## SHS Tech WO View
Asset # | WO # | Type Code | Status Code | Sub Status Code | Date Created | Completed Date | Description | Completion Comments | Requester Remarks
28877 | 192146 | PM | CANCL | CANCL | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
50362 | 192147 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
50363 | 192148 | PM | CANCL | CANCL | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
50981 | 192149 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Testers, Electrical Safety
52268 | 192150 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Testers, Electrical Safety
60022 | 192151 | PM | CANCL | CANCL | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
62124 | 192152 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
62270 | 192153 | PM | CANCL | CANCL | 2023-01-01 02:18 | 2024-06-10 13:03 | PM-MAJOR - Annual PM For: Simulators, Multiparameter
62532 | 192154 | PM | CANCL | CANCL | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: LASERS, SURGICAL, CARBON DIOXIDE
62616 | 192155 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Testers, Electrical Safety
62631 | 192156 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-02-08 17:09 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/08/2023 05:08:37 PM, Rendon Lopez , Elizabeth) Located in LVL 7 OR PM completed, all tests passed. Calibrated scale
62632 | 192157 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-02-01 11:11 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/31/2023 11:10:17 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
63277 | 192158 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
63507 | 192159 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-01-17 13:57 | PM-MAJOR - Annual PM For: Holmium Laser System
63525 | 192160 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
63526 | 192161 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
63786 | 192162 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-06-02 14:31 | PM-MAJOR - Annual PM For: Simulators, Multiparameter | (TC 06/02/2023 02:29:26 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration Attached
63872 | 192163 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-11-30 09:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/29/2023 02:42:12 PM, Rhuland, Courtney) Replaced hood Completed latch recall Calibrated LV and main voltage Calibrated and verified scale Replaced porthole seal Replaced corner grommet Replaced probe shield (TC 11/30/2023 08:27:48 AM, Rhuland, Courtney) PM complete Functional check - 35C, 75% humidity Read 35C, 74.8% humidity Electrical safety - 84.2uA, 189mOhms
63893 | 192164 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Radiofrequency Generators
64964 | 192165 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-06-02 14:44 | PM-MAJOR - Annual PM For: Testers, External Cardiac Pacemaker | (TC 06/02/2023 02:42:27 PM, Chukri, Samir) Completed PM Performed by Geting. Certificate of Calibration attached
64972 | 192166 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Testers, Electrical Safety
64996 | 192167 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-02-13 11:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 02/10/2023 11:36:27 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
65134 | 192168 | PM | CMPLT | CMPLT | 2023-01-01 02:18 | 2023-08-28 11:13 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 08/28/2023 11:12:15 AM, Reid, Emmerson ) PM Complete, see proceedure for details. NOTE: the dose cord works but does not light up.
65194 | 192169 | PM | CANCL | CANCL | 2023-01-01 02:18 | 2024-06-25 13:56 | PM-MAJOR - Annual PM For: Microscopes, Light, Examination
65736 | 192170 | PM | ACTIV | ISSUE | 2023-01-01 02:18 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
65884 | 192171 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-31 12:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/31/2023 12:34:53 PM, Rhuland, Courtney) Replaced front and rear case Replaced pressure sensor due to repeated channel error 354.6690 Replaced mylar gasket Flashed, calibrated and performed PM All tests passed Refer to ASM for details
65890 | 192172 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-08 11:45 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/08/2023 11:44:07 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65914 | 192173 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-06 14:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/06/2023 02:47:33 PM, Rhuland, Courtney) Front and rear case replaced - recall PM complete All tests passed Refer to ASM for details
65951 | 192174 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-07 09:08 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/07/2023 09:08:12 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65958 | 192175 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65991 | 192176 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65998 | 192177 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-05-30 07:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/30/2023 07:31:05 AM, Rhuland, Courtney) Rear case replaced (recall) Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
66034 | 192178 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-05-19 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/19/2023 09:47:30 AM, Rhuland, Courtney) Rear and front cases replaced (recall) Female IUI connector replaced PM complete All tests passed Refer to ASM for details
66112 | 192179 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM - Annual
66498 | 192180 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-26 10:36 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 01/25/2023 10:36:18 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
66503 | 192181 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-04 08:54 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 01/03/2023 08:53:47 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
66524 | 192182 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-08 17:04 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 02/08/2023 05:03:44 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
66525 | 192183 | PM | CANCL | DUP | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM changed to 2 years
66530 | 192184 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-11 10:06 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 01/10/2023 10:05:18 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
66681 | 192185 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-11 10:47 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/10/2023 02:44:18 PM, Rhuland, Courtney) Recall parts replaced (TC 07/11/2023 10:43:05 AM, Rhuland, Courtney) Calibrated RH Sensor (TC 07/11/2023 10:46:20 AM, Ferreira, Qwynn) PM Passed Calibrated scale Calibrated Line voltage Calibrated main vltage Electrical Safety - OK Functional Check - OK
66704 | 192186 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-04-18 13:49 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/18/2023 01:46:10 PM, Ferreira, Qwynn) Tested the scale and it was off by 10gs Janet mentioned the scale was off by more, but checkd with a 5.000kg weight and it read 4.990kgs The scale cable was however tucked under the matress which could cause the heavier weight reading PM Passed Calibrated RH Calibrated line votlage Calibrated main voltage Calibrated system voltage Calibrated scale Replaced RH sensor Replaced 3 grommets Electrical Safety - OK Functional Test - OK
66777 | 192187 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66786 | 192188 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM MAJOR ANNUAL: X2 Monitors, Physiologic, Patient Transport
66843 | 192189 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | BP Check
67227 | 192190 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-02 16:02 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/02/2023 04:00:54 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Replaced blender with overhauled, tested-OK
67405 | 192191 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | 4505
67406 | 192192 | PM | CANCL | CANCL | 2023-01-01 02:19 | 2024-06-11 12:40 | PM-MAJOR - Annual PM For: Testers, ESU
67410 | 192193 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-11 10:00 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/10/2023 10:00:10 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
67411 | 192194 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-11 10:03 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/10/2023 10:03:02 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
67413 | 192195 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-26 10:54 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/25/2023 10:53:23 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. missing diaphgram, replaced
67416 | 192196 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-08 17:06 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 02/08/2023 05:05:58 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
67425 | 192197 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-01 11:29 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/31/2023 11:28:26 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67426 | 192198 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-18 10:34 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/17/2023 10:33:24 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67428 | 192199 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-08 17:05 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 02/08/2023 05:04:38 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Replaced diaphgram
67894 | 192200 | PM | CMPLT | PMFAIL | 2023-01-01 02:19 | 2023-01-09 14:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/09/2023 02:57:52 PM, Rendon Lopez , Elizabeth) Damaged DSOS seal, to be sent out for repair
67917 | 192201 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-06 13:42 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/06/2023 10:58:47 AM, Ferreira, Qwynn) Replaced Hood Replaced two grommets Cleaned out bottom tray Installed Recall Parts Calibrated scale Calibrated LV Calibrated Main voltage Calibrated system voltage (TC 07/06/2023 01:41:18 PM, Ferreira, Qwynn) Functional Test - Passed Electrical Safety - Passed PM Passed
67922 | 192202 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-25 13:41 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 01/25/2023 01:38:45 PM, Ferreira, Qwynn) Note on the device said scale not working Two things were noticed, the pin at the back was pushed out and the latch that tilts the bed was loose Tightened the scale latch - OK Pushed pin back in the for the scale - OK PM Passed Replaced Battery- Yes Calibrated scale Calibrated system voltage Calibrated line voltage Calibrated main voltage Electrical Safety - OK Functional Test - OK
67950 | 192203 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-10-19 13:38 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 10/18/2023 02:52:25 PM, Rhuland, Courtney) Replaced broken hood Replaced broken humidity probe Replaced recall latches Replaced broken porthole door (TC 10/19/2023 01:31:56 PM, Rhuland, Courtney) Tightened base plate to level incubator Calibrated main and LV voltage Calibrated and verified scale Replaced rechargeable battery Functional check - OK Humidity & temp set - 75%, 35C Check - 74.5%, 35C Leakage current - 80.3uA Resistance - 68mOhms
67956 | 192204 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-06-14 13:37 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/14/2023 01:34:58 PM, Rhuland, Courtney) Hood replaced Recall parts replaced Porthole door replaced Power cord replaced PM complete Line and main voltage calibrated Functional check - OK Leakage current - 81.1uA Ground resistance - 45mOhms
67959 | 192205 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-09-28 12:50 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/11/2023 11:42:13 AM, Campbell, Connor) All steps performed under the supervision of Samir The incubator housing roof is cracked still, an order has been placed but tests will need to be re-run when the new housing is installed. Performed a PM of the machine. Refer to Giraffe Omnibed service manual. System Calibration: complete Line voltage calibration: complete Humidifier calibration: complete Scale calibration: complete Tmeperature and Humidity verification: complete/passed Leakage current 55.7uA: passed Ground resistance 126 miliOhm: passed Replaced the power failure alarm battery. Previously a non-rechargeable battery was installed, replaced it with a 150mAh 9V Ni-MH Accu Ultra rechargeable battery Replaced the light as the previous light has housing damage on the base and the neck of the light.
67970 | 192206 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-04-28 14:57 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/28/2023 02:55:55 PM, Ferreira, Qwynn) PM Passed Replaced recall parts Fixed scale cable as one pin was pushed back Calibrated RH Calibrated scale Calibrated main voltage Calibrated system voltage Calibrated line voltage Functional Check - OK Electrical Safety - OK
68356 | 192207 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | not used anymore. replaced with X3
68540 | 192208 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2024-11-22 14:22 | PM Major - Delivery Module SEVO | (TC 11/22/2024 02:22:14 PM, Scero, Tomlyn) 3 year PM parts replaced and tests passed. Results documented on 3 year PM for 1058058
68552 | 192209 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68601 | 192210 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-11-14 10:52 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 11/14/2023 10:51:57 AM, Campbell, Connor) PM complete/passed, refer to ASM
68605 | 192211 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-23 09:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/23/2023 09:38:10 AM, Reid, Emmerson ) Needed to run the PM several times because my computer crashed. Downloaded and cleared log hisotry. Ran ASM PM, all tests passed. PM Complete.
68606 | 192212 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
68610 | 192213 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-23 10:09 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/23/2023 10:08:08 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68627 | 192214 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68631 | 192215 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-04-26 09:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/26/2023 09:45:29 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 200 mohms Leakage current: 69.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68645 | 192216 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-04-05 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2023 09:22:06 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68646 | 192217 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-25 13:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 01:25:32 PM, Canjimavadivel, Kamallesh ) PM tests pass
68704 | 192218 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68718 | 192219 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-08 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/07/2023 02:48:12 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 304 mohms Leakage current: 57.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 02/08/2023 11:23:22 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3498 mamphr New battery capacity: 4048 mamphr PM completed.
68722 | 192220 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-06-09 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 02:11:49 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 244 mohms Leakage current: 76.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 06/09/2023 02:42:20 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3961 mamphr PM completed.
68751 | 192221 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2024-03-26 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68763 | 192222 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-18 14:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/18/2023 02:00:37 PM, Cortez, John) Kam cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4359 mAhs
68770 | 192223 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68780 | 192224 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-05-10 08:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/08/2023 03:21:02 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 221 mohms Leakage current: 70.3 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. Power cord was damaged; replaced power cord (TC 05/10/2023 08:21:58 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3420 mamphr New battery capacity: 4008 mamphr PM completed.
68787 | 192225 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2024-05-08 15:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/18/2023 11:42:26 AM, Cortez, John) Kam cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4146 mAhs
68795 | 192226 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-09-01 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/31/2023 11:26:17 AM, Rendon Lopez , Elizabeth) Replaced battery, PM completed, all tests passed
68808 | 192227 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-30 08:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/29/2023 03:54:46 PM, Cortez, John) Cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioning battery overnight (TC 08/29/2023 08:17:43 AM, Cortez, John) 4105 mAhs
68813 | 192228 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-02 09:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/02/2023 09:56:56 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 10 mohms Leakage current: 72.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68819 | 192229 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68821 | 192230 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68834 | 192231 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-06-16 08:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/15/2023 08:09:59 AM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4111mAh
68836 | 192232 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-12-01 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 12/01/2023 09:47:05 AM, Campbell, Connor) PM passed/complete, refer to ASM refer to procedure for ESA
68859 | 192233 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68863 | 192234 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-07 10:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/07/2023 10:42:37 AM, Reid, Emmerson ) Ran optimal battery conditioning test, new battery capacity 4056 mamphrs, good. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68939 | 192235 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68954 | 192236 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-06-20 08:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/20/2023 08:47:18 AM, Reid, Emmerson ) Replaced the battery. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68959 | 192237 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68975 | 192238 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-09-08 07:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 07:36:00 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 107 mohms Leakage current: 74.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 09/08/2023 07:36:29 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3982 mamphr PM completed.
69005 | 192239 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69009 | 192240 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-12 17:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 04:59:14 PM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4140 mAhs
69019 | 192241 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-09-07 06:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 07:19:45 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 67 mohms Leakage current: 61.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 09/07/2023 06:36:08 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4051 mamphr PM completed.
69022 | 192242 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-06-26 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/26/2023 02:47:21 PM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69080 | 192243 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-24 08:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/21/2023 08:55:05 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69084 | 192244 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-23 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 02:24:38 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69095 | 192245 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-25 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 02:57:56 PM, Canjimavadivel, Kamallesh ) Performed rate calibration PM tests pass
69154 | 192246 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69165 | 192247 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69187 | 192248 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-16 10:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 01/16/2023 10:46:20 AM, Baghirzade, Nara) Originally, unit came down due to channel error. Downloaded error log and noted one error 240.4150 LOG ONLY. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69193 | 192249 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-02 09:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 09:38:01 AM, Rhuland, Courtney) Rate calibrated PM complete All tests passed post calibration Refer to ASM for details
69210 | 192250 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69241 | 192251 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-10-05 10:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 10:39:15 AM, Campbell, Connor) PM passed/complete, refer to ASM
69264 | 192252 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69275 | 192253 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69284 | 192254 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-05-26 10:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/26/2023 09:59:50 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69291 | 192255 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69305 | 192256 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69311 | 192257 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-09-20 08:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 08:42:00 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
69325 | 192258 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69363 | 192259 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-10-25 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/25/2023 10:45:34 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69366 | 192260 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-11-14 11:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/14/2023 11:02:58 AM, Campbell, Connor) PM complete/passed, refer to ASM
69387 | 192261 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-02 13:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 01:46:16 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69396 | 192262 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69406 | 192263 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-09-20 08:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/20/2023 08:40:00 AM, Campbell, Connor) PM complete/passed, refer to ASM
69410 | 192264 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-04-26 12:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/26/2023 12:34:11 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69450 | 192265 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-07 09:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/07/2023 09:41:13 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69455 | 192266 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-10-24 08:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/24/2023 07:59:48 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69479 | 192267 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-17 09:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/04/2023 09:07:29 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69483 | 192268 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-02 11:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 11:19:14 AM, Rhuland, Courtney) Rate calibrated PM complete All tests passed post calibration Refer to ASM for details
69643 | 192269 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-23 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 10:32:17 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69666 | 192270 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69681 | 192271 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
32540 | 192272 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-03-06 20:08 | PM - Monthly | PM completed Mechanical PM done
32560 | 192273 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-31 15:00 | PM - Monthly | PM Major done NRA/AIR cal S FOV done Slice count - 1121053 TOS phantom a(-991.4, 9.7) b(333.6, 14.3) c(132.7, 13.4) d(99.1, 13.7) e(-104.2, 11.5) f(0.4, 14.2) -System passed all tests
32580 | 192274 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-01 15:28 | PM - Monthly | PM Major done Mechanical done Air cal done for 120 kV on all FOV Slice count - 250436 TOS phantom a(-997.7, 10.4) b(334.3, 14.6) c(135.3, 15.1) d(98.0, 14.9) e(-106.7, 12.9) f(1.5, 14.7) -System passed all tests
500320 | 192275 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-07 11:12 | PM - Semi-Annual | (TC 07/07/2023 11:11:40 AM, Green, Paul) Not used and only used for backup.
501440 | 192276 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-22 18:49 | PM - Semi-Annual | A complete PM was done.
501440 | 192277 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-21 20:51 | HARP - Semi-Annual | HARP was done.
502044 | 192278 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-28 14:58 | PM - Semi-Annual | (TC 02/28/2023 02:57:34 PM, Green, Paul) PM as per procedure. Greased up/down motion.
502492 | 192279 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-10-27 14:10 | PM-Major, annual PM | (TC 10/27/2023 02:07:08 PM, Green, Paul) Moved spare down to MSH 2B so MRX2B could be removed. Performed PM as per procedure. Full calibrations done and fully functional tests done. Reinstalled back in room and took spare injector up to 5. Had a call from MRT John and he was having some problems with the injector. Went over and looked at several issues that he was having. Was able to resolve these issues.
541287 | 192280 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM
1046652 | 192281 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1047101 | 192282 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-27 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 01/27/2023 10:37:43 AM, Reid, Emmerson ) Ran Optimal Battery Conditioning test, new battery capacity 4055 msmphr. Ran ASM MSH Pcu, all tests passed. PM Complete.
1047842 | 192283 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-12 09:39 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/12/2023 09:38:51 AM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1048602 | 192284 | PM | CANCL | CANCL | 2023-01-01 02:19 | 2024-06-11 09:10 | PM-MAJOR - Annual PM
1048603 | 192285 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-06-21 10:46 | PM-MAJOR - Annual PM | (TC 06/21/2023 10:20:26 AM, Reid, Emmerson ) Upgraded firmware to V040201. Pushed Palliative - Dec 16/15 library. PM complete, see proceedure for details.
1048606 | 192286 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-07-12 10:01 | PM-MAJOR - Annual PM | (TC 07/11/2023 10:00:29 AM, Cortez, John) PM passed (see procedure for results)
1048627 | 192287 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048637 | 192288 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048639 | 192289 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048643 | 192290 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048651 | 192291 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048654 | 192292 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048655 | 192293 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048663 | 192294 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048666 | 192295 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048676 | 192296 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048677 | 192297 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048686 | 192298 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048688 | 192299 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048691 | 192300 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048694 | 192301 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048697 | 192302 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048700 | 192303 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049658 | 192304 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-18 10:39 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/17/2023 10:38:16 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1050045 | 192305 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-08-03 08:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 08:53:49 AM, Rendon Lopez , Elizabeth) Rear case cracked, replaced under recall. PM completed, all tests passed
1050089 | 192306 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Scales, Infant
1050143 | 192307 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | X2 inactive. replaced with X3
1050260 | 192308 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-09-08 11:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/08/2023 11:34:32 AM, Campbell, Connor) PM complete, refer to ASM
1050476 | 192309 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-04 08:57 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 01/03/2023 08:55:52 AM, Rendon Lopez , Elizabeth) PM parts replaced, all tests passed. Expiratory cassette membrane replaced SN: 102767.
1050478 | 192310 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-17 12:01 | PM-MAJOR- Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | Performed PM Replaced cable guide Earth leakage 20 uA Resistance 131 mOhm Pre check pass
1050480 | 192311 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-03 11:34 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | Replaced filters and nozzles Pre check pass
1050481 | 192312 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-16 12:05 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | Performed PM Replaced cable guide Earth leakage 20 uA Resistance 130 mOhm Pre check pass
1050482 | 192313 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-16 12:03 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | Performed PM Earth leakage 21 uA Resistance 129 mOhm Pre check pass
1051379 | 192314 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Scales, Infant
1051733 | 192315 | PM | CANCL | CANCL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Transport
1053033 | 192316 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM-MAJOR-Annual PM For : Aspirator unit battery
1053447 | 192317 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-06-13 14:23 | PM MAJOR-Stretchers, Mobile, Hospital, Patient Transfer | (TC 06/13/2023 02:20:39 PM, Chukri, Samir) Completed PM by Stryker Tech see attached (TC 06/13/2023 02:29:43 PM, Ferreira, Qwynn) Supervised and assisted the tech with the PM
1053489 | 192318 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-27 14:33 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult | (TC 02/27/2023 02:24:14 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 9862 hrs Total vent time: 5141 hrs
1053490 | 192319 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-28 13:55 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult | (TC 02/28/2023 01:50:49 PM, Canjimavadivel, Kamallesh ) Replaced expieratory cassette membrane Replaced filters and nozzles Pre check pass Total op time: 13824 hrs Total vent time: 8362 hrs
1053491 | 192320 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-28 09:48 | PM-MAJOR - Annual PM For: Ventilators | (TC 02/28/2023 09:20:37 AM, Canjimavadivel, Kamallesh ) Replaced expiratory cassette membrane Replaced 2x batteries Replaced filters and nozzles Pre check pass Total op time: 13116 hrs Total vent time: 6541 hrs
1053492 | 192321 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-03-20 09:52 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/20/2023 09:49:20 AM, Canjimavadivel, Kamallesh ) Replaced battery Replaced cable guide Replaced filters and nozzles Pre check pass Total op time: 18114 hrs Total vent time: 10480 hrs
1053493 | 192322 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-02-28 09:05 | PM-MAJOR - Annual PM For: Ventilators | (TC 02/28/2023 09:00:24 AM, Canjimavadivel, Kamallesh ) Replaced expiratory cassette membrane Replaced filters and nozzles Pre check pass Total op time: 10976 hrs Total vent time: 6135 hrs
1053728 | 192323 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-12 12:16 | PM-MAJOR- Annual PM for: Pacemakers, Cardiac, External | (TC 01/12/2023 12:15:13 PM, Scero, Tomlyn) PM checks complete. Device safety and technical checks passed.
1053995 | 192324 | PM | ACTIV | ISSUE | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: Insufflators
1054460 | 192325 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-04 07:55 | Lubrication cycle | Refer to WO#190321
1054462 | 192326 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-04 07:57 | Lubrication cycle | Refer to WO#192326
1054468 | 192327 | PM | CMPLT | CMPLT | 2023-01-01 02:19 | 2023-01-04 07:58 | Lubrication cycle | Refer to WO#192327
1055180 | 192328 | PM | CNP | ENL | 2023-01-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055562 | 192329 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-03-21 11:28 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/21/2023 11:24:16 AM, Canjimavadivel, Kamallesh ) Replaced cable guide Replaced filters and nozzles Pre check pass Total op time: 14938 hrs Total vent time: 7458 hrs
1057764 | 192330 | PM | CANCL | CANCL | 2023-01-01 02:20 |  | PM MAJOR-Monitors, Physiologic, Central Station | Storage
1057892 | 192331 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-03-22 11:14 | PM MAJOR - ANNUAL PM for: High Flow, Warmer | (TC 03/22/2023 11:08:18 AM, Reid, Emmerson ) General Alarm Test - OK Low Recirculating Fluid Alarm Test - OK Over Temp Alarm Test - OK Disposable Alarm Test - OK Fluid Warmer Alarm Test - OK Drained old fluid and added new recirculating fluid Performed Functional Test - OK Temperature 40.8 C Replaced O-rings Cleaned fan filter Refilled with new recirculating fluid Electrical Safety - Current Leakage - 73.3 uA - Ground Resistance - 43 mohms PM Comlplete
1058144 | 192332 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-03-15 09:02 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 03/08/2023 09:01:14 AM, Rendon Lopez , Elizabeth) Incoming inspection from PM with ventilator 1058146.
1058145 | 192333 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-03-15 09:02 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 03/08/2023 09:02:13 AM, Rendon Lopez , Elizabeth) Incoming inspection from PM with ventilator 1058146.
1058146 | 192334 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-03-15 09:01 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | (TC 02/09/2023 02:11:29 PM, Rendon Lopez , Elizabeth) Sent out ventilator for PM RA#S2023-1165 (TC 03/08/2023 08:59:49 AM, Rendon Lopez , Elizabeth) Incoming inspection, tested-OK
1058277 | 192335 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-07-14 13:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/14/2023 10:01:23 AM, Canjimavadivel, Kamallesh ) PM tests all pass, see logs for details Optimal battery conditioning started (TC 07/14/2023 01:08:59 PM, Canjimavadivel, Kamallesh ) Battery conditioning failed multiple times, replaced battery.
1058542 | 192336 | PM | CANCL | CANCL | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible |  | Every year needs dust cleaning, every two years the battery is changed.
1058543 | 192337 | PM | CANCL | CANCL | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible |  | Every year needs dust cleaning, every two years the battery is changed.
1058544 | 192338 | PM | CANCL | CANCL | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible |  | Every year needs dust cleaning, every two years the battery is changed.
1058545 | 192339 | PM | CANCL | CANCL | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible |  | Every year needs dust cleaning, every two years the battery is changed.
1058546 | 192340 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-12-18 11:11 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 12/18/2023 11:10:32 AM, Reid, Emmerson ) Replace the battery, PM complete. | Every year needs dust cleaning, every two years the battery is changed.
1058552 | 192341 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-12-18 08:25 | PM Major Annual: Power Systems, Uninterruptible | (TC 12/15/2023 08:24:17 AM, Reid, Emmerson ) Replaced the battery. PM Complete.
1058605 | 192342 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-01-13 08:33 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 01/12/2023 08:32:44 AM, Reid, Emmerson ) Vendor came in and performed the PM.
1058606 | 192343 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-01-13 08:35 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 01/12/2023 08:34:51 AM, Reid, Emmerson ) Vendor came in and performed the PM.
1059058 | 192344 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059134 | 192345 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059135 | 192346 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059136 | 192347 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059137 | 192348 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059138 | 192349 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059139 | 192350 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059140 | 192351 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059141 | 192352 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059142 | 192353 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-12-18 11:15 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 12/18/2023 11:14:09 AM, Reid, Emmerson ) Replaced the battery.
1059143 | 192354 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059144 | 192355 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059145 | 192356 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059146 | 192357 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059147 | 192358 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-01-30 09:27 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 01/30/2023 09:23:33 AM, Reid, Emmerson ) Switched this unit with a spare. The unit was giving an SpO2 Malfunction and SpO2 was not working with multiple sensors. I powered it down and left overnight. This morning it is working now. Will keep as a spare.
1059148 | 192359 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059149 | 192360 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059150 | 192361 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059151 | 192362 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059152 | 192363 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059153 | 192364 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059154 | 192365 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059155 | 192366 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059158 | 192367 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059159 | 192368 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059160 | 192369 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059161 | 192370 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059162 | 192371 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059163 | 192372 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059164 | 192373 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-12-18 11:17 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 12/18/2023 11:16:30 AM, Reid, Emmerson ) Replaced the battery
1059165 | 192374 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059166 | 192375 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059168 | 192376 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059169 | 192377 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059170 | 192378 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059171 | 192379 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059172 | 192380 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059173 | 192381 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059174 | 192382 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059175 | 192383 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059176 | 192384 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059177 | 192385 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059178 | 192386 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059179 | 192387 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059180 | 192388 | PM | CANCL | CANCL | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059181 | 192389 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059182 | 192390 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059183 | 192391 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059184 | 192392 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059185 | 192393 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059187 | 192394 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059188 | 192395 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1061154 | 192396 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-01-24 16:45 | PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 01/24/2023 04:44:31 PM, Chau, Ken) Closing WO as system is not under UHN list
1061364 | 192397 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-02-13 20:52 | HARP - Semi Annual | HARP
1061364 | 192398 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-02-28 20:54 | PM - Semi Annual | PM was done.
1061867 | 192399 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Waste-Disposal Units
1063181 | 192400 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-02-21 09:09 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Adult/Pediatric | (TC 02/21/2023 09:07:34 AM, Canjimavadivel, Kamallesh ) Ahmad from Bomimed performed PM and upgraded software to Ver 3.0.3. see attached report.
1063182 | 192401 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-02-21 09:10 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Adult/Pediatric | (TC 02/21/2023 09:09:28 AM, Canjimavadivel, Kamallesh ) Ahmad from Bomimed performed PM and upgraded software to Ver 3.0.3. see attached report.
1063310 | 192402 | PM | ACTIV | ISSUE | 2023-01-01 02:20 |  | PM-MAJOR - ANNUAL PM For: CO2 Laser
1063799 | 192403 | PM | CMPLT | CMPLT | 2023-01-01 02:20 | 2023-12-28 15:26 | PM-Major, annual PM | (TC 11/03/2023 03:51:04 PM, Green, Paul) Replaced pots. Performed position calibration. Had problems with pressure calibration. (TC 11/08/2023 12:13:27 PM, Green, Paul) Performed pressure calibration, ok. Performed functional testing, and problem with B side not engaging. Replaced B side docking pin sensor cable. Problem persists. (TC 11/08/2023 12:13:27 PM, Green, Paul) Performed pressure calibration, ok. Performed functional testing, and problem with B side not engaging. Replaced B side docking pin sensor cable. Problem persists. (TC 12/28/2023 03:24:19 PM, Green, Paul) Calibrated pressure with new cal syringes. Passes. Performed final test, ok. Will reinstall back in room when time permits.
1064331 | 192404 | PM | CANCL | CANCL | 2023-01-01 02:20 |  | PM-MAJOR - Annual PM For: Scales, Infant
| 192405 | PA | CMPLT | CMPLT | 2023-01-03 08:41 | 2023-01-09 08:09 | sterile water for altrix
1048927 | 192406 | CM | CMPLT | CMPLT | 2023-01-03 09:29 | 2023-01-09 14:24 | Equipment malfunction | (TC 12/29/2022 09:29:13 AM, Rendon Lopez , Elizabeth) Sent out for repair (TC 01/09/2023 02:24:10 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1057509 | 192407 | CM | CMPLT | CMPLT | 2023-01-03 09:30 | 2023-01-09 14:25 | Not detected | (TC 12/29/2022 09:30:13 AM, Rendon Lopez , Elizabeth) Sent out for repair (TC 01/09/2023 02:25:13 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1053432 | 192408 | CM | CMPLT | CMPLT | 2023-01-03 09:32 | 2023-01-09 14:15 | Not detected | (TC 12/29/2022 09:31:55 AM, Rendon Lopez , Elizabeth) Sent out for repair (TC 01/09/2023 02:15:00 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1051391 | 192409 | CM | CMPLT | CMPLT | 2023-01-03 09:33 | 2023-01-09 14:23 | Physical damage | (TC 12/29/2022 09:33:21 AM, Rendon Lopez , Elizabeth) Sent out for repair (TC 01/09/2023 02:21:52 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1048003 | 192410 | CM | CANCL | CANCL | 2023-01-03 09:34 |  | Malfunction
1046276 | 192411 | CM | CMPLT | CMPLT | 2023-01-03 11:00 | 2023-02-06 14:19 | MSH CFRH 7flr APLIO 300 SYSTEM WENT BLANK MONDAY WHEN CASE STARTED, | (TC 01/03/2023 11:00:55 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 SYSTEM WENT BLANK MONDAY WHEN CASE STARTED PROBE PVT-781VTE SN# FSC1934521 MFG MARCH 2019 (TC 01/03/2023 08:02:00 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 SYSTEM WENT BLANK MONDAY WHEN CASE STARTED PROBE PVT-781VTE SN# FSC1934521 MFG MARCH 2019 (TC 01/03/2023 08:02:00 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 SYSTEM WENT BLANK MONDAY WHEN CASE STARTED PROBE PVT-781VTE SN# FSC1934521 MFG MARCH 2019 | MSH CFRH 7flr APLIO 300 SYSTEM WENT BLANK MONDAY WHEN CASE STARTED
| 192412 | CM | CMPLT | CMPLT | 2023-01-03 11:33 | 2023-07-11 07:28 | MSH NICU REQUESTING TO PURCHASE A ZONARE ZS3 WITH FOUNDATION SUPPORT | (TC 01/03/2023 11:33:30 AM, Whittington, Scott) MSH NICU REQUESTING TO PURCHASE A ZONARE ZS3 WITH FOUNDATION SUPPORT PROVIDE JOHN BALCOM , MGR OF NICU WITH INFO TO PROCEED WITH CAPITAL ACQUISITION FORM (TC 01/04/2023 03:54:34 PM, Whittington, Scott) MSH NICU REQUESTING TO PURCHASE A ZONARE ZS3 WITH FOUNDATION SUPPORT (TC 01/12/2023 12:32:03 PM, Whittington, Scott) MSH NICU REQUESTING TO PURCHASE A ZONARE ZS3 WITH FOUNDATION SUPPORT | MSH NICU REQUESTING TO PURCHASE A ZONARE ZS3 WITH FOUNDATION SUPPORT
1054523 | 192413 | CM | CMPLT | CMPLT | 2023-01-03 11:42 | 2023-01-03 12:59 | V4 lead broken | (TC 01/03/2023 12:53:36 PM, Canjimavadivel, Kamallesh ) Replaced damaged V4 lead set. Replaced V1-3 set as it was affected under recall. Repalced damaged V7 lead as well
67961 | 192414 | CM | CMPLT | CMPLT | 2023-01-03 11:54 | 2023-01-03 11:59 | door falling off and large moisture build up | (TC 01/03/2023 11:54:28 AM, Ferreira, Qwynn) the side door rail had some minor corrosion/chemical build up hindering the door and springs. Removed the rail, cleaned the rail, springs, and slots of the rail. Reinstalled the rail, springs, side panel, and door Tested - OK Ran the incubator to test the humidity No more build up than normal was seen on the hood of the incubator. Tested at 85% RH Functiona test - OK
1065432 | 192415 | CM | CMPLT | CMPLT | 2023-01-03 13:15 | 2023-01-03 13:16 | Initial Accept | (TC 01/03/2023 01:15:20 PM, Samanian, Kamran) performed electrical safety test GR=30mOHM, LC=55.6uA system self test passed
1065418 | 192416 | CM | CMPLT | CMPLT | 2023-01-03 13:22 | 2023-01-04 11:37 | opened a ticket to helpdesk for adding MAC to NAC | (TC 01/03/2023 01:22:19 PM, Samanian, Kamran) reviewed the process what to do. Sent email to help desk to open a ticket. #N-IM039695936. waiting for mirza to add it. (TC 01/04/2023 11:36:32 AM, Samanian, Kamran) Done by Mirza!, tested and passed.
1056430 | 192417 | CM | CMPLT | CMPLT | 2023-01-03 14:10 | 2023-01-03 14:12 | "BP not reading" | (TC 01/03/2023 02:10:04 PM, Cortez, John) Confirmed issue Hose defective, replaced with spare All parameters tested ok
| 192418 | CM | CMPLT | CMPLT | 2023-01-03 14:30 | 2023-02-03 15:47 | BD Alaris plastics parts order (recall parts) | (TC 01/03/2023 02:30:16 PM, Cortez, John) Requesting order for (no charge): 10x syringe front cases 10x syringe rear cases 10x syringe left handles (TC 01/03/2023 02:37:22 PM, Cortez, John) SC# 1002848990
501072 | 192419 | CM | CMPLT | CMPLT | 2023-01-03 15:24 | 2023-07-11 07:27 | MSH CEOU-3224(Rm5) NOT NETWORKING NW JACK PUNCHED IN | (TC 01/03/2023 03:25:10 PM, Whittington, Scott) CEOU-3224(Rm5) NOT NETWORKING NW JACK PUNCHED IN REPAIRED NW JACK MOUNT FOUND SYSTEM NOT CONNECTING MOVED TO CONSULT ROOM, CONNECTED IMMEDIATELY , ABLE TO UPLOAD STUDIES TO PACS AND DOWNLOAD MWL BACK IN ROOM 5, A SERIES OF REBOOT, NW CABLE RECONNECTION, AND SYSTEM CONNECTED TO NW WILL ADVISE MIRZA | CEOU-3224(Rm5) NOT NETWORKING NW JACK PUNCHED IN
1056125 | 192420 | PM | CMPLT | CMPLT | 2023-01-04 08:00 | 2023-01-04 08:01 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/03/2023 08:00:15 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
66503 | 192421 | CM | CMPLT | CMPLT | 2023-01-04 08:52 | 2023-01-04 08:54 | Recorder not printing | (TC 01/03/2023 08:52:38 AM, Rendon Lopez , Elizabeth) Removed from room, replaced recorder with spare. Tested-OK (TC 01/03/2023 08:52:38 AM, Rendon Lopez , Elizabeth) Cleaned damaged recorder
52252 | 192422 | CM | CMPLT | CMPLT | 2023-01-04 14:55 | 2023-01-04 14:57 | Bed cannot straighten and stuck wheel | (TC 01/04/2023 02:55:36 PM, Ferreira, Qwynn) repaired the tilt latch as it was missing a screw added grease to the tilt assembly Lifted incubator and inspected wheel Seemed dry with no lubrication Put bearing oil in the wheel and is working fine Functional Check - OK
1054468 | 192423 | CM | CMPLT | CMPLT | 2023-01-04 15:21 | 2023-01-13 14:50 | Flow slows down halfway | (TC 01/05/2023 10:51:32 AM, Canjimavadivel, Kamallesh ) Performed testing scenarios to try and recreate issue. Meanwhile reached out to vendor for additional info (TC 01/06/2023 10:52:15 AM, Canjimavadivel, Kamallesh ) Received reply from vendor with recall notice related to issue. Informed ICU and waiting recommendation from vendor as repait kit not available at this time. (TC 01/09/2023 12:58:16 PM, Canjimavadivel, Kamallesh ) Device removed from service until kit is available to repair but will remain in ICU . Recommendation is to use manual pressure bags in the meantime.
69595 | 192424 | CM | CMPLT | CMPLT | 2023-01-04 15:56 | 2023-01-04 15:58 | "not working" | (TC 01/04/2023 03:56:56 PM, Cortez, John) Cleaned, Downloaded error/event logs, Patient-side pressure sensor errors were logged, High pressure sensor replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
| 192425 | CM | CMPLT | CMPLT | 2023-01-05 07:45 | 2023-01-04 07:45 | Checking the IBE ports with Adeel | (TC 01/04/2023 07:44:59 AM, Samanian, Kamran) done!
| 192426 | CM | CMPLT | CMPLT | 2023-01-05 07:46 | 2023-01-04 07:46 | Checking RS232 with carol about CIM monitor | (TC 01/04/2023 07:46:08 AM, Samanian, Kamran) Done!
1065418 | 192427 | CM | CMPLT | CMPLT | 2023-01-05 07:47 | 2023-01-04 07:48 | filled out the NAC form and sent to IS security | (TC 01/04/2023 07:48:07 AM, Samanian, Kamran) sent out the NAC form. Also checked the device connection with server and all is good.
| 192428 | CM | CMPLT | CMPLT | 2023-01-05 08:30 | 2023-01-26 14:22 | Panda warmer blender overhaul | Received
| 192429 | CM | CMPLT | CMPLT | 2023-01-05 10:15 | 2023-02-08 14:11 | Order Battery Module for servo U/N Ventilator | (TC 01/05/2023 10:18:21 AM, Chukri, Samir) Quote Number 000257102.1 order 4 battery module for Servo U. (TC 02/08/2023 02:09:37 PM, Chukri, Samir) Received 4 batteries waiting for the 10 Batteries. charge it to R.T department (TC 01/05/2023 10:18:21 AM, Chukri, Samir) Quote Number 000257102.1 order 4 battery module for Servo U. | Quote # 000257102.1
1064202 | 192430 | CM | CMPLT | CMPLT | 2023-01-05 10:35 | 2023-01-09 12:58 | O2 concentration reading low | (TC 01/05/2023 10:35:21 AM, Canjimavadivel, Kamallesh ) Verified alarm in logs. Performed pre-check tests and did not identify any issues. O2 related issues occurred in June shortly after incoming. Reached out to vendor to investigate further. (TC 01/09/2023 12:55:35 PM, Canjimavadivel, Kamallesh ) Alex from Getinge came to take a look. No major issues identified, discovered pre-check was not performed prior to use and believes this caused issue. See attachment for findings. Device returned to service.
1056544 | 192431 | CM | CMPLT | CMPLT | 2023-01-05 11:02 | 2023-01-05 11:04 | Low battery, does not check in to the dashboard | (TC 01/05/2023 11:02:23 AM, Baghirzade, Nara) Battery was down to 24%; device could not check in to the dashboard. When turning the device on, it indicates "device failed". Changed the AED battery; checked network settings and ensured that access points list is correct. Manually sent check-in to the dashboard. After powering off the device, the checkmark was present in the window.
67830 | 192432 | CM | CMPLT | CMPLT | 2023-01-05 12:22 | 2023-01-05 13:16 | PM yearly | (TC 01/05/2023 01:15:33 PM, Scero, Tomlyn) CO2 Calibration complete. Flow checked. PM complete.
| 192433 | PA | CMPLT | CMPLT | 2023-01-05 12:36 | 2023-05-30 12:00 | GCX rail order for POD 2 AGM mock ups | (TC 01/05/2023 12:52:37 PM, Scero, Tomlyn) Received quote from BOMImed and ordered on ereq (TC 01/31/2023 08:13:53 AM, Scero, Tomlyn) parts received (TC 05/30/2023 11:58:54 AM, Scero, Tomlyn) Complete on all POD2 AGMs 68537 1058056 68374 68368 68369 68377 68554
65916 | 192434 | CM | CMPLT | CMPLT | 2023-01-05 13:28 | 2023-01-05 13:35 | Channel error | (TC 01/05/2023 01:33:29 PM, Baghirzade, Nara) Downloaded error logs; noted that the pressure sensor failure is a repeated error. In the last repair, the pressure sensor was re-seated. Since it is a repeatable error, the pressure sensor has been replaced. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. | When plugging in to the PCU, error 354.6770 appears indicating a faulty pressure sensor
1059828 | 192435 | CM | CMPLT | CMPLT | 2023-01-05 14:09 | 2023-01-05 14:10 | Cleaning | (TC 01/05/2023 02:09:50 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
| 192436 | CM | CMPLT | CMPLT | 2023-01-05 17:32 | 2023-01-05 17:32 | Site coverage and adding software to new computer
| 192437 | CM | CMPLT | CMPLT | 2023-01-05 17:41 | 2023-01-05 17:41 | Site coverage and new laptop software installed
65381 | 192438 | CM | CMPLT | CMPLT | 2023-01-06 07:40 | 2023-01-06 07:40 | QC Checklist | (TC 01/06/2023 07:40:02 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 North.
67230 | 192439 | PM | CMPLT | CMPLT | 2023-01-06 09:11 | 2023-01-06 09:13 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/05/2023 09:11:35 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed Blender replaced
67391 | 192440 | CM | CMPLT | CMPLT | 2023-01-06 09:15 | 2023-01-06 09:31 | Bed assembly cracked | (TC 01/04/2023 09:15:19 AM, Rendon Lopez , Elizabeth) Bed assembly cracked and missing bottom panel. Replaced bed assembly and bottom panel. Tested all tilt functions-OK
1055097 | 192441 | CM | CMPLT | CMPLT | 2023-01-06 09:24 | 2023-01-06 09:26 | Bottle side pressure sensor is defective | (TC 01/06/2023 09:23:46 AM, Baghirzade, Nara) Error logs downaloded; recurring pressure sensor error noted in logs. Did an analog sensor test and bottle side voltage would stay at 4.99V and would not change. Replaced pressure sensor. Calibrated pressure and rate; check-in tests passed.
| 192442 | CM | CMPLT | CMPLT | 2023-01-06 10:09 | 2023-02-06 08:18 | Meeting in ENDO for move date | (TC 01/06/2023 10:09:21 AM, Samanian, Kamran) done! Discussed about what to move (TC 01/18/2023 08:09:58 AM, Samanian, Kamran) done! (TC 02/06/2023 08:18:05 AM, Samanian, Kamran) moved on JAN 20. done!
| 192443 | CM | CMPLT | CMPLT | 2023-01-06 10:11 | 2023-01-06 10:11 | rebuilding PHYSIO3 server. MSHPHIIXPHY3 due to crash caused by OS patch | (TC 01/05/2023 10:11:16 AM, Samanian, Kamran) done. tested and passed
1061116 | 192444 | CM | CMPLT | CMPLT | 2023-01-06 12:38 | 2023-01-06 12:39 | power cord plug is bent | (TC 01/06/2023 12:38:25 PM, Samanian, Kamran) verified the problem, repaired the plug, tested the bed, BAT needs to be charged. bed is functional
1063746 | 192445 | CM | CMPLT | CMPLT | 2023-01-06 14:09 | 2023-01-06 14:13 | Cleaning | (TC 01/06/2023 02:12:13 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1059938 | 192446 | CM | CMPLT | CMPLT | 2023-01-06 15:26 | 2023-01-06 15:27 | Leads not working | (TC 01/06/2023 03:25:47 PM, Canjimavadivel, Kamallesh ) Replaced RL, LL, LA leads as they were damaged
1063193 | 192447 | CM | CMPLT | CMPLT | 2023-01-06 16:50 | 2023-01-06 16:50 | Arm swings too much. |  | Repaired monitor arm so half of it would stay locked in place.
500869 | 192448 | CM | CMPLT | CMPLT | 2023-01-06 16:56 | 2023-01-06 16:56 | Monitor up/down cable broken. Database full. |  | Replaced cable and rebuilt the database.
1050069 | 192449 | CM | CMPLT | CMPLT | 2023-01-09 08:17 | 2023-01-09 14:41 | Not turning on | (TC 01/09/2023 02:29:44 PM, Canjimavadivel, Kamallesh ) Battery and power cord tested okay on different device, suspected issue with power supply. Swapped power supply from a retired unit, turning on and charging. Installed bp hose and cuff
67153 | 192450 | CM | CMPLT | CMPLT | 2023-01-09 08:18 | 2023-01-09 08:19 | BP not working | (TC 01/09/2023 08:18:37 AM, Samanian, Kamran) verified the problem. NBP hose was loose. Tighetned and tested all functions on the monitor and passed.
67135 | 192451 | CM | CMPLT | CMPLT | 2023-01-09 09:06 | 2023-01-09 10:09 | Inflates but doesn't take reading | (TC 01/09/2023 10:07:32 AM, Canjimavadivel, Kamallesh ) Leak in cuff prevented from reaching target pressure. Replaced cuff and tested with simulator, tested ok
1060001 | 192452 | CM | CMPLT | CMPLT | 2023-01-09 10:15 | 2023-01-09 10:56 | No signal, unable to get reading | (TC 01/09/2023 10:55:08 AM, Canjimavadivel, Kamallesh ) Replaced damaged LA and V9 leads. Tested ok with simulator
1056464 | 192453 | CM | CMPLT | CMPLT | 2023-01-09 10:27 | 2023-01-09 10:27 | BPO not working | replaced adult cuff. tested BP, SpO2 and temp. Functional.
1055946 | 192454 | CM | CMPLT | CMPLT | 2023-01-09 10:28 | 2023-01-09 10:28 | BP not working | Repaired cuff hose. replaced adult cuff. tested. BP, SpO2 and temp. functional
| 192455 | CM | CMPLT | CMPLT | 2023-01-09 12:44 | 2023-01-16 08:39 | OR HDMI-DVI cable order - spare parts | (TC 01/09/2023 12:50:32 PM, Samanian, Kamran) S/C# 1002852999 (TC 01/09/2023 12:50:32 PM, Samanian, Kamran) S/C# 1002852999 Received PO# 7202140789 (TC 01/09/2023 12:50:32 PM, Samanian, Kamran) S/C# 1002852999 Received PO# 7202140789
67935 | 192456 | CM | CMPLT | CMPLT | 2023-01-09 12:58 | 2023-01-09 12:59 | Broken Door and not turning on | (TC 01/09/2023 12:58:10 PM, Ferreira, Qwynn) replaced two door latches replaced one port door grommet Switch at the back was turned off causing the device to not turn on Functional Test - OK
1065409 | 192457 | CM | CMPLT | CMPLT | 2023-01-09 13:48 | 2023-01-09 13:48 | long NBP hose needed | (TC 01/09/2023 01:47:48 PM, Samanian, Kamran) long NBP hose was made to be macthed with the monitor and cuff for regular function,. Taken to the IVF and given to Wendy.
| 192458 | CM | CMPLT | CMPLT | 2023-01-09 13:49 | 2023-01-09 13:49 | OR team meeting | (TC 01/09/2023 01:49:18 PM, Samanian, Kamran) done!
65367 | 192459 | CM | CMPLT | CMPLT | 2023-01-09 14:01 | 2023-01-09 14:14 | Will not power on | (TC 01/04/2023 02:02:57 PM, Cortez, John) After troubleshooting, discovered that power button is faulty. Swapped button with button from retired unit. All parameters tested ok
1062136 | 192460 | CM | CMPLT | CMPLT | 2023-01-09 14:08 | 2023-01-25 13:57 | Loud noise coming from head board | (TC 01/09/2023 02:09:39 PM, Ferreira, Qwynn) Checked the bed and noticed no visible damage When powered on and moving, the motor makes a loud grinding noise. No error codes present at the time Reached out to Stryker to inform them and waiting for repair to be completed. (TC 01/25/2023 01:56:44 PM, Ferreira, Qwynn) Tested the bed with Mark, the motor was replaced under warranty Unit was notified
1062256 | 192461 | CM | CMPLT | CMPLT | 2023-01-09 14:11 | 2023-01-09 14:13 | Error code on the display and no functions working | (TC 01/09/2023 02:11:27 PM, Ferreira, Qwynn) Checkedthe display and saw the lock function was turned on which would allow the head, foot, or bed to be raised or lowered. Once the bed was reset, the scale was zeroed and a new patient was added. All functions and features are working correctly.
1051307 | 192462 | CM | CMPLT | CMPLT | 2023-01-09 14:28 | 2023-01-09 14:30 | Not charging or working | (TC 01/09/2023 02:28:32 PM, Rendon Lopez , Elizabeth) Cleaned contacts, charged, tested-OK
63347 | 192463 | CM | CMPLT | NFF | 2023-01-09 14:49 | 2023-01-09 14:51 | Only one lead is showing, regardless of ECG lead placement | (TC 01/09/2023 02:48:51 PM, Ferreira, Qwynn) Cleaned the leads Attached the simulator for the 12 lead ECG was displaying all leads and rthyms that were used. Tried to wiggle the cables to see if any leads were bad or possibly cabeling - None Removed leads V1-V6 to run just RA,RL,LR,LL and all rythms were displaying correctly. Returned to unit
67882 | 192464 | CM | CMPLT | CMPLT | 2023-01-09 14:57 | 2023-02-21 11:11 | Created by Failed PM WO # 190751 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/10/2023 09:11:00 AM, Rendon Lopez , Elizabeth) Sent out for repair RMA # 312794 (TC 02/21/2023 03:53:42 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK. Downloaded configuration and library. | (TC 01/09/2023 02:56:42 PM, Rendon Lopez , Elizabeth) Damaged DSOS seal, to be sent out for repair
67894 | 192465 | CM | CMPLT | CMPLT | 2023-01-09 14:58 | 2023-02-21 16:00 | Created by Failed PM WO # 192200 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/10/2023 09:10:16 AM, Rendon Lopez , Elizabeth) Sent out for repair RMA # 312794 (TC 02/21/2023 03:59:12 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK. Downloaded configuration and library. | (TC 01/09/2023 02:57:52 PM, Rendon Lopez , Elizabeth) Damaged DSOS seal, to be sent out for repair
1059778 | 192466 | FC | CMPLT | CMPLT | 2023-01-10 08:44 | 2023-01-10 08:44 | QC Checklist | (TC 01/10/2023 08:43:55 AM, Baghirzade, Nara) QC Checklist completed. Returned to 11 South
1065478 | 192467 | CM | CMPLT | CMPLT | 2023-01-10 08:48 | 2023-01-10 08:52 | Received loaner from 3M | (TC 01/10/2023 08:48:32 AM, Baghirzade, Nara) Received loaner VAC from 3M in place for the one that is being PM'ed. Attached the drop off receipt and inspection card to this WO. Dropped off the VAC with Sarabeth. | Received loaner VAC from 3M while MSH VAC is out for PM.
502044 | 192468 | CM | CMPLT | CMPLT | 2023-01-10 08:53 | 2023-01-10 09:01 | High humidity Problem | (TC 01/09/2023 09:01:15 AM, Paz, Ramon (Tino)) Replaced head dessicant
1056468 | 192469 | CM | CMPLT | CMPLT | 2023-01-10 09:21 | 2023-01-10 09:22 | Reported blood pressure not working | (TC 01/10/2023 09:20:57 AM, Baghirzade, Nara) Device had depleted battery; recharged the Root. Tested with simulator the BP and SpO2; modules are functional. Checked BP hose and cuffs; all are good. Temperature module is functional. No fault found with device. Most likely the battery was too low to take BP measurements at the time.
1055920 | 192470 | CM | CMPLT | CMPLT | 2023-01-10 09:43 | 2023-01-20 07:57 | Broken basket; ordering new basket | (TC 01/10/2023 09:46:43 AM, Baghirzade, Nara) SC 1002853671 Ordered a basket for the rolling stand (TC 01/19/2023 08:46:52 AM, Baghirzade, Nara) Received basket. PO 7202141268 | Basket is broken off the stand. The power cord prongs are also broken.
1056563 | 192471 | CM | CMPLT | CMPLT | 2023-01-10 13:25 | 2023-01-10 13:28 | Battery replacement | (TC 01/10/2023 01:24:54 PM, Canjimavadivel, Kamallesh ) Replaced battery and upgraded software to version 6.03.006
| 192472 | CM | CMPLT | CMPLT | 2023-01-10 13:31 | 2023-01-10 13:32 | Cable for ICU | (TC 01/10/2023 01:30:48 PM, Canjimavadivel, Kamallesh ) Provided cable from stock | Unit requested SpO2 trunk cable for X3
1065480 | 192473 | ACC | CMPLT | CMPLT | 2023-01-10 13:45 | 2023-01-10 13:47 | Incoming Inspection | (TC 01/10/2023 01:46:18 PM, Reid, Emmerson ) Assembled the stand, completed electrical safety, and asset tagged. Incoming documents attached to WO. Incoming complete.
1065481 | 192474 | ACC | CMPLT | CMPLT | 2023-01-10 13:51 | 2023-01-10 13:54 | Incoming Inspection | (TC 01/10/2023 01:51:32 PM, Reid, Emmerson ) Assembled the stand, completed electrical safety, and asset tagged. Incoming documents attached to WO. Incoming complete.
| 192475 | CM | CMPLT | CMPLT | 2023-01-10 14:33 | 2023-01-10 14:34 | Changing measurement mode to: Ask User | (TC 01/10/2023 02:33:07 PM, Baghirzade, Nara) Changing measurement mode to: Ask User Put MX800 on config mode, then in global settings the measurement mode was changed to Ask User. Configured monitors in Zone B and procedure room. All were successfully completed.
500868 | 192476 | CM | CMPLT | CMPLT | 2023-01-10 15:39 | 2023-01-10 15:39 | MSH 5TH FLOR ROOM 4 STUDY NEEDS TO BE PUSHED | (TC 01/10/2023 03:38:55 PM, Whittington, Scott) MSH 5TH FLOR ROOM 4 STUDY NEEDS TO BE PUSHED | MSH 5TH FLOR ROOM 4 STUDY NEEDS TO BE PUSHED
| 192477 | CM | CMPLT | CMPLT | 2023-01-11 07:53 | 2023-01-09 07:54 | Meeting with Katrina to test Hyland and Philips | (TC 01/09/2023 02:00:00 PM, Samanian, Kamran) performed tests with TEST server Hyland. Philips IBE needs adjustment
| 192478 | CM | CMPLT | CMPLT | 2023-01-11 07:56 | 2023-01-10 07:58 | Upgarded NICU monitor configuration issue- needed change | (TC 01/10/2023 07:55:52 AM, Samanian, Kamran) Verified the issue through both email and testing on the current monitors. MX800 config did NOT have the proper config for X2 if transfers from Resus to Room. adjusted proper configs and tested in the lab. decided to do manually configure all monitors in the NICU rooms. Zone A assigned and done.
| 192479 | CM | CMPLT | CMPLT | 2023-01-11 07:59 | 2023-01-10 07:59 | meeting for VN500-Servo n integration testing | (TC 01/10/2023 07:58:51 AM, Samanian, Kamran) done! decided RT and clinical analysts review the codes and parameters and then we can do some testing
| 192480 | CM | CMPLT | CMPLT | 2023-01-11 08:01 | 2023-01-13 11:16 | IBE ticket in Philips, Case# 0121010340 for ENDO in IBE change | (TC 01/10/2023 08:00:45 AM, Samanian, Kamran) Reviewed what is needed to be done ( currently ENDO is going out through port 9004. we need new ENDO going out from port 9005 ) . opened the case. not able to speak with John Reilly . replied to his email and waiting for his response/call. (TC 01/11/2023 07:42:53 AM, Samanian, Kamran) Spoken with John, he adjusted the filter so that ENDO unit beds are also going out of port 9005. Tested on ENDO_4 bed and worked fine. asked john to close the case (TC 01/13/2023 11:16:18 AM, Samanian, Kamran) Verified the port TEST 9002 and added ENDO to the test port. tested and passed
65988 | 192481 | PM | CMPLT | CMPLT | 2023-01-11 08:12 | 2023-01-13 12:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 01/11/2023 08:12:28 AM, Baghirzade, Nara) Originally device was sent down because of channel error 354.6770. Downloaded error log; according to past WOs, the pressure sensor was re-seated multiple times. Pressure sensor was replaced. Upon replacing the sensor, I found the rear case and left handle to be damaged. Currently no parts available in biomed, waiting for BD order to be sent. (TC 01/13/2023 12:21:28 PM, Baghirzade, Nara) Replaced the left handle and rear case. Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
64239 | 192482 | PM | CMPLT | CMPLT | 2023-01-11 08:55 | 2023-01-11 08:58 | PM-MAJOR - Annual PM For: PATIENT MONITORS, PORTABLE | (TC 01/10/2023 08:56:07 AM, Scero, Tomlyn) PM Complete. Flow and leak test passed. Calibration done.
64243 | 192483 | PM | CMPLT | CMPLT | 2023-01-11 09:04 | 2023-01-11 09:04 | PM-MAJOR - Annual PM For: PATIENT MONITORS, PORTABLE | (TC 01/10/2023 09:03:56 AM, Scero, Tomlyn) PM Complete. Flow and leak test passed. Calibration done.
64245 | 192484 | PM | CMPLT | CMPLT | 2023-01-11 09:05 | 2023-01-11 09:05 | PM-MAJOR - Annual PM For: PATIENT MONITORS, PORTABLE | (TC 01/10/2023 09:04:49 AM, Scero, Tomlyn) PM Complete. Flow and leak test passed. Calibration done.
1056571 | 192485 | CM | CMPLT | CMPLT | 2023-01-11 09:05 | 2023-01-11 09:06 | Cabinet alarming | (TC 01/11/2023 09:05:27 AM, Canjimavadivel, Kamallesh ) Replaced 9V battery in cabinet
64246 | 192486 | PM | CMPLT | CMPLT | 2023-01-11 09:07 | 2023-01-11 09:07 | PM-MAJOR - Annual PM For: PATIENT MONITORS, PORTABLE | (TC 01/10/2023 09:07:11 AM, Scero, Tomlyn) PM Complete. Flow and leak test passed. Calibration done.
1059943 | 192487 | PM | CMPLT | CMPLT | 2023-01-11 09:32 | 2023-01-11 09:33 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/09/2023 09:32:24 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed Calibrated scale
1065499 | 192488 | ACC | CMPLT | CMPLT | 2023-01-11 09:43 | 2023-01-11 09:48 | Incoming Inspection | (TC 01/11/2023 09:42:56 AM, Reid, Emmerson ) Assembled stand and unit, chanrged the batteries, performed self test and asset tagged. Incoming Complete.
1065500 | 192489 | ACC | CMPLT | CMPLT | 2023-01-11 10:01 | 2023-01-11 10:02 | Incoming Inspection | (TC 01/11/2023 10:02:01 AM, Reid, Emmerson ) Assembled stand and unit, chanrged the batteries, performed self test and asset tagged. Incoming Complete.
67393 | 192490 | PM | CMPLT | CMPLT | 2023-01-11 10:02 | 2023-01-11 10:02 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/10/2023 10:02:07 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
502048 | 192491 | CM | CMPLT | CMPLT | 2023-01-11 14:40 | 2023-01-27 15:43 | MSH SPP - IU22- User said no connection | (TC 01/11/2023 02:39:57 PM, Chau, Ken) Fixed walk jack Tested connection and no issue -if it doesnt connect, have to message Mirza from MSH IT to connect machines
1049903 | 192492 | CM | CMPLT | CMPLT | 2023-01-11 17:05 | 2023-01-12 15:11 | MSH R63 Revo Mobile - No battery | (TC 01/11/2023 05:04:37 PM, Chau, Ken) User reported battery only last a short time and turns off -replaced batteries
| 192493 | CM | CMPLT | CMPLT | 2023-01-12 07:56 | 2023-01-11 11:00 | Bi-Weekly Departmental Meetings: Fleet Management, P3A, ICAT & Compliance | (TC 01/11/2023 11:00:00 AM, Baghirzade, Nara) Bi-Weekly Departmental Meetings: Fleet Management, P3A, ICAT & Compliance meeting
| 192494 | CM | CMPLT | CMPLT | 2023-01-12 08:00 | 2023-01-10 15:00 | Biomed device testing touchbase NICU | (TC 01/10/2023 03:00:00 PM, Baghirzade, Nara) biomed device testing touchbase with Kamran and IT
65754 | 192495 | PM | CMPLT | CMPLT | 2023-01-12 08:48 | 2023-01-12 08:49 | PM CO2 Module | (TC 01/12/2023 08:48:22 AM, Scero, Tomlyn) Relocated device to CCU from ICU storage that they weren't using. PM complete. Calibration and leak tests ok.
1051408 | 192496 | CM | CMPLT | CMPLT | 2023-01-12 08:53 | 2023-01-12 08:58 | Making a screeching sound | (TC 01/11/2023 08:52:49 AM, Reid, Emmerson ) Was reported that the unit was making a screeching sound. When I was showed the unit it was working normally but the person said that earlier it was making the sound. There was damage to the case and slightly to the probe end. I used some adhesive to repair the case. I have not head the unit make any strange sounds throughout the repair with periodic testing. Will send back to the unit.
| 192497 | CM | CMPLT | CMPLT | 2023-01-12 10:36 | 2023-01-12 10:39 | Flow Meter repair for Optiflow - airvo | (TC 01/12/2023 10:37:43 AM, Scero, Tomlyn) Flowmeter repaired for RT OR. Needle valve was stuck shut. Samir helped. Returned to OR RT
| 192498 | PA | CMPLT | CMPLT | 2023-01-12 12:54 | 2023-01-19 08:58 | Side bezel for X3 | (TC 01/12/2023 01:00:59 PM, Canjimavadivel, Kamallesh ) Contacted vendor for quote. Placed order for side bezel | Biomed stock, to be charged back as used
32560 | 192499 | CM | CMPLT | CMPLT | 2023-01-12 18:44 | 2023-01-12 18:42 | System boot fault/no boot directory reported | System boot fault/no boot directory reported Service completed troubleshooting Service backed up system disk Service replaced system disk CT rebooted/BIOS tested CT phantom tested OK
65624 | 192500 | CM | CMPLT | CMPLT | 2023-01-13 07:49 | 2023-01-13 08:31 | MX800 monitor in room 1432 is losing connectivity intermittently to the PICIX central station | (TC 01/12/2023 02:30:00 PM, Baghirzade, Nara) MX800 monitor in room 1432 is losing connectivity intermittently to the PICIX central station. Actions taken: 1. Replaced the network cable. 2. Restarted the MX800 monitor. 3. Confirmed network settings for MX800. 4. Replaced the MX800 monitor with the same equipment label, revision, and configuration. 5. In PICIX configuration for SDU unit, unlocked the monitor in the bed assignment and locked it again to the same bed 1432. 6. Changed the equipment label in monitor and central station. 7. Rebooted the PICIX central. 8. Disabled and enabled port switch. 9. Checked the switch panel. 10. Found out the patient was still admitted on a MX800 monitor in PACU; discharged the patient from PACU monitor. 11. Discharged and re-admitted the patient. None of the above actions taken were successful. After we discharged and re-admitted the patient, Kamran and I noticed that the patient had two profiles on the Philips system. When the patient transferred from OR to PACU, the patient was admitted on Philips using their last name, first name, and all demographics fields were filled out. The staff at SDU were expecting the patient after PACU so they also admitted the patient on the philips system but only using their last and first name, creating a confusion for the philips system. We had to discharge the patient on the philips system and then admit her using only part of her last name and the MX800 was successfully connecting to the PICIX. When we tried to put her full last name back, the connectivity issue would occur. Briefed Amanda (SDU PCM) and the nurses that were in SDU at the time on transferring the patient on the philips system rather than admitting them all over again. Will touch base with Amanda again regarding this. (TC 01/12/2023 02:30:00 PM, Samanian, Kamran) MX800 monitor in room 1432 is losing connectivity intermittently to the PICIX central station. Actions taken: 1. Replaced the network cable. 2. Restarted the MX800 monitor. 3. Confirmed network settings for MX800. 4. Replaced the MX800 monitor with the same equipment label, revision, and configuration. 5. In PICIX configuration for SDU unit, unlocked the monitor in the bed assignment and locked it again to the same bed 1432. 6. Changed the equipment label in monitor and central station. 7. Rebooted the PICIX central. 8. Disabled and enabled port switch. 9. Checked the switch panel. 10. Found out the patient was still admitted on a MX800 monitor in PACU; discharged the patient from PACU monitor. 11. Discharged and re-admitted the patient. None of the above actions taken were successful. After we discharged and re-admitted the patient, Kamran and I noticed that the patient had two profiles on the Philips system. When the patient transferred from OR to PACU, the patient was admitted on Philips using their last name, first name, and all demographics fields were filled out. The staff at SDU were expecting the patient after PACU so they also admitted the patient on the philips system but only using their last and first name, creating a confusion for the philips system. We had to discharge the patient on the philips system and then admit her using only part of her last name and the MX800 was successfully connecting to the PICIX. When we tried to put her full last name back, the connectivity issue would occur.
1065517 | 192501 | CM | CMPLT | CMPLT | 2023-01-13 09:19 | 2023-01-13 11:00 | Incoming Inspection and Assembly | (TC 01/13/2023 10:59:36 AM, Ferreira, Qwynn) Assembled the scale Inputted into TMS Stickered Inspected
1065518 | 192502 | CM | CMPLT | CMPLT | 2023-01-13 09:26 | 2023-01-13 10:58 | Incoming Inspection and Assembly | (TC 01/13/2023 10:57:50 AM, Ferreira, Qwynn) Assembled the scale Inputted into TMS Stickered Inspected
67880 | 192503 | PM | CMPLT | CMPLT | 2023-01-13 09:30 | 2023-01-13 09:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/11/2023 09:31:50 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1065513 | 192504 | CM | CMPLT | CMPLT | 2023-01-13 10:49 | 2023-01-13 10:50 | PM | (TC 01/12/2023 10:49:03 AM, Reid, Emmerson ) PM completed by vendor. Added unit to TMS and asset tagged.
1065514 | 192505 | CM | CMPLT | CMPLT | 2023-01-13 10:51 | 2023-01-13 10:51 | PM | (TC 01/12/2023 10:51:12 AM, Reid, Emmerson ) PM completed by vendor. Added unit to TMS and asset tagged.
1047837 | 192506 | CM | CMPLT | CMPLT | 2023-01-13 11:07 | 2023-01-13 11:07 | alarming | tested. functional. returned back to RT.
| 192507 | CM | CMPLT | CMPLT | 2023-01-13 11:18 | 2023-01-13 11:07 | Meeting With Adeel One to One | (TC 01/13/2023 11:29:59 AM, Samanian, Kamran) done!
1046035 | 192508 | CM | CMPLT | CMPLT | 2023-01-13 11:23 | 2023-01-13 11:24 | MSH ER DRX EVO E5-02(ER) DETECTOR BATTERY REQUIRED, PROVIDED FROM PREVIOUSLY PURCHASED LOCAL STOCK ACQUIRED NOV 1 2022 | (TC 01/13/2023 11:23:29 AM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) DETECTOR BATTERY REQUIRED, PROVIDED FROM PREVIOUSLY PURCHASED LOCAL STOCK ACQUIRED NOV 1 2022 | MSH ER DRX EVO E5-02(ER) DETECTOR BATTERY REQUIRED, PROVIDED FROM PREVIOUSLY PURCHASED LOCAL STOCK ACQUIRED NOV 1 2022
| 192509 | CM | CMPLT | CMPLT | 2023-01-13 11:32 | 2023-01-12 11:34 | Start receiving email alerts from FP about Prirmary server due to performance issue , opened CASE# 0121031156 | (TC 01/12/2023 11:32:12 AM, Samanian, Kamran) Verified all functions in PICiX system and all good. checked the logs and noticed connection between Primary and IBE has issue ( connected-disconnected) - Opened the ticket to Philips and was advises FSE needs to come to the site to check. Quote was received to approve. It seemed issue was hospital network as it was resolved by itself at 10AM. no action needed
| 192510 | CM | CMPLT | CMPLT | 2023-01-13 12:39 | 2023-03-09 13:48 | X3 handles- RED for ER | (TC 01/13/2023 12:39:14 PM, Samanian, Kamran) quote attached S/C#1002857633 (TC 03/09/2023 01:42:58 PM, Samanian, Kamran) Received item, PO# 7202143952
| 192511 | CM | CMPLT | CMPLT | 2023-01-13 12:48 | 2023-02-22 07:51 | X3 handles- Purple for ICU | (TC 01/13/2023 12:49:02 PM, Samanian, Kamran) S/C#1002857636 (TC 01/13/2023 12:49:02 PM, Samanian, Kamran) S/C#1002857636 (TC 02/22/2023 07:51:14 AM, Samanian, Kamran) received and installed by Kam
1061217 | 192512 | CM | CMPLT | CMPLT | 2023-01-13 13:26 | 2023-01-13 13:28 | Monitors turned off | (TC 01/13/2023 01:25:54 PM, Reid, Emmerson ) It was reported that the monitors turned off and then on for two seconds and then back off whlile using the olympus scope. I went to check it out but when I got there it was working normally. I sstayed for 20 mins and it was consistently working. I wasn't able to determine exactly what happened but i suspect that the monitors did not actually turn off, rather the camera head light might have accidentaly been turned off making it seem like the monitors turned off. Making this as a note for the future just incase.
1060002 | 192513 | CM | CMPLT | CMPLT | 2023-01-13 14:25 | 2023-01-13 14:42 | V5 lead not working | (TC 01/13/2023 02:41:03 PM, Canjimavadivel, Kamallesh ) Confirmed issue with simulator. Replaced chest lead set, tested ok
1056567 | 192514 | CM | CMPLT | CMPLT | 2023-01-13 14:55 | 2023-01-13 14:56 | Battery replacement | (TC 01/13/2023 02:55:29 PM, Canjimavadivel, Kamallesh ) Replaced battery
1049903 | 192515 | CM | CMPLT | CMPLT | 2023-01-13 17:17 | 2023-01-13 17:17 | detector not communicating |  | Found the detector batteries were not charging. Changed detector contacts and detector charging contacts in the Revolution.
1046035 | 192516 | CM | CMPLT | CMPLT | 2023-01-13 17:20 | 2023-01-13 17:20 | Lead curtain adjustment knob broken. |  | Repair curtain adjustment knob.
| 192517 | CM | CMPLT | CMPLT | 2023-01-13 17:22 | 2023-01-13 17:22 | Site coverage.
| 192518 | PA | CMPLT | CMPLT | 2023-01-16 07:54 | 2023-12-29 10:08 | Sterile water for ED | (TC 01/16/2023 09:04:30 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 02/27/2023 07:59:31 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 04/25/2023 03:02:42 PM, Canjimavadivel, Kamallesh ) Placed order for water (TC 06/05/2023 08:52:06 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 06/23/2023 08:27:48 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 07/31/2023 10:57:09 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 09/28/2023 08:20:20 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 10/25/2023 09:13:44 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 11/01/2023 09:08:48 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 12/18/2023 09:04:13 AM, Canjimavadivel, Kamallesh ) Placed order for sterile water
1065521 | 192519 | ACC | CMPLT | CMPLT | 2023-01-16 08:38 | 2023-01-16 08:40 | Incoming inspection | (TC 01/13/2023 08:38:30 AM, Rendon Lopez , Elizabeth) Incoming inspection.
65909 | 192520 | PM | CMPLT | CMPLT | 2023-01-16 08:40 | 2023-01-16 08:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 01/16/2023 08:46:03 AM, Baghirzade, Nara) Originally device was sent down due to channel error 354.6770; checking previous WOs, the pressure sensor was previously re-seated mutliple times. Pressure sensor disc was replaced. Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65909 | 192521 | PM | CANCL | DUP | 2023-01-16 08:40 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | WO duplicated by mistake.
1065525 | 192522 | ACC | CMPLT | CMPLT | 2023-01-16 08:40 | 2023-01-16 08:42 | Incoming inspection | (TC 01/13/2023 08:41:27 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065526 | 192523 | ACC | CMPLT | CMPLT | 2023-01-16 08:42 | 2023-01-16 08:43 | Incoming inspection | (TC 01/13/2023 08:42:58 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065527 | 192524 | ACC | CMPLT | CMPLT | 2023-01-16 08:44 | 2023-01-16 08:45 | Incoming inspection | (TC 01/13/2023 08:44:49 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065528 | 192525 | ACC | CMPLT | CMPLT | 2023-01-16 08:45 | 2023-01-16 08:47 | Incoming inspection | (TC 01/13/2023 08:46:16 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065522 | 192526 | ACC | CMPLT | CMPLT | 2023-01-16 08:48 | 2023-01-16 08:49 | Incoming inspection | (TC 01/13/2023 08:47:57 AM, Rendon Lopez , Elizabeth) Assembled, entered in TMS, charging batteries. Incoming inspection, OK
1054770 | 192527 | CM | CMPLT | CMPLT | 2023-01-16 09:19 | 2023-06-08 08:34 | Not turning on | (TC 02/08/2023 09:48:14 AM, Canjimavadivel, Kamallesh ) Power button defective. Called vendor for quote and placed order for power button assembly (TC 02/09/2023 02:49:12 PM, Canjimavadivel, Kamallesh ) Installed battery (TC 05/25/2023 02:46:33 PM, Canjimavadivel, Kamallesh ) Found power button from old unit retired for parts and replaced. Discovered touch screen not working. (TC 05/26/2023 01:22:03 PM, Canjimavadivel, Kamallesh ) Swapped screen from old retired unit kept for parts. Touchscreen working now.
69066 | 192528 | FC | CMPLT | CMPLT | 2023-01-16 10:57 | 2023-01-16 10:59 | No yellow tag attached. Functional check only | (TC 01/16/2023 10:58:15 AM, Baghirzade, Nara) Pump was sent without a yellow tag. Good physical condition. Downloaded error logs and did not see any recent errors. Check-in tests passed and completed.
| 192529 | ACC | CMPLT | CMPLT | 2023-01-16 11:33 | 2023-02-06 14:17 | MSH OBSTETRICAL CONFERENCE CANON i700 PRISM ACCEPTANCE | (TC 01/18/2023 04:17:04 PM, Dolinar, John) MSH OBSTETRICAL CONFERENCE CANON i700 PRISM ACCEPTANCE MOVE SYSTEM AND PROBE BOX TO MSH FROM TGH (TC 01/18/2023 01:50:39 PM, Whittington, Scott) MSH OBSTETRICAL CONFERENCE CANON i700 PRISM ACCEPTANCE MSH OBSTETRICAL CONFERENCE CANON i700 PRISM ACCEPTANCE MOVE SYSTEM AND PROBE BOX TO MSH FROM TGH (TC 01/23/2023 03:14:14 PM, Whittington, Scott) MSH OBSTETRICAL CONFERENCE CANON i700 PRISM ACCEPTANCE | MSH OBSTETRICAL CONFERENCE CANON i700 PRISM ACCEPTANCE 125 mOHM 25uA
| 192530 | CM | CMPLT | CMPLT | 2023-01-16 11:49 | 2023-01-17 08:38 | Network cables for ENDO unit for new data drops in ENDO 2 and 3 | (TC 01/16/2023 11:49:04 AM, Samanian, Kamran) ENDO 2, data drop # 6SE-253-D connected to port#12 ENDO 3, data drop # 6SE-265-D connected to port#13 ENDO 4, data drop # 6SE-07-B connected to port#14 All labeled. Currently there are 2 ports assigned for the bed labeles ENDO_2, _3, _4 for switch port mapping. Needs to be tested (TC 01/17/2023 08:38:20 AM, Samanian, Kamran) Tested and passed
1047858 | 192531 | CM | CMPLT | CMPLT | 2023-01-16 12:19 | 2023-01-17 08:15 | Broken pole clamp | (TC 01/17/2023 08:14:11 AM, Baghirzade, Nara) Pole clamp was replaced. Ran functional tests and occlusion tests; all passed.
501173 | 192532 | CM | CMPLT | CMPLT | 2023-01-16 12:56 | 2023-01-16 12:57 | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 | (TC 01/16/2023 12:56:35 PM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 SYSTEM SLOW CLEARED ALL PT EXAMS FROM BROWSER REMAKE DB RUN DFGR 99% HDD SPACE | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173
| 192533 | CM | CMPLT | CMPLT | 2023-01-17 07:22 | 2023-01-17 07:23 | Biomed Huddle | (TC 01/16/2023 02:00:00 PM, Baghirzade, Nara) Biomed Huddle
| 192534 | CM | CMPLT | CMPLT | 2023-01-17 07:23 | 2023-01-17 07:24 | VSM Testing | (TC 01/16/2023 01:00:00 PM, Baghirzade, Nara) VSM testing with Nima and Naz
65909 | 192535 | CM | CMPLT | CMPLT | 2023-01-17 09:42 | 2023-01-17 10:20 | Channel error | (TC 01/17/2023 10:19:42 AM, Baghirzade, Nara) Downloaded error logs and noted that 351.6760 error indicates that the calibration flag was not set, although the unit was calibrated just a few days ago. After speaking to Elizabeth (biomed), it turns out that the software glitches out after calibration if a PM is selected, therefore the way to save the calibration is to perform a check-in after calibrating the pressure sensor disc. Calibrated the pressure sensor disc. Check-in tests passed and completed. Ran a basic infusion; no errors or issues occurred during infusion.
| 192536 | CM | CMPLT | CMPLT | 2023-01-17 09:56 | 2023-01-17 09:58 | Missing adapter for X-series | (TC 01/17/2023 09:56:31 AM, Canjimavadivel, Kamallesh ) Found adapter missing, replaced from stock. Notified PCM
1054466 | 192537 | CM | CMPLT | CMPLT | 2023-01-17 10:02 | 2023-01-18 14:16 | Screen on rail not working | (TC 01/17/2023 10:01:59 AM, Canjimavadivel, Kamallesh ) Created a case with Hillrom to have technician repair bed on Dec 14, 2022. Followed up on Dec 28, 2022. Found repair completed note with date January 13, 2023. Followed up to get service report. (TC 01/18/2023 02:15:04 PM, Canjimavadivel, Kamallesh ) Attached service report
65988 | 192538 | CM | CMPLT | CMPLT | 2023-01-17 10:39 | 2023-01-17 10:40 | Channel error | (TC 01/17/2023 10:38:57 AM, Baghirzade, Nara) Downloaded error logs and noted that 351.6760 error indicates that the calibration flag was not set, although the unit was calibrated just a few days ago. After speaking to Elizabeth (biomed), it turns out that the software glitches out after calibration if a PM is selected, therefore the way to save the calibration is to perform a check-in after calibrating the pressure sensor disc. Calibrated the pressure sensor disc. Check-in tests passed and completed. Ran a basic infusion; no errors or issues occurred during infusion.
1065478 | 192539 | CM | CMPLT | CMPLT | 2023-01-17 10:42 | 2023-01-18 13:56 | Returning loaner unit to 3M | (TC 01/17/2023 10:43:19 AM, Baghirzade, Nara) Called 3M and requested a pickup for the loaner unit. Confirmation #: 12452430 (TC 01/18/2023 01:54:24 PM, Baghirzade, Nara) Removed loaner asset tag. 3M picked up loaner. Attached pick-up receipt.
65901 | 192540 | CM | CMPLT | CMPLT | 2023-01-17 11:05 | 2023-01-17 11:05 | Channel error | (TC 01/17/2023 11:04:48 AM, Baghirzade, Nara) Downloaded error logs and noted that 351.6760 error indicates that the calibration flag was not set, although the unit was calibrated just a few days ago. After speaking to Elizabeth (biomed), it turns out that the software glitches out after calibration if a PM is selected, therefore the way to save the calibration is to perform a check-in after calibrating the pressure sensor disc. Calibrated the pressure sensor disc. Check-in tests passed and completed. Ran a basic infusion; no errors or issues occurred during infusion.
65264 | 192541 | CM | CMPLT | CMPLT | 2023-01-17 11:56 | 2023-01-17 11:56 | Alarming | Presuure sensor calibration valva failed exp vave failed safety valve failled Rex replaced exp valve . tested. Functional. passed pre-use check.
1061416 | 192542 | CM | CMPLT | CMPLT | 2023-01-17 11:58 | 2023-01-17 11:58 | not working | reset the unit. tested. functional. informed user
| 192543 | CM | CMPLT | CMPLT | 2023-01-17 12:25 | 2023-01-17 12:34 | NICU Testing | (TC 01/17/2023 12:33:03 PM, Baghirzade, Nara) Philips system testing with Kamran in NICU. Checked that data is going through to Cerner from centrals in Zone A, B, and C after the architectural change has been applied to the Philips systems. Checked that alarms are going through to the Ascom phones from Zone A, B, and C. Testing was completed and successful.
| 192544 | PA | CMPLT | CMPLT | 2023-01-17 12:57 | 2023-01-20 07:54 | Ordering thermometers for Rheumatology | (TC 01/17/2023 12:59:36 PM, Baghirzade, Nara) SC 1002860133 (TC 01/20/2023 07:53:03 AM, Baghirzade, Nara) PO 7202145687
1056536 | 192545 | CM | CMPLT | CMPLT | 2023-01-17 13:05 | 2023-01-17 13:05 | Battery replacement | (TC 01/17/2023 01:05:09 PM, Canjimavadivel, Kamallesh ) Replaced battery
68028 | 192546 | CM | CMPLT | CMPLT | 2023-01-17 13:16 | 2023-01-17 13:16 | Calibration required | (TC 01/17/2023 01:15:45 PM, Canjimavadivel, Kamallesh ) Performed calibration, passed
1059823 | 192547 | CM | CMPLT | CMPLT | 2023-01-17 13:20 | 2023-03-06 10:29 | Battery/Charger Issue | (TC 01/17/2023 01:20:32 PM, Reid, Emmerson ) Went over to check the battery charger. One of the batteries was showing an amber coloured light in both charging slots. The unit was saying that the battery had been doing this, however I said for them to let it charge overnight and let me know if it is still showing this tomorrow.
69544 | 192548 | PM | CMPLT | CMPLT | 2023-01-17 13:42 | 2023-01-17 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 01/17/2023 01:42:24 PM, Baghirzade, Nara) Pump was sent to biomed because the door was broken, and would not stay closed. Front case and door cover broken; both were replaced. PM passed and completed. IUI connectors and device cleaned.
501105 | 192549 | CM | CMPLT | CMPLT | 2023-01-17 14:02 | 2023-01-18 13:47 | MSH 5 US IU22 Error 609, system cannot start up | (TC 01/17/2023 02:01:41 PM, Fan, Tak) Could not start up system. Placed call to Philips. Cody will do SW reload tomorrow. (TC 01/18/2023 01:45:35 PM, Whittington, Scott) MSH 5 US IU22 Error 609, system cannot start up JILLIAN RELOADED SW, REPORTED OK | MSH 5 US IU22 Error 609, system cannot start up
1064771 | 192550 | CM | CMPLT | CMPLT | 2023-01-17 14:09 | 2023-01-17 14:12 | Moved to OR 8 Boom | (TC 01/17/2023 02:10:34 PM, Reid, Emmerson ) Moved from a mobile cart to the Boom in OR 8. Tested with connected gas, and was working.
1049902 | 192551 | CM | CMPLT | CMPLT | 2023-01-17 16:25 | 2023-01-17 16:26 | MSH R62 Mobile - Detector communication issue | (TC 01/17/2023 04:25:09 PM, Chau, Ken) Detector wouldnt communicate with system -looked into logs and no errors -low drop on detector -removed battery and let detector reset, tested okay after
1056545 | 192552 | CM | CMPLT | CMPLT | 2023-01-18 07:35 | 2023-01-18 07:36 | Replaced battery | (TC 01/18/2023 07:34:45 AM, Baghirzade, Nara) A call came in during my on-call about this AED; after they used it, it would report unit failed because the battery was very depleted. Replaced battery. Battery testing passed. Checked the battery capacity and it was at 95%. Checked device network and all good.
| 192553 | RDS | CMPLT | CMPLT | 2023-01-18 07:37 | 2023-01-18 07:37 | NICU Huddle | (TC 01/17/2023 03:40:00 PM, Baghirzade, Nara) NICU huddle
1059778 | 192554 | FC | CMPLT | CMPLT | 2023-01-18 07:59 | 2023-01-18 08:00 | QC Checklist | (TC 01/18/2023 07:59:06 AM, Baghirzade, Nara) QC Checklist completed. Returned to 11 South.
| 192555 | PA | CMPLT | CMPLT | 2023-01-18 07:59 | 2023-01-25 12:04 | sterile water for altrix
| 192556 | PA | CMPLT | CMPLT | 2023-01-18 08:05 | 2023-02-22 09:12 | HotLine Parts Order - To Charge back as used
69394 | 192557 | PM | CMPLT | CMPLT | 2023-01-18 08:24 | 2023-01-18 08:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 01/18/2023 08:48:21 AM, Baghirzade, Nara) Pump was originally sent down because of channel error. Downloaded error log and noted a repeated error 211.2000. In the service manual, it refers to an unspecified communication error and to replace the logic board. Logic board was replaced. Flashed the pump and updated the serial number for the board. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. After PM, ran a basic infusion at different rates and no errors occurred during infusion. Pump is fully functional.
67943 | 192558 | CM | CMPLT | CMPLT | 2023-01-18 08:28 | 2023-01-18 08:32 | Scale does not work - icon disappeared | (TC 01/18/2023 08:28:52 AM, Ferreira, Qwynn) Unit came down with note sayingt he scale icon does not appear/scale does not work. Tested the unit - no scale icon appeared Noticed when playing with the scale cable that there was a flicker in the display screen and a relay switching. Tested with a new scale - similar issue was present Replaced the scale cable running to the control board. Tested with old scale - no realy noises but still no scale icon Tested with new scale - no flickering, scale icon is present, working. Functional Test - OK
| 192559 | CM | CMPLT | CMPLT | 2023-01-18 08:30 | 2023-01-17 08:31 | Distributed mode - PICiX-IBE connection | (TC 01/17/2023 08:30:14 AM, Samanian, Kamran) Converted centrlzied mode to distributed mode in PICiX @11:17AM.. tested two rooms in NICU for vital data charting and ASCOM and all good. Checked ENDO and PACU and OR and data charting is good and passed. Followed up with NICU in the afternoon and all good.
66680 | 192560 | CM | CMPLT | CMPLT | 2023-01-18 08:32 | 2023-01-18 08:35 | Broken, not working correctly | (TC 01/18/2023 08:33:26 AM, Ferreira, Qwynn) three issues were present 1: Hood was not seated flush to the chassis which prohibited the use of the latches so doors would not lock. 2: Two port door latches were cracked/chipped - Replaced the latches 3: the main switch at the back was turned off. Functional Test - OK
| 192561 | CM | CMPLT | CMPLT | 2023-01-18 08:33 | 2023-01-17 08:37 | ENDO Philips bed label testing with Katrina | (TC 01/17/2023 08:33:16 AM, Samanian, Kamran) Added bed labels ENDO_1 to ENDO_4 to the PICiX for ENDO. added equipment label such as ENDOTEST for testing IBE. (TC 01/17/2023 08:34:06 AM, Samanian, Kamran) met with Katrina on the floor and tested all bed labels by moving the monitor to each room. System is designed based on switch port mapping.
| 192562 | CM | CMPLT | CMPLT | 2023-01-18 08:35 | 2023-01-16 08:35 | Huddle meeting | (TC 01/16/2023 08:35:34 AM, Samanian, Kamran) done!
| 192563 | CM | CMPLT | CMPLT | 2023-01-18 08:36 | 2023-01-16 08:37 | Philips ICU configuration meeting | (TC 01/16/2023 08:36:34 AM, Samanian, Kamran) Done!
67919 | 192564 | CM | CMPLT | CMPLT | 2023-01-18 08:37 | 2023-01-18 08:42 | Broken Door and Clip | (TC 01/18/2023 08:41:09 AM, Ferreira, Qwynn) Replaced the broken hinge clip Replaced the broken door Functional Check - OK
| 192565 | CM | CMPLT | CMPLT | 2023-01-18 10:09 | 2023-01-18 10:11 | Ophthalmoscope functional check and lamp replacement | (TC 01/18/2023 10:09:29 AM, Baghirzade, Nara) Checked battery on the handle; battery must be replaced in July 2023. Macroview ophthalmoscope lamp was replaced. Checked another ophthalmscope with the handle and it was also functional. Touched base with Mara re: battery replacement in July.
1059942 | 192566 | PM | CMPLT | CMPLT | 2023-01-18 10:29 | 2023-01-18 10:29 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/17/2023 10:29:07 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed Calibrated scale
| 192567 | PA | CMPLT | CMPLT | 2023-01-18 10:41 | 2023-01-18 10:42 | Bunnell HFV contract charge back | Bunnell contract 2023-2024
1053391 | 192568 | CM | CMPLT | CMPLT | 2023-01-18 10:58 | 2023-01-18 10:59 | Functional check after patient use | (TC 01/18/2023 10:58:25 AM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 51 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
| 192569 | PA | CMPLT | CMPLT | 2023-01-18 11:56 | 2023-02-14 13:00 | Ordering sheath assembly for PFT systems | (TC 01/18/2023 11:59:51 AM, Baghirzade, Nara) SC 1002861218 (TC 02/14/2023 12:56:18 PM, Baghirzade, Nara) PO 7202146721
1056227 | 192570 | CM | CMPLT | CMPLT | 2023-01-18 12:58 | 2023-01-17 08:05 | Distributed mode - PICiX-IBE connection | (TC 01/17/2023 08:04:19 AM, Samanian, Kamran) Converted Centralzied mode to Distributed mode. Tested ENDO,PACU, NICU, OR and data integration is fine. Also tested ASCOM phones in NICU and all good.
| 192571 | CM | CMPLT | CMPLT | 2023-01-18 13:26 | 2023-02-03 15:46 | Ordering spare fuses for general biomed use | (TC 01/18/2023 02:10:19 PM, Cortez, John) Ordering fuse kit (fast/slow), 240 pcs for biomed stock. Also ordering 5 15A fuses
1065550 | 192572 | ACC | CMPLT | CMPLT | 2023-01-18 13:28 | 2023-01-18 14:12 | Incoming inspection | (TC 01/18/2023 02:11:31 PM, Canjimavadivel, Kamallesh ) Added to TMS and asset tagged stretcher
1065549 | 192573 | ACC | CMPLT | CMPLT | 2023-01-18 13:49 | 2023-01-18 13:51 | Incoming inspection and assembly | (TC 01/18/2023 01:49:30 PM, Baghirzade, Nara) Assembled new bladder scanner on rolling styand for 60M Urology. Adjusted the time on the bladder scanner. Attached instructions on a ring. Charged both batteries to 100%. Ran a self-test; all passed. Dropped off the bladder scanner with Leia (CCM) at 60M Urology.
1054046 | 192574 | PM | CMPLT | CMPLT | 2023-01-18 13:55 | 2023-01-18 13:56 | PM-MAJOR-SENTEC DMS | (TC 01/18/2023 01:55:16 PM, Ferreira, Qwynn) PM Passed Self Test- OK Electrical Safety - 64.6uA, 130mOhms
1057547 | 192575 | CM | CMPLT | CMPLT | 2023-01-18 14:39 | 2023-01-23 11:17 | Battery malfunction | (TC 01/18/2023 02:40:17 PM, Scero, Tomlyn) Replaced old (2020) battery with a new one. Left to charge and X3 was working fine. (TC 01/18/2023 02:40:52 PM, Scero, Tomlyn) Used X3 to test SPO2 sensors from OR (RT) that they were having issues with.
1059828 | 192576 | CM | CMPLT | CMPLT | 2023-01-18 14:46 | 2023-01-19 07:46 | Cleaning | (TC 01/18/2023 02:46:58 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
67240 | 192577 | CM | CMPLT | NFF | 2023-01-18 15:29 | 2023-01-18 15:32 | error message on screen | (TC 01/18/2023 03:29:13 PM, Cortez, John) Ran pump for 30 mins, no issues found Volume accuracy within error range 39.7 g volume after 30 mins at 75ml/hr
1047847 | 192578 | CM | CMPLT | CMPLT | 2023-01-18 15:45 | 2023-01-18 15:50 | battery needs to be replaced | (TC 01/12/2023 03:45:38 PM, Cortez, John) Replaced battery
1057015 | 192579 | CM | CMPLT | CMPLT | 2023-01-19 07:57 | 2023-01-19 13:37 | Broken side bezel | (TC 01/19/2023 01:35:24 PM, Canjimavadivel, Kamallesh ) Replaced side bezel
| 192580 | CM | CMPLT | CMPLT | 2023-01-19 08:07 | 2023-01-18 08:09 | Hyland meeting- | (TC 01/18/2023 08:06:41 AM, Samanian, Kamran) discussed about current status of existing MDR, SW Rev of PACSgear Core and MDR. upgrade plan on both Sw or HW
| 192581 | CM | CMPLT | CMPLT | 2023-01-19 08:42 | 2023-01-19 08:43 | Novus meeting | (TC 01/18/2023 11:40:00 AM, Baghirzade, Nara) Novus meeting re: PFT and CPET parts, support, and training.
1055920 | 192582 | CM | CMPLT | CMPLT | 2023-01-19 08:48 | 2024-03-15 08:17 | Ordering root handle | (TC 01/19/2023 08:48:43 AM, Baghirzade, Nara) Received the basket from Masimo today; assembled it on to the rolling stand. Waiting for the handle. Request part number and price from Masimo. (TC 05/10/2023 08:03:29 AM, Baghirzade, Nara) Have not received a response from Masimo regarding part number and price for a handle. Re-sent the request to their customer service email. (TC 01/10/2024 02:59:00 PM, Baghirzade, Nara) SC 1003144252 (TC 03/15/2024 08:16:43 AM, Baghirzade, Nara) PO 7202346440
67816 | 192583 | PM | CMPLT | CMPLT | 2023-01-19 09:35 | 2023-01-19 13:35 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 01/19/2023 01:32:45 PM, Ferreira, Qwynn) PM Passed Replaced Battery - Yes ECG @100 = 100 spO2 @99 = 99 Resp Rate @40 = 40 Temp @37 = 37 Static @200 = 199 Dynamic @120/80 = 117/80 Accuracy @280 =279 Leakage 4.7 in 60 seconds Linearity @150 = 150 Valve 5.2 in 5 seconds
| 192584 | PA | CMPLT | CMPLT | 2023-01-19 09:46 | 2023-02-16 14:45 | PM kit for servo ventilators | (TC 01/19/2023 09:51:08 AM, Canjimavadivel, Kamallesh ) Placed order for PM kit (TC 02/16/2023 02:44:38 PM, Canjimavadivel, Kamallesh ) Received PM kits
1063746 | 192585 | CM | CMPLT | CMPLT | 2023-01-19 10:23 | 2023-01-19 10:24 | Cleaning | (TC 01/19/2023 10:23:32 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
| 192586 | CON | CMPLT | CMPLT | 2023-01-19 10:51 | 2023-01-19 10:52 | MSH IVF APLIO 300 SERVICE REPORTS, COMPILE AND SEND TO WENDY WANG | (TC 01/19/2023 10:52:07 AM, Whittington, Scott) MSH IVF APLIO 300 SERVICE REPORTS, COMPILE AND SEND TO WENDY WANG | MSH IVF APLIO 300 SERVICE REPORTS, COMPILE AND SEND TO WENDY WANG
68229 | 192587 | PM | CMPLT | CMPLT | 2023-01-19 13:26 | 2023-01-19 13:30 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 01/19/2023 01:26:54 PM, Ferreira, Qwynn) PM Passed Replaced Battery - Yes ECG @100BPM = 100bpm spO2 @99 = 99 Resp Rate @40 = 40 Temp @37 = 37 Static @ 200 = 201 Dynamic @ 120/80 = 118/79 Accuracy @280 = 281.1 Leakage 4.9 in 60 seconds Linearity @150 = 150 Valve 5.1 in 5 seconds Labeled for the NICU and added to the NICU pool
1056125 | 192588 | PM | CMPLT | CMPLT | 2023-01-19 14:03 | 2023-01-19 15:14 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/19/2023 03:13:07 PM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
502044 | 192589 | CM | CMPLT | CMPLT | 2023-01-19 14:14 | 2023-01-31 17:59 | Acquisition aborting due to errors | (TC 01/19/2023 02:14:21 PM, Paz, Ramon (Tino)) Errors while acquiring. Head 0 has 7 missing modules. USB device 85 has failed to reset. Another error about aquisition software must be closed. Adjusted 5 volts to PSB to 5.2 volts. was at 5.00 v. Customer still had problems. Moved USB cable to head to another port. (TC 01/20/2023 11:56:31 AM, Paz, Ramon (Tino)) Replacement PSB board came in but DOA. USB device could not be found message. Replacement part did not work. Ordered a new one but there was a delay in arrival. Scheduled delivery was Monday, February 23. (TC 01/25/2023 12:03:19 PM, Paz, Ramon (Tino)) Part has not arrived yet since package was held in customs due to missing broker information. In the meantime, we reseated module 85 and module 5 as well. Tested with a QC. (TC 01/31/2023 05:51:05 PM, Paz, Ramon (Tino)) Replaced PSB (Passport System Board). One replacement did not work. It came up with DETECTOR ERRORS AND MODULE ERRORS. Swapped with the second board that Digirad sent and that seem to work. Also installed PEM standoffs om module 85 and 67. Had issues with Module 84 (corner) Ran a test flood, Passed QC.
66847 | 192590 | CM | CMPLT | CMPLT | 2023-01-19 14:26 | 2023-01-19 14:27 | Upgraded Version | (TC 01/19/2023 02:26:18 PM, Ferreira, Qwynn) Upgraded to Rev M
66817 | 192591 | CM | CMPLT | CMPLT | 2023-01-19 14:27 | 2023-01-19 14:28 | upgraded version | (TC 01/19/2023 02:27:26 PM, Ferreira, Qwynn) Upgraded to Rev M
531060 | 192592 | CM | CMPLT | CMPLT | 2023-01-19 14:41 | 2023-01-19 14:44 | Does not turn on | (TC 01/19/2023 02:41:33 PM, Baghirzade, Nara) Replaced the power supply. Put all the connections back. Replaced one of the zip-ties at the back of the power supply case. Pump in good physical condition. Functional checks passed. @1 % = 28 mmHg @32% = 100 mmHg @57% = 153 mmHg @77% = 195 mmHg @100% = 248 mmHg Ran pump for a little while and no errors or malfunctions occured. | Pump does not turn on; contacted mother's choice to order a power supply replacement
1064502 | 192593 | CM | CMPLT | CMPLT | 2023-01-19 14:47 | 2023-01-19 14:47 | HARP Documentation | (TC 01/19/2023 02:47:22 PM, Paz, Ramon (Tino)) Filled out UHN HARP/REDACT documentation for the GE iDXA unit at MSH.
1062714 | 192594 | CM | CMPLT | CMPLT | 2023-01-19 14:58 | 2023-01-19 14:57 | Could not calibrate the detector. | Replaced the detector, calibrated the system, tested O.K.
1065556 | 192595 | CM | CMPLT | CMPLT | 2023-01-20 08:09 | 2023-01-20 08:11 | Incoming and asset tagging | (TC 01/20/2023 08:09:18 AM, Baghirzade, Nara) Incoming inspection and functional test passed. Asset tagged the thermometer and attached calibration certificate to this WO. Dropped it off with Dieter at 60M, 2nd fl.
1065557 | 192596 | CM | CMPLT | CMPLT | 2023-01-20 08:15 | 2023-01-20 08:15 | Incoming inspection and asset tagging | (TC 01/20/2023 08:14:49 AM, Baghirzade, Nara) Incoming inspection and functional test passed. Asset tagged the thermometer and attached calibration certificate to this WO. Dropped it off with Dieter at 60M, 2nd fl.
501067 | 192597 | CM | CMPLT | CMPLT | 2023-01-20 09:03 | 2023-08-24 08:54 | MSH CEOU RM#1 - Control arm dropping | (TC 01/20/2023 09:03:26 AM, Chau, Ken) Control arm has been ordered -ETA 2 weeks (TC 05/08/2024 02:59:35 PM, Fan, Tak) Replaced operator panel gas spring assembly. | MSH CEOU RM#1 - Control arm dropping
501107 | 192598 | CM | CMPLT | CMPLT | 2023-01-20 09:04 | 2023-06-06 16:15 | MSH CEOU RM#3 - Front casters broken, requires replacement | (TC 01/20/2023 09:04:19 AM, Chau, Ken) Front casters ordered (TC 02/08/2023 10:42:39 AM, Chau, Ken) Registry deleted by Philips -tested network and working properly | MSH CEOU RM#3 - Front casters broken, requires replacement
| 192599 | CM | CMPLT | CMPLT | 2023-01-20 11:07 | 2023-01-24 12:28 | Alaris PCU LCD screen | (TC 01/20/2023 11:21:36 AM, Rendon Lopez , Elizabeth) BD wrong part delivered, new LCD assembly ordered and returned wrong part with waived restocking fee
| 192600 | CM | CMPLT | CMPLT | 2023-01-20 12:29 | 2023-01-31 07:50 | Order three Suction unit | (TC 02/01/2023 07:50:18 AM, Chukri, Samir) Received three suction regulator delivered to Naz to R.T
| 192601 | PA | CMPLT | CMPLT | 2023-01-20 14:01 | 2023-01-20 14:13 | NICU Neonatal BP Cuffs Order - 4 Sizes | (TC 01/25/2023 09:56:27 AM, Chukri, Samir) Received NIBP cuff from Spectrom and Masimo and deliverd to NICU
501584 | 192602 | CM | CMPLT | CMPLT | 2023-01-20 16:03 | 2023-01-23 16:10 | MSH OPG SPP IU22 Space bar on qwerty keyboard not working. |  | Space bar on qwerty keyboard not working. Called Philips. Ref# 0121076935 Replaced keyboard
501618 | 192603 | CM | CMPLT | CMPLT | 2023-01-20 16:05 | 2023-01-23 16:09 | MSH OPG 3-521 E1 BPP IU22 qwerty keyboard unresponsive intermittently. |  | qwerty keyboard unresponsive intermittently. Called Philips. Ref# 0121076948 Replaced qwerty cable Replace keyboard
32510 | 192604 | CM | CMPLT | CMPLT | 2023-01-20 16:09 | 2023-01-20 16:10 | Tube head handle loose |  | Tube head handle loose. Tightened tube head handle.
502044 | 192605 | CM | CMPLT | CMPLT | 2023-01-20 16:11 | 2023-01-20 16:11 | Interface board need replacing. |  | Helped Barry and Tino replace interface board.
1046196 | 192606 | CM | CMPLT | CMPLT | 2023-01-22 16:03 | 2023-01-22 16:01 | xray tube error/tube arc reported | xray tube error/tube arc reported Service completed troubleshooting Service stored tube onsite when delivered Service/SIEMENS replaced xray tube/calibrated/backup CT phantom tested OK
| 192607 | PA | CMPLT | CMPLT | 2023-01-23 07:53 | 2023-02-09 14:29 | Hood Order
1065121 | 192608 | CM | CMPLT | CMPLT | 2023-01-23 08:47 | 2023-01-23 08:56 | P01/P03/P05 Error message | (TC 01/23/2023 08:55:01 AM, Ferreira, Qwynn) Tested the device but was not turning on Replaced batteries Only showed error p05 Replaced the temp prob - read correctly and no errors
1059150 | 192609 | CM | CMPLT | NFF | 2023-01-23 09:04 | 2023-01-23 09:06 | "spo2 sensor not working" | (TC 01/20/2023 09:03:56 AM, Cortez, John) Unit only provided X3 module, SPO2 tested ok with simulator+spare sensor
1065560 | 192610 | CM | CMPLT | CMPLT | 2023-01-23 09:13 | 2023-01-23 09:14 | Inspection | (TC 01/23/2023 09:13:37 AM, Ferreira, Qwynn) Installed batteries added to tms Temp tested @37 = 37
1065561 | 192611 | CM | CMPLT | CMPLT | 2023-01-23 09:18 | 2023-01-23 09:19 | Inspection | (TC 01/23/2023 09:18:07 AM, Ferreira, Qwynn) added batteries added to tms tested temperature @37 = 37
1060001 | 192612 | CM | CMPLT | CMPLT | 2023-01-23 09:34 | 2023-01-23 11:26 | Leads V4, V6 not working | (TC 01/23/2023 11:04:49 AM, Canjimavadivel, Kamallesh ) Confirmed issue with simulator. Replaced chest lead set, tested ok
1065563 | 192613 | ACC | CMPLT | CMPLT | 2023-01-23 10:53 | 2023-01-23 11:06 | Incoming inspection | (TC 01/23/2023 11:05:08 AM, Baghirzade, Nara) Incoming inspection completed. Asset tagged the device and created asset on TMS for inventory purposes only. Electrical safety results: power cord resistance = 11 mohms leakage current = 9.1 uA Returned device to Dentistry clinic front desk.
1056386 | 192614 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 10:28 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:27:58 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056535 | 192615 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 09:53 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:52:28 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056536 | 192616 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-15 08:27 | Software upgrade to version 06.03.006.7404 | (TC 01/17/2023 08:26:11 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056537 | 192617 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-15 08:28 | Software upgrade to version 06.03.006.7404 | (TC 01/17/2023 08:28:04 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056538 | 192618 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 10:42 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:41:35 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056539 | 192619 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 10:57 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:57:12 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056540 | 192620 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 11:08 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:07:39 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056541 | 192621 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 11:20 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:19:50 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056542 | 192622 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 09:03 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:02:07 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056543 | 192623 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 09:21 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:20:25 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056544 | 192624 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 09:24 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:23:58 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056545 | 192625 | CON | CMPLT | CMPLT | 2023-01-23 11:05 | 2023-03-16 09:26 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:24:55 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056546 | 192626 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:29 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:26:35 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056547 | 192627 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 11:09 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:08:54 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056548 | 192628 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:31 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:31:10 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056549 | 192629 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:33 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:32:19 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056550 | 192630 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:39 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:37:32 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056551 | 192631 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:40 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:39:42 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056552 | 192632 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:47 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:45:43 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056553 | 192633 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-15 08:39 | Software upgrade to version 06.03.006.7404 | (TC 02/16/2023 08:38:14 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056554 | 192634 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 11:00 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:58:55 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056555 | 192635 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:54 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:53:54 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056556 | 192636 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-15 08:39 | Software upgrade to version 06.03.006.7404 | (TC 01/25/2023 08:38:59 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056557 | 192637 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:55 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:54:55 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056558 | 192638 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:57 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:56:24 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056559 | 192639 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:58 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:57:24 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056560 | 192640 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-15 08:30 | Software upgrade to version 06.03.006.7404 | (TC 01/04/2023 08:29:09 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056561 | 192641 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:09 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:07:23 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056562 | 192642 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:12 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:10:42 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056563 | 192643 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-15 08:40 | Software upgrade to version 06.03.006.7404 | (TC 01/11/2023 08:39:57 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056564 | 192644 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-15 08:36 | Software upgrade to version 06.03.006.7404 | (TC 01/04/2023 08:34:34 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056565 | 192645 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:50 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:45:16 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056566 | 192646 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 11:18 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:18:11 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056567 | 192647 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-15 08:36 | Software upgrade to version 06.03.006.7404 | (TC 01/13/2023 08:36:07 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056568 | 192648 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-15 08:37 | Software upgrade to version 06.03.006.7404 | (TC 01/13/2023 08:37:00 AM, Canjimavadivel, Kamallesh ) Upgrade complete
1056569 | 192649 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:27 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:26:37 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056570 | 192650 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:51 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:50:26 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056571 | 192651 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:54 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:53:28 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056572 | 192652 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:56 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:55:30 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056573 | 192653 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 11:15 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:14:45 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056574 | 192654 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:37 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:36:19 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056575 | 192655 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:13 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:12:34 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056576 | 192656 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:21 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:20:42 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056577 | 192657 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:22 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:21:22 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056578 | 192658 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:23 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:23:12 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056579 | 192659 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:36 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:35:24 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056580 | 192660 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:30 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:29:46 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056581 | 192661 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 11:19 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:18:49 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056582 | 192662 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 11:21 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:20:42 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056583 | 192663 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 11:04 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 11:03:35 AM, Canjimavadivel, Kamallesh ) Update completed by Zoll
1056584 | 192664 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:19 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:19:06 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056585 | 192665 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 10:20 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 10:15:20 AM, Canjimavadivel, Kamallesh ) Upgrade completed by Zoll
1056586 | 192666 | CON | CMPLT | CMPLT | 2023-01-23 11:06 | 2023-03-16 09:22 | Software upgrade to version 06.03.006.7404 | (TC 03/16/2023 09:21:51 AM, Canjimavadivel, Kamallesh ) Upgraded completed by Zoll
1056587 | 192667 | CON | ACTIV | ISSUE | 2023-01-23 11:06 |  | Software upgrade to version 06.03.006.7404
1065564 | 192668 | CM | CMPLT | CMPLT | 2023-01-23 11:07 | 2023-01-23 11:12 | Incoming Inspection | (TC 01/23/2023 11:11:12 AM, Ferreira, Qwynn) Set the device to have connectors for theunit to use to perform research Only to be used in the NICU and for the purposes of research only Stickered for use at Sinai.
1054043 | 192669 | PM | CMPLT | CMPLT | 2023-01-23 12:53 | 2023-01-23 12:54 | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO | (TC 01/23/2023 12:53:23 PM, Ferreira, Qwynn) PM Passed Self Test - Ok Electrical Safety - 154mOhm, 68.4uA
68279 | 192670 | CM | CMPLT | CMPLT | 2023-01-23 12:56 | 2023-01-23 12:58 | CMOS battery replacement | (TC 01/23/2023 12:55:42 PM, Baghirzade, Nara) Vinod mentioned that when the PC boots up, there are issues loading the BIOS settings and that the CMOS battery is running low. Replaced the CMOS battery; when we booted the PC, the prior BIOS settings were not retained. We set the date and time settings, and looked it up online, which mentions to leave the computer running for 2 days so that the CMOS battery can retain the settings. Vinod will turn off the PC on wednesday and then turn it back on to check if the BIOS settings were retained.
1055801 | 192671 | CM | CMPLT | CMPLT | 2023-01-23 14:10 | 2023-01-30 13:19 | Flat ECG line while trying to Cardio vert | (TC 01/23/2023 02:53:57 PM, Scero, Tomlyn) Replaced ECG leads and trunk. Tested using simulators for leads and pads. Everything working fine. returned to CCU crash cart.
65959 | 192672 | CM | CMPLT | CMPLT | 2023-01-23 14:17 | 2023-02-09 10:42 | "does not program" | (TC 01/23/2023 10:39:55 AM, Cortez, John) Unable to replicate issue. Rear case damaged. Cleaned unit, Replaced rear case (no charge-recall) Downloaded error/event logs, Calibration tasks completed in ASM, PM completed successfully, see ASM for results Ran test infusions at different rates (99ml/hr, 50 ml/hr and 1ml/hr) for 30 mins. Didn't come across any errors
| 192673 | CM | CMPLT | CMPLT | 2023-01-23 16:15 | 2023-01-23 16:15 | Site coverage.
| 192674 | CM | ACTIV | ISSUE | 2023-01-24 08:33 |  | ordered the spare parts for OR table
1062779 | 192675 | CM | CMPLT | CMPLT | 2023-01-24 08:58 | 2023-01-24 09:10 | Remote troubleshooting with Medisoft | (TC 01/24/2023 08:57:51 AM, Baghirzade, Nara) Smain from Medisoft connected to the CPET computer for ECG + Expair troubleshooting. He made a copy of the current version (1.34) of Expair and has upgraded the software to 1.35. We then went through the sequence of steps for doing a CPET and Smain suggested that the techs turn on all devices before starting spirometry. He will connect remotely for one of the tests to observe the ECG freezing during patient test.
| 192676 | PA | CMPLT | CMPLT | 2023-01-24 09:04 | 2024-01-03 13:37 | Sterile water for ICU | (TC 01/24/2023 09:09:58 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 02/03/2023 12:52:13 PM, Canjimavadivel, Kamallesh ) Placed order for water (TC 02/21/2023 12:57:47 PM, Canjimavadivel, Kamallesh ) Placed order for water (TC 03/20/2023 01:26:14 PM, Cortez, John) Placed order for water (TC 06/13/2023 03:58:38 PM, Cortez, John) Placed order for water (TC 07/25/2023 08:30:12 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 08/08/2023 04:01:16 PM, Cortez, John) Placed order for water (TC 08/21/2023 09:33:11 AM, Cortez, John) Ordered (TC 09/05/2023 01:20:47 PM, Cortez, John) Placed order for water (TC 10/02/2023 10:07:25 AM, Cortez, John) Placed order (TC 10/18/2023 02:31:45 PM, Cortez, John) Placed order (TC 10/30/2023 09:56:36 AM, Cortez, John) Placed order for water (TC 11/01/2023 09:14:10 AM, Cortez, John) Placed water (TC 12/12/2023 10:17:57 AM, Cortez, John) Placed order for water (TC 12/29/2023 10:19:58 AM, Cortez, John) Placed order
65908 | 192677 | CM | CMPLT | CMPLT | 2023-01-24 09:18 | 2023-02-01 07:43 | Ordering plunger assembly | (TC 01/24/2023 09:21:28 AM, Baghirzade, Nara) SC 1002866011 (TC 02/01/2023 07:42:49 AM, Baghirzade, Nara) PO 7202149618 | Plunger assembly is broken; does not move properly.
501063 | 192678 | CM | CMPLT | CMPLT | 2023-01-24 09:32 | 2023-01-24 11:34 | MSH CEOU RM#9 - IU22 - Error 511 SW, Monitor replacement | (TC 01/20/2023 09:32:19 AM, Chau, Ken) Error 511 resolved by FSE, deleting database in windows -waiting for monitor to be replaced (TC 01/24/2023 11:33:24 AM, Chau, Ken) Replaced monitor Reloaded SW Reconfigured Network Settings
1065567 | 192679 | CM | CMPLT | CMPLT | 2023-01-24 10:06 | 2023-01-30 13:11 | Leak in breathing system | (TC 01/25/2023 01:09:45 PM, Scero, Tomlyn) Took the parts apart and cleaned the safety valve and soda lime holder. Tested on AGM in OR 12, no leaks present.
| 192680 | PA | CMPLT | CMPLT | 2023-01-24 10:22 | 2023-01-31 08:17 | Ordering Cal 1 test gas for PACU | (TC 01/31/2023 08:16:17 AM, Reid, Emmerson ) Recieved PO 7202149928
1054034 | 192681 | CM | CMPLT | CMPLT | 2023-01-24 10:27 | 2023-02-16 09:39 | probe does not pass CaliScan test | probe does not pass CaliScan test. Shipping to vendor for warranty repair. Vendor replaced main PCBA and damaged scan cable. Tested. Functional.
501888 | 192682 | CM | CMPLT | CMPLT | 2023-01-24 15:28 | 2023-01-24 15:28 | MSH US RM 2 APLIO 500 SERVICE LICENCE INSTALL | (TC 01/24/2023 03:28:10 PM, Whittington, Scott) MSH US RM 2 APLIO 500 SERVICE LICENCE INSTALL | MSH US RM 2 APLIO 500 SERVICE LICENCE INSTALL
| 192683 | CM | CMPLT | CMPLT | 2023-01-24 17:17 | 2023-01-24 17:17 | Site coverage.
65381 | 192684 | CM | CMPLT | CMPLT | 2023-01-25 07:32 | 2023-01-25 07:33 | QC Checklist | (TC 01/25/2023 07:32:17 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
| 192685 | PA | CMPLT | CMPLT | 2023-01-25 08:49 | 2023-01-27 14:57 | Ordering rechargeable handle batteries | (TC 01/25/2023 08:57:39 AM, Baghirzade, Nara) SC 1002867455 (TC 01/27/2023 02:56:08 PM, Baghirzade, Nara) PO 7202150557
68913 | 192686 | CM | CMPLT | CMPLT | 2023-01-25 09:21 | 2023-03-29 15:52 | Damaged screen | (TC 01/24/2023 09:22:58 AM, Rendon Lopez , Elizabeth) Received replacement LCD, replaced. Logic board is also damaged, ordered part. (TC 03/27/2023 03:49:31 PM, Rendon Lopez , Elizabeth) Replaced logic board, network board, keypady (recall). Flashed PCU, downloaded config. Tested, all tests passed
| 192687 | CM | CMPLT | CMPLT | 2023-01-25 09:28 | 2023-03-21 09:25 | Spare LVP parts | Received
1053378 | 192688 | CM | CMPLT | CMPLT | 2023-01-25 09:35 | 2023-01-25 09:36 | Expiratory valve membrane needs replacement | (TC 01/24/2023 09:34:49 AM, Rendon Lopez , Elizabeth) Replaced membrane, reset hours, completed checks-OK Expiratory valve SN: 103075
69208 | 192689 | CM | CMPLT | CMPLT | 2023-01-25 09:36 | 2023-01-25 09:52 | Channel error | (TC 01/24/2023 09:51:08 AM, Rendon Lopez , Elizabeth) Downloaded error log, error 242.4010. Motor encoder, defective AIL assy. Replaced AIL assembly, front door and LCD. Calibrated-OK
1056553 | 192690 | CM | CMPLT | CMPLT | 2023-01-25 09:43 | 2023-02-16 08:48 | Device fail status | (TC 01/24/2023 09:43:15 AM, Canjimavadivel, Kamallesh ) Attempted to reload configuration to resolve issue, still no language files. Contacted Zoll tech support. Advised to send in unit as well as USB used for upgrade to investigate issue with upgrade. Zoll will be providing a loaner as well. SR# 971917 (TC 01/25/2023 01:49:26 PM, Canjimavadivel, Kamallesh ) Packed up device with USBs for investigation and shipped via KJV (TC 02/14/2023 08:46:11 AM, Canjimavadivel, Kamallesh ) Received from vendor, uploaded config and wifi settings. (TC 02/16/2023 08:46:59 AM, Canjimavadivel, Kamallesh ) Packaged loaner and sent back via shipping label provided by vendor | During upgrade, language files were corrupted. Keep saying unit failed
66461 | 192691 | CM | CMPLT | CMPLT | 2023-01-25 11:10 | 2023-02-01 09:15 | O2 sensor cal failure | "test skipped due to missing precondition" O2/Air supply missing. ordered PATO SA. Received and installed PATO SA. Tested. Functional.
1061419 | 192692 | CM | CMPLT | CMPLT | 2023-01-25 11:37 | 2023-01-25 11:37 | not moving | reset the bed. unlocked the controls. tested. functional.
62644 | 192693 | CM | CMPLT | CMPLT | 2023-01-25 11:47 | 2023-01-25 11:47 | alarming | tested. functional.
66444 | 192694 | CM | CMPLT | CMPLT | 2023-01-25 11:48 | 2023-01-25 11:48 | alarming | tested. functional.
62980 | 192695 | PM | CMPLT | CMPLT | 2023-01-25 11:57 | 2023-01-25 11:58 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 01/25/2023 11:57:33 AM, Baghirzade, Nara) PM passed and completed. No physical damage. @30.5cm distance between the surface and the light, the light intensity is: high = 31 uw/cm2/nm (30-35 uw/cm2/nm) low = 13.9 uw/cm2/nm (12-15uw/cm2/nm) Target light is good. Hours: 37.8 Cleaned the mesh filters and vacuumed the filters. One rubber foot was replaced as it was missing. all 5 rubber feet are in good condition. Electrical safety: Power cord resistance = 285 mohms Current leakage = 49.2 uA
| 192696 | PA | CMPLT | CMPLT | 2023-01-25 12:05 | 2023-01-26 12:55 | sterile water for altrix
1048018 | 192697 | CM | CMPLT | CMPLT | 2023-01-25 12:20 | 2023-01-25 12:23 | not charging | (TC 01/25/2023 12:20:22 PM, Ferreira, Qwynn) tested the unit and there was no charging indicator Perfromed basic troubleshooting to move cable to see if it is a cable issue or internal issue Noticed that the pin of the charging cable was lodged in the main charging port Removed the charging pin and stripped the charging cable to expose the inner cable assembly Resoldered the cable ensuring the grounding connection does not short with the hot end Tested the device and the device was charging normally and functioning correctly
| 192698 | CM | CMPLT | CMPLT | 2023-01-25 14:43 | 2023-01-25 14:44 | Exam Chair not working - Clinic | (TC 01/25/2023 02:43:07 PM, Ferreira, Qwynn) Foot pedal was not working - tested the foot pedal was working when connected to RJ45 tester Power cord had power going to it Ensured all connections were secure powered the device on and everything was working, problem was most likely loose cable that was fixed during testing
| 192699 | CM | CMPLT | CMPLT | 2023-01-25 14:59 | 2023-07-11 07:28 | MSH BIUS GEL WARMER | (TC 01/25/2023 03:07:19 PM, Whittington, Scott) MSH BIUS GEL WARMER | MSH BIUS GEL WARMER
1060700 | 192700 | CM | CMPLT | CMPLT | 2023-01-25 15:13 | 2023-02-01 08:25 | Overflow canisters filled | (TC 02/01/2023 08:24:06 AM, Reid, Emmerson ) Jon from MDRD cleaned the device and I returned it to service.
1063808 | 192701 | CM | CMPLT | CMPLT | 2023-01-25 15:13 | 2023-02-06 14:16 | MSH 15 SPP GE VOLUSON REMOTE VIDEO OLYMPUS SYSTEM | (TC 01/25/2023 03:13:21 PM, Whittington, Scott) MSH 15 SPP GE VOLUSON REMOTE VIDEO OLYMPUS SYSTEM RESEATED SVHS CABLE IN BOOM I/P COLOUR RETURNED. CHROMA CONNECTION MUST HAVE NOT BEEN SUFFICIENT | MSH 15 SPP GE VOLUSON REMOTE VIDEO OLYMPUS SYSTEM
| 192702 | CM | CMPLT | CMPLT | 2023-01-25 15:26 | 2023-07-11 07:28 | MSH BI HOLOGIC SAFETY ALERT | (TC 01/25/2023 03:28:32 PM, Whittington, Scott) MSH BI HOLOGIC SAFETY ALERT | MSH BI HOLOGIC SAFETY ALERT
1061204 | 192703 | CM | CMPLT | CMPLT | 2023-01-26 09:41 | 2023-01-26 09:42 | Boom Brake line popped off again | (TC 01/26/2023 09:41:45 AM, Reid, Emmerson ) The purple brake line popped off again on the foot boom due to not enough slack in the arm. Last time I pulled out some slack but from movement this causesd it to tighten up again. I am reaching out to olympus to find a better solution.
| 192704 | CM | CMPLT | CMPLT | 2023-01-26 12:18 | 2023-01-26 12:19 | Philips infrastructure meeting | (TC 01/26/2023 12:18:37 PM, Baghirzade, Nara) Philips infrastructure meeting
1048598 | 192705 | CM | CMPLT | CMPLT | 2023-01-26 13:25 | 2023-01-26 13:28 | Drug set stuck on pump | (TC 01/26/2023 01:26:19 PM, Reid, Emmerson ) The drug set was stuck on the pump and the key was not working. I noticed the pump batteries were dead and the pump would not turn on so I changed the batteries. I turned on the pump, and went into manual program mode, then I was able to get the set off. Working now.
1053347 | 192706 | CM | CMPLT | CMPLT | 2023-01-26 14:17 | 2023-01-27 10:15 | MSH BI 1270 H1 12NW stack migration to Cisco 172.20.220.136 | (TC 01/26/2023 02:17:12 PM, Whittington, Scott) MSH BI 1270 H1 12NW stack migration to Cisco 172.20.220.136 (TC 01/27/2023 10:14:21 AM, Whittington, Scott) MSH BI 1270 H1 12NW stack migration to Cisco 172.20.220.136 FOLLOWUP ALL MODALITIES ON MSH BI PULLING MWL | From: Whittington, Scott Sent: January 26, 2023 2:14 PM To: Riaz, Mirza <mirza.riaz@dxc.com> Cc: Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca>; Vaillancourt, Melanie - Sinai Health System <Melanie.Vaillancourt@sinaihealth.ca>; Kim, So Young - Sinai Health System <SoYoung.Kim@sinaihealth.ca>; Hewlett, Rose - Sinai Health System <Rose.Hewlett@sinaihealth.ca>; Potvin, Rochelle - Sinai Health System <Rochelle.Potvin@sinaihealth.ca>; Li, Chao - Sinai Health System <Chao.Li@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: RE: [External] 12NW stack migration to Cisco The only IP I can identify from our database is 172.20.220.136 which is H1 Thanks, Scott From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: January 26, 2023 1:46 PM To: Riaz, Mirza <mirza.riaz@dxc.com> Cc: Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca>; Vaillancourt, Melanie - Sinai Health System <Melanie.Vaillancourt@sinaihealth.ca>; Kim, So Young - Sinai Health System <SoYoung.Kim@sinaihealth.ca>; Hewlett, Rose - Sinai Health System <Rose.Hewlett@sinaihealth.ca>; Potvin, Rochelle - Sinai Health System <Rochelle.Potvin@sinaihealth.ca>; Li, Chao - Sinai Health System <Chao.Li@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] 12NW stack migration to Cisco Hi Mirza Thanks for the heads up. I will check the IP. We will check all of the modalities tomorrow morning to make sure they have connected to the network Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Riaz, Mirza <mirza.riaz@dxc.com> Sent: Thursday, January 26, 2023 1:32:54 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] 12NW stack migration to Cisco Hi Scott, We are moving devices from 12NW breast imaging to Cisco switches tonight. I see following devices are connect and will moved over ifc5 (Slot: 1 Port: 5) (5) Windows 8c:dc:d4:43:4c:94 (Hewlett Packard) 172.20.223.81 ifc17 (Slot: 1 Port: 17) (17) Not Profiled ac:1f:6b:b2:ae:64 172.20.220.136 Medical imaging H1 ifc129 (Slot: 2 Port: 1) (129) Windows b0:7b:25:1b:80:43 172.20.220.148 Medical imaging ifc130 (Slot: 2 Port: 2) (130) PXE Client 54:bf:64:7b:62:a3 172.20.220.133 Medical imaging ifc131 (Slot: 2 Port: 3) (131) Windows 6c:3b:e5:16:ce:ab (Hewlett Packard) 172.20.222.1 Medical imaging ifc135 (Slot: 2 Port: 7) (135) Not Profiled e4:54:e8:6a:93:33 172.20.220.135 Medical imaging Nac Disabled Thanks, Mirza DXC Technology
500801 | 192707 | CM | CMPLT | CMPLT | 2023-01-26 14:23 | 2023-01-27 10:22 | MSH BIUS STUDY FROM JAN 25 WILL NOT SEND TO JDMIPACS | (TC 01/26/2023 02:22:44 PM, Whittington, Scott) MSH BIUS STUDY FROM JAN 25 WILL NOT SEND TO JDMIPACS (TC 01/26/2023 02:22:44 PM, Whittington, Scott) MSH BIUS STUDY FROM JAN 25 WILL NOT SEND TO JDMIPACS INSTALLED NEW SERVICE LICENCE 2023 PUSH STUDY TO JDMIPACS-MSH OK | MSH BIUS STUDY FROM JAN 25 WILL NOT SEND TO JDMIPACS
65640 | 192708 | CM | CMPLT | CMPLT | 2023-01-26 14:28 | 2023-01-26 14:28 | alarming | tested. functional.
1055948 | 192709 | CM | CMPLT | CMPLT | 2023-01-26 14:29 | 2023-01-26 14:29 | BP cuff not working | Tested BP temp and SpO2. Functional
1055574 | 192710 | CM | CMPLT | CMPLT | 2023-01-26 14:30 | 2023-01-26 14:30 | BP cuff not working | Tested BP, temp and SpO2. Functional.
1056471 | 192711 | CM | CMPLT | CMPLT | 2023-01-26 14:31 | 2023-01-26 14:31 | BP cuff not working | Tested BP, temp and SpO2. Functional.
1054537 | 192712 | PM | CMPLT | CMPLT | 2023-01-26 14:39 | 2023-01-26 14:40 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/26/2023 02:39:48 PM, Canjimavadivel, Kamallesh ) Wiped outside surface. Cleaning cycle complete
| 192713 | CM | CMPLT | CMPLT | 2023-01-26 14:42 | 2023-02-02 08:35 | CADD pump spare pole mount adapter | Received
1063000 | 192714 | PM | CMPLT | CMPLT | 2023-01-26 15:01 | 2023-01-26 15:02 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/25/2023 03:01:25 PM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
1064771 | 192715 | CM | CMPLT | CMPLT | 2023-01-27 09:16 | 2023-01-30 14:51 | Incident# 54544 | (TC 01/27/2023 09:29:54 AM, Reid, Emmerson ) Went up to the OR with Jessica. Spoke with Edwina and Maureen about this incident. They told us that the nurses were not using the device correctly but had figured it out. We also noticed with the newer UHI-4 models the startup is different from the older ones. The newer ones require that the cavity mode button be pressed before the unit can be run. This was what confused the staff and caused the incident. Jessica has reached out to Jaron to determine why this new ones are like this, if it can be changed, and if we can standardise all the UHI-4 models to be the same. (TC 01/30/2023 02:50:34 PM, Reid, Emmerson ) Jaron was able to send instuctions to change the settings. On startup hold the cavity button, then change from setting 1 to setting 2. Working now.
500800 | 192716 | CM | CMPLT | CMPLT | 2023-01-27 10:24 | 2023-01-27 10:27 | MSH BIUS APLIO XG INSTALL 2023 SERVICE LICENCE, UPDATE ARCHIVE NODES, INCLUDING JDMIPACS-MSH | (TC 01/27/2023 10:24:20 AM, Whittington, Scott) MSH BIUS APLIO XG INSTALL 2023 SERVICE LICENCE, UPDATE ARCHIVE NODES, INCLUDING JDMIPACS-MSH | MSH BIUS APLIO XG INSTALL 2023 SERVICE LICENCE, UPDATE ARCHIVE NODES, INCLUDING JDMIPACS-MSH
500802 | 192717 | CM | CMPLT | CMPLT | 2023-01-27 10:25 | 2023-01-27 10:28 | MSH BIUS APLIO XG INSTALL 2023 SERVICE LICENCE, UPDATE ARCHIVE NODES, INCLUDING JDMIPACS-MSH | (TC 01/27/2023 10:25:07 AM, Whittington, Scott) MSH BIUS APLIO XG INSTALL 2023 SERVICE LICENCE, UPDATE ARCHIVE NODES, INCLUDING JDMIPACS-MSH | MSH BIUS APLIO XG INSTALL 2023 SERVICE LICENCE, UPDATE ARCHIVE NODES, INCLUDING JDMIPACS-MSH
68713 | 192718 | PM | CANCL | CANCL | 2023-01-27 10:53 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68020 | 192719 | CM | CMPLT | CMPLT | 2023-01-27 12:16 | 2023-06-16 08:16 | processor defective, not reading scopes | (TC 01/27/2023 12:15:56 PM, Samanian, Kamran) Tried to clean the sensor by Olympus rep in the unit but no success. advised by Olympus no charge PO to be sent out for repair. S/C#1002870095 (TC 01/31/2023 09:44:29 AM, Samanian, Kamran) filled out the repair form RMA#23015285 (TC 06/16/2023 08:14:56 AM, Samanian, Kamran) device was returned after repair and installed by Rob. tested and passed (TC 01/27/2023 12:15:56 PM, Samanian, Kamran) Tried to clean the sensor in CLV-190 by Olympus rep in the unit but no success. advised by Olympus no charge PO to be sent out for repair. repair approved by Anita- quote attached (TC 06/16/2023 08:14:56 AM, Samanian, Kamran) device was returned after repair and installed by Rob. tested and passed PO# 7202152599
500846 | 192720 | CM | CMPLT | CMPLT | 2023-01-27 12:16 | 2023-01-27 12:17 | MSH US APLIO XG 2023 SRVC LIC & DFRG | (TC 01/27/2023 12:15:56 PM, Whittington, Scott) MSH US APLIO XG 2023 SRVC LIC & DFRG | MSH US APLIO XG 2023 SRVC LIC & DFRG
500847 | 192721 | CM | CMPLT | CMPLT | 2023-01-27 12:17 | 2023-01-27 12:17 | MSH US APLIO XG 2023 SRVC LIC & DFRG | (TC 01/27/2023 12:17:43 PM, Whittington, Scott) MSH US APLIO XG 2023 SRVC LIC & DFRG | MSH US APLIO XG 2023 SRVC LIC & DFRG
| 192722 | CM | CANCL | CANCL | 2023-01-27 12:18 |  | Power bar order
1056453 | 192723 | CM | CMPLT | CMPLT | 2023-01-27 14:58 | 2023-01-27 15:03 | Missing accessories and functional check | (TC 01/27/2023 03:02:12 PM, Baghirzade, Nara) Replaced all accessories as they were missing; functional check ok. Tested all modules. Unit is fully functional. Spoke to Sri regarding their hose, spo2, cables, etc. stock and they have none. Followed up with Nima re: stock issue.
| 192724 | MEP | ACTIV | ISSUE | 2023-01-27 15:18 |  | Miscellaneous | (TC 01/27/2023 03:17:51 PM, Canjimavadivel, Kamallesh ) Meeting to discuss daily check procedure with defibs (TC 01/27/2023 03:18:25 PM, Canjimavadivel, Kamallesh ) Discussion with Zoll for dashboard feedback (TC 02/06/2023 08:09:01 AM, Canjimavadivel, Kamallesh ) Picking up AED batteries from UHN (TC 02/06/2023 08:08:49 AM, Canjimavadivel, Kamallesh ) Huddle (TC 02/10/2023 11:30:21 AM, Canjimavadivel, Kamallesh ) Mandatory e-learning (TC 02/10/2023 03:01:57 PM, Canjimavadivel, Kamallesh ) Cleaned up device dashboard (TC 02/16/2023 01:27:20 PM, Canjimavadivel, Kamallesh ) Patient monitoring committee (TC 02/21/2023 12:58:10 PM, Canjimavadivel, Kamallesh ) Prepared document with device dashboard instructions. (TC 02/21/2023 01:00:27 PM, Canjimavadivel, Kamallesh ) Huddle (TC 02/28/2023 03:08:21 PM, Canjimavadivel, Kamallesh ) Huddle (TC 02/28/2023 03:08:32 PM, Canjimavadivel, Kamallesh ) Patient monitoring committee (TC 03/01/2023 02:05:41 PM, Canjimavadivel, Kamallesh ) Capital planning meeting (TC 03/13/2023 11:39:38 AM, Canjimavadivel, Kamallesh ) Changed time on defibs in ICU, CCU, ED for daylight savings (TC 05/01/2023 02:50:46 PM, Canjimavadivel, Kamallesh ) PICiX setup lab (TC 05/25/2023 02:43:23 PM, Canjimavadivel, Kamallesh ) Patient monitoring group (TC 05/25/2023 02:43:44 PM, Canjimavadivel, Kamallesh ) Philips bi-weekly meeting (TC 05/25/2023 02:44:40 PM, Canjimavadivel, Kamallesh ) PM updates (TC 05/26/2023 03:13:08 PM, Canjimavadivel, Kamallesh ) 1 on 1 meeting (TC 05/29/2023 03:03:23 PM, Canjimavadivel, Kamallesh ) Philips MPV config meeting (TC 05/31/2023 10:27:59 AM, Canjimavadivel, Kamallesh ) Creating procedure for X-series PM (TC 05/31/2023 03:30:00 PM, Canjimavadivel, Kamallesh ) Hillrom support meeting (TC 06/01/2023 01:26:59 PM, Canjimavadivel, Kamallesh ) Planning software upgrade for 950 humidifiers (TC 06/01/2023 02:02:28 PM, Canjimavadivel, Kamallesh ) Updated on-call book (TC 06/02/2023 09:18:13 AM, Canjimavadivel, Kamallesh ) Scheduled patch updates on FocalPoint (TC 06/02/2023 09:50:22 AM, Canjimavadivel, Kamallesh ) TMS cleanup of older Hillrom bed WOs (TC 06/02/2023 12:35:55 PM, Canjimavadivel, Kamallesh ) Wound care discussion (TC 06/02/2023 03:18:49 PM, Canjimavadivel, Kamallesh ) Rounds (TC 06/05/2023 10:21:17 AM, Canjimavadivel, Kamallesh ) Installed under desk keyboard tray (TC 06/05/2023 02:06:14 PM, Canjimavadivel, Kamallesh ) Huddle (TC 06/05/2023 02:43:58 PM, Canjimavadivel, Kamallesh ) Unboxing weights and put away in storage (TC 06/06/2023 08:05:14 AM, Canjimavadivel, Kamallesh ) Updating orders on TMS (TC 06/06/2023 02:58:26 PM, Canjimavadivel, Kamallesh ) Helped with alarm testing in ICU (TC 06/07/2023 11:21:17 AM, Canjimavadivel, Kamallesh ) Helped Nigel clear out RT workshop (TC 06/07/2023 01:51:08 PM, Canjimavadivel, Kamallesh ) Equipment cleanup and disposal (TC 06/07/2023 02:41:47 PM, Canjimavadivel, Kamallesh ) Swapped out X-series batteries for conditioning (TC 06/08/2023 12:38:11 PM, Canjimavadivel, Kamallesh ) Cleaning out equipment in RT room (TC 06/09/2023 09:07:11 AM, Canjimavadivel, Kamallesh ) Went up to ICU to look at ultrasound machine. Forwarded info JDMI team to assess. (TC 06/09/2023 03:03:02 PM, Canjimavadivel, Kamallesh ) Discussion with tech from Baxter regarding bed repairs (TC 06/13/2023 03:20:14 PM, Canjimavadivel, Kamallesh ) wound vac prep (TC 06/13/2023 03:20:41 PM, Canjimavadivel, Kamallesh ) went to materials management to locate missing package (TC 06/14/2023 01:02:02 PM, Canjimavadivel, Kamallesh ) Prepared document for wound vac meeting (TC 06/14/2023 03:11:24 PM, Canjimavadivel, Kamallesh ) Preparing presentation for wound vac meeting (TC 06/15/2023 09:37:59 AM, Canjimavadivel, Kamallesh ) Cleaning up boxes in front area (TC 06/15/2023 01:33:07 PM, Canjimavadivel, Kamallesh ) wound vac presentation prep (TC 06/15/2023 01:33:23 PM, Canjimavadivel, Kamallesh ) Presentation with wound care team (TC 06/15/2023 02:56:36 PM, Canjimavadivel, Kamallesh ) PM updates (TC 06/19/2023 08:34:59 AM, Canjimavadivel, Kamallesh ) Put request for porter to remove recliner (TC 06/19/2023 08:44:23 AM, Canjimavadivel, Kamallesh ) Checked out ultrasound machine that was sent down and emailed details to JDMI team (TC 06/19/2023 10:35:12 AM, Canjimavadivel, Kamallesh ) Updated defib PM schedules with procedure (TC 06/19/2023 02:43:26 PM, Canjimavadivel, Kamallesh ) Huddle (TC 06/21/2023 11:32:05 AM, Canjimavadivel, Kamallesh ) Education committee (TC 06/21/2023 12:50:39 PM, Canjimavadivel, Kamallesh ) Creating PM proceudres for progressa beds and updating procedure for X-series (TC 06/21/2023 03:09:45 PM, Canjimavadivel, Kamallesh ) Capital plan meeting (TC 06/21/2023 03:10:04 PM, Canjimavadivel, Kamallesh ) Compiling alarm data from PICiX (TC 06/22/2023 10:57:44 AM, Canjimavadivel, Kamallesh ) Bi-weekly meeting with Philips (TC 06/22/2023 02:14:06 PM, Canjimavadivel, Kamallesh ) Added PM schedule to Progressa beds (TC 06/22/2023 02:46:16 PM, Canjimavadivel, Kamallesh ) Validate daily check document (TC 06/23/2023 02:23:44 PM, Canjimavadivel, Kamallesh ) Cleaning workshop (TC 06/26/2023 08:31:32 AM, Canjimavadivel, Kamallesh ) Swapping back PICiX in 2 nursing stations (TC 06/29/2023 11:10:14 AM, Canjimavadivel, Kamallesh ) Creating SBAR for wound vac follow up (TC 06/29/2023 03:00:03 PM, Canjimavadivel, Kamallesh ) OneRecall meeting (TC 07/04/2023 10:44:41 AM, Canjimavadivel, Kamallesh ) Follow-up on POs with Plexxus (TC 07/05/2023 01:29:26 PM, Canjimavadivel, Kamallesh ) Capital biweekly (TC 07/05/2023 02:42:19 PM, Canjimavadivel, Kamallesh ) Helping Qwynn install shelf (TC 07/06/2023 09:49:09 AM, Canjimavadivel, Kamallesh ) SES introduction meeting (TC 07/06/2023 10:44:27 AM, Canjimavadivel, Kamallesh ) Swapped out batteries from 2 rooms with batt malfunction errors. Started recalibration (TC 07/06/2023 10:52:51 AM, Canjimavadivel, Kamallesh ) Philips biweekly meeting (TC 07/10/2023 02:12:34 PM, Canjimavadivel, Kamallesh ) Huddle (TC 07/11/2023 02:14:50 PM, Canjimavadivel, Kamallesh ) TMS cleanup servo-u PM (TC 07/12/2023 11:16:24 AM, Canjimavadivel, Kamallesh ) Moving monitors to L6 storage (TC 07/14/2023 08:07:33 AM, Canjimavadivel, Kamallesh ) Follow up with Plexxus on old orders (TC 07/14/2023 09:05:26 AM, Canjimavadivel, Kamallesh ) cleaned up front area (TC 07/14/2023 02:09:03 PM, Canjimavadivel, Kamallesh ) Looked into making a cable needed for Alaris pump maintenance. Found alternative that can be ordered. Asked Samir to order from Amazon (TC 07/17/2023 11:01:05 AM, Canjimavadivel, Kamallesh ) Capital meeting (TC 07/17/2023 11:01:54 AM, Canjimavadivel, Kamallesh ) Adjusting PM schedule for pumps (TC 07/17/2023 02:08:33 PM, Canjimavadivel, Kamallesh ) Huddle (TC 07/17/2023 03:19:58 PM, Canjimavadivel, Kamallesh ) Building shelves for shop and cleaning up (TC 07/18/2023 03:15:32 PM, Canjimavadivel, Kamallesh ) Cleaning up shop and moving out cabinets (TC 07/19/2023 08:00:47 AM, Canjimavadivel, Kamallesh ) Following up on orders with Plexxus (TC 07/19/2023 01:09:01 PM, Canjimavadivel, Kamallesh ) SES internal prep meeting (TC 07/20/2023 12:56:53 PM, Canjimavadivel, Kamallesh ) Systems committee (TC 07/20/2023 02:40:44 PM, Canjimavadivel, Kamallesh ) Cleaning (TC 07/21/2023 01:47:01 PM, Canjimavadivel, Kamallesh ) Made cable for Alaris pump software (TC 07/24/2023 11:22:09 AM, Canjimavadivel, Kamallesh ) Installed latest version of SupportTool (TC 07/24/2023 01:39:31 PM, Canjimavadivel, Kamallesh ) Huddle (TC 07/24/2023 02:41:37 PM, Canjimavadivel, Kamallesh ) Went to 6th floor to inventory PICiX hardware (TC 07/24/2023 01:39:31 PM, Canjimavadivel, Kamallesh ) Huddle (TC 07/25/2023 08:41:22 AM, Canjimavadivel, Kamallesh ) Verifying TMS data for wound VAC (TC 07/25/2023 10:45:45 AM, Canjimavadivel, Kamallesh ) Cleaning LVPs (TC 07/25/2023 03:00:01 PM, Canjimavadivel, Kamallesh ) PM update (TC 07/25/2023 03:00:17 PM, Canjimavadivel, Kamallesh ) Looking at training options for Draeger vents (TC 07/31/2023 01:36:46 PM, Canjimavadivel, Kamallesh ) Huddle (TC 07/31/2023 03:05:57 PM, Canjimavadivel, Kamallesh ) Looking up manual for servo heater (TC 08/01/2023 10:55:02 AM, Canjimavadivel, Kamallesh ) Following up with older orders with Plexxus (TC 08/01/2023 01:38:04 PM, Canjimavadivel, Kamallesh ) Drug library update on pumps in biomed (TC 08/01/2023 02:45:31 PM, Canjimavadivel, Kamallesh ) Replaced batteries on Alaris pumps and started conditioning (TC 08/02/2023 08:09:51 AM, Canjimavadivel, Kamallesh ) Cleaning up WOs and replying to emails (TC 08/03/2023 02:44:32 PM, Canjimavadivel, Kamallesh ) Cleaning (TC 08/04/2023 09:17:11 AM, Canjimavadivel, Kamallesh ) On-call daily check (TC 08/08/2023 08:09:57 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 08/08/2023 03:22:37 PM, Canjimavadivel, Kamallesh ) Helped load up waste into Samir's car (TC 08/09/2023 07:56:48 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 08/10/2023 08:32:13 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 08/14/2023 10:08:45 AM, Canjimavadivel, Kamallesh ) Adding procedures to PM schedules (TC 08/14/2023 10:47:39 AM, Canjimavadivel, Kamallesh ) Moved microstreams to storage (TC 08/14/2023 11:11:25 AM, Canjimavadivel, Kamallesh ) Follow up woth KJV regarding older order (TC 08/14/2023 03:13:28 PM, Canjimavadivel, Kamallesh ) Huddle (TC 08/16/2023 09:03:25 AM, Canjimavadivel, Kamallesh ) completed onerecall task for ventilator alert (TC 08/16/2023 11:11:28 AM, Canjimavadivel, Kamallesh ) Patient monitoring meeting (TC 08/17/2023 08:33:28 AM, Canjimavadivel, Kamallesh ) Looking at Hillrom bed data on TMS (TC 08/17/2023 10:16:26 AM, Canjimavadivel, Kamallesh ) Helped with issue in OR 13 (TC 08/17/2023 11:02:15 AM, Canjimavadivel, Kamallesh ) Philips bi-weekly (TC 08/17/2023 03:08:29 PM, Canjimavadivel, Kamallesh ) PM update (TC 08/17/2023 03:13:48 PM, Canjimavadivel, Kamallesh ) Cleaning upmps for PM (TC 08/18/2023 09:07:29 AM, Canjimavadivel, Kamallesh ) ICU bed data analysis in TMS (TC 08/18/2023 03:07:59 PM, Canjimavadivel, Kamallesh ) Cleaning up items that are to be disposed from the workshop (TC 08/22/2023 10:44:03 AM, Canjimavadivel, Kamallesh ) CMBES meeting (TC 08/23/2023 03:01:13 PM, Canjimavadivel, Kamallesh ) Itinerary planning for Draeger training (TC 08/25/2023 11:10:40 AM, Canjimavadivel, Kamallesh ) 1 on 1 meeting (TC 08/25/2023 11:11:29 AM, Canjimavadivel, Kamallesh ) Cleaning up after leak found (TC 08/28/2023 02:46:35 PM, Canjimavadivel, Kamallesh ) Huddle (TC 08/28/2023 02:47:34 PM, Canjimavadivel, Kamallesh ) Unlocking tablets (TC 08/29/2023 03:51:26 PM, Canjimavadivel, Kamallesh ) Unlocking tablets (TC 08/31/2023 08:24:02 AM, Canjimavadivel, Kamallesh ) Cleaning up procedures for progressa beds (TC 08/31/2023 10:05:12 AM, Canjimavadivel, Kamallesh ) CMBES planning (TC 08/31/2023 10:49:28 AM, Canjimavadivel, Kamallesh ) Philips biweekly (TC 08/31/2023 11:08:43 AM, Canjimavadivel, Kamallesh ) Booking arrangements for training (TC 08/31/2023 03:11:48 PM, Canjimavadivel, Kamallesh ) TMS data cleanup (TC 09/01/2023 01:12:35 PM, Canjimavadivel, Kamallesh ) Discussion with Andrew about Zoll RescueNet (TC 09/05/2023 02:23:08 PM, Canjimavadivel, Kamallesh ) Helping Maisha identify equipment on shelves (TC 09/05/2023 02:23:37 PM, Canjimavadivel, Kamallesh ) CMBES conference prep (TC 09/06/2023 09:52:21 AM, Canjimavadivel, Kamallesh ) Unlocking tablets (TC 09/06/2023 07:54:19 AM, Canjimavadivel, Kamallesh ) Testing scenario for competency training (TC 09/07/2023 01:02:25 PM, Canjimavadivel, Kamallesh ) CMBES prep meeting (TC 09/08/2023 01:11:08 PM, Canjimavadivel, Kamallesh ) Zoll product feedback with team (TC 09/12/2023 09:59:59 AM, Canjimavadivel, Kamallesh ) Closed OneRecall alert (TC 09/13/2023 11:23:21 AM, Canjimavadivel, Kamallesh ) CMBES prep meeting (TC 09/14/2023 12:58:27 PM, Canjimavadivel, Kamallesh ) CMBES meeting (TC 09/15/2023 12:53:11 PM, Canjimavadivel, Kamallesh ) Full team meeting (TC 09/15/2023 12:53:25 PM, Canjimavadivel, Kamallesh ) Wound VAC discussion with Adeel (TC 09/15/2023 12:53:11 PM, Canjimavadivel, Kamallesh ) Full team meeting (TC 09/18/2023 08:01:44 AM, Canjimavadivel, Kamallesh ) Huddle (TC 09/18/2023 08:02:03 AM, Canjimavadivel, Kamallesh ) CMBES planning (TC 09/18/2023 08:02:34 AM, Canjimavadivel, Kamallesh ) Creating poll for CMBES (TC 09/19/2023 03:15:33 PM, Canjimavadivel, Kamallesh ) Wound care discussion (TC 09/19/2023 03:16:52 PM, Canjimavadivel, Kamallesh ) Patient monitoring scenarios (TC 09/20/2023 03:18:50 PM, Canjimavadivel, Kamallesh ) Capital projects update (TC 09/20/2023 03:19:16 PM, Canjimavadivel, Kamallesh ) CMBES poll update (TC 09/20/2023 03:21:59 PM, Canjimavadivel, Kamallesh ) Patient monitoring scenarios (TC 09/21/2023 12:53:23 PM, Canjimavadivel, Kamallesh ) CMBES conference planning (TC 09/21/2023 12:53:40 PM, Canjimavadivel, Kamallesh ) Patient monitoring scenarios (TC 09/25/2023 07:51:47 AM, Canjimavadivel, Kamallesh ) Daily check update (TC 09/26/2023 08:07:40 AM, Canjimavadivel, Kamallesh ) Daily check update (TC 09/26/2023 03:14:24 PM, Canjimavadivel, Kamallesh ) Systems committee update (TC 09/27/2023 01:57:36 PM, Canjimavadivel, Kamallesh ) Department meeting (TC 09/27/2023 01:58:14 PM, Canjimavadivel, Kamallesh ) Patient monitoring scenarios (TC 09/27/2023 02:11:24 PM, Canjimavadivel, Kamallesh ) Daily check update (TC 09/28/2023 08:29:16 AM, Canjimavadivel, Kamallesh ) Daily check update (TC 09/28/2023 10:41:31 AM, Canjimavadivel, Kamallesh ) CMBES committee meeting (TC 09/28/2023 10:41:53 AM, Canjimavadivel, Kamallesh ) Philips bi-weekly meeting (TC 09/29/2023 03:11:04 PM, Canjimavadivel, Kamallesh ) Patient monitoring scenarios (TC 09/29/2023 03:11:34 PM, Canjimavadivel, Kamallesh ) PM update (TC 09/29/2023 03:15:07 PM, Canjimavadivel, Kamallesh ) Zoll dashboard feedback (TC 10/02/2023 02:34:08 PM, Canjimavadivel, Kamallesh ) Huddle (TC 10/02/2023 02:37:39 PM, Canjimavadivel, Kamallesh ) CMBES poll follow-up analysis (TC 10/03/2023 03:07:49 PM, Canjimavadivel, Kamallesh ) Zoll testing with Jonathon (TC 10/04/2023 08:04:50 AM, Canjimavadivel, Kamallesh ) Disposing batteries to e-waste (TC 10/12/2023 11:26:14 AM, Canjimavadivel, Kamallesh ) CMBES committee (TC 10/12/2023 02:19:29 PM, Canjimavadivel, Kamallesh ) Summarized data from CMBES poll and provided to group (TC 10/12/2023 02:20:02 PM, Canjimavadivel, Kamallesh ) Discussed TMS clean up plan with Samir (TC 10/19/2023 01:00:44 PM, Canjimavadivel, Kamallesh ) TMS cleanup (TC 10/20/2023 08:37:30 AM, Canjimavadivel, Kamallesh ) Cleaning up incoming boxes area (TC 10/20/2023 01:52:43 PM, Canjimavadivel, Kamallesh ) Full team meeting (TC 10/24/2023 02:24:58 PM, Canjimavadivel, Kamallesh ) Huddle (TC 10/25/2023 02:48:01 PM, Canjimavadivel, Kamallesh ) Testing case push from AED3 with Jonathan (TC 10/26/2023 02:32:00 PM, Canjimavadivel, Kamallesh ) Meeting with wound care to finalize qc process for VAC Ulta (TC 10/27/2023 01:19:04 PM, Canjimavadivel, Kamallesh ) 1 on 1 meeting (TC 10/31/2023 08:44:33 AM, Canjimavadivel, Kamallesh ) Daily check (TC 11/01/2023 08:44:48 AM, Canjimavadivel, Kamallesh ) Daily check (TC 11/01/2023 03:03:06 PM, Canjimavadivel, Kamallesh ) CMBES planning meeting (TC 11/01/2023 03:03:26 PM, Canjimavadivel, Kamallesh ) Huddle (TC 11/02/2023 07:45:39 AM, Canjimavadivel, Kamallesh ) Daily check update (TC 11/02/2023 10:16:12 AM, Canjimavadivel, Kamallesh ) Created briefing note document for wound care process (TC 11/03/2023 03:22:03 PM, Canjimavadivel, Kamallesh ) Draeger training online (TC 11/06/2023 11:30:00 PM, Canjimavadivel, Kamallesh ) Health check update with Philips (TC 11/07/2023 02:29:37 PM, Canjimavadivel, Kamallesh ) Updating defib info in dashboard (TC 11/07/2023 02:30:12 PM, Canjimavadivel, Kamallesh ) working with IT to try and resolve issue with e-req (TC 11/08/2023 09:50:03 AM, Canjimavadivel, Kamallesh ) Updating defib dashboard (TC 11/17/2023 11:21:24 AM, Canjimavadivel, Kamallesh ) CMBES conference planning (TC 11/17/2023 11:21:50 AM, Canjimavadivel, Kamallesh ) Full team meeting (TC 11/17/2023 11:22:05 AM, Canjimavadivel, Kamallesh ) ED monthly meeting (TC 11/20/2023 02:50:21 PM, Canjimavadivel, Kamallesh ) Huddle (TC 11/20/2023 03:19:46 PM, Canjimavadivel, Kamallesh ) Getting costs for servo repair for RT team (TC 11/21/2023 03:08:11 PM, Canjimavadivel, Kamallesh ) Contacting vendors for CMBES conf (TC 11/22/2023 08:03:19 AM, Canjimavadivel, Kamallesh ) Cleaning up old POs and WOs (TC 11/22/2023 10:53:27 AM, Canjimavadivel, Kamallesh ) TMS cleanup and inventory (TC 11/22/2023 02:04:37 PM, Canjimavadivel, Kamallesh ) Sim Man training (TC 11/22/2023 03:01:08 PM, Canjimavadivel, Kamallesh ) Systems committee (TC 11/24/2023 03:21:07 PM, Canjimavadivel, Kamallesh ) CMBES comittee update (TC 11/29/2023 01:33:10 PM, Canjimavadivel, Kamallesh ) Huddle (TC 11/30/2023 02:58:49 PM, Canjimavadivel, Kamallesh ) Meeting with support services to discuss new cleaning plan for C. diff. (TC 12/01/2023 10:36:43 AM, Canjimavadivel, Kamallesh ) CMBES commitee (TC 12/07/2023 01:10:11 PM, Canjimavadivel, Kamallesh ) Huddle (TC 12/08/2023 10:05:29 AM, Canjimavadivel, Kamallesh ) CMBES committee (TC 12/08/2023 10:51:07 AM, Canjimavadivel, Kamallesh ) Following up with vendors for CMBES conference (TC 12/11/2023 01:02:10 PM, Canjimavadivel, Kamallesh ) Skin and wound care follow up meeting (TC 12/11/2023 01:02:39 PM, Canjimavadivel, Kamallesh ) Wound VAC follow up with past WOs to identify QC fails (TC 12/12/2023 07:55:19 AM, Canjimavadivel, Kamallesh ) Email Stryker to sort out wrong part delivery (TC 12/12/2023 03:30:00 PM, Canjimavadivel, Kamallesh ) Helping John swap out light source for ENDO (TC 12/13/2023 08:35:54 AM, Canjimavadivel, Kamallesh ) Emailing vendor for CMBES conference (TC 12/13/2023 01:56:38 PM, Canjimavadivel, Kamallesh ) Huddle (TC 12/15/2023 12:43:11 PM, Canjimavadivel, Kamallesh ) Mandatory e-learning (TC 12/15/2023 12:43:11 PM, Canjimavadivel, Kamallesh ) CMBES conference planning (TC 12/15/2023 12:43:11 PM, Canjimavadivel, Kamallesh ) Full team meeting (TC 12/15/2023 02:54:34 PM, Canjimavadivel, Kamallesh ) IBE solution review (TC 12/18/2023 11:19:12 AM, Canjimavadivel, Kamallesh ) Working on paper for CMBES conference (TC 12/19/2023 11:04:28 AM, Canjimavadivel, Kamallesh ) Pulling alarm logs from PICiX for Nima (TC 12/19/2023 11:05:15 AM, Canjimavadivel, Kamallesh ) CMBES committee planning (TC 12/20/2023 08:50:15 AM, Canjimavadivel, Kamallesh ) Updating asset location info for extra X3 power supplies (TC 12/22/2023 09:05:34 AM, Canjimavadivel, Kamallesh ) Updated documents for CMBES committee (TC 12/22/2023 02:00:40 PM, Canjimavadivel, Kamallesh ) CMBES planning committee (TC 12/22/2023 02:01:12 PM, Canjimavadivel, Kamallesh ) Systems committee meeting (TC 01/04/2024 02:50:07 PM, Canjimavadivel, Kamallesh ) TMS clean-up (TC 01/05/2024 08:20:49 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 01/08/2024 08:08:39 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 01/08/2024 08:09:42 AM, Canjimavadivel, Kamallesh ) Meeting with GoEngineer team to discuss details for CMBES conference (TC 01/10/2024 09:26:05 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 01/11/2024 08:12:52 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 01/11/2024 02:50:47 PM, Canjimavadivel, Kamallesh ) Paper for CMBES conference (TC 01/12/2024 10:17:47 AM, Canjimavadivel, Kamallesh ) 1 on 1 Meeting (TC 01/15/2024 10:30:00 AM, Canjimavadivel, Kamallesh ) OR training session (TC 01/15/2024 03:00:00 PM, Canjimavadivel, Kamallesh ) Capital meeting (TC 01/16/2024 01:48:05 PM, Canjimavadivel, Kamallesh ) Monday huddle (TC 01/19/2024 11:09:42 AM, Canjimavadivel, Kamallesh ) Full team meeting (TC 01/19/2024 11:09:58 AM, Canjimavadivel, Kamallesh ) CMBES conference update (TC 01/19/2024 02:34:18 PM, Canjimavadivel, Kamallesh ) Finalizing draft for CMBES paper (TC 01/22/2024 03:08:08 PM, Canjimavadivel, Kamallesh ) Huddle (TC 01/22/2024 03:08:25 PM, Canjimavadivel, Kamallesh ) Editing paper for CMBES submission (TC 01/23/2024 10:33:32 AM, Canjimavadivel, Kamallesh ) Helped Naz load boxes onto cart (TC 01/24/2024 03:19:57 PM, Canjimavadivel, Kamallesh ) Helped Qwynn with segmenting case (TC 01/25/2024 10:13:21 AM, Canjimavadivel, Kamallesh ) Follow-up with vendors for CMBES conference. (TC 01/25/2024 02:11:53 PM, Canjimavadivel, Kamallesh ) Updated TMS for HL-90 PM procedure (TC 01/25/2024 03:32:53 PM, Canjimavadivel, Kamallesh ) Segmenting on 3D slicer (TC 01/26/2024 10:33:54 AM, Canjimavadivel, Kamallesh ) Post-print processing of model (TC 01/26/2024 01:30:00 PM, Canjimavadivel, Kamallesh ) Picked up 3D printer from receiving for Qwynn (TC 01/29/2024 08:46:10 AM, Canjimavadivel, Kamallesh ) Updating info for CMBES committee (TC 01/29/2024 02:00:00 PM, Canjimavadivel, Kamallesh ) Huddle (TC 01/30/2024 08:11:25 AM, Canjimavadivel, Kamallesh ) Replying to emails and TMS activities (TC 02/05/2024 02:00:00 PM, Canjimavadivel, Kamallesh ) Huddle (TC 02/06/2024 11:23:05 AM, Canjimavadivel, Kamallesh ) Follow-up on old shopping carts (TC 02/12/2024 03:06:25 PM, Canjimavadivel, Kamallesh ) Huddle (TC 02/15/2024 09:00:10 AM, Canjimavadivel, Kamallesh ) Updating presenter info for CMBES conference (TC 02/15/2024 01:49:21 PM, Canjimavadivel, Kamallesh ) Cleaned up biomed shop (TC 02/16/2024 10:15:30 AM, Canjimavadivel, Kamallesh ) Follow up for CMBES planning committee (TC 02/16/2024 01:04:45 PM, Canjimavadivel, Kamallesh ) Full team meeting (TC 02/16/2024 02:30:21 PM, Canjimavadivel, Kamallesh ) Cleaned up shop (TC 02/26/2024 10:58:15 AM, Canjimavadivel, Kamallesh ) BD recall meeting (TC 02/26/2024 11:10:27 AM, Canjimavadivel, Kamallesh ) Helped Courtney with X-series upgrade for NICU defibs (TC 02/26/2024 02:13:42 PM, Canjimavadivel, Kamallesh ) Huddle (TC 02/27/2024 02:57:24 PM, Canjimavadivel, Kamallesh ) PM meeting (TC 02/28/2024 01:29:56 PM, Canjimavadivel, Kamallesh ) Capital meeting (TC 02/29/2024 03:09:49 PM, Canjimavadivel, Kamallesh ) Servo-u 4.4 overview with Getinge (TC 03/01/2024 10:49:31 AM, Canjimavadivel, Kamallesh ) CMBES committee updates (TC 03/04/2024 03:16:40 PM, Canjimavadivel, Kamallesh ) Huddle (TC 03/08/2024 01:47:33 PM, Canjimavadivel, Kamallesh ) Completing performance review (TC 03/12/2024 10:15:50 AM, Canjimavadivel, Kamallesh ) Closing pump recall WOs (TC 03/13/2024 03:21:22 PM, Canjimavadivel, Kamallesh ) Closing BD pump WOs (TC 03/14/2024 09:06:52 AM, Canjimavadivel, Kamallesh ) Closing BD recall WOs (TC 03/14/2024 11:40:58 AM, Canjimavadivel, Kamallesh ) Covering a call from 17M regarding PSN ADT issue. Discovered Naz was also looking into issue and handed it off. Informed unit we will connect once resolved from Masimo end. (TC 03/14/2024 03:26:55 PM, Canjimavadivel, Kamallesh ) Closing BD recall WOs (TC 03/15/2024 08:18:50 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 03/15/2024 11:31:42 AM, Canjimavadivel, Kamallesh ) Booking accomodations for RSNA conf (TC 03/15/2024 11:32:11 AM, Canjimavadivel, Kamallesh ) Registering for CMBES conf (TC 03/15/2024 11:32:31 AM, Canjimavadivel, Kamallesh ) CMBES conf planning meeting update (TC 03/15/2024 03:16:35 PM, Canjimavadivel, Kamallesh ) Closing BD recall WOs (TC 03/18/2024 08:26:09 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 03/18/2024 01:55:40 PM, Canjimavadivel, Kamallesh ) Huddle (TC 03/18/2024 03:25:17 PM, Canjimavadivel, Kamallesh ) Closing BD recall wOs (TC 03/19/2024 08:43:20 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 03/19/2024 03:11:56 PM, Canjimavadivel, Kamallesh ) Closing BD recall WOs (TC 03/21/2024 08:11:32 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 03/20/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Showinf Rutvik how to complete Zimmer ATS PM (TC 03/21/2024 11:02:57 AM, Canjimavadivel, Kamallesh ) Closing BD recall WOs (TC 03/22/2024 11:10:13 AM, Canjimavadivel, Kamallesh ) Moving ventilators from docks to RT room (TC 03/22/2024 11:31:29 AM, Canjimavadivel, Kamallesh ) Closing alerts in OneRecall (TC 03/27/2024 03:01:28 PM, Canjimavadivel, Kamallesh ) PM updates (TC 03/27/2024 03:02:15 PM, Canjimavadivel, Kamallesh ) Closing BD recall WOs (TC 04/02/2024 10:20:39 AM, Canjimavadivel, Kamallesh ) Following up on outstanding PO orders with Mohawk Medbuy (TC 04/03/2024 01:21:06 PM, Canjimavadivel, Kamallesh ) 1 on 1 meeting (TC 04/03/2024 01:33:10 PM, Canjimavadivel, Kamallesh ) Pulling up and sorting alarm data for ICU PICiX and forwarded to Kamran (TC 04/04/2024 02:23:36 PM, Canjimavadivel, Kamallesh ) Reviewing hamilton vent alarms with Kamran (TC 04/04/2024 02:28:38 PM, Canjimavadivel, Kamallesh ) Moving equipment to 6th fl storage (TC 04/05/2024 09:54:33 AM, Canjimavadivel, Kamallesh ) CMBES conference planning update (TC 04/05/2024 09:57:51 AM, Canjimavadivel, Kamallesh ) Following up on outstanding orders (TC 04/05/2024 09:58:13 AM, Canjimavadivel, Kamallesh ) Reached out to vendor for PM parts quote (TC 04/05/2024 02:30:00 PM, Canjimavadivel, Kamallesh ) Meeting with GoEngineer to discuss details for CMBES presentation (TC 04/09/2024 02:59:47 PM, Canjimavadivel, Kamallesh ) Cleaning up after BD recall work on 14 (TC 04/16/2024 01:31:48 PM, Canjimavadivel, Kamallesh ) Huddle (TC 04/24/2024 02:19:29 PM, Canjimavadivel, Kamallesh ) PM updates (TC 04/25/2024 02:25:48 PM, Canjimavadivel, Kamallesh ) Spoke with Renee from Zoll regarding some changes in dashboard (TC 04/26/2024 12:59:28 PM, Canjimavadivel, Kamallesh ) Cleaning up AED data on TMs and dashboard. Validated locations of AEDs at MSH (TC 04/26/2024 01:00:18 PM, Canjimavadivel, Kamallesh ) Following up on emails (TC 04/29/2024 01:54:04 PM, Canjimavadivel, Kamallesh ) Huddle (TC 04/30/2024 01:06:41 PM, Canjimavadivel, Kamallesh ) Servo recertification with Samir (TC 05/13/2024 03:49:38 PM, Canjimavadivel, Kamallesh ) Huddle (TC 05/14/2024 01:29:47 PM, Canjimavadivel, Kamallesh ) Updated support tool to latest version (TC 05/16/2024 09:50:51 AM, Canjimavadivel, Kamallesh ) Followup on on old orders, went to receiving to ask about missing boxes (TC 05/16/2024 10:30:00 AM, Canjimavadivel, Kamallesh ) CMBES conference planning follow-up meeting (TC 05/16/2024 01:53:58 PM, Canjimavadivel, Kamallesh ) Worked on submission for CMBES conference (TC 05/24/2024 08:54:43 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 05/24/2024 10:50:59 AM, Canjimavadivel, Kamallesh ) Prepared presentation for CMBES conference (TC 05/27/2024 08:02:10 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 05/27/2024 11:01:36 AM, Canjimavadivel, Kamallesh ) Updated info in Zoll dashboard (TC 05/27/2024 02:26:29 PM, Canjimavadivel, Kamallesh ) Huddle (TC 05/31/2024 09:56:44 AM, Canjimavadivel, Kamallesh ) Worked with Steve Arai to get cables required for RD capability on vital signs monitors. (TC 06/04/2024 01:21:41 PM, Canjimavadivel, Kamallesh ) Tally up feedback forms from the CMBES conference. Follow up with presenters (TC 06/05/2024 11:26:36 AM, Canjimavadivel, Kamallesh ) Cleaned up AED data on TMS and forwarded list to Nigel for inventory check (TC 06/10/2024 10:55:05 AM, Canjimavadivel, Kamallesh ) Swapping out X-series batteries (TC 06/11/2024 11:17:45 AM, Canjimavadivel, Kamallesh ) Swapping defib batteries (TC 06/11/2024 03:16:56 PM, Canjimavadivel, Kamallesh ) Validating PICiX hardware info (TC 06/12/2024 11:00:00 AM, Canjimavadivel, Kamallesh ) Capital projects meeting (TC 06/12/2024 01:49:20 PM, Canjimavadivel, Kamallesh ) Validating PICiX hardware in NICU (TC 06/19/2024 03:23:02 PM, Canjimavadivel, Kamallesh ) Huddle (TC 06/26/2024 03:31:56 PM, Canjimavadivel, Kamallesh ) PM updates (TC 07/22/2024 09:51:02 AM, Canjimavadivel, Kamallesh ) Configuring email alerts for device dashboard (TC 07/22/2024 02:39:49 PM, Canjimavadivel, Kamallesh ) Huddle (TC 07/29/2024 10:14:41 AM, Canjimavadivel, Kamallesh ) Moving stryker PM parts to bed shop (TC 07/29/2024 10:15:06 AM, Canjimavadivel, Kamallesh ) Closing old open WOs (TC 07/29/2024 01:04:39 PM, Canjimavadivel, Kamallesh ) Capital update (TC 07/29/2024 01:55:34 PM, Canjimavadivel, Kamallesh ) Removing alerts from defib dashboard (TC 08/21/2024 02:41:29 PM, Canjimavadivel, Kamallesh ) Capital projects update (TC 08/21/2024 02:47:04 PM, Canjimavadivel, Kamallesh ) Discussion of recall for Level 1 fast flow device (TC 08/26/2024 02:00:00 PM, Canjimavadivel, Kamallesh ) Huddle (TC 09/09/2024 02:00:00 PM, Canjimavadivel, Kamallesh ) Huddle (TC 09/16/2024 01:59:03 PM, Canjimavadivel, Kamallesh ) Huddle (TC 09/23/2024 09:01:39 AM, Canjimavadivel, Kamallesh ) Huddle (TC 10/07/2024 02:32:12 PM, Canjimavadivel, Kamallesh ) Huddle (TC 10/21/2024 03:20:20 PM, Canjimavadivel, Kamallesh ) Cleaning up dialysis PM schedules in TMS (TC 10/28/2024 11:11:11 AM, Canjimavadivel, Kamallesh ) Stryker lifenet info session (TC 11/04/2024 02:34:23 PM, Canjimavadivel, Kamallesh ) Huddle (TC 11/18/2024 02:32:30 PM, Canjimavadivel, Kamallesh ) Huddle (TC 11/25/2024 02:39:48 PM, Canjimavadivel, Kamallesh ) Huddle (TC 11/27/2024 02:16:57 PM, Canjimavadivel, Kamallesh ) Dialysis overview with team (TC 11/27/2024 02:20:06 PM, Canjimavadivel, Kamallesh ) Picking up boxes from receiving (TC 11/28/2024 10:11:58 AM, Canjimavadivel, Kamallesh ) Capital meeting updates (TC 11/28/2024 11:01:09 AM, Canjimavadivel, Kamallesh ) Following up on recall notice for wound VAC
| 192725 | CM | CMPLT | CMPLT | 2023-01-30 10:27 | 2023-02-27 13:24 | Order Radical batteries for Becky | (TC 01/30/2023 10:27:18 AM, Chukri, Samir) Ordered Batteries for Radical 7 no charge. (TC 02/27/2023 01:24:42 PM, Chukri, Samir) Received batteries.
1062140 | 192726 | CM | CMPLT | CMPLT | 2023-01-30 11:01 | 2023-01-30 13:09 | switches not working | (TC 01/30/2023 01:07:57 PM, Scero, Tomlyn) Went to 12S to check the bed in the hallway. Turned the main switch off and pifted the foot panel and reseated it. Plugged in the power in and turned the main switch back on to reset the bed. All the button now working fine on the keypad. Notified NUA at the desk. | "Can you please look into it. None of the switches (keypad) wokring. please assist. thanks.
1051085 | 192727 | CM | CMPLT | CMPLT | 2023-01-30 11:02 | 2023-02-01 10:45 | Battery not holding a charge | (TC 01/30/2023 11:18:50 AM, Scero, Tomlyn) Battery malfunctioning. Tried to recondition, failed. Replaced battery and changed the power cord to a longer one because they asked.
1056547 | 192728 | CM | CMPLT | CMPLT | 2023-01-30 11:09 | 2023-01-30 11:13 | battery replacement | (TC 01/30/2023 11:09:05 AM, Canjimavadivel, Kamallesh ) Replaced battery
1056587 | 192729 | CM | CMPLT | CMPLT | 2023-01-30 11:20 | 2023-01-30 11:21 | Battery replacement | (TC 01/30/2023 11:20:04 AM, Canjimavadivel, Kamallesh ) Replaced battery
1065595 | 192730 | CM | ACTIV | ISSUE | 2023-01-30 12:40 |  | Acceptance
| 192731 | PA | CMPLT | CMPLT | 2023-01-30 12:54 | 2023-02-07 14:07 | Lead sets for TC70 | (TC 01/30/2023 08:10:27 AM, Canjimavadivel, Kamallesh ) Placed order for leads (TC 02/07/2023 02:06:20 PM, Canjimavadivel, Kamallesh ) Received leads, placed on shelf | Biomed stock
501105 | 192732 | CM | CMPLT | CMPLT | 2023-01-30 14:23 | 2023-01-31 16:23 | System would not boot up | (TC 01/31/2023 11:53:28 AM, Chau, Ken) Replaced 3 of the HDD (TC 02/02/2023 01:31:17 PM, Fan, Tak) System still had issues. Main monitor also shuts off intermittently. Replaced UMB. Updated MAC address with IT team. (TC 02/03/2023 09:34:35 AM, Whittington, Scott) MSH 5 US IU22 WITH UMB REPLACED AND IT NOTIFIED OF MAC CHANGE, STILL NOT NETWORKING FOLLOWUP WITH MIRZA TO HAVE NEW MAC WHITELISTED IN UTILITY ROOM WHERE SYSTEM NORMALLY UPLOADS IMAGES, BOTTOM NW JACK SET TO 10 MB MIDDLE JACK IS OK AT AUTO/AUTO SYSTEM NETOWKING OK (TC 02/02/2023 09:34:35 AM, Whittington, Scott) MSH 5 US IU22 WITH UMB REPLACED AND IT NOTIFIED OF MAC CHANGE, STILL NOT NETWORKING FOLLOWUP WITH MIRZA TO HAVE NEW MAC WHITELISTED IN UTILITY ROOM WHERE SYSTEM NORMALLY UPLOADS IMAGES, BOTTOM NW JACK SET TO 10 MB MIDDLE JACK IS OK AT AUTO/AUTO SYSTEM NETOWKING OK | System stops at software load #49 giving a 609 error code. Called Philips. Ref# 0121121673 Replaced all three hard drives. new mac 00:D0:68:19:9F:0F
501584 | 192733 | CM | CMPLT | CMPLT | 2023-01-30 14:26 | 2023-01-30 14:26 | Remote monitor not working. |  | Rewired video cable. Tested OK
1064773 | 192734 | CM | CMPLT | CMPLT | 2023-01-30 14:49 | 2023-01-30 14:50 | Changed the settings | (TC 01/30/2023 02:48:49 PM, Reid, Emmerson ) After the incident with the staff not able to start the insufflator Jaron sent instructions on how to change this setting. Changed this and working now.
| 192735 | CM | CMPLT | CMPLT | 2023-01-30 17:14 | 2023-01-30 17:14 | Site coverage
| 192736 | CM | CMPLT | CMPLT | 2023-01-31 07:34 | 2023-01-27 07:35 | ENDO Go live on 6th floor- all day in the unit | (TC 01/23/2023 07:34:05 AM, Samanian, Kamran) present in the unit to resolve any issue (TC 01/24/2023 07:34:46 AM, Samanian, Kamran) present in the unit to resolve any issue (TC 01/25/2023 07:34:58 AM, Samanian, Kamran) present in the unit to resolve any issue (TC 01/26/2023 07:35:12 AM, Samanian, Kamran) present in the unit to resolve any issue (TC 01/27/2023 07:35:28 AM, Samanian, Kamran) present in the unit to resolve any issue
1065567 | 192737 | CM | CMPLT | CMPLT | 2023-01-31 09:07 | 2023-01-31 10:23 | 2 Year PM | (TC 01/25/2023 10:22:24 AM, Scero, Tomlyn) All PM parts replaced. Breathing system cleaned and tested. No leaks. Passed Self test.
1059828 | 192738 | CM | CMPLT | CMPLT | 2023-01-31 09:36 | 2023-01-31 09:37 | Cleaning | (TC 01/31/2023 09:36:40 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1054957 | 192739 | CM | CMPLT | CMPLT | 2023-01-31 11:15 | 2023-01-31 11:17 | Reading low | (TC 01/31/2023 11:16:39 AM, Canjimavadivel, Kamallesh ) Calibration passed on tester after one try. | Staff reported they believe thermometer reading about 1 degree lower than it should
1062149 | 192740 | CM | CMPLT | CMPLT | 2023-01-31 11:19 | 2023-01-31 13:25 | Constantly reading between 34.5-34.7 | (TC 01/31/2023 01:23:30 PM, Canjimavadivel, Kamallesh ) Passed calibration after second attempt
501064 | 192741 | CM | CMPLT | CMPLT | 2023-01-31 11:49 | 2023-02-07 16:34 | MSH CEOU RM# 7 - Main monitor display is off | (TC 01/31/2023 11:49:35 AM, Chau, Ken) System monitor is displaying, its intermittent. Was advised they restarted several times -keeping WO open until end of week -closig work order, been 7 days and no reported issue
502047 | 192742 | CM | CMPLT | CMPLT | 2023-01-31 11:51 | 2023-01-31 11:52 | OPG SPP RM#1 - Display isnt working on extended monitors | (TC 01/31/2023 11:51:17 AM, Chau, Ken) Connections was mixed up -placed video connections on proper spot and tested okay
1059778 | 192743 | FC | CMPLT | CMPLT | 2023-01-31 15:43 | 2023-01-31 15:54 | QC Checklist | (TC 01/31/2023 03:53:23 PM, Baghirzade, Nara) QC Checklist completed. Returned to 11 South.
| 192744 | PA | CMPLT | CMPLT | 2023-01-31 15:45 | 2023-05-15 08:07 | Ordering exam lights for NICU | (TC 01/31/2023 03:49:55 PM, Baghirzade, Nara) SC 1002872931 (TC 05/15/2023 01:07:44 PM, Rhuland, Courtney) Lights received and unboxed Serial numbers: 1119636024 - date sticker added for warranty 1120145021 - date sticker added for warranty 1119636019 1119641015 - date sticker added for warranty 1119636015 - date sticker added for warranty 1119641023 1119636008 1119636021 1119636027 1119641016 1119636010 - date sticker added for warranty 1119636004 - date sticker added for warranty 1119636017 1119641020 1119636011 1118620036 1119636013 1119641014 - date sticker added for warranty 1119636020 1119638018 - date sticker added for warranty Received lights., PO#7202154903(8). We received total 20 x lights. PO#7202154905(7) and 7202157465(5)
| 192745 | CON | CMPLT | CMPLT | 2023-01-31 15:54 | 2023-01-31 15:55 | Secondary alerting for Masimo PSN | (TC 01/31/2023 03:54:30 PM, Baghirzade, Nara) Secondary alerting for Masimo PSN with John, Naz, Agata, and Becky on 11th fl.
| 192746 | CM | CMPLT | CMPLT | 2023-01-31 16:27 | 2023-01-31 16:27 | Site Coverage
26115 | 192747 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26116 | 192748 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26122 | 192749 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26123 | 192750 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2024-09-19 13:03 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 09/19/2024 01:03:24 PM, Rhuland, Courtney) Shipped out to Spectramedx for PM Returned and all tests passed Verified O2 values 21% - 20.9% 40% - 38% 60% - 59% 80% - 80.9% 100% - 100% Spectramedx report attached
26132 | 192751 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26142 | 192752 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26147 | 192753 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
51743 | 192754 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2024-01-12 14:23 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 01/12/2024 02:22:39 PM, Rhuland, Courtney) PM complete See attached checklist for details
52117 | 192755 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2024-01-26 12:04 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | See WO #201218
60025 | 192756 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
60026 | 192757 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: ELECTROSURGICAL UNITS, SPECIALTY
62500 | 192758 | PM | CANCL | CANCL | 2023-02-01 02:19 | 2024-07-12 13:41 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport
62610 | 192759 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-02-14 08:25 | Multiple PM Work Order for 62610 | (TC 02/14/2023 08:25:45 AM, Samanian, Kamran) closed, replaced with new one- device retired
62611 | 192760 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
62618 | 192761 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-05-23 11:54 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/23/2023 11:44:44 AM, Rhuland, Courtney) Power cord replaced Heater doors would not close when lowering canopy Removed canopy top and tightened screw holding the door cable Issue resolved PM complete RH calibrated All tests passed Electrical safety - OK
62728 | 192762 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency
62729 | 192763 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency
62934 | 192764 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
62942 | 192765 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: High Flow Insufflator
63058 | 192766 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63523 | 192767 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Scales, Infant
63937 | 192768 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Pulse-Free Flow Peristaltic Pump
64240 | 192769 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM Major Annual: Microstream CO2 extension | Unable to located equipment in the CCU
64241 | 192770 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM Major Annual: Microstream CO2 extension
65117 | 192771 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-26 10:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/26/2023 10:22:57 AM, Reid, Emmerson ) PM complete, see proceedure for details.
65123 | 192772 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-04-21 11:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/21/2023 11:23:15 AM, Reid, Emmerson ) PM Complete, see proceedure for details.
65130 | 192773 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-30 09:54 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/30/2023 09:53:07 AM, Reid, Emmerson ) PM Complete, see proceedure for details.
65153 | 192774 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-10-31 11:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 10/31/2023 11:32:02 AM, Reid, Emmerson ) PM Complete, see procedure for details.
65155 | 192775 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-11-20 12:38 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 11/20/2023 12:37:15 PM, Reid, Emmerson ) PM Complete, see proceedure for details.
65172 | 192776 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-30 11:30 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/30/2023 11:29:06 AM, Reid, Emmerson ) PM complete, see proceedure for details
65881 | 192777 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-08 10:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/08/2023 10:22:58 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65904 | 192778 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65905 | 192779 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-09-22 12:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/22/2023 12:21:40 PM, Campbell, Connor) PM complete/passed, refer to ASM
65912 | 192780 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-07-11 09:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/16/2023 09:31:12 AM, Rhuland, Courtney) Front case replaced Mylar gasket replaced Top disk holder replaced LED backlight cable replaced (spare) Waiting on rear cases (backorder) (TC 07/11/2023 09:39:02 AM, Rhuland, Courtney) Rear case replaced PM failed clamp open test Calibration and check in performed All tests passed PM complete Refer to ASM for details
65913 | 192781 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-07-11 08:20 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/06/2023 12:51:24 PM, Rhuland, Courtney) Front and rear case replaced Left IUI connector replaced Left handle replaced Mylar gasket replaced (TC 07/11/2023 08:19:22 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
65919 | 192782 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65922 | 192783 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-29 11:00 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/29/2023 11:00:08 AM, Baghirzade, Nara) Originally was sent down due to channel error when attempting to setup for use. Checked error logs and noted a repeated pressure sensor error; checked previous WOs and sensor has already been reseated. Pressure sensor disc was replaced. Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65927 | 192784 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65936 | 192785 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-21 13:49 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2023 01:47:39 PM, Baghirzade, Nara) Replaced plunger assembly. Full calibration of barrel clamp, pressure sensor disc, and plunger accuracy and force was completed. Check-in tests passed and completed. IUIs and device cleaned.
65939 | 192786 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65950 | 192787 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65955 | 192788 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-04-20 09:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/20/2023 09:32:47 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
65961 | 192789 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-07-04 16:13 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/04/2023 04:11:17 PM, Rhuland, Courtney) Plunger claws closing very slowly Opened up top of plunger assembly and cleaned claws and reseated spring to proper orientation Left side handle replaced PM complete All tests passed Refer to ASM for details
65972 | 192790 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-04-19 09:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/19/2023 09:47:14 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65986 | 192791 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-07-13 14:49 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/13/2023 02:48:16 PM, Rhuland, Courtney) Rear case replaced Left side handle replaced PM complete All tests passed Refer to asm for details
65994 | 192792 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-09 13:02 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/09/2023 12:56:30 PM, Baghirzade, Nara) Sent to biomed due to channel error. When reviewing the error log, noted a repeated 354.6770 error. Pressure sensor was replaced. Replaced the front case as it was damaged. Reprinted the NICU barcode sticker. During PM, the barrel clamp was failing the functional test; replaced the barrel clamp with a spare one in biomed. Successfully completed a full calibration of the plunger, barrel clamp, and pressure sensor disc. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
66036 | 192793 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66058 | 192794 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-05-16 12:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/16/2023 12:48:01 PM, Rhuland, Courtney) Cracked front and rear cases replaced (recall) Female IUI connector replaced Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
66059 | 192795 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-07-18 06:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/17/2023 08:36:32 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 58 mohms Leakage current: 78.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 07/18/2023 06:56:12 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4033 mamphr PM completed.
66068 | 192796 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66462 | 192797 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-04-26 08:36 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. NO PS500. Tested. Functional. SW.2.42 Leakage: 71uA and 84 ohm.
66472 | 192798 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-05-01 07:55 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. NO PS500. Tested. Functional. SW.2.42 Leakage: 75uA and 83 ohm.
66783 | 192799 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | No longer in use, replaced with X3
66803 | 192800 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | BP Check
66804 | 192801 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | BP Check
66806 | 192802 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | BP Check
67377 | 192803 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67378 | 192804 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67769 | 192805 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-23 14:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/23/2023 02:43:25 PM, Reid, Emmerson ) PM complete, see proceedure for details.
67888 | 192806 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-19 08:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/13/2023 08:45:13 AM, Rendon Lopez , Elizabeth) Updated software and library. PM completed, all tests passed
67915 | 192807 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-30 13:09 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/30/2023 01:07:07 PM, Ferreira, Qwynn) PM Passed Replaced two port hole latches for the recall as the doors were swapped Calibrated ssytem voltage Calibrated main voltage Calibrated line voltage Scale was accurate Functional Test - OK Electrical Safety - OK
67916 | 192808 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-10-02 11:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 10/02/2023 11:00:40 AM, Baghirzade, Nara) PM completed. Porthole door hinge was repaired. Bottom clips on door hinge were replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C. Functional tests passed. Electrical safety results - Leakage current = 85.0 uA Power cord resistance = 154 mohms
67929 | 192809 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-08 09:48 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/08/2023 09:46:33 AM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C on the analyzer. Functional tests passed. Electrical safety results - Leakage current = 33.0 uA Power cord resistance = 246 mohms
67943 | 192810 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-22 13:57 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 03/22/2023 01:55:26 PM, Ferreira, Qwynn) PM Passed Came down with no temp visible on screen Went through troubleshooting steps Found it to be the air probe Replaced air probe and tested - Values visible on screen. Replaced recall parts Calibrated scale Calibrated line voltage calibrated main voltage calibrated system voltage Electrical Safety - OK Functional Test - OK
67946 | 192811 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-11-09 11:05 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/09/2023 10:24:13 AM, Campbell, Connor) Porthole door was cracked, replaced it all latches replaced as per recall PM Passed Calibrated System - OK Calibrated Line Voltage - OK Calibrated Scale - OK Leakage Current Off - 77.4 uA Leakage Current ON - 82.6 uA Ground Resistance Off - 66 mOhms Ground Resistance ON - 63 mOhms Functional Test - Pass - Temp = 33.0 deg - Humidity = 75%
67965 | 192812 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-09-12 12:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/12/2023 11:59:25 AM, Campbell, Connor) PM Passed Calibrated system voltage Calibrated main voltage Calibrated scale Electrical Safety - OK * refer to procedure Functional Test - OK changed the north east door lock to be the correct one because the door is on the wrong way meaning the locks were not oriented correctly before.
68022 | 192813 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-05-02 08:28 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. PS500 last year. PM GS500 this year. Replaced left door. Replaced tank holder(lug). Tested. Functional. SW.2.51 Leakage: 81 uA and 90 ohm.
68023 | 192814 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-05-04 07:49 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. PS500 last year. PM GS500 this year. Tested. Functional. SW.2.51 Leakage: 75 uA and 88 ohm.
68296 | 192815 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-01 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/01/2023 11:20:49 AM, Rhuland, Courtney) Rate calibrated PM complete All tests passed post calibration Refer to ASM for details
68415 | 192816 | PM | ACTIV | READY | 2023-02-01 02:19 |  | PM MAJOR - Delivery Module DESO
68416 | 192817 | PM | ACTIV | READY | 2023-02-01 02:19 |  | PM MAJOR - Delivery Module DESO
68555 | 192818 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-05-17 09:28 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 05/08/2023 03:56:33 PM, Rendon Lopez , Elizabeth) Brought unit down. 6 year PM parts replacement started (TC 05/11/2023 09:25:44 AM, Rendon Lopez , Elizabeth) Pm parts replaced, tasks started (TC 05/15/2023 09:25:44 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed (TC 05/16/2023 09:25:44 AM, Rendon Lopez , Elizabeth) Device returned to L&D OR as spare. See attached results sheet
68623 | 192819 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68736 | 192820 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-04-13 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/13/2023 02:32:20 PM, Reid, Emmerson ) Replaced the battery. Ran ASM MSH PM PCU, all tests passed.
68743 | 192821 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68753 | 192822 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68756 | 192823 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68766 | 192824 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-09-08 11:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/08/2023 11:36:49 AM, Campbell, Connor) Battery replaced PM passed and complete
68790 | 192825 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-05-24 13:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/24/2023 01:37:16 PM, Rhuland, Courtney) Bent power cord replaced (spare) Missing cord retainer replaced Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
68848 | 192826 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68889 | 192827 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68918 | 192828 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68920 | 192829 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-11-06 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/30/2023 08:50:21 AM, Rendon Lopez , Elizabeth) Run battery conditioning (TC 10/31/2023 08:50:21 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
68964 | 192830 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-02-08 10:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/07/2023 01:04:21 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 300 mohms Leakage current: 74.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 02/08/2023 10:31:56 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4086 mamphr PM completed.
68998 | 192831 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-02 11:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 11:46:36 AM, Baghirzade, Nara) Replaced front case. Replaced power cord. PCU was missing power cord holder/cover and power cord strap. Placed holder and cord strap. PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 76 mohms Leakage current: 69.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
69001 | 192832 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-03 07:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 12:37:27 PM, Baghirzade, Nara) Replaced power cord. PM passed and completed. Electrical Safety - Ground resistance: 102 mohms Leakage current: 60.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/03/2023 07:38:08 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4121 mamphr PM completed.
69006 | 192833 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-12 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/12/2023 02:18:11 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 146 mohms Leakage current: 70.3 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
69089 | 192834 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-02 09:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 09:40:16 AM, Rhuland, Courtney) Rate calibrated PM complete All tests passed post calibration Refer to ASM for details
69122 | 192835 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69134 | 192836 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-06-09 14:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:57:11 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69184 | 192837 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69231 | 192838 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-01 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/01/2023 01:39:14 PM, Rhuland, Courtney) Rate calibrated PM complete All tests passed post calibration Refer to ASM for details
69233 | 192839 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69244 | 192840 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69297 | 192841 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-04-03 08:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2023 08:15:31 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69304 | 192842 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69380 | 192843 | PM | CANCL | CANCL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69391 | 192844 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-08-18 07:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/18/2023 07:53:05 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69512 | 192845 | PM | CNP | ENL | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69585 | 192846 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-11-06 09:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/01/2023 09:10:09 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69616 | 192847 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-07-19 13:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/19/2023 01:13:14 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
65316 | 192848 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-04-24 15:46 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 04/21/2023 03:37:21 PM, Cortez, John) Physical inspection passed, cleaned unit, Functional test passed, Mains leakage (uA), used BC safety analyzer: NC, SFC - HI 0.7, 7.3 - LO 0.9, 6.3 - P/P 0.6, 9.6 - BI 0.1, 0.5 - handswitch 0.5, 0.7 all within safety limits RF output passed, see attached document RF leakage test passed -35 HI 54 mA -35 BI1 34 mA. BI2 13 mA
32540 | 192849 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-06 20:09 | PM - Monthly | PM completed NRA/AIR cal Small FOV done sc-582487 TOS phantom a(-992.8,9.9) b(332.0,14.5) c(132.7,13.1) d(97.4,13.5) e(-104.3,12.7) f(0.6,14.2) -passed phantom tests
32560 | 192850 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-02-28 15:00 | PM - Monthly | PM completed Mechanical PM done Faulty heat exchanger replaced SC-1138341 TOS phantom a(-994.5,9.7) b(332.7,14.3) c(132.8,14.3) d(95.9,13.3) e(-106.4,13.2) f(0.1,12.4) -passed all phantom tests
32580 | 192851 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-01 15:20 | PM - Monthly | PM completed Mechanical PM done NRA/AIR Cal done-S FOV sc-265674 TOS phantom a(-993.7,10.0) b(334.5,15.5) c(132.6,14.7) d(97.0,15.0) e(-106.9,13.2) f(0.7,13.8) -passed all phantom tests
541124 | 192852 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541125 | 192853 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541127 | 192854 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541128 | 192855 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541205 | 192856 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541206 | 192857 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541207 | 192858 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541208 | 192859 | PM | ACTIV | ISSUE | 2023-02-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1046273 | 192860 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-02 16:11 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 02/22/2023 02:11:40 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/02/2023 04:11:24 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046274 | 192861 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-02 16:10 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 02/22/2023 02:13:00 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/02/2023 04:09:59 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046275 | 192862 | PM | CMPLT | CMPLT | 2023-02-01 02:19 | 2023-03-02 16:10 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 02/23/2023 07:17:43 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/02/2023 04:10:42 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046276 | 192863 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-02 16:12 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 02/24/2023 07:22:44 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 02/23/2023 07:22:44 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/02/2023 04:11:57 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046294 | 192864 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-31 11:52 | PM-MAJOR - Annual PM for: Electrosurgical Units | (TC 03/30/2023 11:50:50 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached results. using for L&D spare
1046407 | 192865 | PM | ACTIV | ISSUE | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1046408 | 192866 | PM | ACTIV | ISSUE | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047159 | 192867 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | Annual PM | WO created semi annually by accident.
1047844 | 192868 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-07-14 10:04 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/14/2023 10:03:23 AM, Baghirzade, Nara) Battery replaced. PM completed. See procedure document for PM results. PM schedule updated.
1047854 | 192869 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-05-11 09:30 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/11/2023 09:28:43 AM, Baghirzade, Nara) PM completed. See PM procedure for results.
1047858 | 192870 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-05-10 11:40 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/10/2023 11:39:21 AM, Rhuland, Courtney) PM complete All tests passed Refer to procedure for details (TC 05/10/2023 02:07:32 PM, Rhuland, Courtney) Noticed housing was out of place Snapped case back into position Ran functional check - OK
1048472 | 192871 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-08 11:49 | HARP - Semi Annual |  | Harp completed.
1048472 | 192872 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-08 11:51 | PM |  | PM completed. Tube covers melted. Ordered new ones.
1048527 | 192873 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-05-19 07:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/19/2023 07:39:44 AM, Rhuland, Courtney) Cracked rear and front cases replaced (recall) Mylar gasket replaced PM complete All tests passed Refer to ASM for details
1048537 | 192874 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-08 17:08 | HARP - Semi-Annual | HARP done.
1048537 | 192875 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-08 17:18 | PM - Semi-Annual | PM was done.
1048629 | 192876 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048632 | 192877 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048636 | 192878 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048641 | 192879 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048647 | 192880 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048648 | 192881 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048649 | 192882 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048650 | 192883 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048652 | 192884 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048656 | 192885 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048659 | 192886 | PM | CNP | ENL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048667 | 192887 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048669 | 192888 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048671 | 192889 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048679 | 192890 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048682 | 192891 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048687 | 192892 | PM | CNP | ENL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048690 | 192893 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048692 | 192894 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048698 | 192895 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048699 | 192896 | PM | CNP | ENL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049416 | 192897 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-05-31 08:30 | PM-MAJOR - Annual PM For: Scale | (TC 05/29/2023 08:29:29 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Refer to WO 195011
1051372 | 192898 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051373 | 192899 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051378 | 192900 | PM | CMPLT | PMFAIL | 2023-02-01 02:20 | 2023-10-05 09:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 09:54:23 AM, Campbell, Connor) replaced battery attempted to connect to ASM, not working faulty network board waiting for order of replacements
1053034 | 192901 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-28 10:44 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 03/28/2023 10:42:46 AM, Ferreira, Qwynn) Completed PM Tested the X3 and stickered Functional test - OK Sunction battery replaced in 2022 Sent back to unit
1053348 | 192902 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-08 16:54 | MSH BI 1272 H3 HARP MARCH 7 TRGT SCHEDULED SEMI-ANNUAL MSH_1272_MG | (TC 03/08/2023 04:54:07 PM, Dolinar, John) MSH BI 1272 H3 HARP MARCH 7 TRGT SCHEDULED SEMI-ANNUAL MSH_1272_MG (TC 03/08/2023 04:54:30 PM, Whittington, Scott) MSH BI 1272 H3 HARP MARCH 7 TRGT SCHEDULED SEMI-ANNUAL MSH_1272_MG (TC 03/08/2023 05:54:00 PM, Whittington, Scott) MSH BI 1272 H3 HARP MARCH 7 TRGT SCHEDULED SEMI-ANNUAL MSH_1272_MG | MSH BI 1272 H3 HARP MARCH 7 TRGT SCHEDULED SEMI-ANNUAL MSH_1272_MG
1053378 | 192903 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-17 12:27 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | Completed annual PM. Tested. Functional.
1054037 | 192904 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM | PM not required
1054108 | 192905 | PM | CANCL | DUP | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every 2 years
1054109 | 192906 | PM | CANCL | DUP | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every 2 years
1054111 | 192907 | PM | CANCL | DUP | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every 2 years
1054112 | 192908 | PM | CANCL | DUP | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every 2 years
1054113 | 192909 | PM | CANCL | DUP | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every 2 years
1054114 | 192910 | PM | CANCL | DUP | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every 2 years
1054115 | 192911 | PM | CANCL | DUP | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every 2 years
1054441 | 192912 | PM | ACTIV | ISSUE | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM
1057680 | 192913 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-02-09 13:13 | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO | (TC 02/09/2023 01:12:28 PM, Ferreira, Qwynn) PM Passed Self - Test - OK Electrical Safety - OK 132.2uA 36mOhm
1057926 | 192914 | PM | ACTIV | ISSUE | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1058056 | 192915 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-31 14:42 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 03/27/2023 02:40:25 PM, Scero, Tomlyn) PM started. 2 year Maintenance parts replaced. (TC 03/29/2023 02:41:12 PM, Scero, Tomlyn) Continued with PM (TC 03/31/2023 02:41:27 PM, Scero, Tomlyn) All PM tests complete and passed. Results sheet attached .
1058057 | 192916 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-09-08 09:24 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 08/31/2023 02:37:55 PM, Scero, Tomlyn) 2 year PM parts replaced and IPM-L started. O2 concentration wouldn't rise above 94%. Back and forth with tech support on how to troubleshoot and fix issue. Read documentation on how to use the Visia2 program. Found RS232 cable and eventually conected AGM to the program on my laptop. Used the program to calibrate the gas sensors so they are now reading properly. Will continue with PM procedures. (TC 09/05/2023 09:11:47 AM, Scero, Tomlyn) Continued with PM proceedures (TC 09/07/2023 09:21:41 AM, Scero, Tomlyn) All tests and proceedures complete and passed. Results sheet attached. Side cover right (MK04158) and 3 Snap lock connector (MK027799) replaced becaues they were broken.
1058058 | 192917 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Anesthesia Units | AGM not available
1058059 | 192918 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Anesthesia Units | AGM not available
1058183 | 192919 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-06-21 14:12 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/21/2023 02:11:22 PM, Reid, Emmerson ) PM complete, see proceedure for details.
1058190 | 192920 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-06-23 10:09 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/23/2023 10:08:21 AM, Reid, Emmerson ) PM Complete, see proceedure for details.
1058191 | 192921 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-06-06 10:07 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/06/2023 10:06:46 AM, Reid, Emmerson ) PM Complete, see proceedure for details.
1058192 | 192922 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-04-21 10:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/21/2023 10:56:31 AM, Reid, Emmerson ) PM complete, see proceedure for details.
1058193 | 192923 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-06-28 14:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/28/2023 02:23:01 PM, Reid, Emmerson ) PM complete, see proceedure for details.
1058196 | 192924 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-11-06 14:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 11/06/2023 02:53:56 PM, Reid, Emmerson ) PM Complete, see procedure for details.
1058197 | 192925 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-12-05 14:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 10/27/2023 09:53:44 AM, Reid, Emmerson ) Requested a PO for a repair to be done due to error 45985 showing on the screen of the pump (TC 12/05/2023 02:42:56 PM, Reid, Emmerson ) Pushed library and completed PM, see procedure for details.
1058255 | 192926 | PM | CNP | ENL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059202 | 192927 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | 3 Year PM was decided. Changed pm schedule to reflect that.
1059204 | 192928 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | 3 Year PM was decided. Changed pm schedule to reflect that.
1059206 | 192929 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | 3 Year PM was decided. Changed pm schedule to reflect that.
1059207 | 192930 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | 3 Year PM was decided. Changed pm schedule to reflect that.
1059208 | 192931 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | 3 Year PM was decided. Changed pm schedule to reflect that.
1059209 | 192932 | PM | CANCL | CANCL | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | 3 Year PM was decided. Changed pm schedule to reflect that.
1059223 | 192933 | PM | CNP | ENL | 2023-02-01 02:20 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide | Looked for equipment multiple times. May have been taken to another unit.
1059655 | 192934 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-02-01 11:43 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 01/31/2023 11:42:57 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1059795 | 192935 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-17 17:38 | HARP | HARP
1059795 | 192936 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-16 17:39 | PM - Semi-Annual | PM was done.
1059796 | 192937 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-09 14:44 | HARP | HARP
1059796 | 192938 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-04-27 17:53 | PM - Semi-Annual
1059810 | 192939 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-07 16:38 | HARP | HARP
1059810 | 192940 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-07 16:24 | PM - Semi-Annual | PM was done.
1063401 | 192941 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-05-31 14:13 | PM-MAJOR - Annual PM For: Ventilators | (TC 05/31/2023 01:37:44 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 8364 hrs Total vent time: 4203 hrs Replaced cable guide
1063402 | 192942 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-02-28 14:28 | PM-MAJOR - Annual PM For: Ventilators | (TC 02/28/2023 02:24:38 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 328 hrs Total vent time: 220 hrs
1063403 | 192943 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-21 14:11 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/21/2023 02:08:58 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 7541 hrs Total vent time: 3459 hrs
1063404 | 192944 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-20 11:07 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/20/2023 11:06:13 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 6516 hrs Total vent time: 4116 hrs
1063405 | 192945 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-06-01 10:37 | PM-MAJOR - Annual PM For: Ventilators | (TC 06/01/2023 10:32:00 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 8723 hrs Total vent time: 4620 hrs
1063406 | 192946 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-22 10:24 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/22/2023 10:24:33 AM, Chukri, Samir) Completed and replaced PM-5000 H kit. Passed pre-user and patient circuit leak .
1063407 | 192947 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-20 10:32 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/20/2023 10:27:30 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 7428 hrs Total vent time: 4388 hrs
1063408 | 192948 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-06 10:06 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/06/2023 10:00:43 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 307 hrs Total vent time: 0 hrs
1063409 | 192949 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-05-01 09:56 | PM-MAJOR - Annual PM For: Ventilators | (TC 05/01/2023 09:48:22 AM, Canjimavadivel, Kamallesh ) Cleaned fan filter Replaced filters and nozzles Pre check pass Total op time: 1892 hrs Total vent time: 661 hrs
1063746 | 192950 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-04-07 14:00 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 04/07/2023 02:00:00 PM, Baghirzade, Nara) Replaced the back filters Replaced battery Replaced air filter Replaced port/connector covers Assembled the unit Unit was disinfected, rinsed, and refilld with sterile water and sent back to unit. (TC 04/07/2023 02:00:00 PM, Ferreira, Qwynn) Replaced the back filters Replaced Battery Replaced air filter Replaced port/connector covers Assembled the unit Unit was disinfected, rinsed, and refilld with sterile water and sent back to unit.
1064201 | 192951 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-14 15:03 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/14/2023 03:02:40 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 6290 hrs Total vent time: 2956 hrs
1064202 | 192952 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-22 10:41 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/22/2023 10:40:07 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 6875 hrs Total vent time: 3972 hrs
1064203 | 192953 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-21 12:30 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/21/2023 12:30:22 PM, Chukri, Samir) Replaced PM-5000 kit. Exp.Membrane 85% performed functional and electrical safety Procuder 5466 attached
1064204 | 192954 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-08-18 09:52 | PM-MAJOR - Annual PM For: Ventilators | (TC 08/18/2023 09:51:17 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 4165 hrs Total vent time: 3156 hrs
1064205 | 192955 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-06-01 09:39 | PM-MAJOR - Annual PM For: Ventilators | (TC 06/01/2023 09:33:03 AM, Canjimavadivel, Kamallesh ) Replaced cable guide Replaced filters and nozzles Pre check pass Total op time: 9312 hrs Total vent time: 5744 hrs
1064206 | 192956 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-06 10:36 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/06/2023 10:35:40 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 1 hrs Total vent time: 0 hrs
1064207 | 192957 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-06 11:05 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/06/2023 11:02:34 AM, Canjimavadivel, Kamallesh ) Replaced 2x battery Replaced filters and nozzles Pre check pass Total op time: 1 hrs Total vent time: 0 hrs
1064208 | 192958 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-22 11:17 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/22/2023 11:14:36 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 6982 hrs Total vent time: 3017 hrs
1064209 | 192959 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-20 11:40 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/20/2023 11:37:30 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 7159 hrs Total vent time: 3624 hrs
1064210 | 192960 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-21 13:39 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/21/2023 01:32:46 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 2763 hrs Total vent time: 756 hrs
1064211 | 192961 | PM | CMPLT | CMPLT | 2023-02-01 02:20 | 2023-03-22 15:44 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/22/2023 03:43:17 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 4664 hrs Total vent time: 2548 hrs
1064212 | 192962 | PM | CNP | EIU | 2023-02-01 02:20 |  | PM-MAJOR - Annual PM For: Ventilators | (TC 08/24/2023 08:03:08 AM, Canjimavadivel, Kamallesh ) Attempt to PM multiple times, always found while on patient. PM will be performed when next available
| 192963 | PA | CMPLT | CMPLT | 2023-02-01 08:13 | 2023-05-15 08:09 | Giraffe LED Lights | (TC 05/15/2023 12:50:06 PM, Rhuland, Courtney) Lights received and unboxed Serial numbers: 1119636024 - date sticker added for warranty 1120145021 - date sticker added for warranty 1119636019 1119641015 - date sticker added for warranty 1119636015 - date sticker added for warranty 1119641023 1119636008 1119636021 1119636027 1119641016 1119636010 - date sticker added for warranty 1119636004 - date sticker added for warranty 1119636017 1119641020 1119636011 1118620036 1119636013 1119641014 - date sticker added for warranty 1119636020 1119638018 - date sticker added for warranty Received lights., PO#7202154905(7). We received total 20 x lights. PO#7202154903(8) and 7202157465(5)
68415 | 192964 | CM | CMPLT | CMPLT | 2023-02-01 09:10 | 2023-04-11 14:05 | Factory PM TSB 259 | (TC 02/01/2023 10:47:29 AM, Scero, Tomlyn) Shipped to Draeger using KJV (TC 04/11/2023 02:02:08 PM, Scero, Tomlyn) Returned form factory service/pm at draeger.
1046024 | 192965 | CM | CMPLT | CMPLT | 2023-02-01 09:12 | 2023-04-26 10:37 | Factory PM TSB 259 | (TC 02/01/2023 10:46:45 AM, Scero, Tomlyn) Shipped to Draeger using KJV (TC 04/26/2023 10:35:05 AM, Scero, Tomlyn) device returned from draeger. service report attached.
52181 | 192966 | PM | CMPLT | CMPLT | 2023-02-01 09:52 | 2023-02-01 09:55 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 02/01/2023 09:52:36 AM, Ferreira, Qwynn) Came down with broken wheel Replaced wheel Noticed the latch for the tilt assembly was not working - Reseated the latch and tightened Installed the Latch Recall Parts for the ports, latches, and locks. PM Passed Calibrated System voltage Calibrated Main voltage calibrated line voltage Calibrated scale Electrical Safety - OK Functional Check - OK
1063746 | 192967 | CM | CMPLT | CMPLT | 2023-02-01 10:12 | 2023-02-01 10:12 | Cleaning | (TC 02/01/2023 10:12:12 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
65940 | 192968 | CM | CMPLT | CMPLT | 2023-02-01 11:34 | 2023-02-01 11:36 | "Channel error" | (TC 02/01/2023 11:33:58 AM, Baghirzade, Nara) When connecting to PCU, error "354.6770" appears. Downloaded error log and noted that the error is a "log only". Re-seated the pressue sensor. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1053348 | 192969 | CM | CMPLT | CMPLT | 2023-02-01 13:47 | 2023-02-07 16:33 | MSH Hologic H3 - Cannot Send Images to PACs | (TC 02/01/2023 01:48:33 PM, Chau, Ken) User reported only certain images are sent over to PACs. Other images they have to send to downtime to be able to get images over. -contacted PACs for settings required to be able to send images over -user having issues only with Tomo-projection and tomo-combo -16/20 images can be sent over but 4 is blocked -verified AET (MSH_1272_IMG) -verified storage commit is disabled -PACs disable storage commit on their end -waiting for users to confirm if images can be sent to PACs without restriction (TC 02/02/2023 05:13:51 PM, Chau, Ken) Images still causing issue with tomo-projection -case created with PACs and they are looking into issue -case created with Hologic as they performed an update a month ago and where issue started to arise (TC 02/03/2023 01:51:50 PM, Chau, Ken) Request for SOP UID Class: provided dicom conformance to Daniel Reyes at JDMIPACs Contacted Hologic; Denis have editted system to be able to send approperiate studies -2D For processing -2D for presentation -3D projects (SCO) -Notices (GSPS) -3D Slices (BTO) -Generated 2D (BTO) -Generated 2D (For presentation) -I-View (For Presentation) (TC 02/07/2023 04:32:37 PM, Chau, Ken) System is operational now -Hologic made modification with image sending -copied configuration from H1 and H2 -verified with MRT that all 20/20 tomo-combo images are received on PACs | User reported only certain images are sent over to PACs. Other images they have to send to downtime to be able to get images over.
1056451 | 192970 | CM | CMPLT | CMPLT | 2023-02-01 14:22 | 2023-03-20 07:30 | Sending for replacement | (TC 02/01/2023 02:29:52 PM, Baghirzade, Nara) SC 1002873901 (TC 03/20/2023 07:28:52 AM, Baghirzade, Nara) PO 7500183430 | RIB MCU comms failure - must be sent in to Masimo for a replacement unit.
502044 | 192971 | CM | CMPLT | CMPLT | 2023-02-01 14:58 | 2023-02-01 14:58 | Part Return to Digirad | (TC 02/01/2023 02:57:53 PM, Paz, Ramon (Tino)) Arranged parts to be shipped back to Digirad. 3 PSB assemblies.
1061316 | 192972 | CM | CMPLT | CMPLT | 2023-02-01 15:21 | 2023-02-01 15:29 | Software doesn't start | (TC 02/01/2023 03:21:20 PM, Fan, Tak) IT deleted local files on Dr Jain's PC since his PC was full. This causes a dissociation between the server and client. GE can remote in and clear the server side. Jose or GE will need to adjust allocation and watermark if users have storage issues on other PCs. Faith also wanted to install echopac on laptops. Installed on one. She will possibly install on others.
1047839 | 192973 | CM | CMPLT | CMPLT | 2023-02-01 15:27 | 2023-02-01 15:29 | "pump stops running" | (TC 02/01/2023 03:27:23 PM, Cortez, John) Unable to replicate issue Completed PM successfully -Power up test passed, 7566:00 on time 3690:25 infusion time 5079000 motor revs -Force sensor check: 14.92 lbs -Small syringe size: 0.271 in Position: 0.704 in -Large syringe size: 1.254 in Position: 4.683 in -Plunger travel test passed, -Motor drive and occlusion passed, -Flow delivery accuracy: Start: 4.688 in End: 1.157 in Total travel: 3.531 in -Line leakage 15.4 uA
1047843 | 192974 | CM | CMPLT | CMPLT | 2023-02-01 15:34 | 2023-02-01 15:39 | "Had issues programming desired infusion" | (TC 02/01/2023 03:34:38 PM, Cortez, John) Nurse stated that she tried to infuse at 26ml per 30 mins, and 52 ml per hr, but was unable to. Nurse also stated that the 30 ml syringe wasn't recognized by the pump. Unable to replicate issue. PM completed successfully. Date/time corrected -Power up test passed, 7566:00 on time 3690:25 infusion time 5079000 motor revs -Force sensor check: 14.38 lbs -Small syringe size: 0.267 in Position: 0.701 in -Large syringe size: 1.253 in Position: 4.717 in -Plunger travel test passed, -Motor drive and occlusion passed, -Flow delivery accuracy: Start: 4.367 in End: 0.818 in Total travel: 3.549 in -Line leakage 18.5 uA
65918 | 192975 | CM | CMPLT | CMPLT | 2023-02-02 08:22 | 2023-02-02 08:25 | Channel error | (TC 02/02/2023 08:22:37 AM, Baghirzade, Nara) Pump was sent down for "channel error"; when connecting to the PCU, there were no errors. During check-in, the barrel clamp was failing the "closed" position; the sizer sensor assembly was replaced under recall. Barrel clamp was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1059162 | 192976 | CM | CMPLT | NFF | 2023-02-02 09:16 | 2023-02-02 09:18 | BP reading low | (TC 02/02/2023 09:16:23 AM, Reid, Emmerson ) Unit called that the BP reading was low. I brought it to biomed and tested, read 121/81 on 120/80 settings. NFF.
1054769 | 192977 | CM | CMPLT | CMPLT | 2023-02-02 13:36 | 2023-02-02 13:40 | Audit Inspection | (TC 02/02/2023 01:38:51 PM, Ferreira, Qwynn) See Procedure for details
1065159 | 192978 | CM | CMPLT | CMPLT | 2023-02-02 13:42 | 2023-02-02 13:45 | Audit Inspection | (TC 02/02/2023 01:44:18 PM, Ferreira, Qwynn) See procedure for results
1065162 | 192979 | CM | CMPLT | CMPLT | 2023-02-02 13:47 | 2023-02-02 13:49 | Audit Inspection | (TC 02/02/2023 01:48:53 PM, Ferreira, Qwynn) see procedure for details
1054774 | 192980 | CM | CMPLT | CMPLT | 2023-02-02 13:51 | 2023-02-02 13:53 | Audit Inspection | (TC 02/02/2023 01:53:04 PM, Ferreira, Qwynn) See procedure for details
1065157 | 192981 | CM | CMPLT | CMPLT | 2023-02-02 13:55 | 2023-02-02 13:59 | Audit Inspection | (TC 02/02/2023 01:58:46 PM, Ferreira, Qwynn) See procedure for details
1065158 | 192982 | CM | CMPLT | CMPLT | 2023-02-02 14:01 | 2023-02-02 14:03 | Audit Inspection | (TC 02/02/2023 02:02:48 PM, Ferreira, Qwynn) See procedure for details
| 192983 | CM | CMPLT | CMPLT | 2023-02-02 14:03 | 2023-02-02 14:03 | IVF PM | (TC 02/02/2023 02:03:08 PM, Samanian, Kamran) Exam rooms VSM and scales were done
| 192984 | CM | CMPLT | CMPLT | 2023-02-02 14:04 | 2023-02-01 14:04 | Meeting with Dr. Poorva about MIB | (TC 02/01/2023 02:04:09 PM, Samanian, Kamran) done!
1053728 | 192985 | FC | CMPLT | CMPLT | 2023-02-02 14:17 | 2023-02-02 14:18 | Verification check | (TC 02/02/2023 02:17:28 PM, Scero, Tomlyn) Functional check complete after use.
1065161 | 192986 | CM | CMPLT | CMPLT | 2023-02-02 14:22 | 2023-02-02 14:25 | Audit Inspection | (TC 02/02/2023 02:24:06 PM, Ferreira, Qwynn) See procedure for results
1054771 | 192987 | CM | CMPLT | CMPLT | 2023-02-02 14:26 | 2023-02-02 14:28 | Audit Inspection | (TC 02/02/2023 02:27:10 PM, Ferreira, Qwynn) See procedure for results
1054377 | 192988 | CM | CMPLT | CMPLT | 2023-02-02 14:31 | 2023-02-02 14:33 | Audit Inspection | (TC 02/02/2023 02:32:13 PM, Ferreira, Qwynn) see procedure for results
1065163 | 192989 | CM | CMPLT | CMPLT | 2023-02-02 14:39 | 2023-02-02 14:41 | Audit Inspection | (TC 02/02/2023 02:40:35 PM, Ferreira, Qwynn) see procedure for details
68871 | 192990 | PM | CMPLT | CMPLT | 2023-02-02 14:40 | 2023-02-02 15:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/02/2023 03:15:10 PM, Reid, Emmerson ) PM Complete see proceedure
1054768 | 192991 | CM | CMPLT | CMPLT | 2023-02-02 14:45 | 2023-02-02 14:48 | Audit Inspections | (TC 02/02/2023 02:47:54 PM, Ferreira, Qwynn) see procedure for details
1065160 | 192992 | CM | CMPLT | CMPLT | 2023-02-02 14:54 | 2023-02-02 14:57 | Audit Inspection | (TC 02/02/2023 02:56:03 PM, Ferreira, Qwynn) See procedure for details
1054775 | 192993 | CM | CMPLT | CMPLT | 2023-02-02 15:01 | 2023-02-02 15:05 | Audit Inspection | (TC 02/02/2023 03:04:24 PM, Ferreira, Qwynn) See procedure for results
68670 | 192994 | PM | CMPLT | CMPLT | 2023-02-02 15:01 | 2023-02-02 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/02/2023 03:00:52 PM, Reid, Emmerson ) PM Complete
26152 | 192995 | CM | CMPLT | CMPLT | 2023-02-02 15:51 | 2023-02-02 15:53 | Oxygen filter broke off | (TC 02/02/2023 03:50:47 PM, Ferreira, Qwynn) Removed the broken brass fitting from the filter Removed the broken brass fitting from the blender Dissambled the blender to remove the filter mount Added new brass fitting with white sealing tape Replaced the filter assembly and tested the blender.
32520 | 192996 | CM | CMPLT | CMPLT | 2023-02-03 09:40 | 2023-03-22 14:56 | MSH DR2 BETA TUBE SWIVEL NOT HOLDING | (TC 02/03/2023 09:40:39 AM, Whittington, Scott) MSH DR2 BETA TUBE SWIVEL NOT HOLDING (TC 02/02/2023 09:40:39 AM, Whittington, Scott) MSH DR2 BETA TUBE SWIVEL NOT HOLDING (TC 03/14/2023 07:55:00 PM, Whittington, Scott) MSH DR2 BETA TUBE SWIVEL REORTED NOT HOLDING | MSH DR2 BETA TUBE SWIVEL NOT HOLDING
1065602 | 192997 | CM | CMPLT | CMPLT | 2023-02-03 09:45 | 2023-02-03 10:01 | oncoming inspection | loaner unit incoming inspection. 0.024 ohm. leakage 91.5 uA
1065603 | 192998 | CM | CMPLT | CMPLT | 2023-02-03 09:57 | 2023-02-03 09:57 | incoming inspection | incoming inspection. 0.076 ohm. leakage 104.3 uA
1063182 | 192999 | CM | CMPLT | CMPLT | 2023-02-03 10:18 | 2023-02-10 08:16 | 1 Year, MR1 PM kit for two MR1 Ventilator | (TC 02/03/2023 10:25:09 AM, Chukri, Samir) Ordered two PM kits for MR1 ventilator
66431 | 193000 | CM | CMPLT | CMPLT | 2023-02-03 11:48 | 2023-02-03 12:07 | turning on and off | (TC 02/03/2023 12:05:29 PM, Scero, Tomlyn) replaced a defective/worn out battery adaptor. ran device on the simulator for a while to be sure the connection wasn't lost and unit remained on.
66462 | 193001 | CM | CMPLT | CMPLT | 2023-02-03 11:59 | 2023-02-24 12:07 | ordering parts for VN500 vents | received parts for VN500 vents. PO#7202156702
1055799 | 193002 | CM | CMPLT | CMPLT | 2023-02-03 12:08 | 2023-02-03 14:40 | squished under strecher during transport. | (TC 02/03/2023 02:31:35 PM, Scero, Tomlyn) Nurse in CCU asked me to check the defib after it was accidently squished under a strecher during transport. Everything was checked out and looks ok. 30J test complete and no visible signs of damage were present. Returned the device to CCU
| 193003 | CM | CMPLT | CMPLT | 2023-02-03 12:46 | 2023-02-10 08:17 | ordered 8X 1Year PM kit for vent, Hamilton-T1 | (TC 02/03/2023 12:51:04 PM, Chukri, Samir) Ordered 8 1 Year PM kit for Hamilton -T1 ventilator
1056576 | 193004 | CM | CMPLT | CMPLT | 2023-02-03 13:24 | 2023-02-03 13:25 | Battery Replacement | (TC 02/03/2023 01:24:16 PM, Reid, Emmerson ) Replaced the battery
| 193005 | CM | CMPLT | CMPLT | 2023-02-03 15:49 | 2023-05-12 16:59 | Masimo Roots- temperature reading issues on 10S | (TC 01/25/2023 03:49:34 PM, Cortez, John) Informed by Barb (PCM) from 10S of issues that staff were experiencing when using the Masimo Root monitors: -Wrong cuffs on BP monitors, -Various difficulties with getting a temperature reading ("probe off", long wait times, set to continuous monitoring rather than spot check mode) Barb asked for an audit on all of the monitors. Checked 14 monitors on the first day (TC 01/26/2023 03:58:39 PM, Cortez, John) Checked 12 monitors (TC 02/02/2023 03:59:07 PM, Cortez, John) Joined 10S huddle to discuss findings and gain feedback from nurses (TC 02/07/2023 10:14:20 AM, Cortez, John) Checked remaining monitors, joined 10S huddle
1054549 | 193006 | CM | CMPLT | CMPLT | 2023-02-06 08:58 | 2023-02-02 11:00 | Keyboard cover damaged | (TC 02/02/2023 11:00:00 AM, Baghirzade, Nara) Keyboard cover had rips in it; cover was replaced.
1056550 | 193007 | CM | CMPLT | CMPLT | 2023-02-06 10:02 | 2023-02-06 10:02 | battery alarm | replaced battery. tested. functional.
502088 | 193008 | PM | CMPLT | CMPLT | 2023-02-06 10:20 | 2023-06-15 19:52 | PM - Semi-Annual | (TC 06/15/2023 07:50:41 PM, Paz, Ramon (Tino)) Scheduled PM 3 performed as per OEM intructions. Performed CT Constancy as well
502089 | 193009 | PM | CMPLT | CMPLT | 2023-02-06 10:24 | 2023-06-27 14:53 | PM - Semi-Annual | (TC 05/23/2023 12:46:35 PM, Paz, Ramon (Tino)) PM rescheduled due to lack of personnel. (TC 06/27/2023 02:51:56 PM, Green, Paul) PM3 performed as per CBDOCS.
66472 | 193010 | CM | CMPLT | CMPLT | 2023-02-06 11:20 | 2023-06-06 08:25 | Ordering PS500 batteries | received batteries. PO#7202157303
1061333 | 193011 | FC | CMPLT | CMPLT | 2023-02-06 12:48 | 2023-02-06 12:51 | QC Checklist | (TC 02/06/2023 12:47:42 PM, Baghirzade, Nara) Sarabeth mentioned that the device was alarming for blockage even though there was no blockage. Ran the QC checklist and all tests passed. Did not encounter any error or alarm. Put it on therapy mode and ran the device for a few minutes without any errors. Did another QC checklist and all passed. Dropped it off with Sarabeth and let her know that I could not find or replicate the error. She will keep an extra vac in case this causes a problem again. Device was recently PM'ed so if error occurs again, then will be sending the device to 3M.
501105 | 193012 | CM | CMPLT | CMPLT | 2023-02-06 13:58 | 2023-03-08 12:32 | MSH 5 US IU22 ROAM POWERS DOWN SPONTANEOUSLY FIVE TIMES IN NICU | (TC 02/06/2023 01:59:04 PM, Whittington, Scott) MSH 5 US IU22 ROAM POWERS DOWN SPONTANEOUSLY FIVE TIMES IN NICU (TC 02/06/2023 01:59:04 PM, Whittington, Scott) MSH 5 US IU22 ROAM POWERS DOWN SPONTANEOUSLY FIVE TIMES IN NICU GLEAN LOGS, SEND TO CODY (TC 02/28/2023 11:53:17 AM, Whittington, Scott) MSH 5 US IU22 ROAM POWERS DOWN SPONTANEOUSLY FIVE TIMES IN NICU CODY HAS CLOSED THE PHILIPS CASE SYSTEM STOPPED FUNCTIONING AND WENT INTO SERVICE LOGIN SCREEN REBOOT AND FUNCTIONING AGAIN (TC 03/08/2023 12:31:13 PM, Whittington, Scott) MSH 5 US IU22 ROAM POWERS DOWN SPONTANEOUSLY FIVE TIMES IN NICU CODY RELOADED SW RETURNED TO CLINICAL USE AS BACKUP TO TWH SYSTEM | MSH 5 US IU22 ROAM POWERS DOWN SPONTANEOUSLY FIVE TIMES IN NICU SITE 553760
500868 | 193013 | CM | CMPLT | CMPLT | 2023-02-06 14:27 | 2023-02-06 14:27 | msh 5 us apli xg srvc lic & dfrg | (TC 02/06/2023 02:26:57 PM, Whittington, Scott) msh 5 us apli xg srvc lic & dfrg | msh 5 us apli xg srvc lic & dfrg
1046477 | 193014 | CM | CMPLT | CMPLT | 2023-02-06 15:11 | 2023-02-06 15:12 | No synced | (TC 02/06/2023 03:11:06 PM, Reid, Emmerson ) I ran the sync proceedure for the smoke evac. Working now
501260 | 193015 | CM | CMPLT | CMPLT | 2023-02-07 07:18 | 2023-02-07 07:25 | Not staying on | (TC 02/06/2023 11:24:00 AM, Fan, Tak) Checked system, not charging. Shuts off shortly. Checked PS, no voltage on any pins. Reconnected power conenctors and reset battery. System charging now. Tested, OK.
| 193016 | CM | CMPLT | CMPLT | 2023-02-07 07:42 | 2023-02-06 07:42 | ICU for ASCOM training | (TC 02/06/2023 07:42:34 AM, Samanian, Kamran) done!
62618 | 193017 | CM | CMPLT | CMPLT | 2023-02-07 08:03 | 2023-02-07 08:09 | Door Broken and scale not working | (TC 02/07/2023 08:06:18 AM, Ferreira, Qwynn) replaced the main door latch tested the scale and no icon was displaying on the screen. Notcied the brass pin at the back of the scale connector was pushed back. Reseated the brass pin and tested the scale - OK Functional Test - OK
69515 | 193018 | CM | CMPLT | CMPLT | 2023-02-07 08:42 | 2023-02-07 08:43 | Functional check/check-in | (TC 02/07/2023 08:42:52 AM, Baghirzade, Nara) Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
69211 | 193019 | PM | CMPLT | CMPLT | 2023-02-07 08:46 | 2023-02-07 08:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/07/2023 08:46:09 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68882 | 193020 | PM | CMPLT | CMPLT | 2023-02-07 08:49 | 2023-02-07 08:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/07/2023 08:55:19 AM, Baghirzade, Nara) PM completed. Battery replaced. See procedure for results.
1065075 | 193021 | FC | CMPLT | CMPLT | 2023-02-07 09:27 | 2023-02-07 09:29 | Functional check/check-in | (TC 02/07/2023 09:29:08 AM, Baghirzade, Nara) Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
68839 | 193022 | PM | CMPLT | CMPLT | 2023-02-07 09:48 | 2023-02-07 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/07/2023 09:53:46 AM, Baghirzade, Nara) PM completed. Battery replaced. See procedure for results.
68172 | 193023 | CM | CMPLT | CMPLT | 2023-02-07 09:51 | 2023-02-08 14:18 | unit caught fire | (TC 02/08/2023 02:17:02 PM, Scero, Tomlyn) unit completely destroyed and melted. looks like water got into the unit and fried the PCB. threw it out.
62702 | 193024 | CM | CMPLT | CMPLT | 2023-02-07 09:52 | 2023-02-07 10:23 | broken limb leads LA and RL | (TC 02/07/2023 10:21:32 AM, Scero, Tomlyn) RL and LA limb leads broken. Replaced and unit tested ok. Returned device to CCU
501107 | 193025 | CM | CMPLT | CMPLT | 2023-02-07 10:01 | 2023-02-08 11:30 | Error 511 on startup | (TC 02/07/2023 10:01:37 AM, Fan, Tak) Checked system. Can't clear with reset. FSE will be paged for either SW reload or registry editing. (TC 02/07/2023 01:58:43 PM, Fan, Tak) Jillian will be in tomorrow. Retrieved logs for her. Cody came in instead.
1053429 | 193026 | CM | CMPLT | CMPLT | 2023-02-07 10:17 | 2023-02-22 13:18 | not detected | (TC 02/07/2023 10:16:54 AM, Rendon Lopez , Elizabeth) Confirmed to be sent out for repair. (TC 02/21/2023 01:17:23 PM, Rendon Lopez , Elizabeth) Notified broken beyond repair. Retire from service
1055668 | 193027 | CM | CMPLT | CMPLT | 2023-02-07 10:18 | 2023-02-24 09:07 | not detected | (TC 02/07/2023 10:19:12 AM, Rendon Lopez , Elizabeth) Confirmed to be sent out for repair (TC 02/22/2023 09:07:14 AM, Rendon Lopez , Elizabeth) Received, incoming inspection-OK
1047075 | 193028 | CM | CMPLT | CMPLT | 2023-02-07 10:20 | 2023-02-24 09:07 | Doesn't pick up FHR | (TC 02/07/2023 10:20:21 AM, Rendon Lopez , Elizabeth) Confirm issue. Repair under warranty, sent out (TC 02/22/2023 09:06:37 AM, Rendon Lopez , Elizabeth) Received, incoming inspection-OK
1057134 | 193029 | CM | CMPLT | CMPLT | 2023-02-07 10:22 | 2023-02-24 09:06 | Piece inside, noisy | (TC 02/07/2023 10:22:03 AM, Rendon Lopez , Elizabeth) Confirmed. Repair under warranty, sent out (TC 02/22/2023 09:05:49 AM, Rendon Lopez , Elizabeth) Received, incoming inspection-OK
1065604 | 193030 | ACC | CMPLT | CMPLT | 2023-02-07 10:41 | 2023-02-07 10:42 | Incoming inspection | (TC 02/06/2023 10:41:33 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065605 | 193031 | ACC | CMPLT | CMPLT | 2023-02-07 10:42 | 2023-02-07 10:42 | Incoming inspection | (TC 02/06/2023 10:42:02 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065606 | 193032 | ACC | CMPLT | CMPLT | 2023-02-07 10:42 | 2023-02-07 10:43 | Incoming inspection | (TC 02/06/2023 10:42:31 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065607 | 193033 | CM | CMPLT | CMPLT | 2023-02-07 10:43 | 2023-02-07 10:43 | Incoming inspection | (TC 02/06/2023 10:43:03 AM, Rendon Lopez , Elizabeth) Incoming inspection
1065608 | 193034 | ACC | CMPLT | CMPLT | 2023-02-07 10:44 | 2023-02-07 10:44 | Incoming inspection | (TC 02/06/2023 10:43:40 AM, Rendon Lopez , Elizabeth) Incoming inspection
69011 | 193035 | CM | CMPLT | CMPLT | 2023-02-07 11:08 | 2023-02-07 11:10 | Functional test/checkin | (TC 02/07/2023 11:09:17 AM, Baghirzade, Nara) Check-in tests passed and completed. Refer to procedures for results.
| 193036 | CM | CMPLT | CMPLT | 2023-02-07 11:42 | 2023-02-07 11:44 | Functional check | (TC 02/07/2023 11:42:31 AM, Baghirzade, Nara) Received thermometer from the senior wellness clinic in Scarborough. The rechargeable battery is defective; Kamran will be ordering new batteries and will drop off the thermometer when ready. SN: 00328785000252
1059778 | 193037 | FC | CMPLT | CMPLT | 2023-02-07 11:57 | 2023-02-07 11:57 | QC checklist | (TC 02/07/2023 11:57:05 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 S
1059828 | 193038 | CM | CMPLT | CMPLT | 2023-02-07 11:58 | 2023-02-06 15:00 | Sent to biomed without indicating the issue | (TC 02/06/2023 03:00:00 PM, Baghirzade, Nara) In service mode, noted the error logs; most errors were pertaining to user interface. Ran the altrix for more than an hour at different temperatures (15, 25, and 35). No errors occurred and device is fully functional.
68781 | 193039 | PM | CMPLT | CMPLT | 2023-02-07 12:58 | 2023-02-08 10:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/07/2023 12:59:45 PM, Baghirzade, Nara) PM passed and completed. Power cord was replaced due to damaged prongs. Electrical Safety - Ground resistance: 67 mohms Leakage current: 89.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 02/08/2023 10:13:58 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4095 mamphr PM completed.
1051031 | 193040 | CM | CMPLT | CMPLT | 2023-02-07 14:11 | 2023-02-08 11:29 | Faith would like to disable sending for research studies. | (TC 02/07/2023 02:11:27 PM, Fan, Tak) Faith would like to disable wifi and network to do research studies at sickkids. Might set up a "Offline Research" dataflow rather than have users enable and disable wireless settings. (TC 02/08/2023 11:27:07 AM, Fan, Tak) Tested options: remove wireless adapter, disable wireless, offline dataflow. Setup "Local Archive - Offline (research)" dataflow, which doesn't send to viewpoint.
| 193041 | CM | CMPLT | CMPLT | 2023-02-07 14:48 | 2023-02-07 14:48 | rebuilding MSHBIOMED1S | (TC 02/07/2023 02:48:08 PM, Samanian, Kamran) done!
66106 | 193042 | CM | CMPLT | CMPLT | 2023-02-07 14:57 | 2023-02-07 14:58 | QC checklist | (TC 02/07/2023 02:57:40 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 North
| 193043 | CM | CMPLT | CMPLT | 2023-02-08 07:46 | 2023-02-10 08:51 | Batteries for testing scales | Received, for BIomed
1056455 | 193044 | CM | CMPLT | CMPLT | 2023-02-08 09:15 | 2023-03-28 10:51 | Not charging. no power | Received new unit with sn.2000160183. tested. Functional
69008 | 193045 | CM | CMPLT | CMPLT | 2023-02-08 09:22 | 2023-02-08 09:24 | Power cord prongs broken | (TC 02/08/2023 09:24:05 AM, Baghirzade, Nara) Replaced power cord. Check-in tests passed and completed.
1065612 | 193046 | ACC | CMPLT | CMPLT | 2023-02-08 09:47 | 2023-02-09 10:00 | Incoming Inspection | (TC 02/09/2023 09:59:18 AM, Reid, Emmerson ) Incoming complete, electrical safety in the proceedure.
62703 | 193047 | CM | CMPLT | NFF | 2023-02-08 10:26 | 2023-02-08 10:29 | Reported no leads picking up | (TC 02/08/2023 10:26:21 AM, Canjimavadivel, Kamallesh ) Tested with simulator, all leads pick up signal. Replaced chest lead set as it was affected by recall. | Sent down with printout showing no leads
65381 | 193048 | CM | CMPLT | CMPLT | 2023-02-08 10:29 | 2023-02-08 10:30 | QC checklist | (TC 02/08/2023 10:29:37 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
501071 | 193049 | CM | CMPLT | CMPLT | 2023-02-08 10:41 | 2023-02-13 10:31 | MSH CEOU RM#12 - System Shut off, System Top and Bottom Monitor Blank | (TC 02/08/2023 10:41:24 AM, Chau, Ken) System shut off by itself once but it didnt shut down by itself again -top monitor and bottom monitor black -case#0121175422 for a UMB replacement (TC 02/13/2023 10:15:34 AM, Fan, Tak) Replaced UMB. Submitted MAC address change. Tested network.
1065613 | 193050 | CM | CMPLT | CMPLT | 2023-02-08 10:44 | 2023-02-08 10:58 | Incoming inspection | (TC 02/08/2023 10:44:31 AM, Canjimavadivel, Kamallesh ) Pump was sent down with issue but no asset tag. No fault found. Added to database and asset tagged. All functions good
69289 | 193051 | FC | CMPLT | NFF | 2023-02-08 10:46 | 2023-02-08 10:47 | Sent with no tag; functional check | (TC 02/08/2023 10:46:47 AM, Baghirzade, Nara) Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
| 193052 | PA | CMPLT | CMPLT | 2023-02-08 10:59 | 2023-02-14 13:15 | Ordering Masimo SpO2 cables and sensors for biomed stock | (TC 02/08/2023 11:30:59 AM, Baghirzade, Nara) SC 1002879310 (TC 02/14/2023 01:14:31 PM, Baghirzade, Nara) PO 7202159248
66099 | 193053 | CM | CMPLT | CMPLT | 2023-02-08 11:20 | 2023-02-08 11:21 | "battery not working, flashes red" | (TC 02/08/2023 11:19:43 AM, Cortez, John) Unable to replicate issue. Charged battery overnight. Possibly an issue with the outlet? Port A tube has damaged connector. Unit to replace. Functional test ok
500858 | 193054 | CM | CMPLT | CMPLT | 2023-02-08 12:50 | 2023-02-08 12:50 | msh 5 us apli xg srvc lic & dfrg | (TC 02/08/2023 12:50:19 PM, Whittington, Scott) msh 5 us apli xg srvc lic & dfrg | msh 5 us apli xg srvc lic & dfrg
63876 | 193055 | CM | CMPLT | CMPLT | 2023-02-08 12:52 | 2023-02-08 12:57 | System error 25 and reading in Fahrenheit | (TC 02/08/2023 12:54:37 PM, Ferreira, Qwynn) System error 25 was displayed Replaced the fan assembly - OK Tested both temperature settings and both were displayed in Celcius not Fahrenheit. Replaced latches with recall parts. Repaired the hood with UV10Med resin Functional Test - OK Electrical Safety - OK
500857 | 193056 | CM | CMPLT | CMPLT | 2023-02-08 12:53 | 2023-02-08 12:53 | msh 5 us apli xg srvc lic & dfrg | (TC 02/08/2023 12:52:35 PM, Whittington, Scott) msh 5 us apli xg srvc lic & dfrg | msh 5 us apli xg srvc lic & dfrg
65930 | 193057 | CM | CMPLT | CMPLT | 2023-02-08 13:24 | 2023-02-09 12:53 | Channel error; calibration check | (TC 02/09/2023 12:51:31 PM, Baghirzade, Nara) Performed full calibration. Check-in tests passed. Ran a basic infusion on various rates; no errors.
1061333 | 193058 | FC | CMPLT | CMPLT | 2023-02-08 13:33 | 2023-02-08 13:42 | QC checklist | (TC 02/08/2023 01:41:02 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1059550 | 193059 | CM | CMPLT | CMPLT | 2023-02-08 13:48 | 2023-02-08 13:58 | Reconfigured for NICU | (TC 02/08/2023 01:48:01 PM, Ferreira, Qwynn) Worked with Kamran This was being used for the biomed shop Renamed the pc from biomed to nicu overview 2 Tested to make sure it was working - OK Added the 8 displays for the rooms in zone A (TC 02/08/2023 01:57:13 PM, Samanian, Kamran) Worked with Qwynn to change the Biomed PC to NICUO2
1065165 | 193060 | CM | CMPLT | CMPLT | 2023-02-08 14:00 | 2023-02-08 14:02 | Audit Inspection | (TC 02/08/2023 02:01:49 PM, Ferreira, Qwynn) Functional Test for Audit and inspection. See procedure for results.
1062151 | 193061 | CM | CMPLT | CMPLT | 2023-02-08 14:04 | 2023-02-08 14:37 | Not working | (TC 02/08/2023 02:36:04 PM, Canjimavadivel, Kamallesh ) Cleaned earwax from lense. Passed calibration after one attempt.
68577 | 193062 | CM | CMPLT | CMPLT | 2023-02-08 14:04 | 2023-02-08 14:05 | Audit Inspection | (TC 02/08/2023 02:04:57 PM, Ferreira, Qwynn) See procedure for Results.
1065164 | 193063 | CM | CMPLT | CMPLT | 2023-02-08 14:11 | 2023-02-08 14:12 | Audit Inspection | (TC 02/08/2023 02:11:50 PM, Ferreira, Qwynn) See procedure for results
1057538 | 193064 | CM | CMPLT | CMPLT | 2023-02-08 14:12 | 2023-02-08 14:14 | broken bezel | (TC 02/08/2023 02:13:28 PM, Scero, Tomlyn) went to OR RT and replaced broken bezel
1065517 | 193065 | CM | CMPLT | CMPLT | 2023-02-08 14:13 | 2023-02-08 14:17 | Audit Inspection | (TC 02/08/2023 02:16:12 PM, Ferreira, Qwynn) see procedure for details
1057559 | 193066 | CM | CMPLT | CMPLT | 2023-02-08 14:15 | 2023-02-08 14:16 | missing/broken bezel | (TC 02/08/2023 02:15:42 PM, Scero, Tomlyn) replaced missing/broken bezel
1065518 | 193067 | CM | CMPLT | CMPLT | 2023-02-08 14:18 | 2023-02-08 14:19 | Audit Inspection | (TC 02/08/2023 02:19:05 PM, Ferreira, Qwynn) see procedure for details
1062149 | 193068 | CM | CMPLT | CMPLT | 2023-02-08 14:48 | 2023-02-08 15:18 | Reading low temp result | (TC 02/08/2023 03:12:13 PM, Canjimavadivel, Kamallesh ) Performed calibration, passed after one attempt
1059544 | 193069 | CM | CMPLT | CMPLT | 2023-02-08 15:25 | 2023-02-08 15:27 | Re-install philips software | (TC 02/08/2023 03:25:19 PM, Ferreira, Qwynn) Worked with Kamran, re-installed the philips software onto this PC Came from NICU and would not allow the philips application to open - red error message Renamed the PC and configured it for the Biomed shop and added the IP. Let run for a day in the Biomed shop with demo mode and no interruptions or unexpected errors. Functional.
67433 | 193070 | CM | CMPLT | CMPLT | 2023-02-08 17:57 | 2023-02-16 08:37 | Not powering on | (TC 02/08/2023 06:02:03 PM, Ferreira, Qwynn) Replaced the batteries as the old ones were dead. Tested the scale and it was able to turn on, but when bumped or moved it would shut off. Saw the terminals in the battery compartment were cracked causing an interruption of the power. Waiting to see solutions. (TC 02/16/2023 08:30:08 AM, Ferreira, Qwynn) Cut the rivets that were holding the C cell holder in place Soldered the new C cell holder to the new connectors Added double sided tape to hold the holder down Assembled the board and attached the scale display back Tested the scale and everything was working Tested with 5kg weight = 5kg
67948 | 193071 | CM | CMPLT | CMPLT | 2023-02-08 18:35 | 2023-02-15 14:31 | Missing inside door, missing rubber, broken latch | (TC 02/08/2023 06:35:28 PM, Ferreira, Qwynn) Hood was broken on top of the missing door, broken latch, and missing grommet. Removed old hood and components Replaced the hood with new hood. Work in progress. (TC 02/15/2023 02:29:21 PM, Ferreira, Qwynn) Removed old parts and added the recall parts. Noted the serial number for records. Tested the scale as there was some noted corrosion but was cleaned off. Functional Check - OK
502044 | 193072 | CM | CMPLT | CMPLT | 2023-02-08 20:02 | 2023-05-16 08:07 | Arm Cable Broken | (TC 02/08/2023 08:03:23 PM, Paz, Ramon (Tino)) Brake cable broken. Replaced with a new one. Upon replacement, we noticed that the Extended Arm assembly was loose by the carriage. We attempted to tighten the two long bolts that attaches the rest of the arm. Attempted to remove the bolts and the brake aseembly to access the bolts. We could not tighten it. We decided to reassemble everything.. Also tightened the brake handle assembly. UPON POWER UP, module 32 was not working. Opened Head and resested #32 and also added the missing PEM standoffs. Spoke to DIGIRAD and they are sending us retrofit kits.
| 193073 | CM | CMPLT | CMPLT | 2023-02-09 07:54 | 2023-02-08 07:54 | Working on CESO presentation | (TC 02/08/2023 07:54:03 AM, Samanian, Kamran) prepared.
| 193074 | PA | CMPLT | CMPLT | 2023-02-09 09:08 | 2023-02-15 14:42 | Formlabs order
1056548 | 193075 | CM | CMPLT | CMPLT | 2023-02-09 09:41 | 2023-02-09 09:41 | Battery replacement | (TC 02/09/2023 09:41:08 AM, Canjimavadivel, Kamallesh ) Replaced battery
1054537 | 193076 | CM | CMPLT | CMPLT | 2023-02-09 10:13 | 2023-02-09 10:15 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/08/2023 10:13:59 AM, Cortez, John) Completed cleaning cycle
1063000 | 193077 | CM | CMPLT | CMPLT | 2023-02-09 10:15 | 2023-02-09 10:16 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/08/2023 10:15:40 AM, Cortez, John) Completed cleaning cycle
1054679 | 193078 | CM | CMPLT | CMPLT | 2023-02-09 10:45 | 2023-02-09 10:46 | "broken plug" | (TC 01/20/2023 10:44:59 AM, Cortez, John) Replaced missing plugs on power adapter
1055906 | 193079 | CM | CMPLT | CMPLT | 2023-02-09 10:53 | 2023-02-09 10:54 | "doesn't take temperature" | (TC 01/31/2023 10:53:10 AM, Cortez, John) Swapped probe with spare
65367 | 193080 | CM | CMPLT | CMPLT | 2023-02-09 11:03 | 2023-02-09 11:05 | "BP cuff leaking air" | (TC 01/26/2023 11:03:35 AM, Cortez, John) Flexiport connector on hose damaged. Replaced hose
1054543 | 193081 | CM | CMPLT | CMPLT | 2023-02-09 11:11 | 2023-02-09 11:16 | LA lead damaged | (TC 02/03/2023 11:15:20 AM, Cortez, John) Replaced LA limb lead
| 193082 | PA | CMPLT | CMPLT | 2023-02-09 11:12 | 2023-02-15 13:06 | Ordering PCU batteries and sensors | (TC 02/09/2023 11:32:55 AM, Baghirzade, Nara) SC 1002880421 (TC 02/15/2023 01:05:07 PM, Baghirzade, Nara) PO 7202160088
68923 | 193083 | PM | CMPLT | CMPLT | 2023-02-09 11:58 | 2023-02-09 12:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/09/2023 11:58:09 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 69 mohms Leakage current: 78.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68899 | 193084 | PM | CMPLT | CMPLT | 2023-02-09 12:14 | 2023-02-09 12:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/09/2023 12:14:02 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 13 mohms Leakage current: 67.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
1051897 | 193085 | CM | CMPLT | CMPLT | 2023-02-09 12:57 | 2023-02-09 12:59 | Parents reported to nurse that pump is broken | (TC 02/09/2023 12:59:20 PM, Baghirzade, Nara) No physical damaged inspected. Ran the breast pump for a full day at various suction levels and speeds; no errors or malfunctions. At 45 mins, pump times out and shut off; passed. No fault found with device. Functional test results: @1% = 28mmHg @32 = 100 mmHg @57% = 158 mmHg @77% = 204 mmHg @100% = 257 mmHg Sent back to NICU with attention to Tammy and Erin (lactation specialists)
1054781 | 193086 | CM | CMPLT | CMPLT | 2023-02-09 13:15 | 2023-02-09 13:17 | Audit Inspection | (TC 02/09/2023 01:17:05 PM, Ferreira, Qwynn) see procedure for results
| 193087 | PA | CMPLT | CMPLT | 2023-02-09 14:43 | 2023-02-15 08:20 | Digi Key Order
| 193088 | CM | CMPLT | CMPLT | 2023-02-09 15:05 | 2023-02-09 15:05 | Meeting with Andrew re: ophthalmology lasers | (TC 02/09/2023 03:04:49 PM, Baghirzade, Nara) Meeting with Andrew re: ophthalmology lasers
| 193089 | CM | CMPLT | CMPLT | 2023-02-09 15:05 | 2023-02-06 14:00 | Monday Huddle | (TC 02/06/2023 02:00:00 PM, Baghirzade, Nara) Monday biomed huddle
| 193090 | CM | CMPLT | CMPLT | 2023-02-09 15:06 | 2023-02-07 15:30 | NICU huddle | (TC 02/07/2023 03:30:00 PM, Baghirzade, Nara) NICU huddle
| 193091 | CM | CMPLT | CMPLT | 2023-02-10 08:01 | 2023-02-21 13:19 | Zoll input cable | (TC 02/21/2023 01:17:47 PM, Scero, Tomlyn) Arrived. Given to Cath lab and installed on GE system
69345 | 193092 | FC | CMPLT | CMPLT | 2023-02-10 10:18 | 2023-02-10 10:19 | Check-in/functional | (TC 02/10/2023 10:18:32 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
1051880 | 193093 | CM | CMPLT | CMPLT | 2023-02-10 10:19 | 2023-02-10 10:25 | Power cord connector was missing screws | (TC 02/10/2023 10:23:40 AM, Baghirzade, Nara) Replaced the screws and secured the power cord connector to the case. While the device was open, i was able to find the one missing screw inside the case. Functional test results: @ 1% = 30 mmHg @ 32% = 104 mmHg @ 57% = 157 mmHg @ 77% = 200 mmHg @ 100% = 245 mmHg
| 193094 | CM | CMPLT | CMPLT | 2023-02-10 10:30 | 2023-02-10 10:30 | Incubators meeting | (TC 02/10/2023 10:29:48 AM, Baghirzade, Nara) Incubators meeting with Adeel, Olcay, and Qwynn.
1055917 | 193095 | CM | CMPLT | CMPLT | 2023-02-10 11:49 | 2023-02-10 11:51 | Temperature module not working | (TC 02/10/2023 11:49:41 AM, Baghirzade, Nara) Replaced temperature well and probe; fully functional. Checked SpO2 and BP; all functional.
1053348 | 193096 | CM | CMPLT | CMPLT | 2023-02-10 12:37 | 2023-02-27 14:49 | Console table movement buttons broken | (TC 02/10/2023 12:37:03 PM, Fan, Tak) Reassembled and taped together. Contacted Jon for replacement.
1049905 | 193097 | CM | CMPLT | CMPLT | 2023-02-10 12:39 | 2023-02-14 10:21 | "Drives too fast" | (TC 02/10/2023 12:38:53 PM, Fan, Tak) Right strain gauge way off (1000 range). Tested and recalibrated. Tested range, back is OK but forward right goes to 200 when forward right goes to 20. Drives OK, will replace strain gauges and recal. (TC 02/14/2023 10:20:09 AM, Fan, Tak) Strain gauge values were off again. Replaced strain gauges. Calibrated.
1061435 | 193098 | CM | CMPLT | CMPLT | 2023-02-10 13:41 | 2023-02-10 13:43 | Bed controls not responsive | (TC 02/10/2023 01:41:34 PM, Baghirzade, Nara) Bed controls were not working, both on user and patient side. Reset the bed and power cycled by turning off the battery and unplugging. Turned battery back on and plugged bed to power; all bed controls (both patient and user side) are functional and all functions of the bed have been restored.
1062389 | 193099 | CM | CMPLT | CMPLT | 2023-02-10 13:43 | 2023-02-10 13:43 | Bed controls not working | (TC 02/10/2023 01:43:13 PM, Baghirzade, Nara) Bed controls were not working, both on user and patient side. Reset the bed and power cycled by turning off the battery and unplugging. Turned battery back on and plugged bed to power; all bed controls (both patient and user side) are functional and all functions of the bed have been restored.
1059778 | 193100 | FC | CMPLT | CMPLT | 2023-02-10 14:37 | 2023-02-10 14:39 | QC Checklist | (TC 02/10/2023 02:38:20 PM, Baghirzade, Nara) QC Checklist completed. Returned to 11S
| 193101 | CM | CMPLT | CMPLT | 2023-02-10 15:14 | 2023-02-10 07:50 | CESO-Powershell presentation | (TC 02/10/2023 07:49:44 AM, Samanian, Kamran) presented
1047230 | 193102 | CM | CMPLT | CMPLT | 2023-02-13 10:04 | 2023-02-13 10:05 | Not accurate | (TC 02/10/2023 10:04:07 AM, Rendon Lopez , Elizabeth) Cleaned contacts, tested-OK
67088 | 193103 | CM | CMPLT | CMPLT | 2023-02-13 11:40 | 2023-02-13 11:40 | BP not working | (TC 02/13/2023 11:39:42 AM, Rendon Lopez , Elizabeth) BP set to neonatal. Changed to adult, tested all functions-OK
1064351 | 193104 | CM | CMPLT | CMPLT | 2023-02-13 13:05 | 2023-02-13 13:05 | MSH ICU SPARQ ULTRASOUND NEW SYSTEMS STORED IN NEW ICU ROOM 5-299 | (TC 02/13/2023 01:05:32 PM, Whittington, Scott) MSH ICU SPARQ ULTRASOUND NEW SYSTEMS STORED IN NEW ICU ROOM 5-299 | MSH ICU SPARQ ULTRASOUND NEW SYSTEMS STORED IN NEW ICU ROOM 5-299
1064352 | 193105 | CM | CMPLT | CMPLT | 2023-02-13 13:06 | 2023-02-13 13:06 | MSH ICU SPARQ ULTRASOUND NEW SYSTEMS STORED IN NEW ICU ROOM 5-299 | (TC 02/13/2023 01:06:29 PM, Whittington, Scott) MSH ICU SPARQ ULTRASOUND NEW SYSTEMS STORED IN NEW ICU ROOM 5-299 | MSH ICU SPARQ ULTRASOUND NEW SYSTEMS STORED IN NEW ICU ROOM 5-299
500858 | 193106 | CM | CMPLT | CMPLT | 2023-02-13 13:39 | 2023-02-13 13:40 | MSH 5 US RM 7 SYSTEM FROZE UP ON CARLOS LAST WEEK | (TC 02/13/2023 01:39:34 PM, Whittington, Scott) MSH 5 US RM 7 SYSTEM FROZE UP ON CARLOS LAST WEEK OPERATING CORRECTLY | MSH 5 US RM 7 SYSTEM FROZE UP ON CARLOS LAST WEEK
500847 | 193107 | CM | CMPLT | CMPLT | 2023-02-13 13:41 | 2023-02-13 13:41 | MSH 5 US RM 1 SYSTEM FROZE UP ON CARLOS LAST WEEK | (TC 02/13/2023 01:41:08 PM, Whittington, Scott) MSH 5 US RM 1 SYSTEM FROZE UP ON CARLOS LAST WEEK | MSH 5 US RM 1 SYSTEM FROZE UP ON CARLOS LAST WEEK
1055799 | 193108 | CM | CMPLT | CMPLT | 2023-02-13 13:50 | 2023-02-13 13:55 | battery failure | (TC 02/13/2023 01:53:13 PM, Scero, Tomlyn) email notification from Zoll saying that the defib failed do to the battery. I went to CCU and tested the defib, confirmed the issue and replaced the battery. put the battery on the charger so it could recondition it.
53076 | 193109 | CM | CMPLT | CMPLT | 2023-02-13 14:38 | 2023-02-13 14:41 | Died on transport - no vitals | (TC 02/13/2023 02:38:39 PM, Ferreira, Qwynn) tested the transport noticed the charging cable for the X2 dock was not correctly seated. Changed the power cables from a 90deg bend to straight plug in. Added foam to apply pressure to keep the cable tucked in. secured with tape temporarily. Replaced x2 battery Tested the X2 - OK Tested the transport - OK Tightened the wheel straps as they were incorrectly placed.
| 193110 | PA | CMPLT | CMPLT | 2023-02-13 14:42 | 2023-02-16 13:12 | Sterile Water
1056493 | 193111 | CM | CMPLT | CMPLT | 2023-02-14 08:07 | 2023-02-14 08:08 | Radical displaying "no SD card" | (TC 02/14/2023 08:07:23 AM, Baghirzade, Nara) Re-seated the SD card. Functional tests passed; checked SpO2, BP, and temp.
1056436 | 193112 | CM | CMPLT | NFF | 2023-02-14 08:14 | 2023-02-14 08:15 | Temperature probe not working | (TC 02/14/2023 08:14:11 AM, Baghirzade, Nara) Measured temperature multiple times; was unable to see any errors. Ensured that the connector was sitting in properly and temperature module is fully functional. No fault found. Checked SpO2 and BP; all functional.
1051825 | 193113 | CM | CMPLT | CMPLT | 2023-02-14 09:51 | 2023-02-14 09:57 | removing from IT closet on 4th floor used with ENDO | (TC 02/13/2023 09:51:09 AM, Samanian, Kamran) removed from IT closet (TC 02/14/2023 09:51:27 AM, Samanian, Kamran) copied the configuration to biomed laptop. erased the configuration from switch. just default loaded.
1065686 | 193114 | ACC | CMPLT | CMPLT | 2023-02-14 10:55 | 2023-05-12 09:32 | Incoming Inspection
1065687 | 193115 | ACC | CMPLT | CMPLT | 2023-02-14 11:08 | 2023-05-12 09:32 | Incoming inspection
1063808 | 193116 | CM | CMPLT | CMPLT | 2023-02-14 11:10 | 2023-07-11 07:29 | MSH 15 SPP VOLUSON 10 QWERTY KYBD , SOME KEYS SHOWING RESISTANCE: REPORTED BY LIZ & INGRID | (TC 02/14/2023 11:10:01 AM, Whittington, Scott) MSH 15 SPP VOLUSON 10 QWERTY KYBD , SOME KEYS SHOWING RESISTANCE: REPORTED BY LIZ & INGRID WILL LOOKUP GE PN# AND ORDER UNDER CONTRACT ALL KEYS ARE WORKING | MSH 15 SPP VOLUSON 10 QWERTY KYBD , SOME KEYS SHOWING RESISTANCE: REPORTED BY LIZ & INGRID
65924 | 193117 | CM | CMPLT | CMPLT | 2023-02-14 11:12 | 2023-02-14 11:17 | Channel Error | (TC 02/14/2023 11:12:33 AM, Reid, Emmerson ) Downloaded log history, had many previous issues with the pressure sensor. Replaced the pressure sensor, and the back case as it broke when opening the pump. Ran Calibration, all tests passed. Ran ASM PM, all tests passed. Working now.
1049878 | 193118 | FC | CMPLT | CMPLT | 2023-02-14 12:26 | 2023-02-14 12:28 | Checking safety goggles before use | (TC 02/14/2023 12:26:32 PM, Baghirzade, Nara) Checked 4 safety goggles'; all have indications for wavelengths 970 nm to 1470 nm, with OD 5+ All goggles are in perfect physical condition; no scratches, fogging, or damage on the lenses. Returned to Kim Ngo.
52259 | 193119 | CM | CMPLT | CMPLT | 2023-02-14 12:38 | 2023-02-14 12:40 | Porthole hinge broken | (TC 02/09/2023 12:38:27 PM, Baghirzade, Nara) Replaced all latches under recall. Noticed that only one castor has a working break, rest of the 3 breaks on castors are broken. Replaced the SW castor. Incubator now has 2 working castors, one on each corner.
52259 | 193120 | PM | CMPLT | CMPLT | 2023-02-14 12:41 | 2023-02-14 12:55 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 02/14/2023 12:52:56 PM, Baghirzade, Nara) During functional check, the RH sensor was inaccurately displaying the humidity. The humidity on the incubator was also taking longer than normal to change on the display. After trying to calibrate the RH sensor twice and both times it failed, I replaced the RH sensor. Calibrated the RH sensor again and it failed. Used salt to calibrate the RH sensor and it worked. PM completed. VRef, Mains V, LV, and RH were calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 74% Set temp to 35C, measured 35C approximately on all 4 corner sensors on the analyzer. All functions of the incubator are good. Electrical safety results - Leakage current = 34.9 uA Power cord resistance = 342 mohms
1054987 | 193121 | CM | CMPLT | CMPLT | 2023-02-14 13:11 | 2023-02-14 13:11 | moving from 11S room to 11N room locations | (TC 02/14/2023 01:10:59 PM, Samanian, Kamran) done!
1053348 | 193122 | CM | CMPLT | CMPLT | 2023-02-14 14:19 | 2023-07-11 07:29 | MSH BI 1272 H3 12th Floor - North Side COMPUTER UPGRADE FOR CONTRAST ENHANCED BIOPSY | (TC 02/14/2023 02:19:01 PM, Whittington, Scott) MSH BI 1272 H3 12th Floor - North Side COMPUTER UPGRADE FOR CONTRAST ENHANCED BIOPSY NOTE CURRENT COMPUTER MAC CONNECTION 2: 3C:EC:EF:00:06:49 (TC 02/15/2023 09:54:51 AM, Whittington, Scott) MSH BI 1272 H3 12th Floor - North Side COMPUTER UPGRADE FOR CONTRAST ENHANCED BIOPSY MIRZA STATED NEW MAC WHITE LISTED The new MAC is 3C:EC:EF:B9:68:45 The old MAC was 3C:EC:EF:00:06:49 From: HELP Desk (MSH) <helpdesk.MSH@sinaihealth.ca> Sent: February 15, 2023 12:24 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Ultrasound Support <Ultrasound.Support@uhn.ca>; Moe - Dynamic Imaging <moe@dynamicimaging.ca>; Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca>; Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca> Subject: [External] N-IM039814972-RE: MSH 12 Breast Imaging room H3 new MAC Hi Scott, A ticket has been logged for this email. Ticket number : N-IM039814972, sent to Network Team. If you have any questions or concerns, please contact us at ext4357 and reference the above case number. Best regards, Patrick Brunet MSH Service Desk Phone : 416 586 4800 x 4357 Toll Free:1-866-244-9909(MSH) Email: toronto.carecentre@dxc.com From: Whittington, Scott [mailto:Scott.Whittington@uhn.ca] Sent: Tuesday, February 14, 2023 8:40 PM To: HELP Desk (MSH) <helpdesk.MSH@sinaihealth.ca> Cc: Ultrasound Support <Ultrasound.Support@uhn.ca>; Moe - Dynamic Imaging <moe@dynamicimaging.ca>; Fisico, Jason <Jason.Fisico@sinaihealth.ca>; Schwietz, Matthew <Matthew.Schwietz@sinaihealth.ca> Subject: MSH 12 Breast Imaging room H3 new MAC CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Good evening The above system has had its NIC card replaced The new MAC is 3C:EC:EF:B9:68:45 The old MAC was 3C:EC:EF:00:06:49 Please whitelist the new MAC Thanks Scott 416-402-2462 | MSH BI 1272 H3 12th Floor - North Side COMPUTER UPGRADE FOR CONTRAST ENHANCED BIOPSY
1065738 | 193123 | ACC | CMPLT | CMPLT | 2023-02-14 14:31 | 2023-02-14 14:33 | Incoming inspection | (TC 02/14/2023 02:31:00 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Leakage current = 62.4 uA Power cord resistance = 15 mohms Quick in service with the vendor and troubleshooting tips for biomed.
1049941 | 193124 | CM | CMPLT | CMPLT | 2023-02-15 07:43 | 2023-02-15 07:43 | Panel doesn't stay in place | (TC 02/14/2023 11:45:00 AM, Fan, Tak) Readjusted frogleg brakes.
1049940 | 193125 | CM | CMPLT | CMPLT | 2023-02-15 07:44 | 2023-02-16 15:20 | Sticky unresponsive buttons. | (TC 02/16/2023 03:19:57 PM, Fan, Tak) Replaced panel. Returned to vendor.
502366 | 193126 | CM | CMPLT | CMPLT | 2023-02-15 07:44 | 2023-02-16 15:21 | Sticky unresponsive buttons. | (TC 02/15/2023 03:20:33 PM, Fan, Tak) Replaced panel. Returned to vendor.
1061333 | 193127 | FC | CMPLT | CMPLT | 2023-02-15 11:28 | 2023-02-15 11:37 | QC Checklist | (TC 02/15/2023 11:31:03 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1054986 | 193128 | CM | CMPLT | CMPLT | 2023-02-15 12:28 | 2023-02-15 12:34 | Consumables for respiratory monitoring | (TC 02/15/2023 12:28:39 PM, Baghirzade, Nara) Met with Sarabeth to discuss the consumables required for respiratory monitoring. They require tapes and sensors; dropped it off with Sarabeth.
| 193129 | TEDU | CMPLT | CMPLT | 2023-02-15 13:01 | 2023-02-15 13:02 | Procuity and MV3 PM training | (TC 02/15/2023 01:01:23 PM, Baghirzade, Nara) Procuity and MV3 PM training
| 193130 | CM | CMPLT | CMPLT | 2023-02-15 14:51 | 2023-02-15 14:51 | working in ENDO closet to get information | (TC 02/15/2023 02:51:35 PM, Samanian, Kamran) done!
| 193131 | CM | CMPLT | CMPLT | 2023-02-15 14:52 | 2023-02-15 14:53 | meeting with Abdel | (TC 02/15/2023 02:52:24 PM, Samanian, Kamran) discussed about FP and OS patch, IBE alert from overview, Network diagram
32510 | 193132 | CM | CMPLT | CMPLT | 2023-02-15 15:08 | 2023-04-24 11:59 | MSH DR1 SYSTEM REPORTED "DOWN"; AT SITE SYSTEM OPERATING; FURTHER INVESTIGATION REQ | (TC 02/15/2023 03:08:31 PM, Whittington, Scott) MSH DR1 SYSTEM REPORTED "DOWN"; AT SITE SYSTEM OPERATING; FURTHER INVESTIGATION REQ | MSH DR1 SYSTEM REPORTED "DOWN"; AT SITE SYSTEM OPERATING; FURTHER INVESTIGATION REQ
32520 | 193133 | CM | CMPLT | CMPLT | 2023-02-15 15:10 | 2023-02-24 07:52 | MSH DR2 SYSTEM REPORTED "DOWN"; AT SITE SYSTEM OPERATING; FURTHER INVESTIGATION REQ | (TC 02/15/2023 03:10:14 PM, Whittington, Scott) MSH DR2 SYSTEM REPORTED "DOWN"; AT SITE SYSTEM OPERATING; FURTHER INVESTIGATION REQ (TC 02/15/2023 07:51:00 PM, Whittington, Scott) MSH DR2 SYSTEM REPORTED "DOWN"; AT SITE SYSTEM OPERATING; FURTHER INVESTIGATION REQ FOUND OGP ETRROR AND CORRECTED | MSH DR2 SYSTEM REPORTED "DOWN"; AT SITE SYSTEM OPERATING; FURTHER INVESTIGATION REQ
1055911 | 193134 | CM | CMPLT | CMPLT | 2023-02-15 15:11 | 2023-03-29 15:43 | Mounting plate damaged, can't attach Root to stand | (TC 02/15/2023 03:46:06 PM, Cortez, John) Ordered quick connect kit (TC 02/24/2023 03:52:39 PM, Cortez, John) Part arrived. installed on rollstand (TC 03/29/2023 03:42:12 PM, Cortez, John) Replaced temp probe All parameters tested ok
1062025 | 193135 | CM | CMPLT | NFF | 2023-02-15 15:55 | 2023-02-15 15:59 | "fall risk alarm not working" | (TC 02/14/2023 03:55:33 PM, Cortez, John) Power cycled bed, unable to replicate error Confirmed functionality of bed exit alarm with calibrated weights
1062006 | 193136 | CM | CMPLT | CMPLT | 2023-02-15 16:00 | 2023-02-15 16:02 | "unable to set bed exit alarm" | (TC 02/15/2023 04:00:08 PM, Cortez, John) "Limited motion" error code Power cycling bed seems to resolve issue Nurse will re weigh patient at another time and call me back if issue remains
1047849 | 193137 | PA | CMPLT | CMPLT | 2023-02-16 08:21 | 2023-02-27 10:53 | Missing plunger lever; ordering plunger assembly kit | (TC 02/16/2023 08:30:29 AM, Baghirzade, Nara) SC 1002886260 (TC 02/27/2023 10:51:38 AM, Baghirzade, Nara) PO 7202163871
1059828 | 193138 | CM | CMPLT | CMPLT | 2023-02-16 09:32 | 2023-02-16 09:32 | Cleaning | (TC 02/16/2023 09:32:05 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1063746 | 193139 | CM | CMPLT | CMPLT | 2023-02-16 11:22 | 2023-02-16 11:22 | Cleaning | (TC 02/16/2023 11:21:59 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1055594 | 193140 | CM | CMPLT | CMPLT | 2023-02-16 11:28 | 2023-02-16 11:29 | BP not working | (TC 02/16/2023 11:28:11 AM, Baghirzade, Nara) Checked SpO2, BP, and temperature; all modules are functional. No fault found with the BP. Noticed that there was no configuration loaded to the Root; loaded the MSH clinical configuration and network configuration. Checked that the HR low limit is 45. Device is fully functional.
1061333 | 193141 | FC | CMPLT | CMPLT | 2023-02-16 11:30 | 2023-02-16 11:32 | QC checklist | (TC 02/16/2023 11:31:45 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
1050481 | 193142 | CM | CMPLT | CMPLT | 2023-02-16 12:02 | 2023-02-24 10:21 | Techinal error TE: 7 | Techical error 55, 16 and 7 on the logs. Sent it to Getinge. Passed pre-use check. Running with test lung. We got suggestiond form Alex(Getinge): 1- Make sure the control cable going from ventilator to user interface is securely connected. Check that both ends of the cable are fully screwed in. 2- Check the control cable for damage along it’s full length 3- Remove the battery modules and re-insert them (do this for all modules in the plug and play compartments). Make sure that the locking latch is securely engaged on each battery and each sensor module. 4- Restart the Servo and perform 2 pre-use checks (back-to-back without powering off the device). 5- Perform test ventilation using Getinge neonatal test lung for 4 hours and observe for technical alarms or errors. 6- If you find no further errors, you can return the Servo to clinical use. We did all and fully functional, no error.
1061432 | 193143 | CM | CMPLT | CMPLT | 2023-02-16 12:07 | 2023-02-15 12:06 | broken power plug | repaired power plug. tested. Functional.
66457 | 193144 | CM | CMPLT | CMPLT | 2023-02-16 12:08 | 2023-02-15 12:08 | alarming | tested. functional.
65263 | 193145 | CM | CMPLT | CMPLT | 2023-02-16 12:09 | 2023-02-15 12:09 | C500 dropped on its pole | re-installed C500. tested. Functional.
1061456 | 193146 | CM | CMPLT | CMPLT | 2023-02-16 12:11 | 2023-02-15 12:10 | not working | tested. Functional.
1047846 | 193147 | FC | CMPLT | CMPLT | 2023-02-16 13:12 | 2023-02-16 13:13 | Adjusted time and functional check | (TC 02/16/2023 01:11:51 PM, Baghirzade, Nara) Connected the syringe pump the pharmguard, adjust the time. One of the plunger levers was loose; re-seated the springs and gears; lever is functional, and claws fully open and close. Functional tests passed. Tested the pump with 30 mL and 60 mL syringes; all passed.
67935 | 193148 | CM | CMPLT | CMPLT | 2023-02-16 13:13 | 2023-02-16 13:14 | Air probe failure | (TC 02/16/2023 01:13:16 PM, Ferreira, Qwynn) Tested the unit and after 20 minutes there was an air probe failure. Replaced the air probe Installed the recall latches Calibrated the unit and RH sensor Functional Test - OK
1066063 | 193149 | ACC | CMPLT | CMPLT | 2023-02-16 13:59 | 2023-05-12 09:38 | Incoming inspection | (TC 02/16/2023 09:35:29 AM, Cortez, John) Unboxed and assembled Visual/functional check ok Waiting for Jason from MRI to schedule training for staff before putting lift into service
1066064 | 193150 | ACC | CMPLT | CMPLT | 2023-02-16 14:03 | 2023-05-12 09:39 | Incoming inspection | (TC 02/16/2023 09:38:37 AM, Cortez, John) Unboxed and assembled Visual/functional check ok Waiting for Jason from MRI to schedule training for staff before putting lift into service
1047846 | 193151 | CM | CMPLT | CMPLT | 2023-02-16 14:23 | 2023-02-16 14:24 | Configuration updated | (TC 02/16/2023 02:23:09 PM, Baghirzade, Nara) Updated the configuration to the latest: MARCH 2019
1064825 | 193152 | CM | CMPLT | CMPLT | 2023-02-16 14:33 | 2023-02-24 10:17 | SpO2 malfunction alarm | (TC 02/16/2023 02:42:50 PM, Canjimavadivel, Kamallesh ) Confirmed issue and swapped out with spare. Brought to biomed to recreate and isolate issue (TC 02/23/2023 10:15:58 AM, Canjimavadivel, Kamallesh ) Unable to recreate issue, will continue to monitor for reoccurance.
1047840 | 193153 | PM | CMPLT | CMPLT | 2023-02-16 14:40 | 2023-08-21 14:40 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 08/09/2023 02:39:48 PM, Rhuland, Courtney) Rebuilt clutch assembly (TC 08/10/2023 03:17:50 PM, Rhuland, Courtney) Rebuilt device, working on issue with plunger flippers not registering (TC 08/16/2023 02:29:54 PM, Rhuland, Courtney) Disassembled plunger assemby to remove plunger cable Plunger cable found to be damaged Plunger cable on order (TC 08/21/2023 02:36:42 PM, Rhuland, Courtney) Replaced plunger cable Rebuilt device Replaced top case Performed PM, refer to procedure for details
1055526 | 193154 | CM | CMPLT | CMPLT | 2023-02-16 15:00 | 2023-02-16 15:01 | not working | (TC 02/16/2023 03:00:44 PM, Ferreira, Qwynn) tested and had error p03 replaced the temp probe tested at 37deg = 37 sent to unit
501065 | 193155 | CM | CMPLT | CMPLT | 2023-02-16 15:24 | 2023-02-22 07:29 | Monitor not displaying anything | (TC 02/15/2023 05:23:00 PM, Fan, Tak) Working. Swapped video cards, stopped working. Swapped cards back, still not working. Used spare monitor, display shows on ext monitor but not main monitor. Ordered a new monitor. (TC 02/21/2023 02:28:00 PM, Fan, Tak) Replaced monitor.
1063808 | 193156 | CM | CMPLT | CMPLT | 2023-02-16 15:31 | 2023-02-16 15:31 | MSH 15 SPP VOLUSON 10 REMOTE VIDEO NOT WORKING AFTER DEEP CLEAN | (TC 02/16/2023 03:31:17 PM, Whittington, Scott) MSH 15 SPP VOLUSON 10 REMOTE VIDEO NOT WORKING AFTER DEEP CLEAN SVIDEO CABLE HAD BEEN DISCONNECTED FROM BOOM RESTORED, OK | MSH 15 SPP VOLUSON 10 REMOTE VIDEO NOT WORKING AFTER DEEP CLEAN
501173 | 193157 | CM | CMPLT | CMPLT | 2023-02-16 15:33 | 2023-02-16 15:44 | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 REMOTE VIDEO NOT AVAILABLE AFTER DEEP CLEAN | (TC 02/16/2023 03:33:24 PM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 REMOTE VIDEO NOT AVAILABLE AFTER DEEP CLEAN RESTORED CONNECTION OK | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 REMOTE VIDEO NOT AVAILABLE AFTER DEEP CLEAN
| 193158 | PA | CMPLT | CMPLT | 2023-02-17 08:39 | 2023-03-09 08:12 | PM Kits for servo-u | (TC 02/17/2023 08:45:29 AM, Canjimavadivel, Kamallesh ) Placed order for PM kits
| 193159 | CM | CMPLT | CMPLT | 2023-02-17 11:45 | 2023-02-28 09:32 | Alaris PCU RS-232 Board | Received
1046035 | 193160 | CM | CMPLT | CMPLT | 2023-02-17 11:59 | 2023-02-17 12:00 | MSH ER - System Revision Downgrade | (TC 02/17/2023 11:58:39 AM, Chau, Ken) Downgrade from 2.23 to 2.14 -tube cal -no more mA errors caused (TC 02/17/2023 11:58:39 AM, Whittington, Scott) MSH ER - System Revision Downgrade (TC 02/17/2023 11:58:39 AM, Whittington, Scott) MSH ER - System Revision Downgrade Downgrade from 2.23 to 2.14 -tube cal -no more mA errors caused (TC 02/17/2023 11:58:39 AM, Dolinar, John) MSH ER - System Revision Downgrade Downgrade from 2.23 to 2.14 -tube cal -no more mA errors caused | MSH ER - System Revision Downgrade
1053348 | 193161 | CM | CMPLT | CMPLT | 2023-02-17 13:14 | 2023-02-17 13:18 | MSH BI 1272 H3 DURING QC SPOT COMPRESSION PADDLE STOPPED MOVING IN COMPRESSION ERROR PMC38:41 | MSH BI 1272 H3 DURING QC SPOT COMPRESSION PADDLE STOPPED MOVING IN COMPRESSION From: Hewlett, Rose <Rose.Hewlett@sinaihealth.ca> Sent: February 17, 2023 9:11 AM To: Veeneman, Jon <Jon.Veeneman@hologic.com>; Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca> Cc: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] RE: Error on Hologic unit - No compression So it seems that it was the initial use of any of the spot and mag compression paddle that was causing the problem. Once it’s used initially, it seems to work. I tried all the paddles and so far it’s ok Will let you know if we have any issues during the day Rose From: Veeneman, Jon [mailto:Jon.Veeneman@hologic.com] Sent: February 17, 2023 8:51 AM To: Hewlett, Rose <Rose.Hewlett@sinaihealth.ca>; Fisico, Jason <Jason.Fisico@sinaihealth.ca> Cc: Whittington, Scott - UHN - University Health Network <Scott.Whittington@uhn.ca> Subject: RE: Error on Hologic unit - No compression CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. I understand it was one paddle only, can you tell me which paddle was giving you the trouble? And can you try the same paddle from a different room so we can determine if it’s the software or the paddle. Thanks Jon Veeneman Service Lead Eastern Canada M 289.407.6748 E Jon.veeneman@hologic.com L 2400 Skymark Ave Unit 7A Mississauga ON L4W 5K5 Hologic.com Follow Us Achiever – Ideation – Communication – Arranger - Restorative From: Hewlett, Rose <Rose.Hewlett@sinaihealth.ca> Sent: Friday, February 17, 2023 8:39 AM To: Fisico, Jason <Jason.Fisico@sinaihealth.ca>; Veeneman, Jon <Jon.Veeneman@hologic.com> Cc: Whittington, Scott - UHN - University Health Network <Scott.Whittington@uhn.ca> Subject: RE: Error on Hologic unit - No compression External Mail The up/down foot pedal control for the whole system is working, but the compression is not From: Fisico, Jason Sent: February 17, 2023 8:33 AM To: Veeneman, Jon <Jon.Veeneman@hologic.com> Cc: Hewlett, Rose <Rose.Hewlett@sinaihealth.ca>; Whittington, Scott - UHN - University Health Network <Scott.Whittington@uhn.ca> Subject: RE: Error on Hologic unit - No compression Hi Jon. My understanding is there is no compression at all right now. Scott had said he was going to see if Mo is still in town before checking it out himself. Jason From: Veeneman, Jon [mailto:Jon.Veeneman@hologic.com] Sent: February 17, 2023 8:30 AM To: Fisico, Jason <Jason.Fisico@sinaihealth.ca> Cc: Hewlett, Rose <Rose.Hewlett@sinaihealth.ca> Subject: Re: Error on Hologic unit - No compression CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Ok thanks Jason. Was it a one time issue or will it not compress at all anymore? If rose can give me a time I’ll download the logs to see what’s going on. Jon Veeneman Area Service Lead - Eastern Canada Field Service Engineer Hologic Canada 289-407-6748 (cell) jon.veeneman@hologic.com ________________________________________ From: Fisico, Jason <Jason.Fisico@sinaihealth.ca> Sent: Friday, February 17, 2023 8:28:24 AM To: Veeneman, Jon <Jon.Veeneman@hologic.com> Cc: Hewlett, Rose <Rose.Hewlett@sinaihealth.ca> Subject: Error on Hologic unit - No compression External Mail Hi Jon. Rose reported an error this morning where our H3 unit wouldn’t compress after the paddle was switched. I’ve cc’d her so she can provide an error number. I’ll see if Scott here can take a look but just for your awareness for now. Thanks, Jason Fisico M.MedRadSc(MRI), MRT(N)(MR) MRI Supervisor – Interim Breast Imaging Supervisor Medical Imaging Research Co-Lead JDMI Practice Council Co-Chair jason.fisico@SinaiHealth.ca T 416-586-4800 x1591 F 416-586-8321 (TC 02/21/2023 10:13:09 AM, Whittington, Scott) MSH BI 1272 H3 DURING QC SPOT COMPRESSION PADDLE STOPPED MOVING IN COMPRESSION ERROR PMC38:41 PROBLEM REOCURRED TUESDAY FEB 21 JON V LOGGING IN REMOTELY TO CHECK | MSH BI 1272 H3 DURING QC SPOT COMPRESSION PADDLE STOPPED MOVING IN COMPRESSION
1059656 | 193162 | CM | CMPLT | CMPLT | 2023-02-17 13:23 | 2023-03-15 14:59 | Damaged head - Detached from stand | (TC 02/16/2023 01:23:11 PM, Rendon Lopez , Elizabeth) Device is under warranty, requested replacement from Welch Allyn. No charge replacement on route. (TC 03/15/2023 02:58:08 PM, Rendon Lopez , Elizabeth) Received new device, installed, tested-OK. Returned damaged light
1046196 | 193163 | CM | CMPLT | CMPLT | 2023-02-17 13:45 | 2023-02-17 13:45 | PM | PM completed
1049362 | 193164 | CM | CMPLT | CMPLT | 2023-02-17 13:51 | 2023-02-17 15:03 | PM - Yearly | Injector PM done -Flex Ring calibrated -Potentiometer calibrated -Pressure calibrated -Touchscreen calibrated -System passed all functional tests
1049941 | 193165 | CM | CMPLT | CMPLT | 2023-02-17 17:01 | 2023-02-17 17:01 | Patient not showing up on workstation. |  | Patient as being sent to downtime.
67930 | 193166 | CM | CMPLT | CMPLT | 2023-02-21 08:23 | 2023-02-21 08:23 | Latch recall | (TC 02/17/2023 07:30:00 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
67889 | 193167 | CM | CMPLT | CMPLT | 2023-02-21 08:36 | 2023-02-21 08:37 | Keeps saying "downstream occlusion" | (TC 02/15/2023 08:36:23 AM, Rendon Lopez , Elizabeth) Running for several hours. No fault found
1058237 | 193168 | CM | CMPLT | CMPLT | 2023-02-21 08:39 | 2023-02-21 08:40 | Error | (TC 02/15/2023 08:39:20 AM, Rendon Lopez , Elizabeth) Downloaded error log, no logs recorded. Performed check in tests, all passed
68673 | 193169 | CM | CMPLT | CMPLT | 2023-02-21 08:48 | 2023-02-21 08:49 | Red battery alarm | (TC 02/15/2023 08:48:22 AM, Rendon Lopez , Elizabeth) Replaced battery, PM completed, all tests passed
66530 | 193170 | CM | CMPLT | CMPLT | 2023-02-21 08:52 | 2023-02-21 08:54 | Recorder not printing | (TC 02/16/2023 08:52:44 AM, Rendon Lopez , Elizabeth) Confirmed issue, removed monitor from room. Replaced recorder with spare repaired one. Tested all functions-OK
1054161 | 193171 | CM | CMPLT | CMPLT | 2023-02-21 08:54 | 2023-02-21 08:55 | Missing button | (TC 02/16/2023 08:54:03 AM, Rendon Lopez , Elizabeth) Replaced button with spare, tested-OK
63072 | 193172 | CM | CMPLT | CMPLT | 2023-02-21 11:45 | 2023-02-21 11:46 | clips fell off - Strap for the legs | (TC 02/21/2023 11:45:06 AM, Ferreira, Qwynn) repalced the clip Secured the strap to the clip
66106 | 193173 | CM | CMPLT | CMPLT | 2023-02-21 11:51 | 2023-02-21 11:52 | QC checklist | (TC 02/21/2023 11:51:17 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
1059778 | 193174 | FC | CMPLT | CMPLT | 2023-02-21 12:20 | 2023-02-21 12:20 | QC checklist | (TC 02/21/2023 12:20:03 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
| 193175 | CM | CMPLT | CMPLT | 2023-02-21 12:47 | 2023-02-21 12:47 | meeting - Huddle | (TC 02/21/2023 12:47:29 PM, Samanian, Kamran) done!
| 193176 | CM | CMPLT | CMPLT | 2023-02-21 12:48 | 2023-02-21 12:49 | upgrading 3 x X2s for C12 option and 2 x X2 to M Rev | (TC 02/21/2023 12:48:40 PM, Samanian, Kamran) done!
66692 | 193177 | PM | CMPLT | CMPLT | 2023-02-21 13:52 | 2023-02-21 14:13 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 02/21/2023 01:52:45 PM, Ferreira, Qwynn) PM Passed System calibration - OK Main calibration - ok Line Voltage Cal- OK Electrical Safety - OK Functional Check - ok Unit came down with broken door Repaired the port door and tested the device RH was off by 10% Calibration was performed but did not solve the issue Tested the humidifier and this was not the issue Replaced the SSR and this solved the issue Tested the device to ensure accuracy. Sent back to unit
1048498 | 193178 | CM | CMPLT | CMPLT | 2023-02-21 15:12 | 2023-02-22 08:33 | Cuts in and out | (TC 02/21/2023 03:12:05 PM, Reid, Emmerson ) Opened up the device because it was cutting in and out. Found that the issue was coming from the volume dial due to the corner of the board being broken off. Looks like the device had been opened up before as the tamper seal was broken and I could not find the broken piece inside of the device. (TC 02/22/2023 08:31:54 AM, Reid, Emmerson ) Cannot repair, retiring.
| 193179 | CM | CMPLT | CMPLT | 2023-02-21 17:23 | 2023-02-21 17:23 | Site coverage
1047847 | 193180 | CM | CMPLT | CMPLT | 2023-02-22 07:48 | 2023-02-22 07:55 | Stopped working; no other info provided | (TC 02/22/2023 07:54:08 AM, Baghirzade, Nara) Pump was intermittently switching on/off; found the battery connector was not properly fitted. Re-seated the battery connector. When switching on the pump, a system failure message appeared: "Primary audible alarm BGND test". After looking into the error log, also found a "ACU watchdog failure"; as per the manual, to restart the pump and check if error occurs again. All functional tests passed. Ran the pump on empty syringe and no errors occured during this; also the occlusion and empty syringe alarm were also functional. Pump is fully functional.
| 193181 | CM | CMPLT | CMPLT | 2023-02-22 07:49 | 2023-02-21 07:49 | sinai course online mandatory | (TC 02/21/2023 07:49:02 AM, Samanian, Kamran) done!
1047847 | 193182 | CM | CMPLT | CMPLT | 2023-02-22 07:55 | 2023-02-22 07:56 | Updated time and date | (TC 02/22/2023 07:55:30 AM, Baghirzade, Nara) Date and time was adjusted. Confirmed that the correct MSH NICU config is uploaded.
1056538 | 193183 | FC | CMPLT | CMPLT | 2023-02-22 08:11 | 2023-02-22 08:19 | Functional Check/Verification | (TC 02/22/2023 08:14:18 AM, Baghirzade, Nara) No electrical safety test as the device is battery powered. The AED is battery powered and operated, and conducts self-tests on a daily and weekly basis. When AED is switched on, the AED conducts another self-test. The AED does not require a calibration or energy delivery verification.
| 193184 | MEP | CMPLT | CMPLT | 2023-02-22 08:31 | 2023-06-13 16:00 | Masimo site survey, viewstation patches | (TC 02/21/2023 08:31:25 AM, Cortez, John) Assisted Mirza in verifying if selected datajacks are live (TC 03/03/2023 03:48:10 PM, Cortez, John) Retrieved monitors from 18 (TC 03/02/2023 03:49:15 PM, Cortez, John) Placed roots in rooms on 18 (TC 03/01/2023 03:50:20 PM, Cortez, John) Assisted Masimo reps (TC 02/27/2023 03:50:50 PM, Cortez, John) Assisted Masimo reps
| 193185 | CM | CMPLT | CMPLT | 2023-02-22 08:47 | 2023-04-19 08:48 | XDS to excel export project for Dr. Poorva - NICU | (TC 02/22/2023 08:47:25 AM, Samanian, Kamran) connected the cable between port 16-018BD patch panel to switch DIST A port 18 . Port 18 is set to VLAN 212 XDS laptop IP: 172.31.18.50 with subnet mask 255.255.254.0 (TC 03/20/2023 09:51:45 AM, Samanian, Kamran) Tried to install SQl 2022 and issue. tried to install SQL 2019 and issue. Searched for known problem between SQL and windows 11. had to change the registry key and added a key in Registry key. Also problem with Windows 11 home edition and firewall. had to change the GP to disable the firewall. Also disabled WLAN card. tested application and all passed. (TC 03/27/2023 01:50:33 PM, Samanian, Kamran) worked with sean to check the export on NIB52 today (TC 03/29/2023 08:46:15 AM, Samanian, Kamran) wokring with Philips on deleteing data and clearing the cache for data in database
1057106 | 193186 | CM | CMPLT | CMPLT | 2023-02-22 09:03 | 2023-02-13 09:08 | X3 powered off during transport | (TC 02/13/2023 09:06:37 AM, Ferreira, Qwynn) The 90 degree power cord continues to falls out which stops the charging of the X3 monitor. Swapped for a straight cable, added a piece of support foam to help keep the cable in place and taped the foam in temporarily to stop the cable from unplugging. Charging and working now. More secure.
1050672 | 193187 | CM | CMPLT | CMPLT | 2023-02-22 09:05 | 2023-02-22 09:07 | Lead V4 not picking up | (TC 02/22/2023 09:06:19 AM, Canjimavadivel, Kamallesh ) Replaced both chest leads set as they were affected by recall. Tested ok with simulator.
| 193188 | CM | ACTIV | ISSUE | 2023-02-22 09:13 |  | mistrail warming units warranty replacement
1057867 | 193189 | PM | CMPLT | CMPLT | 2023-02-22 09:35 | 2023-02-22 11:28 | Annual PM | (TC 02/22/2023 09:36:07 AM, Reid, Emmerson ) Replaced the enclosure cover. Replaced O-rings and greased. Replaced circulating fluid. See procedure for details.
64078 | 193190 | CM | CMPLT | CMPLT | 2023-02-22 10:36 | 2023-04-28 08:01 | Hood not lifting | (TC 02/22/2023 10:36:07 AM, Ferreira, Qwynn) Came down not able to raise the canopy Tested the device and looked like a broken or damaged rail. REmoved the canopy and checked the rail and noticed it was chain so the rail was replaced recently. Continued to troubleshoot and found the coupler between the motor and the chain had been damaged. This looks like a force was applied to the canopy during use and the motor spun through the coupler. Ordering new couplers. (TC 04/20/2023 02:24:21 PM, Ferreira, Qwynn) Finished the work order under the PM work order.
1055580 | 193191 | CM | CMPLT | CMPLT | 2023-02-22 13:10 | 2023-02-16 13:09 | NIBP needs calibration | Kam got password for NIBP calibration menu: 4258. Did calibration, tested, functional.
1049901 | 193192 | CM | CMPLT | CMPLT | 2023-02-22 13:30 | 2023-02-22 13:31 | MSH R61 - System Battery issues | (TC 02/22/2023 01:30:13 PM, Chau, Ken) Changed system batteries
1049905 | 193193 | CM | CMPLT | CMPLT | 2023-02-22 15:45 | 2023-02-22 15:47 | Batteries depleted (system batt) | (TC 02/22/2023 03:45:17 PM, Fan, Tak) Replaced system batteries.
| 193194 | PA | CMPLT | CMPLT | 2023-02-23 08:39 | 2023-05-12 13:09 | Gearbox Coupler
1059828 | 193195 | PM | CMPLT | CMPLT | 2023-02-23 09:42 | 2023-03-31 09:58 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 03/31/2023 09:48:10 AM, Ferreira, Qwynn) PM Passed Tested alarms Tested overtemp Tested flow rate Replaced two sheet filters replaced 9V Replaced air exhaust filter and tubing Cleaned some built up fluid in the altrix. It was a blueish greenish colour. Reassembled and performed the cleaning cycle to disinfect, rinse, and refilled with sterile water.
| 193196 | CM | CMPLT | CMPLT | 2023-02-23 10:04 | 2023-02-23 10:04 | MSH El-learning course | (TC 02/23/2023 10:04:28 AM, Samanian, Kamran) finished!
1065609 | 193197 | CM | CMPLT | CMPLT | 2023-02-23 10:06 | 2023-02-23 14:24 | Wrong WO | (TC 02/23/2023 10:06:22 AM, Samanian, Kamran) wrong WO
1057535 | 193198 | CM | CMPLT | CMPLT | 2023-02-23 11:47 | 2023-02-23 11:47 | upgrade | (TC 02/23/2023 11:46:47 AM, Samanian, Kamran) upgraded to rev n.00.04-108 and configured for ICU
1057522 | 193199 | CM | CMPLT | CMPLT | 2023-02-23 11:52 | 2023-02-23 11:53 | upgrade | (TC 02/23/2023 11:52:34 AM, Samanian, Kamran) upgraded to Rev N.00.04-108 and configured for ICU
67921 | 193200 | CM | CMPLT | CMPLT | 2023-02-23 12:01 | 2023-02-23 12:13 | Does not turn on | (TC 02/23/2023 12:11:58 PM, Ferreira, Qwynn) Checked the device and it presented a system error 25 Checked the fan assembly and the grey gromment was not correctly placed Replaced the gromment and assembled the fan functional Test - OK Returned to unit
1057534 | 193201 | CM | CMPLT | CMPLT | 2023-02-23 12:21 | 2023-02-23 12:22 | upgrade | (TC 02/23/2023 12:21:44 PM, Samanian, Kamran) upgraded to Rev N.00.04-108
| 193202 | CM | CMPLT | CMPLT | 2023-02-23 13:03 | 2023-03-07 09:07 | Ordering small adult Masimo BP cuffs for stock | (TC 03/02/2023 03:51:53 PM, Cortez, John) cuffs arrived (TC 03/02/2023 09:06:02 AM, Cortez, John) Will charge back to units when used
67136 | 193203 | CM | CMPLT | NFF | 2023-02-23 13:22 | 2023-02-23 13:33 | Not reading oxygen or BP | (TC 02/23/2023 01:22:22 PM, Canjimavadivel, Kamallesh ) Tested both BP and SpO2, tested ok
1065309 | 193204 | CM | CANCL | CANCL | 2023-02-23 14:26 |  | Issue with battery and Power switch | (TC 02/23/2023 02:28:57 PM, Samanian, Kamran) power LED stays off even when the BAT charged. contacted WA (1FEB)and explained the problem/ NEW battery was sent . Ref# 8006381178. Charged the new battery, the same issue. (TC 02/23/2023 02:29:06 PM, Samanian, Kamran) spoke with tech and received RMA to send the unit back to repair under warranty. Attached is RMA.
1065309 | 193205 | CM | CMPLT | CMPLT | 2023-02-23 14:30 | 2023-03-31 07:56 | Issue with battery and Power switch | (TC 02/23/2023 02:30:44 PM, Samanian, Kamran) power LED stays off even when the BAT charged. contacted WA (1FEB)and explained the problem/ NEW battery was sent . Ref# 8006381178. Charged the new battery, the same issue. (TC 02/23/2023 02:31:04 PM, Samanian, Kamran) spoke with Tech. RMA is sent for repair under warranty. Attached is RMA (TC 02/23/2023 02:31:04 PM, Samanian, Kamran) spoke with Tech. RMA is sent for repair under warranty. Attached is RMA SC#1002891874 (TC 03/02/2023 07:50:11 AM, Samanian, Kamran) shipped by fedex , tracking#771385877002 (TC 03/24/2023 08:25:53 AM, Samanian, Kamran) Received the monitor after repair. power switch was replaced. repair under warranty attached. Being tested in biomed lab (TC 03/31/2023 07:55:50 AM, Samanian, Kamran) mounted the monitor in ENDO-1 on the floor.
1061316 | 193206 | CM | CMPLT | CMPLT | 2023-02-23 15:31 | 2023-02-23 15:31 | Pediatrician wants to install echopac on personal laptop | (TC 02/24/2023 04:04:00 PM, Fan, Tak) Research studies could not be read. Troubleshot with Amish and Faith. The studies are from various versions. Incompatible. Tried to install versions from Echo but GE would not provide licenses. (TC 02/27/2023 02:42:55 PM, Fan, Tak) Transfer issues. Sends but no analysis. Copied over from system from echo, same thing. Tried on echopac dongle from echo with 203 SW, worked but no strain analysis (not a purchased option on that dongle). Reverted SW back to 202.
1056938 | 193207 | CM | CMPLT | CMPLT | 2023-02-24 09:04 | 2023-02-24 09:05 | Alarming, not working properly | (TC 02/22/2023 09:04:22 AM, Rendon Lopez , Elizabeth) Tested for several hours, no fault found. Returned to RT NICU
32510 | 193208 | CM | CMPLT | CMPLT | 2023-02-24 09:11 | 2023-02-24 09:11 | MSH DR1 REPORTED MONITOR NOT BOOTING | (TC 02/23/2023 09:10:45 AM, Whittington, Scott) MSH DR1 REPORTED MONITOR NOT BOOTING SPARE MONITOR WAS INSTALLED | MSH DR1 REPORTED MONITOR NOT BOOTING
32510 | 193209 | CM | CMPLT | CMPLT | 2023-02-24 09:14 | 2023-02-28 10:23 | MSH DR2 TABLE NOT RESPONDING TO UP/DOW AND TABLE TOP MOVE CMDS | (TC 02/23/2023 09:13:36 AM, Whittington, Scott) MSH DR2 TABLE NOT RESPONDING TO UP/DOW AND TABLE TOP MOVE CMDS REBOOT CABINET, TABLE CONTROL RESPONDING FURTHER INVESTIGATION REQ (TC 02/23/2023 07:48:00 PM, Whittington, Scott) MSH DR2 TABLE NOT RESPONDING TO UP/DOW AND TABLE TOP MOVE CMDS INVESTIGATE, CORRECT | MSH DR2 TABLE NOT RESPONDING TO UP/DOW AND TABLE TOP MOVE CMDS
1061333 | 193210 | FC | CMPLT | CMPLT | 2023-02-24 09:15 | 2023-02-24 09:17 | QC checklist | (TC 02/24/2023 09:15:33 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
| 193211 | CM | CMPLT | CMPLT | 2023-02-24 10:21 | 2023-06-20 08:03 | Biomed test equipment calibration | (TC 02/24/2023 10:26:05 AM, Chukri, Samir) schedule for biomed test equipment calibration . Quotation # 2023MSH (TC 06/16/2023 03:13:07 PM, Chukri, Samir) working on the biomede tester list corrected and updatting the 2023 list (TC 06/20/2023 08:01:31 AM, Chukri, Samir) emailed to Ian the final list of Biomed testers certified. correct the PO final cost .
| 193212 | PA | CMPLT | CMPLT | 2023-02-24 13:18 | 2023-03-14 15:06 | Neonate BP cuffs
502088 | 193213 | RQ | CMPLT | CMPLT | 2023-02-24 13:30 | 2023-02-24 14:56 | message- contact service rep | (TC 02/24/2023 02:53:25 PM, Green, Paul) SLD1 persistent error due to uncommanded motion. Travelled to site and inspected. Was aboe to clear persistent error and no further problems. | gantry down
26152 | 193214 | CM | CMPLT | CMPLT | 2023-02-24 15:23 | 2023-02-24 15:25 | Testing for Use | (TC 02/24/2023 03:23:40 PM, Ferreira, Qwynn) Tested the Blender for use to ensure accuracy Tested the reid alarm Tested 21%,40%, 60%,80% and 100% All within 3%
26100 | 193215 | CM | CMPLT | CMPLT | 2023-02-24 15:25 | 2023-02-24 15:27 | Testing for Use | (TC 02/24/2023 03:25:35 PM, Ferreira, Qwynn) Tested the blender for use to ensure accuracy Tested the reid alarm' Tested at 21%, 40%, 60%, 80% and 100% All within 3%
| 193216 | CM | CMPLT | CMPLT | 2023-02-24 17:03 | 2023-02-24 17:03 | Room move |  | Room move
| 193217 | PA | CMPLT | CMPLT | 2023-02-27 07:55 | 2023-02-28 13:51 | Sterile Water
| 193218 | PA | CMPLT | CMPLT | 2023-02-27 08:07 | 2023-04-14 08:33 | AED training power adapter | (TC 02/27/2023 08:11:44 AM, Canjimavadivel, Kamallesh ) Placed order for training adapter (TC 04/14/2023 08:20:59 AM, Canjimavadivel, Kamallesh ) Received
1065316 | 193219 | CM | CMPLT | CMPLT | 2023-02-27 08:33 | 2023-02-27 08:36 | SPO2 sensor missing - provided | (TC 02/27/2023 08:32:56 AM, Samanian, Kamran) done!
| 193220 | CM | CMPLT | CMPLT | 2023-02-27 08:44 | 2023-04-03 10:20 | Masimo sensor and cable ORDER for Philips Monitors in ENDO | (TC 02/27/2023 08:30:05 AM, Samanian, Kamran) serached for the PN and also contacted Masimo (TC 02/27/2023 08:30:05 AM, Samanian, Kamran) serached for the PN and also contacted Masimo PO# 7202169508 (TC 04/03/2023 10:08:32 AM, Samanian, Kamran) received , checked and delivered to Anita Sensor PO# 7202183210
1061210 | 193221 | CM | CMPLT | CMPLT | 2023-02-27 10:23 | 2023-02-27 10:28 | Wall Monitor | (TC 02/27/2023 10:26:37 AM, Samanian, Kamran) this is the charge related to WO# 186529 which part was missing. According to samir since the month has already passed to charge back then this new WO# 193221 created to be able to charge back to OR. For Ref . quote & invoice attched
1056587 | 193222 | CM | CMPLT | CMPLT | 2023-02-27 10:39 | 2023-04-13 14:27 | UPGRADING SW TO 06.03.006.7404& reapir | (TC 02/24/2023 10:40:28 AM, Samanian, Kamran) STARTED TO UPGRADE THE SW IN AED BUT NOT FINISHED DUE TO LANGUAGE FILE CORRUPTED.OPENED THE CASENO IN ZOLL #01580391 AND PROVIDED INFORMATION. SOMEONE FROM ZOLL WILL CONTANT ME (TC 02/27/2023 10:44:46 AM, Samanian, Kamran) CONTACTED ZOLL AGAIN. UNIT NEEDS TO BE SENT OUT FOR REPAIR. SPOKE WITH ALLEN TO GET A LOANER AND RETURN THE DEFECTIVE ONE. (TC 02/28/2023 09:03:11 AM, Samanian, Kamran) Receive RMA# 980022 to ship the defective one to factory (TC 03/02/2023 09:11:01 AM, Samanian, Kamran) Received the loaner- SN# AX19B024967 ( under SR#980044)- Defective AED is sent to ZOLL for repair SC# 1002897287, No charge PO# (TC 03/08/2023 10:30:44 AM, Samanian, Kamran) Received email from ZOLL and updated with SR980022 (TC 04/13/2023 02:23:21 PM, Samanian, Kamran) Received the ZOLL AED after repair. Attched is the report. Had to load the configuration, Wifi config, root certificates two files and timing. AED passed the test. returned back to unit. and loaner one returned back to ZOLL.
531118 | 193223 | FC | CMPLT | CMPLT | 2023-02-27 10:43 | 2023-02-24 13:14 | Functional Check | (TC 02/24/2023 01:12:33 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 118/79 bpm 80 Set 80/40 bpm 50; measured 80/39 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
1058056 | 193224 | CM | CMPLT | CMPLT | 2023-02-27 11:16 | 2023-03-01 11:11 | sensing N2O when it's turning off | (TC 02/21/2023 10:54:42 AM, Scero, Tomlyn) Dr Teresi was wondering why when they intubated a patient in OR 19 (dental) using O2+N2O then switched to O2+ Air + Sevo, there was still N20 being detected by the gas analyzer up to 1 hr after it was discontinued from being used. I checked over the device and downloaded the logs. The device is working properly and the gas analyzers are working as they should. We think the device was acting normally, as it seems the case was being run in Auto-mode, which causes the patients exhaled breath to be recirculated and any unused gases to be rebreathed. Switching the mode to Freshgas mode until all N20 has been used or sent to scavenging would be ideal. | -high flows -sensing N2O when it's turned off after use, entire case
68370 | 193225 | CM | CMPLT | CMPLT | 2023-02-27 11:17 | 2023-03-02 10:38 | loud powersupply fan | (TC 02/22/2023 10:23:15 AM, Scero, Tomlyn) During PM I noticed the power supply fan was a little loud. I removed it, cleaned it and put it back in. Noticed it helped a little. I ordered one and it could take a while to be shipped. I'll replace it if it gets worse. shopping cart 1002890714 MK04951 $320 (TC 03/02/2023 10:37:28 AM, Scero, Tomlyn) fan arrived. AGM is currently in use, will put part in stock for future.
| 193226 | CM | CMPLT | CMPLT | 2023-02-27 12:48 | 2023-03-21 13:25 | Order Patient simulator PR-SL-8 | (TC 02/27/2023 12:57:19 PM, Chukri, Samir) Contacted MedSet and get a quotation , replaced order (TC 03/21/2023 01:26:01 PM, Chukri, Samir) Recevied the units.
531116 | 193227 | FC | CMPLT | CMPLT | 2023-02-27 13:20 | 2023-02-24 13:20 | Functional Check | (TC 02/24/2023 01:20:02 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 118/80 bpm 80 Set 80/40 bpm 50; measured 80/40 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
531121 | 193228 | FC | CMPLT | CMPLT | 2023-02-27 13:22 | 2023-02-24 13:23 | Functional Check | (TC 02/24/2023 01:23:19 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 119/81 bpm 80 Set 80/40 bpm 50; measured 82/39 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
531113 | 193229 | FC | CMPLT | CMPLT | 2023-02-27 13:26 | 2023-02-24 13:26 | Functional Check | (TC 02/24/2023 01:26:40 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 119/79 bpm 80 Set 80/40 bpm 50; measured 80/41 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
531117 | 193230 | FC | CMPLT | CMPLT | 2023-02-27 13:28 | 2023-02-24 13:28 | Functional check | (TC 02/24/2023 01:28:20 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 117/80 bpm 80 Set 80/40 bpm 50; measured 80/40 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
531120 | 193231 | CM | CMPLT | CMPLT | 2023-02-27 13:29 | 2023-02-24 13:29 | Functional check | (TC 02/24/2023 01:29:47 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 118/80 bpm 80 Set 80/40 bpm 50; measured 81/39 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
1063299 | 193232 | FC | CMPLT | CMPLT | 2023-02-27 13:33 | 2023-02-24 13:34 | Functional Check | (TC 02/24/2023 01:34:30 PM, Baghirzade, Nara) Physical inspection passed. Functional tests passed. Set weight 10lb, measured 10lb Set weight 20lb, measured 20lb No calibration needed. Scale is fully functional.
531119 | 193233 | FC | CMPLT | CMPLT | 2023-02-27 13:37 | 2023-02-24 13:52 | Functional Check | (TC 02/24/2023 01:52:44 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 120/79 bpm 80 Set 80/40 bpm 50; measured 80/39 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
531115 | 193234 | FC | CMPLT | CMPLT | 2023-02-27 13:54 | 2023-02-24 13:54 | Functional Check | (TC 02/24/2023 01:54:35 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 117/79 bpm 80 Set 80/40 bpm 50; measured 79/42 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
1063125 | 193235 | FC | CMPLT | CMPLT | 2023-02-27 13:56 | 2023-02-24 13:56 | Functional Check | (TC 02/24/2023 01:56:03 PM, Baghirzade, Nara) Device is in good physical condition. Set 120/80 bpm 80; measured 119/80 bpm 80 Set 80/40 bpm 50; measured 81/41 bpm 50 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
1063112 | 193236 | FC | CMPLT | CMPLT | 2023-02-27 13:57 | 2023-02-24 13:57 | Functional Check | (TC 02/27/2023 01:57:29 PM, Baghirzade, Nara) Physical inspection passed. Functional tests passed. Set weight 10lb, measured 10lb Set weight 20lb, measured 20lb No calibration needed. Scale is fully functional.
| 193237 | PA | CMPLT | CMPLT | 2023-02-27 14:02 | 2023-04-28 15:13 | Servo u/n membrane | (TC 02/27/2023 02:03:02 PM, Canjimavadivel, Kamallesh ) Placed order for membrane packs (TC 04/28/2023 03:12:58 PM, Canjimavadivel, Kamallesh ) Received packs
1056578 | 193238 | CM | CMPLT | CMPLT | 2023-02-27 14:18 | 2023-02-24 14:17 | Replaced battery | (TC 02/24/2023 02:19:03 PM, Baghirzade, Nara) Replaced battery with a spare one found in the AED cabinet. Battery is at 90% capacity. Verified the MSH network settings are correct. Manually checked in the device.
62659 | 193239 | CM | CMPLT | CMPLT | 2023-02-27 14:20 | 2023-02-27 14:21 | Display and measuring rod is unstable, wobbly | (TC 02/24/2023 02:19:55 PM, Baghirzade, Nara) The display was very loose on the mount; removed the display from mount and reinforced the mount and then tightened the display on to the mount. The measuring rod was loose and wobbling around; rod was tightened using the middle screw. Functional check of the scale passed. Placed 10 lbs, measured 10 lbs.
62658 | 193240 | CM | CMPLT | CMPLT | 2023-02-27 14:22 | 2023-02-24 14:22 | Display is not properly secured | (TC 02/24/2023 02:22:32 PM, Baghirzade, Nara) Secured the display on to the mount, tightened the screws. Functional check of scale passed. Set 10lbs, measured 10 lbs.
1047849 | 193241 | CM | CMPLT | CMPLT | 2023-02-27 14:28 | 2023-02-27 14:29 | Lever on the plunger was missing | (TC 02/27/2023 02:28:00 PM, Baghirzade, Nara) Replaced the lever on the plunger head. Claws are opening and closing without resistance. Functional tests passed. Ran infusion successfully. Occlusion alarm is functional. No errors during infusion. Pump is fully functional.
1047849 | 193242 | CM | CMPLT | CMPLT | 2023-02-27 14:30 | 2023-02-27 14:31 | Updated time, confirmed configuration | (TC 02/27/2023 02:30:14 PM, Baghirzade, Nara) Updated date and time. Confirmed correct configuration MARCH 2019.
1060030 | 193243 | CM | CMPLT | CMPLT | 2023-02-27 14:34 | 2023-02-27 14:37 | Replaced umbilical and repaired arm | (TC 02/27/2023 02:34:34 PM, Baghirzade, Nara) Worked with Samir. Replaced the umbilical/braided sleeve for the boydbox. The arm was very loose, Samir had taken it to biomed and had placed spaces, and greased the joint. Placed the arm back with the umbilical. Helium cell was calibrated. Calibrations passed. Bodybox is functional.
1051031 | 193244 | CM | CMPLT | CMPLT | 2023-02-27 14:47 | 2023-02-27 14:48 | System hung up, cannot resolve with restart | (TC 02/27/2023 09:47:00 AM, Fan, Tak) Users tried too send too many studies at once on the spooler. Spooler clogged. Manually removed all the studies from the spooler. Restarted, OK.
501584 | 193245 | CM | CMPLT | CMPLT | 2023-02-27 14:50 | 2023-02-27 14:50 | Keyboard cable guide broken. | (TC 02/27/2023 02:50:00 PM, Fan, Tak) Replaced keyboard guide.
1065226 | 193246 | CM | CMPLT | CMPLT | 2023-02-27 15:03 | 2023-11-16 15:14 | not working; warranty repair
67246 | 193247 | CM | CMPLT | CMPLT | 2023-02-28 06:32 | 2023-02-28 06:35 | Doppler and probe broken | (TC 02/28/2023 06:33:16 AM, Baghirzade, Nara) Doppler and probe are broken, and some of the electronic components are missing. Contacted Stat Medical to inquire about a repair; the cost of the repair is almost the same price as buying a new doppler. Spoke to Shruti and wants to proceed with purchasing a new doppler and probe to replace this. Will retire this asset. | Doppler was dropped and it resulted in a broken case, some electronic pieces have been lost, and the probe is broken as well. Will be ordering a new doppler and probe to replace this.
| 193248 | CM | CMPLT | CMPLT | 2023-02-28 06:38 | 2023-03-03 07:31 | Ordering doppler and 8MHz probe | (TC 02/28/2023 06:44:33 AM, Baghirzade, Nara) SC 1002894643 (TC 03/03/2023 07:30:10 AM, Baghirzade, Nara) PO 7202170020
1066105 | 193249 | CM | CMPLT | CMPLT | 2023-02-28 09:30 | 2023-03-01 10:20 | 6 YR PM | (TC 02/28/2023 10:19:06 AM, Scero, Tomlyn) 6 YEAR PM parts replaced. PM tests done. No leaks. Passed.
67959 | 193250 | CM | CMPLT | CMPLT | 2023-02-28 11:26 | 2023-02-28 11:34 | Latch Recall | (TC 02/28/2023 11:33:08 AM, Baghirzade, Nara) Completed the latch recall for portholes and NE/NW latches. Updated the latch tracking sheet on t drive.
1059778 | 193251 | FC | CMPLT | CMPLT | 2023-02-28 11:34 | 2023-02-28 11:34 | QC checklist | (TC 02/28/2023 11:34:03 AM, Baghirzade, Nara) QC checklist completed. Returned to 11S.
1053349 | 193252 | CM | CMPLT | CMPLT | 2023-02-28 11:59 | 2023-02-28 12:00 | MSH BI 1274 H2 DOWN TIME PACS | (TC 02/28/2023 12:00:18 PM, Whittington, Scott) MSH BI 1274 H2 DOWN TIME PACS | MSH BI 1274 H2 DOWN TIME PACS
502044 | 193253 | CM | CMPLT | CMPLT | 2023-02-28 15:00 | 2023-03-03 15:21 | Arm gap repair | (TC 02/28/2023 02:59:59 PM, Green, Paul) Performed arm gap repair as per procedure. Upon power up, 2 modules were out 52, 71. Opened detector, and upon inspection found pem nuts missing on 25, 89, 52, 38, 80, 71, 73, 65. Glued pem nuts on these modules, installed and retapped modules in place. Upon power up all modules ok and show green in prica. Performed daily QC however, and found no counts from a module. Troubleshot and then ordered 2 new modules from Digirad. Also, collimator ID is not functioning collimator detection is set= "Automatic". Troubleshot this problem and ordered a PSU board to try and resolve this. Put collimator detection = "manual" for now, and let the department know. (TC 03/02/2023 06:15:34 PM, Green, Paul) Hole in detector flood. Troubleshot with changing module positions. Could not get hole to move, despite moving various module positions. Fixed many pem nuts that had come off modules by epoxying them. Talked to Digirad support and determined that we should try an Energy calibration. This corrected the hole in the flood. Performed Uniformity corrections overnight. (TC 03/03/2023 07:49:17 AM, Green, Paul) Troubleshooting continued. Trying to isolate why there is a whole in the detector flood. Found new Energy map corrected for the whole in the flood. Created the new energy map, and then the uniformity maps were created. Had problems with the settings in setting up the uniformity map, as settings where changed, but when the aquisition was started, looked at settings again, and they were different than previously set. Tried several uniformity maps, and seems like there is a potential software bug. Figured out a work around and then ran another uniformity map over night. (TC 03/02/2023 07:49:17 AM, Green, Paul) Troubleshooting continued. Trying to isolate why there is a whole in the detector flood. Found new Energy map corrected for the whole in the flood. Created the new energy map, and then the uniformity maps were created. Had problems with the settings in setting up the uniformity map, as settings where changed, but when the aquisition was started, looked at settings again, and they were different than previously set. Tried several uniformity maps, and seems like there is a potential software bug. Figured out a work around and then ran another uniformity map over night. (TC 03/03/2023 03:20:41 PM, Green, Paul) resolved collimator ID issues and uniformity correction issues.
1056539 | 193254 | CM | CMPLT | CMPLT | 2023-02-28 15:10 | 2023-02-28 15:11 | Battery replacement | (TC 02/28/2023 03:10:51 PM, Canjimavadivel, Kamallesh ) Nigel replaced battery
50517 | 193255 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-03-31 11:58 | PM-MAJOR - Annual PM For: Vacuum Extractors, Obstetric | (TC 03/31/2023 11:57:55 AM, Rendon Lopez , Elizabeth) Performance verification-OK
52236 | 193256 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-12-13 12:45 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 12/08/2023 02:49:02 PM, Rhuland, Courtney) Replaced chassis Replaced hood Replaced humidifier (TC 12/13/2023 12:41:25 PM, Rhuland, Courtney) Replaced chipped inside door Replaced 2 porthole seals Completed latch recall Replaced rechargeable battery Replaced power cord PM complete Calibrated REF, main and line voltage Calibrated and verified scale Functional check - 35C, 75% humidity Read - 35C, 75.6% Electrical safety - 27mOhms, 60uA
52349 | 193257 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Bladder Scanner | Device no longer support. Replaced with newer model.
53076 | 193258 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-03-31 13:51 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 03/28/2023 11:27:11 AM, Ferreira, Qwynn) Replaced battery Cleaned all dirt and dust from compartment Removed suction regulator ordering batteries for the regulator Will continue the PM once parts are in (TC 03/31/2023 01:50:02 PM, Ferreira, Qwynn) Replaced the suction regulators battery and tested Replaced all the missing parts. Replaced two knobs Cleaned Funtional Test - OK Electrical Safety - OK
62640 | 193259 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | 2 year PM
62781 | 193260 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-07-20 10:13 | PM-MAJOR - Annual PM For: Vacuum Pumps, Central | (TC 07/19/2023 10:11:31 AM, Ferreira, Qwynn) PM Passed Replaced battery and labelled Tested the suction and working - OK Installed the device back in the transport
62915 | 193261 | PM | ACTIV | ISSUE | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
64059 | 193262 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-04-18 07:54 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 04/17/2023 07:53:43 AM, Rendon Lopez , Elizabeth) Replaced battery, cleaned, functions-OK
64073 | 193263 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-04-18 07:53 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 04/17/2023 07:47:03 AM, Rendon Lopez , Elizabeth) Battery replaced, cleaned, functions-OK
65118 | 193264 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-06-26 10:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/26/2023 10:45:50 AM, Reid, Emmerson ) PM Complete, see proceedure for details.
65263 | 193265 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-04-25 08:21 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 78uA and 85 ohm.
65883 | 193266 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-07-19 08:12 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/19/2023 08:11:56 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65887 | 193267 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-05-26 14:19 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/26/2023 02:18:38 PM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
65891 | 193268 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65902 | 193269 | PM | CNP | ENL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65948 | 193270 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-03-21 10:07 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2023 10:04:56 AM, Baghirzade, Nara) Pressure sensor was previously re-seated but error is recurring; replaced the pressure sensor. Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65985 | 193271 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-09-28 12:51 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/26/2023 08:33:50 AM, Campbell, Connor) PM complete/passed, refer to ASM
65993 | 193272 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-09-01 11:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/01/2023 11:38:35 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
66026 | 193273 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-06-08 09:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 09:01:16 AM, Rhuland, Courtney) Cracked rear and front case replaced (recall) Handle O-ring replaced PM complete All tests passed Refer to ASM for details
66035 | 193274 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-07-31 14:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 08:55:08 AM, Rhuland, Courtney) Rear and front cases replaced Battery replaced and conditioning in process (TC 07/31/2023 02:12:23 PM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
66064 | 193275 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-02 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 02:26:08 PM, Cortez, John) Cleaned unit, Downloaded error/event logs, "Backup speaker failure" error found in logs, PM completed successfully however, see ASM for results Conditioned battery overnight, Battery capacity: 4069 mAhs
66354 | 193276 | PM | ACTIV | ISSUE | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
66454 | 193277 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-04-24 08:17 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 75uA and 83 ohm.
66455 | 193278 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-05-05 14:02 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 73 uA and 85ohm.
66710 | 193279 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-25 08:27 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/25/2023 08:22:41 AM, Baghirzade, Nara) Originally sent down due to broken latch and missing multiple snap fasteners. Replaced latches and snap fasteners. PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C. Functional tests passed. Electrical safety results - Leakage current = 75.0 uA Power cord resistance = 115 mohms
66761 | 193280 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66816 | 193281 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | BP Check
66819 | 193282 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | BP Check
67376 | 193283 | PM | ACTIV | ISSUE | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67432 | 193284 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Vacuum Pumps, Central
67570 | 193285 | PM | ACTIV | ISSUE | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
67573 | 193286 | PM | ACTIV | ISSUE | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
67574 | 193287 | PM | ACTIV | ISSUE | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
67870 | 193288 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: External Power Supply Module | SSDU merged with ICU. UPS is not in use. PICIX not in use, stored on the storage cart on level 6.
67883 | 193289 | PM | CMPLT | UNL | 2023-03-01 02:26 | 2023-06-13 12:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | unable to locate
67921 | 193290 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-24 11:26 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/24/2023 10:46:22 AM, Rhuland, Courtney) Tag stated fan was loud During functional check, fan motor started making loud noise Replaced fan motor Replaced tilt platform plunger Replaced broken scale Replaced porthole seal Calibrated LV & Main voltage Scale calibrated Functional check - OK Electrical saftety - OK
67927 | 193291 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-05-24 10:39 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/24/2023 10:36:00 AM, Rhuland, Courtney) Cracked hood replaced Side door and porthole latches replaced (recall) Exam light replaced PM complete All tests passed Calibrated LV Calibrated scale Calibrated main voltage Electrical safety - OK
68034 | 193292 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-09-05 13:12 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/05/2023 01:11:23 PM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C. Functional tests passed. Electrical safety results - Leakage current = 86.5 uA Power cord resistance = 67 mohms
68376 | 193293 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-06-16 14:54 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 06/05/2023 10:47:10 AM, Scero, Tomlyn) 6 Year PM started. Brought AGM downstairs for PM. Printed paperwork and started filling it out. (TC 06/14/2023 02:23:54 PM, Scero, Tomlyn) All 6 year PM parts replaced. (TC 06/15/2023 02:14:10 PM, Scero, Tomlyn) Continued with PM proceedures (TC 06/16/2023 02:52:57 PM, Scero, Tomlyn) All PM test complete and passed. Results sheet attached and next PM schedule for 1 year.
68546 | 193294 | PM | CNP | EIU | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: Anesthesia Units | device in use
68587 | 193295 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-04-13 13:26 | PM-MAJOR-Annual PM for Servo-U Ventilator | (TC 04/13/2023 01:25:09 PM, Canjimavadivel, Kamallesh ) PM performed by vendor, see attached service report
68687 | 193296 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68716 | 193297 | PM | CNP | ENL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68724 | 193298 | PM | CNP | ENL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68727 | 193299 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68730 | 193300 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-03-10 14:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/10/2023 02:16:07 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 10 mohms Leakage current: 70.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68732 | 193301 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68750 | 193302 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-24 07:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 08:43:18 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 54 mohms Leakage current: 79.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/24/2023 07:36:23 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4000 mamphr PM completed.
68775 | 193303 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-05-16 13:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/04/2023 01:12:35 PM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4053 mAhrs
68777 | 193304 | PM | CNP | ENL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68792 | 193305 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-03 06:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 08:15:24 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 55 mohms Leakage current: 74.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/03/2023 06:35:24 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4136 mamphr PM completed.
68816 | 193306 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-15 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/15/2023 02:20:01 PM, Baghirzade, Nara) Battery was replaced. Electrical Safety - Ground resistance: 118 mohms Leakage current: 76.8 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. (TC 08/15/2023 02:20:01 PM, Baghirzade, Nara) Battery was replaced. Electrical Safety - Ground resistance: 42 mohms Leakage current: 65.9 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68825 | 193307 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68846 | 193308 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68853 | 193309 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68866 | 193310 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68868 | 193311 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68877 | 193312 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68886 | 193313 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68891 | 193314 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-24 08:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 02:54:50 PM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioning battery overnight (TC 08/23/2023 08:29:40 AM, Cortez, John) 3971 mAhs
68897 | 193315 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68908 | 193316 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-09-19 15:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 03:27:03 PM, Campbell, Connor) replaced the battery (scheduled) replaced the power cord (old one was tearing at the unit side plug) refer to procedure for ES test PM complete/passed, refer to ASM
68911 | 193317 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-04-05 12:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2023 12:38:21 PM, Cortez, John) Cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4054 mAh
68931 | 193318 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-09-01 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/29/2023 10:46:53 AM, Rendon Lopez , Elizabeth) Battery replaced, charging (TC 08/30/2023 10:46:53 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
68934 | 193319 | PM | CANCL | CANCL | 2023-03-01 02:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68942 | 193320 | PM | CNP | ENL | 2023-03-01 02:26 | 2024-03-25 10:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68950 | 193321 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-10-27 13:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/27/2023 01:27:27 PM, Baghirzade, Nara) PM passed and completed. Power cord was damaged; replaced power cord. Battery was replaced. Electrical Safety - Ground resistance: 58 mohms Leakage current: 70.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68956 | 193322 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-08-24 08:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 08:52:04 AM, Canjimavadivel, Kamallesh ) Replaced battery PM tests complete
68980 | 193323 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-07-19 10:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/19/2023 10:01:26 AM, Cortez, John) Cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4162 mAhs
68985 | 193324 | PM | CMPLT | CMPLT | 2023-03-01 02:26 | 2023-06-12 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/12/2023 02:13:53 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 129 mohms Leakage current: 75.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Female IUI replaced. Power cord strap replaced.
68987 | 193325 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69002 | 193326 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-10-18 08:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/18/2023 08:51:51 AM, Campbell, Connor) PM complete/passed, refer to ASM refer to procedure for ESA
69014 | 193327 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69023 | 193328 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-11-14 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/14/2023 10:33:13 AM, Campbell, Connor) PM complete/passed, refer to ASM refer to procedure for ESA conditioned battery
69026 | 193329 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69042 | 193330 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69058 | 193331 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69059 | 193332 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-10 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/10/2023 01:43:28 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69066 | 193333 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-24 09:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 09:11:20 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69078 | 193334 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69098 | 193335 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69141 | 193336 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69143 | 193337 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69148 | 193338 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69155 | 193339 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69157 | 193340 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69206 | 193341 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69237 | 193342 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-09-22 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/22/2023 02:55:06 PM, Campbell, Connor) Pm complete/passed, refer to ASM
69243 | 193343 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69249 | 193344 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-23 09:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 09:34:24 AM, Cortez, John) Cleaned, Downloaded error/event logs, Replaced LVP display and LED cover Calibrated infusion rate, PM completed successfully, results logged in ASM
69255 | 193345 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-07-31 14:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/31/2023 02:50:41 PM, Rhuland, Courtney) Rate and pressure calibrated PM complete All tests passed post calibration Refer to ASM for details
69310 | 193346 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-10-31 10:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/31/2023 09:51:42 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69321 | 193347 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69343 | 193348 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69349 | 193349 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-09-01 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 02:24:26 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69372 | 193350 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-23 13:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 12:59:52 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69378 | 193351 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-05-16 13:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2023 01:07:43 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69382 | 193352 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69404 | 193353 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69405 | 193354 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-11 15:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/11/2023 03:48:39 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69412 | 193355 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-09-08 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/08/2023 10:29:06 AM, Canjimavadivel, Kamallesh ) Cleaned pump Calibrated rate PM pass
69441 | 193356 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-05-26 10:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/26/2023 10:27:39 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
69500 | 193357 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-24 08:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 08:31:28 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69515 | 193358 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-06-09 15:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 03:00:46 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69529 | 193359 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-09-01 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 10:45:23 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69530 | 193360 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-09-01 11:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 11:13:14 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69532 | 193361 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-09-14 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/14/2023 01:40:27 PM, Cortez, John) Cleaned, Downloaded error/event logs, Patient-side pressure sensor errors were logged, Pressure sensor replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
69552 | 193362 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69560 | 193363 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-17 12:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/17/2023 12:26:58 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69566 | 193364 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-08 11:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/08/2023 11:21:04 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69599 | 193365 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69605 | 193366 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-07-28 13:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 01:40:36 PM, Rhuland, Courtney) Pressure and rate calibrated after failed rate accuracy test PM complete All tests passed post calibration Refer to ASM for details
69608 | 193367 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69614 | 193368 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-07-28 12:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 12:49:12 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69626 | 193369 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69632 | 193370 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-06-08 13:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:43:15 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69669 | 193371 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69672 | 193372 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69677 | 193373 | PM | CMPLT | PMFAIL | 2023-03-01 02:27 | 2023-09-07 09:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 09:23:05 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy, fails occlusion test, bottle pressure sensor error in logs. Waiting for part
68149 | 193374 | PM | ACTIV | ISSUE | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Testers
500846 | 193375 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-16 12:45 | MSH 5 US RM 3 APLIO XG PM - Semi-Annual | (TC 03/16/2023 12:45:10 PM, Whittington, Scott) MSH 5 US RM 3 APLIO XG PM - Semi-Annual | MSH 5 US RM 3 APLIO XG PM - Semi-Annual
500847 | 193376 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-28 08:35 | MSH US APLIO XG PM - Semi-Annual | (TC 03/23/2023 03:20:45 PM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual DFRG (TC 03/28/2023 08:34:51 AM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual | MSH US APLIO XG PM - Semi-Annual
500850 | 193377 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-19 16:08 | MSH IRUS "CLYDE" PM - Semi-Annual | (TC 04/18/2023 12:36:44 PM, Whittington, Scott) MSH IRUS "CLYDE" PM - Semi-Annual (TC 04/19/2023 04:08:45 PM, Whittington, Scott) MSH IRUS "CLYDE" PM - Semi-Annual | MSH IRUS "CLYDE" PM - Semi-Annual
32510 | 193378 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-24 11:02 | MSH DR1 PM - Semi-Annual | (TC 04/24/2023 11:02:05 AM, Whittington, Scott) MSH DR1 PM - Semi-Annual | MSH DR1 PM - Semi-Annual
32510 | 193379 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-24 11:01 | MSH DR1 HARP APRIL 21 TRGT- Semi Annual | (TC 04/21/2023 10:31:55 AM, Whittington, Scott) MSH DR1 HARP APRIL 21 TRGT- Semi Annual pre harp (TC 04/24/2023 11:00:41 AM, Whittington, Scott) MSH DR1 HARP APRIL 21 TRGT- Semi Annual | MSH DR1 HARP APRIL 21 TRGT- Semi Annual
32540 | 193380 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-11 15:08 | PM - Monthly
32560 | 193381 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-02 11:15 | PM - Monthly | PM completed Mechanical PM done NRA/AIR CAL-L FOV completed SC-1159404 TOS phantom a(-989.7,10.6) b(332.4,13.8) c(136.0,14.5) d(94.4,12.6) e(-106.8,12.3) f(0.7,13.7) -passed all phantom tests
32580 | 193382 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-02 11:20 | PM - Monthly | PM completed Mechanical PM done NRA/AIR CAL-L FOV completed SC-281692 TOS phantom a(-993.3,9.5) b(334.6,15.7) c(134.4,13.2) d(97.4,13.0) e(-105.8,14.6) f(0.9,13.8) -passed all phantom tests
500732 | 193383 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-19 11:33 | MSH-20-470 (Sports Clinic) PM- SCHEDULED SEMI-ANNUAL | (TC 04/18/2023 12:39:28 PM, Whittington, Scott) MSH-20-470 (Sports Clinic) PM- SCHEDULED SEMI-ANNUAL (TC 04/19/2023 11:33:30 AM, Whittington, Scott) MSH-20-470 (Sports Clinic) PM- SCHEDULED SEMI-ANNUAL | MSH-20-470 (Sports Clinic) PM- SCHEDULED SEMI-ANNUAL
500800 | 193384 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-20 14:02 | MSH BIUS PM - Semi-Annual | (TC 04/18/2023 12:37:21 PM, Whittington, Scott) MSH BIUS PM - Semi-Annual (TC 04/20/2023 02:02:25 PM, Whittington, Scott) MSH BIUS PM - Semi-Annual | MSH BIUS PM - Semi-Annual
500801 | 193385 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-19 14:45 | MSH BIUS PM - Semi-Annual | (TC 04/18/2023 12:37:43 PM, Whittington, Scott) MSH BIUS PM - Semi-Annual (TC 04/19/2023 02:45:03 PM, Whittington, Scott) MSH BIUS PM - Semi-Annual | MSH BIUS PM - Semi-Annual
500802 | 193386 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-19 14:47 | MSH BIUS PM - Semi-Annual | (TC 04/18/2023 12:38:07 PM, Whittington, Scott) MSH BIUS PM - Semi-Annual (TC 04/19/2023 02:46:55 PM, Whittington, Scott) MSH BIUS PM - Semi-Annual | MSH BIUS PM - Semi-Annual
500857 | 193387 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-28 08:35 | MSH US APLIO XG PM - Semi-Annual | (TC 03/23/2023 03:26:28 PM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual DFRG (TC 03/28/2023 08:35:36 AM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual | MSH US APLIO XG PM - Semi-Annual
500858 | 193388 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-28 08:36 | MSH US APLIO XG PM - Semi-Annual | (TC 03/23/2023 03:25:01 PM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual (TC 03/28/2023 08:36:09 AM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual | MSH US APLIO XG PM - Semi-Annual
500868 | 193389 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-27 10:37 | MSH US APLIO XG PM - Semi-Annual | (TC 03/23/2023 03:21:48 PM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual (TC 03/27/2023 10:37:06 AM, Whittington, Scott) PM COMPLETED TRACKBALL FOAMS REPLACED | MSH US APLIO XG PM - Semi-Annual
500869 | 193390 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-28 08:36 | MSH US APLIO XG PM - Semi-Annual | (TC 03/23/2023 03:22:59 PM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual HDD TOO FULL TO COMPLETE DFRG HDD REPLACEMENT NECESSARY (TC 03/28/2023 08:36:42 AM, Whittington, Scott) MSH US APLIO XG PM - Semi-Annual | MSH US APLIO XG PM - Semi-Annual
501647 | 193391 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-07-07 11:10 | PM - Annual | (TC 07/07/2023 11:09:09 AM, Green, Paul) Not used and only used for backup.
501888 | 193392 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-28 08:38 | MSH 5 US APLIO 500 PM - Semi-Annual | (TC 03/24/2023 01:49:37 PM, Whittington, Scott) MSH 5 US APLIO 500 PM - Semi-Annual | MSH 5 US APLIO 500 PM - Semi-Annual
1046035 | 193393 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-24 12:55 | MSH ER DRX EVO HARP APR 26 TRGT - Semi Annual |  | MSH ER DRX EVO HARP APR 26 TRGT - Semi Annual Harp performed.
1046035 | 193394 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-25 10:55 | PM Major | (TC 04/25/2023 10:54:49 AM, Chau, Ken) PM Okay
1046423 | 193395 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Smoke Evacuation Systems, Surgical
1046448 | 193396 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-31 11:57 | Annual PM | (TC 03/30/2023 11:56:25 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached results.
1046449 | 193397 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-31 11:55 | Annual PM | (TC 03/30/2023 11:55:11 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached results.
1046450 | 193398 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-31 11:54 | Annual PM | (TC 03/30/2023 11:52:37 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
1046569 | 193399 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR _ Annual PM for: smoke evacuation, surgical
1047036 | 193400 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-04 08:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/04/2023 08:04:26 AM, Rhuland, Courtney) Device cleaned PM complete All tests passed Refer to ASM for details
1047830 | 193401 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
1048524 | 193402 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-27 09:00 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/27/2023 08:59:39 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1048529 | 193403 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1048545 | 193404 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-18 06:29 | PM-MAJOR - Annual PM For: Insufflators | (TC 08/18/2023 06:27:14 AM, Reid, Emmerson ) Ran the self test. Ran and checked for occlusion. Did physical inspection. PM Complete.
1048598 | 193405 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048599 | 193406 | PM | ACTIV | ISSUE | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048600 | 193407 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048601 | 193408 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048604 | 193409 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048605 | 193410 | PM | ACTIV | ISSUE | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048628 | 193411 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048630 | 193412 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048631 | 193413 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048633 | 193414 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048635 | 193415 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048638 | 193416 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048640 | 193417 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048644 | 193418 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048657 | 193419 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048661 | 193420 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048665 | 193421 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048672 | 193422 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048673 | 193423 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048675 | 193424 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048680 | 193425 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048683 | 193426 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048684 | 193427 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048685 | 193428 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048689 | 193429 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048696 | 193430 | PM | CANCL | CANCL | 2023-03-01 02:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050250 | 193431 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-09-06 15:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 03:18:23 PM, Campbell, Connor) PM complete, refer to ASM battery replaced power cord replaced
1050252 | 193432 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-07-31 10:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/31/2023 09:58:37 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 80 mohms Leakage current: 63.8 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
1050577 | 193433 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-17 10:56 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | (TC 03/08/2023 01:13:15 PM, Rendon Lopez , Elizabeth) Packed to send out for PM. RA # S2023-1283 (TC 04/13/2023 10:55:13 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK. See attached service report
1050578 | 193434 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-17 10:57 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 03/08/2023 09:03:17 AM, Rendon Lopez , Elizabeth) Sent out for PM (TC 04/13/2023 10:56:45 AM, Rendon Lopez , Elizabeth) Incoming inspection with vent 1050577.
1050623 | 193435 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-07-07 13:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/06/2023 12:56:54 PM, Rhuland, Courtney) Rear case replaced (TC 07/07/2023 01:45:16 PM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
1053349 | 193436 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-19 17:13 | MSH BI 1274 H2 HARP APRIL 19 TRGT SEMI-ANNUAL | (TC 04/19/2023 05:07:00 PM, Whittington, Scott) MSH BI 1274 H2 HARP APRIL 19 TRGT SEMI-ANNUAL (TC 04/19/2023 05:13:02 PM, Dolinar, John) MSH BI 1274 H2 HARP APRIL 19 TRGT SEMI-ANNUAL | MSH BI 1274 H2 HARP APRIL 19 TRGT SEMI-ANNUAL HARP completed.
1053381 | 193437 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-21 09:04 | PM-MAJOR - Annual PM For: Ventilators | performed annual PM. Passed pre-use check. Functional. | 2 years warranty covers PMs and repairs
1053452 | 193438 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-06-13 14:26 | PM MAJOR-Stretchers, Mobile, Hospital, Patient Transfer | (TC 06/13/2023 02:24:07 PM, Chukri, Samir) Completed PM by Stryker report attached. (TC 06/13/2023 02:30:57 PM, Ferreira, Qwynn) supervised and assisted the tech in performing the PM
1054411 | 193439 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-04-26 13:46 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 04/26/2023 01:45:07 PM, Canjimavadivel, Kamallesh ) Verified accuracy, calibration not needed
1054448 | 193440 | PM | CNP | ENL | 2023-03-01 02:27 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 08/30/2023 11:37:05 AM, Cortez, John) see WO 186102 | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054460 | 193441 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-10 09:29 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 03/10/2023 09:24:29 AM, Canjimavadivel, Kamallesh ) Replaced pressure chambers as per recall Replaced and lubricated o-rings All tests pass PM complete
1054462 | 193442 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-10 09:01 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 03/09/2023 09:00:35 AM, Cortez, John) Replaced pressure chambers as per recall Replaced o-rings All tests pass PM complete
1054464 | 193443 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-08-31 14:16 | PM-MAJOR- Factory PM for: ICU Bed Elecric |  | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054465 | 193444 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-11-22 14:27 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 11/22/2023 02:26:29 PM, Cortez, John) Bed+Intellidrive batteries were replaced 60 mOhms, 21 uAmps Pendant missing clip PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054466 | 193445 | PM | CMPLT | PMFAIL | 2023-03-01 02:27 | 2023-11-21 13:25 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 11/21/2023 01:12:40 PM, Cortez, John) Bed+Intellidrive batteries were replaced 70 mOhms, 25 uAmps Pendant missing clip PM failed, buttons on bottom right rail not working during chair/flat test | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054467 | 193446 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-10-13 09:06 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 10/12/2023 09:03:31 AM, Cortez, John) Bed+Intellidrive batteries were replaced 90 mOhms, 23 uAmps Right head side rail slightly bent PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054468 | 193447 | PM | CMPLT | CMPLT | 2023-03-01 02:27 | 2023-03-09 14:34 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 03/09/2023 02:33:34 PM, Canjimavadivel, Kamallesh ) Replaced pressure chambers as per recall Replaced o-rings All tests pass PM complete
1054469 | 193448 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-11-09 13:49 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 11/09/2023 01:47:37 PM, Cortez, John) Bed+Intellidrive batteries were replaced 70 mOhms, 25 uAmps Pendant missing PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054470 | 193449 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-08-15 13:34 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 08/15/2023 01:31:13 PM, Cortez, John) Replaced bed+intellidrive batteries 70 mOhms, 26 uAmps Strain relief missing for AUX power cord Pendant without clip PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054471 | 193450 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-11-08 16:08 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 11/08/2023 04:07:32 PM, Cortez, John) Bed+Intellidrive batteries were replaced 70 mOhms, 24 uAmps Pendant missing clip PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054473 | 193451 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-11-28 16:37 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 11/28/2023 04:35:59 PM, Cortez, John) Bed+Intellidrive batteries were replaced 80 mOhms, 22 uAmps Pendant missing PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054474 | 193452 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-08-16 15:22 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 08/16/2023 03:13:09 PM, Cortez, John) Bed+Intellidrive batteries were replaced 90 mOhms, 22 uAmps Headboard, pendant missing, Cover for hose on bed frame dislodged, Left foot break cover missing, Head angle display reading slightly off by a couple of degrees, attempted to calibrate twice, but no change (TC 08/16/2023 03:13:09 PM, Cortez, John) Bed+Intellidrive batteries were replaced 90 mOhms, 22 uAmps Headboard, pendant missing, Cover for hose on bed frame dislodged, Left foot break cover missing, Head angle display reading slightly off by a couple of degrees, attempted to calibrate twice, but no change | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054475 | 193453 | PM | CMPLT | PMFAIL | 2023-03-01 02:28 | 2023-08-15 13:28 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 08/15/2023 01:23:51 PM, Cortez, John) Replaced bed+intellidrive batteries Right head rail, bent Pendant physical damage 60 mOhms, 25 uAmps Bed unable to lower completely due to bent siderail. Will place service request | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054479 | 193454 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-09-29 15:46 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/28/2023 03:41:22 PM, Cortez, John) Bed+Intellidrive batteries were replaced 60 mOhms, 24 uAmps Pendant missing clip, headboard missing Right head rail slightly bent, will monitor PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1054537 | 193455 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-07-10 09:22 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 03/13/2023 08:06:26 AM, Canjimavadivel, Kamallesh ) Placed order for PM kit (TC 04/28/2023 07:52:23 AM, Canjimavadivel, Kamallesh ) Reached out to vendor for ETA of PM kit, earliest delivery is end of May (TC 07/10/2023 09:20:27 AM, Canjimavadivel, Kamallesh ) Replaced filters knob rings from PM kit. Completed PM with Courtney
1054541 | 193456 | PM | ACTIV | ISSUE | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
1054806 | 193457 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-05-18 12:56 | PM-MAJOR - Annual PM For: Lifts, Patient Transfer | (TC 05/18/2023 12:55:21 PM, Canjimavadivel, Kamallesh ) Functional tests pass
1055122 | 193458 | PM | CNP | ENL | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055178 | 193459 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-09-07 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 09:21:22 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
1055642 | 193460 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-09-26 09:07 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/25/2023 09:06:30 AM, Cortez, John) Bed+Intellidrive batteries were replaced 80 mOhms, 21 uAmps Pendant missing clip PM passed (TC 09/25/2023 09:11:55 AM, Cortez, John) Side panel on intellibridge housing cracked Will look to swap with spare bed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1055643 | 193461 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-09-26 09:21 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/25/2023 09:20:38 AM, Cortez, John) Bed+Intellidrive batteries were replaced 60 mOhms, 23 uAmps Pendant missing clip PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1055644 | 193462 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-12-28 09:42 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 12/28/2023 09:40:41 AM, Canjimavadivel, Kamallesh ) PM complete, see procedure | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1055645 | 193463 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-08-15 15:39 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 08/15/2023 03:36:50 PM, Cortez, John) Replaced bed+intellidrive batteries 70 mOhms, 24 uAmps Missing headboard Missing pendant clip | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1055646 | 193464 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-09-27 15:48 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/27/2023 03:45:30 PM, Cortez, John) Bed+Intellidrive batteries were replaced 50 mOhms, 24 uAmps Pendant missing clip PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1055647 | 193465 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-10-13 15:34 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 10/13/2023 03:32:47 PM, Cortez, John) Bed+Intellidrive batteries were replaced 100 mOhms, 21 uAmps Pendant missing clip PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1057846 | 193466 | PM | ACTIV | ISSUE | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Insufflators
1057847 | 193467 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-11-03 09:06 | PM-MAJOR - Annual PM For: Insufflators | (TC 11/03/2023 09:05:29 AM, Reid, Emmerson ) Turned on and tested flow and occlusion. PM Complete.
1057892 | 193468 | PM | CANCL | CANCL | 2023-03-01 02:28 |  | Lubrication cycle
1058180 | 193469 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2024-05-06 10:08 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058235 | 193470 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-04-17 11:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2023 11:41:33 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1058242 | 193471 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-08-09 09:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 09:57:59 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1058248 | 193472 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-07-27 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/27/2023 11:23:14 AM, Rhuland, Courtney) During PM rate test failed During calibration, pressure calibration failed Patient side pressure sensor replaced Pressure and rate calibrated PM complete All tests passed post calibration Refer to ASM for details
1058261 | 193473 | PM | CANCL | CANCL | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058275 | 193474 | PM | CNP | ENL | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059524 | 193475 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-04-17 10:58 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 03/08/2023 09:03:57 AM, Rendon Lopez , Elizabeth) Sent out for PM (TC 04/13/2023 10:57:41 AM, Rendon Lopez , Elizabeth) Incoming inspection with vent 1050577.
1059701 | 193476 | PM | ACTIV | ISSUE | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1059703 | 193477 | PM | ACTIV | ISSUE | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1059706 | 193478 | PM | ACTIV | ISSUE | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1059869 | 193479 | PM | ACTIV | ISSUE | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
1059874 | 193480 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-06-08 09:00 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/08/2023 08:35:11 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1059875 | 193481 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-09-28 08:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/28/2023 08:39:18 AM, Campbell, Connor) PM complete/passed, refer to ASM
1059876 | 193482 | PM | CNP | ENL | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1059878 | 193483 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-05-24 13:28 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/24/2023 12:56:20 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
1059879 | 193484 | PM | CNP | ENL | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1059883 | 193485 | PM | CNP | ENL | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059884 | 193486 | PM | CNP | ENL | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059944 | 193487 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-08-25 18:45 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
1060030 | 193488 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-03-13 14:44 | PM | (TC 03/13/2023 02:44:00 PM, Baghirzade, Nara) PM completed by vendor (Medisoft) Smain Amoura assisted by distributor (Novus Medical). (TC 03/29/2023 08:38:59 AM, Baghirzade, Nara) Attached service report
1060031 | 193489 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-03-14 12:29 | PM | (TC 03/14/2023 12:29:21 PM, Baghirzade, Nara) PM completed by vendor (Medisoft) Smain Amoura assisted by distributor (Novus Medical). (TC 03/29/2023 08:35:41 AM, Baghirzade, Nara) Attached service report
1060032 | 193490 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-03-22 12:31 | PM | (TC 03/23/2023 12:31:15 PM, Baghirzade, Nara) PM completed by vendor (Medisoft) Smain Amoura assisted by distributor (Novus Medical). (TC 03/29/2023 08:22:21 AM, Baghirzade, Nara) Attached service report
1060164 | 193491 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-03-01 02:28 | Sonesta 3 Exam Table
1060325 | 193492 | PM | ACTIV | READY | 2023-03-01 02:28 |  | Annual PM
1063193 | 193493 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-04-24 12:32 | MSH 5 RPCU BONNIE PM - Semi-Annual | (TC 04/18/2023 12:38:50 PM, Whittington, Scott) MSH 5 RPCU BONNIE PM - Semi-Annual (TC 04/21/2023 04:12:51 PM, Whittington, Scott) MSH 5 RPCU BONNIE PM - Semi-Annual | MSH 5 RPCU BONNIE PM - Semi-Annual
1063372 | 193494 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-09-29 15:37 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/29/2023 03:31:27 PM, Cortez, John) Bed+Intellidrive batteries were replaced 100 mOhms, 22 uAmps Pendant/headboard missing Power cord strap missing, replaced PM passed | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1064964 | 193495 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-04-14 16:05 | HARP | HARP
1064964 | 193496 | PM | CMPLT | CMPLT | 2023-03-01 02:28 | 2023-04-13 16:06 | PM - Semi-Annual | PM
1065463 | 193497 | PM | ACTIV | ISSUE | 2023-03-01 02:28 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
67961 | 193498 | FC | CANCL | CANCL | 2023-03-01 08:51 |  | Issue not specified; functional check
502492 | 193499 | CM | CMPLT | CMPLT | 2023-03-01 09:28 | 2023-03-01 09:32 | Critical Error, SymCode FAIL, Errnum -2506 | (TC 03/01/2023 09:28:35 AM, Green, Paul) Talked with Ona and she could not clear error, even after a shutdown. Travelled to site and inspected error. DioAppDockSensor Errnum -2506. Manually advanced the plungers and then loaded the syringes. Now able to move plungers, perform protocol and rebooted. Error did not return. Ona will let us know if problem returns.
1063709 | 193500 | CM | CMPLT | CMPLT | 2023-03-01 09:35 | 2023-03-01 16:19 | MSH OR SPARQ PM BY PHILIPS | (TC 03/01/2023 09:36:45 AM, Whittington, Scott) MSH OR SPARQ PM BY PHILIPS (TC 03/01/2023 09:36:45 AM, Whittington, Scott) MSH OR SPARQ PM BY PHILIPS | MSH OR SPARQ PM BY PHILIPS
1061333 | 193501 | FC | CMPLT | CMPLT | 2023-03-01 09:39 | 2023-03-01 09:40 | QC checklist | (TC 03/01/2023 09:38:50 AM, Baghirzade, Nara) QC checklist completed. Returned to 11N.
1066107 | 193502 | ACC | CMPLT | CMPLT | 2023-03-01 09:40 | 2023-03-22 12:25 | MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE | (TC 03/01/2023 09:40:06 AM, Whittington, Scott) MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE (TC 03/01/2023 09:40:06 AM, Whittington, Scott) MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE (TC 03/02/2023 04:12:58 PM, Whittington, Scott) MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE TONY CANTUTO IN TO SETUP (TC 03/20/2023 09:55:41 AM, Whittington, Scott) MARCH 17 MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE CLINICAL APPLICATIONS (TC 03/22/2023 12:24:14 PM, Whittington, Scott) MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE DELIVER TO CELINE ALLEN , LOCK UP ON 10TH FLORR CCU (TC 03/22/2023 12:24:49 PM, Dolinar, John) MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE DELIVER TO CELINE ALLEN , LOCK UP ON 10TH FLORR CCU (TC 03/27/2023 01:49:12 PM, Whittington, Scott) MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE ECG ADD ON DELIVERED | MSH GIM ULTRASOUND GE VENUE 40 ACCEPTANCE 132 mOHM 49 uA MFG FEB 2023
1062779 | 193503 | CM | CMPLT | CMPLT | 2023-03-01 11:08 | 2023-03-02 12:40 | O2 cell failed during calibration | (TC 03/01/2023 11:07:41 AM, Baghirzade, Nara) When the PFT tech did calibrations, the O2 sensor was not providing any readings or graphs. It was failing immediately. Even when using different gas tanks, it was producing the same results. Samir and I connected with Russ from Novus and Smain from Medisoft; they both remoted to the system and conducted tests which confirmed that the O2 cell was highly saturated and not decreasing in mV. O2 cell is not functional. Will be ordering a new O2 cell. (TC 03/02/2023 12:39:01 PM, Baghirzade, Nara) O2 cell was replaced by Samir and I. Russ from Novus remoted in and Smain from medisoft; Smain walked Samir and I on how to calibrate the O2 cell and we have also slightly adjusted the CO2 cell. All calibration tests are successful and the O2 cell is functional.
1062779 | 193504 | CM | CMPLT | CMPLT | 2023-03-01 11:13 | 2023-03-02 10:18 | Ordering O2 Cell | (TC 03/01/2023 11:16:47 AM, Baghirzade, Nara) SC 1002896348 (TC 03/02/2023 10:16:21 AM, Baghirzade, Nara) PO 7202170889
1057521 | 193505 | CM | CMPLT | CMPLT | 2023-03-01 11:39 | 2023-03-01 11:39 | upgraded software to N.00.04-108 | (TC 03/01/2023 11:39:01 AM, Samanian, Kamran) done!
1057004 | 193506 | CM | CMPLT | CMPLT | 2023-03-01 12:14 | 2023-03-01 12:15 | upgrade to Rev n.00.04-108 | (TC 03/01/2023 12:14:49 PM, Samanian, Kamran) done! after a few tries
1056125 | 193507 | PM | CMPLT | CMPLT | 2023-03-01 13:57 | 2023-03-01 13:59 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/01/2023 01:57:40 PM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
1057006 | 193508 | CM | CMPLT | CMPLT | 2023-03-01 14:49 | 2023-03-01 14:51 | upgrade to Rev N.00.04-108 | (TC 03/01/2023 02:49:44 PM, Samanian, Kamran) upgraded
1047862 | 193509 | CM | CMPLT | CMPLT | 2023-03-01 14:53 | 2023-03-01 14:54 | Functional check/updating time/date/config | (TC 03/01/2023 02:52:47 PM, Baghirzade, Nara) Confirmed the latest configuration (MARCH 2019) is loaded Updated date and time. Ran a basic infusion; completed successfully. Occlusion alarm is functional. No errors noted.
| 193510 | CM | CMPLT | CMPLT | 2023-03-01 14:53 | 2023-03-01 14:52 | capital meeting | (TC 03/01/2023 02:52:43 PM, Samanian, Kamran) done!
1047841 | 193511 | CM | CMPLT | CMPLT | 2023-03-01 14:54 | 2023-03-01 14:55 | Functional check/updating time/date/config | (TC 03/01/2023 02:54:09 PM, Baghirzade, Nara) Confirmed the latest configuration (MARCH 2019) is loaded Updated date and time. Ran a basic infusion; completed successfully. Occlusion alarm is functional. No errors noted.
66106 | 193512 | CM | CMPLT | CMPLT | 2023-03-01 14:55 | 2023-03-01 14:56 | Loose connector and QC | (TC 03/01/2023 02:55:07 PM, Baghirzade, Nara) Julie sent down VACs with loose connectors. Power connector was loose. Samir and I opened the VAC, tightened the connector. E-mailed Julie to update her that connector issue was resolved and is safe for patient use. QC checklist completed. Returned to 14 South.
65381 | 193513 | CM | CMPLT | CMPLT | 2023-03-01 14:57 | 2023-03-01 14:57 | Loose connector/QC | (TC 03/01/2023 02:56:44 PM, Baghirzade, Nara) Julie sent down VACs with loose connectors. Power connector was loose. Opened the VAC and tightened the connector. E-mailed Julie to update her that connector issue was resolved and is safe for patient use. QC checklist completed. Returned to 14 South.
67961 | 193514 | CM | CMPLT | CMPLT | 2023-03-02 07:33 | 2023-03-02 07:34 | Latch Recall | (TC 03/02/2023 07:33:35 AM, Baghirzade, Nara) Replaced all porthole and NW/NE latches on the incubator hood.
1057528 | 193515 | CM | CMPLT | CMPLT | 2023-03-02 08:07 | 2023-03-01 08:08 | upgraded to Rev. N.00.04-108 | (TC 03/01/2023 08:07:54 AM, Samanian, Kamran) done!
1066114 | 193516 | ACC | CMPLT | CMPLT | 2023-03-02 09:15 | 2023-03-02 09:20 | Incoming inspection | (TC 03/02/2023 09:18:56 AM, Baghirzade, Nara) Incoming inspection completed; asset tagged. QC checklist completed.
1066115 | 193517 | ACC | CMPLT | CMPLT | 2023-03-02 09:20 | 2023-03-02 09:21 | Incoming inspection | (TC 03/02/2023 09:20:20 AM, Baghirzade, Nara) Incoming inspection completed; asset tagged. QC checklist completed.
1066116 | 193518 | CM | CMPLT | CMPLT | 2023-03-02 09:22 | 2023-03-02 09:22 | received as loaner to send back the defective one ( 1056587)
67136 | 193519 | CM | CMPLT | CMPLT | 2023-03-02 09:22 | 2023-03-02 09:30 | Cuff not inflating | (TC 03/02/2023 09:22:39 AM, Canjimavadivel, Kamallesh ) Found leak in cuff during testing. Replaced cuff, working fine
1053381 | 193520 | CM | CMPLT | CMPLT | 2023-03-02 09:23 | 2023-03-21 09:10 | Ordering parts for Servo-N vents | received parts. PO#7202171536
1060099 | 193521 | CM | CMPLT | CMPLT | 2023-03-02 09:29 | 2023-03-02 09:33 | Not booting properly; CMOS battery replacement | (TC 03/02/2023 09:29:33 AM, Baghirzade, Nara) BIOS settings were not being retained after each boot up; boot up was also not occurring properly. Replaced the CMOS battery. The old battery had calcification and some residue buildup which could have prevented proper contact between the battery and the contact. PC booted properly, time and date are correct, and was able to get to the Xltek menu.
67148 | 193522 | CM | CMPLT | CMPLT | 2023-03-02 09:56 | 2023-03-02 10:09 | Thermometer doesn't work | (TC 03/02/2023 10:08:48 AM, Canjimavadivel, Kamallesh ) Faulty probe, replaced probe, tested fine
1047852 | 193523 | CM | CMPLT | CMPLT | 2023-03-02 10:02 | 2023-03-02 10:04 | Updated date/time and config | (TC 03/02/2023 10:02:22 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
501065 | 193524 | CM | CMPLT | CMPLT | 2023-03-02 11:22 | 2023-03-08 07:58 | No main display | (TC 03/02/2023 11:22:00 AM, Fan, Tak) Confirmed issue. Swapped monitors, display fine on PC monitor. Swapped cables, cables are fine. Contacted Philips for replacement monitor. REF 0121304065. (TC 03/07/2023 07:57:19 AM, Fan, Tak) Replaced main display.
1047853 | 193525 | CM | CMPLT | CMPLT | 2023-03-02 12:45 | 2023-03-02 14:53 | Updated time/date and config | (TC 03/02/2023 02:52:20 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
| 193526 | CM | CMPLT | CMPLT | 2023-03-02 15:00 | 2023-03-02 15:02 | Meeting with Andrew and Henna | (TC 03/02/2023 03:00:38 PM, Baghirzade, Nara) Meeting with Henna and Andrew regarding Ophthalmology lasers (TC 03/02/2023 03:02:10 PM, Baghirzade, Nara) Meeting with Andrew and Henna regarding Ophthalmology lasers
| 193527 | CM | CMPLT | CMPLT | 2023-03-02 15:03 | 2023-03-02 15:06 | Scales and Ophthalmoloscopes check | (TC 03/02/2023 03:03:06 PM, Baghirzade, Nara) Did a functional check of all infant scales in family medicine as Mara and Sally were complaining about weight inaccuracies. The scale in Triage 1 is not accurate and has broken parts; it was moved to exam room 10. The other 2 scales are functional but old. Andrew also spoke to Mara about potentially replacing the scales either through Capital or their operational budget. Will discuss further once Andrew confirms. (TC 03/02/2023 03:05:04 PM, Baghirzade, Nara) Swapped out the oto-ophthalmoscope with the spare one Mara had. The current oto-ophthalmoscope will not turn on even after changing the bulbs, outlet, and checking the cables. Will open it up and see if I can fix it otherwise I let Mara know that she may have to purchase a new unit.
| 193528 | CM | CMPLT | CMPLT | 2023-03-02 15:48 | 2023-03-02 15:48 | smitha and nephew initial accept forOR | (TC 03/02/2023 03:48:34 PM, Samanian, Kamran) done!
1066118 | 193529 | ACC | CMPLT | CMPLT | 2023-03-03 07:34 | 2023-03-03 07:42 | Incoming inspection and asset tagging | (TC 03/03/2023 07:41:36 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Doppler is functional.
1066120 | 193530 | CM | CMPLT | CMPLT | 2023-03-03 10:07 | 2023-03-03 10:09 | Frayed probe wire | (TC 03/03/2023 10:07:26 AM, Baghirzade, Nara) Probe cable was frayed; replaced the cable. Device was not in TMS, asset tagged and added to TMS.
1047845 | 193531 | CM | CMPLT | CMPLT | 2023-03-03 10:16 | 2023-03-03 10:17 | Update date/time and config | (TC 03/03/2023 10:16:52 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
1047867 | 193532 | CM | CMPLT | CMPLT | 2023-03-03 10:17 | 2023-03-03 10:18 | Updated date/time and config | (TC 03/03/2023 10:17:33 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional. Front display was missing the orange cover; orange display cover was placed.
67959 | 193533 | CM | CMPLT | CMPLT | 2023-03-03 10:22 | 2023-03-03 10:27 | No issue stated on tag | (TC 03/03/2023 10:22:36 AM, Baghirzade, Nara) Noticed that the hood was breaking off on both sides; Qwynn and I glued the hood with UV10 med resin. Allowed 48 hours to cure and dry. Functional tests passed. Set 35C temp and 75%, humidity; measured 35C temp and between 74%-75% humidity. No errors noted. No fault found. Incubator is fully functional.
| 193534 | CM | CMPLT | CMPLT | 2023-03-03 10:40 | 2023-03-20 07:41 | Ordering lifedop doppler cables for biomed stock | (TC 03/15/2023 12:47:37 PM, Baghirzade, Nara) SC 1002909638 (TC 03/20/2023 07:40:33 AM, Baghirzade, Nara) PO 7202179747
| 193535 | CM | CMPLT | CMPLT | 2023-03-03 13:03 | 2023-03-02 20:00 | RDD Error | RDD Error -Error log showed all RDDs not functional Troubleshooting done - RDD 4 disk 121 seems to bring down all the other JBODS RDD 4 Disk 121 Replaced.
| 193536 | CM | CMPLT | CMPLT | 2023-03-03 13:08 | 2023-03-02 13:08 | Software upgrade | -Software upgrade -This is due to a potential cyber security vulnerability -TGH Prime CT injector asset# 39507 and research CT injector asset # 501960
1055543 | 193537 | CM | CMPLT | CMPLT | 2023-03-03 14:06 | 2023-03-29 10:27 | Battery Indicator RED- internal problem. Battery does not get charged | (TC 03/03/2023 02:05:50 PM, Samanian, Kamran) Contacted tech support. devcie needs to be sent for repair. Case RMA # 00234706 ( temporary ) (TC 03/08/2023 10:28:22 AM, Samanian, Kamran) Received PO# 7202172438 for this repair (TC 03/10/2023 03:11:37 PM, Samanian, Kamran) Received RMA# 1030769 and shipped the package by KJV (TC 03/29/2023 10:16:17 AM, Samanian, Kamran) received a root replacement as repaired. so this device will be retired and new asset tag will be created for the new root.
1065114 | 193538 | CM | CMPLT | CMPLT | 2023-03-03 15:18 | 2023-03-03 15:19 | Reading too high | (TC 02/27/2023 03:18:06 PM, Reid, Emmerson ) Was on with tech support and had to run multiple calibrations before reaching an accurate value. Working now.
1065115 | 193539 | CM | CMPLT | CMPLT | 2023-03-03 15:20 | 2023-03-03 15:21 | Computer not turning on | (TC 02/27/2023 03:19:45 PM, Reid, Emmerson ) Innitially was reported that the device was not reading properly. When I arrived the laptop was not turning on normally, I had to hold the power button down for 10 seconds to shut down, then 10 seconds to turn back on. Also ran calibration to check values and it was in range. Working now.
1056452 | 193540 | CM | CMPLT | CMPLT | 2023-03-06 07:22 | 2023-03-06 07:25 | Missing SpO2 cable and sensor | (TC 03/06/2023 07:23:57 AM, Baghirzade, Nara) Replaced SpO2 cable and finger sensor; functional tests passed. BP, temp, and SpO2 all functional.
62623 | 193541 | CM | CMPLT | CMPLT | 2023-03-06 09:13 | 2023-03-06 09:14 | Latch recall | (TC 03/06/2023 09:13:24 AM, Baghirzade, Nara) Latch recall completed; replaced porthole and NE/NW latches.
1059828 | 193542 | CM | CMPLT | CMPLT | 2023-03-06 09:21 | 2023-03-06 09:22 | Cleaning | (TC 03/06/2023 09:21:26 AM, Ferreira, Qwynn) Drained, cleaned, rinsed, and refilled with sterile water and sent back to floor.
| 193543 | PA | CMPLT | CMPLT | 2023-03-06 09:23 | 2023-03-14 14:59 | Transport Incubator Guardrail and IV Pole Clamp
1051298 | 193544 | CM | CMPLT | CMPLT | 2023-03-06 09:51 | 2023-03-06 09:51 | not working | (TC 02/27/2023 09:50:48 AM, Rendon Lopez , Elizabeth) Cleaned contacts, charged, tested-OK
1063187 | 193545 | CM | CMPLT | CMPLT | 2023-03-06 09:51 | 2023-04-11 16:19 | Not reading | (TC 02/27/2023 09:51:38 AM, Rendon Lopez , Elizabeth) Confirmed problem, kept for repair (TC 04/11/2023 04:18:53 PM, Rendon Lopez , Elizabeth) Received from repair, incoming inspection-OK
63030 | 193546 | PM | CMPLT | CMPLT | 2023-03-06 09:56 | 2023-03-06 09:56 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 02/28/2023 09:56:13 AM, Rendon Lopez , Elizabeth) Replaced battery, functional checks OK
62623 | 193547 | CM | CMPLT | CMPLT | 2023-03-06 09:59 | 2023-03-06 10:02 | Missing door, broken porthole hinge | (TC 03/06/2023 09:59:36 AM, Baghirzade, Nara) Porthole hinge was broken; replaced hinge and re-seated the springs in the hinge. Door was missing; week before, a door was sent to biomed on its own, most likely belonging to this unit. Replaced missing door. All functional tests passed. Set temp to 35C and humidity to 75%; measured 35C and 77%. Scale is functional and accurate; tilt function is good. No errors noted when running the incubator.
1050025 | 193548 | CM | CMPLT | CMPLT | 2023-03-06 10:05 | 2023-03-06 10:07 | Extra monitor - biomed loaner to 9S | (TC 03/06/2023 10:06:05 AM, Samanian, Kamran) Added hose, cuff, used SPo2 sensor and cable and tested the monitor. Passed. Given to 9S as loaner until their Masimo Root gets fixed.
65929 | 193549 | CM | CMPLT | CMPLT | 2023-03-06 10:32 | 2023-03-06 10:34 | channel error | (TC 03/06/2023 10:32:56 AM, Baghirzade, Nara) Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. Ran a basic infusion at various rates; no errors or issues noted during infusion. It was completed successfully.
1054538 | 193550 | CM | CMPLT | CMPLT | 2023-03-06 10:37 | 2023-04-14 09:42 | Repalcing Keyboard from WO 188821 | (TC 04/14/2023 09:31:06 AM, Canjimavadivel, Kamallesh ) Replaced keyboard and keyboard cover
65904 | 193551 | CM | CMPLT | CMPLT | 2023-03-06 10:48 | 2023-03-06 10:49 | Barcode sticker missing | (TC 03/06/2023 10:48:25 AM, Baghirzade, Nara) Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. Ran a basic infusion at various rates; no errors or issues noted. Infusion completed successfully.
1056448 | 193552 | CM | CMPLT | CMPLT | 2023-03-06 11:01 | 2023-03-06 11:02 | SpO2 cable defective | (TC 03/06/2023 11:01:18 AM, Baghirzade, Nara) Red LNC-10 cable was defective; replaced cable. Functional tests passed; checked SpO2, BP, and temp; all functional.
1066129 | 193553 | CM | CMPLT | CMPLT | 2023-03-06 11:05 | 2023-03-06 11:07 | Incoming Inspection | (TC 03/02/2023 11:06:19 AM, Reid, Emmerson ) Incoming complete, see proceedure for electrical safety. For EVALUATION
| 193554 | CM | CMPLT | CMPLT | 2023-03-06 11:22 | 2023-03-06 11:26 | Hood Used | (TC 03/06/2023 11:24:45 AM, Ferreira, Qwynn) Replaced the hood
1055925 | 193555 | CM | CMPLT | CMPLT | 2023-03-06 11:43 | 2023-03-06 11:46 | NIBP requires calibration | (TC 03/06/2023 11:43:41 AM, Baghirzade, Nara) Error displaying "NIBP requires calibration"; followed instructions as per the Masimo service manual to calibrate NIBP. Performed a zero cal and span cal. Calibration was successful and measurements are within +/- 1 mmHg. Checked SpO2 and temperature; all modules are functional.
1066133 | 193556 | ACC | CMPLT | CMPLT | 2023-03-06 13:52 | 2023-07-11 07:30 | MSH VASCULITIS KOFFLER BUILDING VASCULTIS GE VENUE FIT ACCEPTANCE | (TC 03/06/2023 03:21:00 AM, Whittington, Scott) MSH VASCULITIS KOFFLER BUILDING VASCULTIS GE VENUE FIT ACCEPTANCE (TC 03/07/2023 11:25:20 AM, Whittington, Scott) MSH VASCULITIS KOFFLER BUILDING VASCULTIS GE VENUE FIT ACCEPTANCE DOUBLE INSULATED (TC 03/09/2023 03:26:04 PM, Whittington, Scott) MSH VASCULITIS KOFFLER BUILDING VASCULTIS GE VENUE FIT ACCEPTANCE MSH IT NAC & GE MDSS | MSH VASCULITIS KOFFLER BUILDING VASCULTIS GE VENUE FIT ACCEPTANCE DOUBLE INSULATED
1061830 | 193557 | CM | CMPLT | CMPLT | 2023-03-06 14:21 | 2023-10-30 14:11 | Neptune was smoking while in use | (TC 03/06/2023 02:21:13 PM, Reid, Emmerson ) Moved the unit to OR12 so it would not be used and contacted the vendor to fix the device. (TC 03/07/2023 09:56:23 AM, Reid, Emmerson ) NO CHARGE PO created for service SC 1002902217 (TC 03/28/2023 11:18:04 AM, Reid, Emmerson ) All Units have been found with the same issue and have been pulled from service. Currently all the Neptune devices are in OR12 with signage stating not to use. The clinical staff have been notified of this. Waiting to here back from Stryker. (TC 03/31/2023 09:22:37 AM, Reid, Emmerson ) Chris called with an minor update. He had an inservice with the PSAs regarding cleaning. He mentioned that the information was well recieved. He also gave notice about avoiding filling the canisters completely. Still waiting for a quote to replace parts and a permanent solution. Will continue to monitor the situation. (TC 04/25/2023 09:38:51 AM, Reid, Emmerson ) Checked all three loaner neptunes and found that one has started to show some residue in the filter. Reported to Nima, Adeel, and Samir. (TC 06/26/2023 10:48:13 AM, Reid, Emmerson ) Answered emails from stryker regarding having this resolved and scheduling them to come in. (TC 07/13/2023 02:08:50 PM, Reid, Emmerson ) Worked with Chris to identify and decontaminate 4 neptune units that are being picked up next week. This unit is being taken away and has been replaced by the new unit.
1065515 | 193558 | PM | CMPLT | CMPLT | 2023-03-06 14:34 | 2023-03-06 14:35 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 03/06/2023 02:33:45 PM, Reid, Emmerson ) Complete by vendor, see attachde PM report - Jan 10th, 2023.
32809 | 193559 | CM | CMPLT | CMPLT | 2023-03-06 17:29 | 2023-03-06 17:29 | Cassette # 1 not closing properly |  | Replace closing mechanism.
62702 | 193560 | CM | CMPLT | CMPLT | 2023-03-07 08:51 | 2023-03-07 09:20 | restarting when trying take an ECG | (TC 03/07/2023 09:18:02 AM, Scero, Tomlyn) When trying to print an ECG the "paper jam" message apeared and the unit froze, turned off, then restarted. I cleared the paper jam and tested the printer funciton. Redid the ECG and printed it twice. Unit no longer restarted or froze.
69242 | 193561 | CM | CMPLT | NFF | 2023-03-07 16:16 | 2023-03-07 16:17 | Sent down with PCU, "cassettes turn off when adjusting setting" | (TC 03/07/2023 04:16:06 PM, Cortez, John) Unable to replicate error Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69459 | 193562 | CM | CMPLT | NFF | 2023-03-07 16:19 | 2023-03-07 16:20 | Sent down with PCU, "cassettes turn off when adjusting setting" | (TC 03/07/2023 04:19:10 PM, Cortez, John) Unable to replicate issue Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1061364 | 193563 | CM | CMPLT | CMPLT | 2023-03-07 16:28 | 2023-03-07 16:27 | Breaks not engaging. | Adjusted the breaks on one wheel, working fine. The other wheel is on order.
65381 | 193564 | CM | CMPLT | CMPLT | 2023-03-08 07:54 | 2023-03-08 07:54 | QC Checklist | (TC 03/08/2023 07:54:12 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
1055121 | 193565 | PM | CMPLT | CMPLT | 2023-03-08 07:56 | 2023-03-08 07:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2023 07:57:15 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 187 mohms Leakage current: 65.9 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
69365 | 193566 | FC | CMPLT | CMPLT | 2023-03-08 08:24 | 2023-03-08 08:25 | No yellow tag; functional check | (TC 03/08/2023 08:24:33 AM, Baghirzade, Nara) Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
| 193567 | CM | CMPLT | CMPLT | 2023-03-08 08:46 | 2023-03-06 14:45 | Biomed huddle | (TC 03/06/2023 02:45:00 PM, Baghirzade, Nara) Biomed huddle
1056570 | 193568 | CM | CMPLT | CMPLT | 2023-03-08 08:47 | 2023-03-08 08:48 | Battery replacement | (TC 03/08/2023 08:47:52 AM, Canjimavadivel, Kamallesh ) Replaced battery
68966 | 193569 | PM | CMPLT | CMPLT | 2023-03-08 09:09 | 2023-03-08 09:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2023 09:09:17 AM, Baghirzade, Nara) PCU was unable to retain its charge and had to be plugged in all the time; noticed that during the last PM, battery was not conditioned. PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 46 mohms Leakage current: 74.3 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66032 | 193570 | CM | CMPLT | CMPLT | 2023-03-08 09:48 | 2023-03-08 09:50 | Does not charge | (TC 03/08/2023 09:48:06 AM, Baghirzade, Nara) Power cord was defective; replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 13 mohms Leakage current: 90.6 uA
1056936 | 193571 | CM | CMPLT | CMPLT | 2023-03-08 09:59 | 2023-07-28 08:27 | Release lock broken | (TC 03/08/2023 11:00:43 AM, Canjimavadivel, Kamallesh ) Found part number in manual for replacement part. Contacted vendor for quote. Placed order for release button. (TC 07/28/2023 08:25:11 AM, Canjimavadivel, Kamallesh ) Installed lock mechanism in slot
| 193572 | CM | CMPLT | CMPLT | 2023-03-08 10:24 | 2023-03-07 10:24 | Masimo Training | (TC 03/07/2023 10:24:21 AM, Samanian, Kamran) done!
1055934 | 193573 | CM | CMPLT | CMPLT | 2023-03-08 10:44 | 2023-03-08 10:46 | BP times out | (TC 03/08/2023 10:44:08 AM, Baghirzade, Nara) There was damaged to the BP hose so it was leaking each time inflation would occur. Replaced the BP hose and tested calibration. Calibrated zero point and span point. BP, temp, and SpO2 are all functional; functional tests passed.
1046274 | 193574 | CM | CMPLT | CMPLT | 2023-03-08 11:53 | 2023-03-08 11:53 | Probe end strain relief cover riped. |  | Taped probe end cover.
1059778 | 193575 | FC | CMPLT | CMPLT | 2023-03-08 12:12 | 2023-03-08 12:13 | QC checklist | (TC 03/08/2023 12:12:38 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1061207 | 193576 | CM | CMPLT | CMPLT | 2023-03-08 12:13 | 2023-03-08 12:14 | PACS PC is not showing up on Olympus | (TC 03/08/2023 12:13:25 PM, Samanian, Kamran) verified the problem, all other video sources were fine to show up. disconnected and reconnected the VDI cable from wall and reconnected and resolved the issue.
501584 | 193577 | CM | CMPLT | CMPLT | 2023-03-08 12:35 | 2023-07-11 07:31 | MSH OPG 3 SPP IU22 PM BY CODY | (TC 03/08/2023 12:35:20 PM, Whittington, Scott) MSH OPG 3 SPP IU22 PM BY CODY ARRANGED, BUT CODY DID NOT HAVE SUFFICIENT TIME TO DO | MSH OPG 3 SPP IU22 PM BY CODY
1054725 | 193578 | CM | CMPLT | CMPLT | 2023-03-08 12:37 | 2023-03-08 12:39 | MSH 5 IR EPIQ US POWER CORD | (TC 03/08/2023 12:37:10 PM, Whittington, Scott) MSH 5 IR EPIQ US POWER CORD JACKET WORN AT CHASSIS END, SHOWING INSULATED CONDUTORS REPLACE FROM STOCK CODY TO STOCK CODY DID ELECTRICAL SAFETY | MSH 5 IR EPIQ US POWER CORD
1046952 | 193579 | CM | CMPLT | CMPLT | 2023-03-08 12:40 | 2023-03-08 12:42 | MSH ICU SPARQ WILLL NOT BOOT UP | (TC 03/08/2023 12:40:27 PM, Whittington, Scott) MSH ICU SPARQ WILLL NOT BOOT UP (TC 03/08/2023 12:41:48 PM, Whittington, Scott) 1046952 (TC 03/07/2023 05:41:00 PM, Whittington, Scott) MSH ICU SPARQ WILLL NOT BOOT UP | MSH ICU SPARQ WILLL NOT BOOT UP
1046276 | 193580 | CM | CMPLT | CMPLT | 2023-03-08 12:44 | 2023-07-11 07:30 | MSH CFRH 7flr APLIO 300 ORDER TRACKBALL PARTS | (TC 03/08/2023 12:44:06 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 ORDER TRACKBALL PARTS | MSH CFRH 7flr APLIO 300 ORDER TRACKBALL PARTS
1055575 | 193581 | CM | CMPLT | CMPLT | 2023-03-08 13:02 | 2023-03-08 13:02 | BP cuff needs calibration | tested cuff on 2 different unit functional. tested unit BP, SpO2, temp functional
1056473 | 193582 | CM | CMPLT | CMPLT | 2023-03-08 13:05 | 2023-03-08 13:05 | BP leaks | replaced BP hose. tested BP, SpO2, temp. Functional.
1066189 | 193583 | ACC | CMPLT | CMPLT | 2023-03-08 13:35 | 2023-03-08 14:39 | Incoming inspection | (TC 03/08/2023 02:37:59 PM, Canjimavadivel, Kamallesh ) Added to database Incoming inspection passed
1066190 | 193584 | ACC | CMPLT | CMPLT | 2023-03-08 13:37 | 2023-03-08 14:39 | Incoming inspection | (TC 03/08/2023 02:39:07 PM, Canjimavadivel, Kamallesh ) Added to database Incoming inspection passed
1066191 | 193585 | ACC | CMPLT | CMPLT | 2023-03-08 13:38 | 2023-03-08 14:40 | Incoming inspection | (TC 03/08/2023 02:39:38 PM, Canjimavadivel, Kamallesh ) Added to database Incoming inspection passed
1066192 | 193586 | ACC | CMPLT | CMPLT | 2023-03-08 13:39 | 2023-03-08 14:41 | Incoming inspection | (TC 03/08/2023 02:41:01 PM, Canjimavadivel, Kamallesh ) Added to database Incoming inspection passed
1066193 | 193587 | ACC | CMPLT | CMPLT | 2023-03-08 13:40 | 2023-03-08 14:42 | Incoming inspection | (TC 03/08/2023 02:41:55 PM, Canjimavadivel, Kamallesh ) Added to database Incoming inspection passed
1066194 | 193588 | CM | CMPLT | CMPLT | 2023-03-08 13:44 | 2023-03-08 14:45 | Incoming inspection | (TC 03/08/2023 02:44:47 PM, Canjimavadivel, Kamallesh ) Added to database Electrical safety pass Incoming inspection passed
1066195 | 193589 | ACC | CMPLT | CMPLT | 2023-03-08 13:46 | 2023-03-08 14:46 | Incoming inspection | (TC 03/08/2023 02:45:36 PM, Canjimavadivel, Kamallesh ) Added to database Electrical safety pass Incoming inspection passed
1066196 | 193590 | ACC | CMPLT | CMPLT | 2023-03-08 13:47 | 2023-03-08 14:48 | Incoming inspection | (TC 03/08/2023 02:48:08 PM, Canjimavadivel, Kamallesh ) Added to database Electrical safety pass Incoming inspection passed
1066197 | 193591 | ACC | CMPLT | CMPLT | 2023-03-08 13:48 | 2023-03-08 14:51 | Incoming inspection | (TC 03/08/2023 02:50:43 PM, Canjimavadivel, Kamallesh ) Added to database Electrical safety pass Incoming inspection passed
1066198 | 193592 | ACC | CMPLT | CMPLT | 2023-03-08 13:50 | 2023-03-08 14:55 | Incoming inspection | (TC 03/08/2023 02:53:54 PM, Canjimavadivel, Kamallesh ) Added to database Electrical safety pass Incoming inspection passed
65907 | 193593 | CM | CMPLT | CMPLT | 2023-03-08 14:03 | 2023-03-08 14:05 | Channel error | (TC 03/08/2023 02:03:54 PM, Baghirzade, Nara) Pressure sensor was defective; replaced the pressure sensor disc. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1061217 | 193594 | CM | CMPLT | CMPLT | 2023-03-08 14:44 | 2023-03-08 14:45 | small microscope- HDMI to DVI cable | (TC 03/08/2023 02:44:38 PM, Samanian, Kamran) HDMI-DVI cable from Microscope to boom broken. replaced it
1046196 | 193595 | CM | CMPLT | CMPLT | 2023-03-09 00:57 | 2023-03-08 00:56 | CT required upgrade | CT required upgrade
66436 | 193596 | CM | CMPLT | CMPLT | 2023-03-09 09:36 | 2023-03-09 09:46 | Pack was dropped, no longer working | (TC 03/09/2023 09:45:27 AM, Scero, Tomlyn) replaced broken battery adaptor and tested unit on the simulator. Working ok.
1063746 | 193597 | CM | CMPLT | NFF | 2023-03-09 09:45 | 2023-03-09 09:51 | Not warming the patient | (TC 03/09/2023 09:50:15 AM, Ferreira, Qwynn) Tested the Altrix and it was warming the fluid. Let the device run at various temperatures and no alarm present. Measured the temp at 25 = 25.2 Measured temp at 35 = 34.4 Temps are witihn range NFF
1064418 | 193598 | CM | CMPLT | CMPLT | 2023-03-09 09:52 | 2024-01-30 14:32 | Broken Door Corner | (TC 03/09/2023 09:52:18 AM, Ferreira, Qwynn) Was informed the MRI transport was broken and a piece had fallen off. The hood cannot be replaced and a quote has been given for the reapir. Waiting to hear from the unit on whether to replace or not. (TC 01/30/2024 02:30:53 PM, Chukri, Samir) completed and replaced during PM .
66430 | 193599 | CM | CMPLT | CMPLT | 2023-03-09 10:55 | 2023-03-09 10:54 | confguration - labeled as TRxBio2 and added to CCU pool | (TC 03/09/2023 10:54:43 AM, Samanian, Kamran) done!
| 193600 | CM | CMPLT | CMPLT | 2023-03-09 12:38 | 2023-03-14 14:19 | hovermat pump hose leak | (TC 03/09/2023 12:39:27 PM, Rendon Lopez , Elizabeth) Confirmed issue, ordering replacement (TC 03/14/2023 02:18:01 PM, Rendon Lopez , Elizabeth) Received new hose, tested-OK
66710 | 193601 | CM | CMPLT | CMPLT | 2023-03-09 12:53 | 2023-03-09 12:53 | Latch Recall | (TC 03/09/2023 12:53:04 PM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
66710 | 193602 | CM | CMPLT | CMPLT | 2023-03-09 12:54 | 2023-03-09 12:56 | When turning on, keeps beeping and displays error | (TC 03/09/2023 12:53:52 PM, Baghirzade, Nara) Error code 25 displays when turning on the incubator; according to service manual, the fan can be displaced or defective. Fan was not seated properly; re-seated the fan. The NW porthole was damaged and porthole cover was too old; replaced the porthole and cover. Functional tests passed. Set temp to 35C and humidity to 75%; measured 35C and 74%.
| 193603 | PA | CMPLT | CMPLT | 2023-03-09 13:05 | 2023-03-20 07:31 | Ordering syringe pressure sensors | (TC 03/09/2023 01:06:51 PM, Baghirzade, Nara) SC 1002905104 (TC 03/20/2023 07:30:47 AM, Baghirzade, Nara) PO 7202176086
| 193604 | CM | CMPLT | CMPLT | 2023-03-09 13:50 | 2023-04-19 08:37 | Met with Medtronic about Invos 7100 to Philips Monitor | (TC 03/09/2023 01:49:36 PM, Samanian, Kamran) connected EC5#104 to Invos 7100 and EC10 with open interface driver. tested and connectivity passed. a bit delay on appearing numbers on MX monitor (TC 04/19/2023 08:36:48 AM, Samanian, Kamran) medtronic brought back invos 7100 to show dr poorva how it works, plus it was tested with XDS DB Access to export data to excel
1057555 | 193605 | CM | CMPLT | CMPLT | 2023-03-09 14:13 | 2023-03-09 14:25 | Not sending data to cerner | (TC 03/09/2023 02:23:18 PM, Scero, Tomlyn) AA was complaining that data wasn't flowing to the EMR from the monitor and thought changing the X3 would help. Tested the unit and device is wokring fine. Data transfer comes from the EC10, not the X3. Used the X3 to test masimo cables for the OR
1066205 | 193606 | CM | CMPLT | CMPLT | 2023-03-09 14:31 | 2023-03-09 14:32 | Asset tagging and drop off at ENT | (TC 03/09/2023 02:31:01 PM, Baghirzade, Nara) Asset tagging completed. Dropped off the video tower at ENT clinic. Olympus rep will be dropping off the processor the week after.
65929 | 193607 | CM | CMPLT | CMPLT | 2023-03-09 14:35 | 2023-03-09 14:37 | Broken rear case and channel error | (TC 03/09/2023 02:35:30 PM, Baghirzade, Nara) Rear case was broken; replaced the case. Pressure sensor disc was failing the functional test; replaced the pressure sensor. Left handle was broken; replaced the handle. Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 193608 | CM | CMPLT | CMPLT | 2023-03-10 08:01 | 2023-03-31 08:24 | Order X2 safety analyzer AS-2010 and X3 AS-2001 from MedSet. | (TC 03/31/2023 08:23:13 AM, Chukri, Samir) Received all tester. Emmerson entred them into our TMS
65936 | 193609 | CM | CMPLT | CMPLT | 2023-03-10 08:34 | 2023-03-20 07:39 | Broken plunger; ordering plunger assembly | (TC 03/10/2023 08:36:18 AM, Baghirzade, Nara) SC 1002905666 (TC 03/20/2023 07:38:12 AM, Baghirzade, Nara) PO 7202176787
1062137 | 193610 | CM | CMPLT | CMPLT | 2023-03-10 09:05 | 2023-03-10 09:06 | limited function functionality error | (TC 03/09/2023 09:05:08 AM, Cortez, John) Power cycled bed, Functional check ok
1062542 | 193611 | CM | CMPLT | CMPLT | 2023-03-10 09:31 | 2023-03-10 10:24 | Failed DAB self test; unable to measure impedance | (TC 03/10/2023 09:31:20 AM, Baghirzade, Nara) The DAB was old, from 2016. Replaced with the spare one in biomed, dated 2021. E-mailed Rex to let him know that the box was replaced and to let us know if they are able to measure impedance on their next study.
66361 | 193612 | CM | CMPLT | CMPLT | 2023-03-10 09:49 | 2023-03-10 09:49 | upgraded to rev N | (TC 03/10/2023 09:49:31 AM, Samanian, Kamran) done!
66361 | 193613 | CM | CMPLT | CMPLT | 2023-03-10 10:25 | 2023-03-10 10:28 | Arm and rack broken | (TC 03/10/2023 10:25:42 AM, Baghirzade, Nara) The arm holding the MX800 was broken and collapsed on top of the rack, causing the rack and its mounting to break. Kamran upgraded the rack to revision N. Replaced the rack. Qwynn and I replaced the arm and placed the MX800 back on to the arm; checked that the monitor connects to central and that it recognizes the rack. All is functional. Qwynn will retire the asset for the broken rack. (TC 03/10/2023 10:27:52 AM, Samanian, Kamran) Upgraded rack to Rev. N. for NICU
69487 | 193614 | CM | CMPLT | CMPLT | 2023-03-10 10:28 | 2023-03-10 10:30 | "malfunction" | (TC 03/09/2023 10:29:12 AM, Cortez, John) Cleaned, Downloaded error/event logs, Bottle-side pressure sensor errors were logged, Pressure sensor replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
| 193615 | PA | CMPLT | CMPLT | 2023-03-10 10:29 | 2023-03-15 12:22 | Ordering sterile water for NICU | (TC 03/10/2023 10:37:05 AM, Baghirzade, Nara) SC 1002905910 (TC 03/15/2023 12:20:51 PM, Baghirzade, Nara) Order # 4150250619
| 193616 | CM | CMPLT | CMPLT | 2023-03-10 10:53 | 2023-05-04 09:04 | Spare WA thermometer probes | Received
1051733 | 193617 | CM | CMPLT | CMPLT | 2023-03-10 11:20 | 2023-02-21 11:21 | Guardrail Coating | (TC 03/10/2023 11:19:50 AM, Ferreira, Qwynn) Found the device on the 15th floor. There was pink tape covering the damaged section of the rail. Found the protective coating had been stripped off. Ordered new rail and IV pole mount Filled out the required incident report 55322
66000 | 193618 | CM | CMPLT | CMPLT | 2023-03-10 14:09 | 2023-03-16 15:48 | Damaged front sreen | (TC 03/10/2023 02:09:26 PM, Cortez, John) Cleaned unit, Replaced front and rear cases, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioning battery overnight (TC 03/16/2023 03:46:58 PM, Cortez, John) 4148 mAh new capacity
1054543 | 193619 | CM | CMPLT | CMPLT | 2023-03-10 14:12 | 2023-03-10 14:15 | Leads V4-V6 not working | (TC 03/10/2023 02:12:18 PM, Baghirzade, Nara) Replaced leads V4-V6. Also provided cleartabs.
501618 | 193620 | CM | CMPLT | CMPLT | 2023-03-10 14:56 | 2023-03-10 14:57 | C5 probe not recognizing. | (TC 03/10/2023 02:56:13 PM, Fan, Tak) Tested on 2 other machines and the original. Connected fine. Johnny also checked during lunch. The pins had packaging foam around some of the pins, may have covered the connection but moved out of the way when reconnected a few times.
| 193621 | CM | CMPLT | CMPLT | 2023-03-10 15:10 | 2023-03-10 15:11 | Working on Masimo to connect to MX monitor | (TC 03/10/2023 03:10:05 PM, Samanian, Kamran) used EC5#104 and EC10 open interface driver, adjsuted baud rate to 19200, used USB to RS232 adapter and was able to export data to MX. tested and passed
1061205 | 193622 | CM | ACTIV | ISSUE | 2023-03-10 15:13 |  | Olympus wall monitor damaged , SN# 722021835461 | (TC 03/10/2023 03:13:38 PM, Samanian, Kamran) switched the monitor with OR12 wall monitor, tested and passed. got RMA # 300009234 from Olympus for sending this for repair (TC 06/13/2023 12:32:49 PM, Samanian, Kamran) Received the quote and it was approved by Reeba. attached quote this WO. (TC 06/13/2023 12:32:49 PM, Samanian, Kamran) Received the quote and it was approved by Reeba. attached quote this WO. S/C#1002979199
| 193623 | CM | CMPLT | CMPLT | 2023-03-10 15:42 | 2023-05-29 09:45 | Bunnell ventilator test equipment | (TC 03/10/2023 03:49:27 PM, Rendon Lopez , Elizabeth) Bunnell ventilator test equipment Received
1054537 | 193624 | CM | CMPLT | CMPLT | 2023-03-10 16:03 | 2023-03-10 16:04 | Cleaning cycle | (TC 03/10/2023 04:03:32 PM, Cortez, John) Completed cleaning cycle
1050672 | 193625 | CM | CMPLT | CMPLT | 2023-03-13 09:20 | 2023-03-13 09:38 | Screen malfunction, glitchy | (TC 03/13/2023 09:37:05 AM, Canjimavadivel, Kamallesh ) Screen working fine, touch screen test passed. Could not replicate issue. Repaired power supply holder on cart.
1060001 | 193626 | CM | CMPLT | CMPLT | 2023-03-13 09:56 | 2023-03-13 10:09 | Broken | (TC 03/13/2023 10:08:20 AM, Canjimavadivel, Kamallesh ) Replaced RA and LL leads as they had exposed wires. Tested okay with simulator.
| 193627 | PA | CMPLT | CMPLT | 2023-03-13 10:06 | 2023-06-12 08:45 | Hoods for Incubator | arriving in june
65846 | 193628 | CM | CMPLT | CMPLT | 2023-03-13 10:10 | 2023-03-13 10:34 | Power cord exposed | (TC 03/13/2023 10:10:12 AM, Reid, Emmerson ) Replaced the power cord. Could not find a grounding pin. Working now.
1059711 | 193629 | CM | CMPLT | CMPLT | 2023-03-13 10:55 | 2023-03-13 11:00 | RA connector pin came off | (TC 03/13/2023 10:59:22 AM, Reid, Emmerson ) Replaced the RA lead and tested. Working now. (TC 03/13/2023 10:59:22 AM, Reid, Emmerson ) Replaced the RA lead and tested. Working now.
63724 | 193630 | CM | CMPLT | CMPLT | 2023-03-13 11:05 | 2023-03-13 11:20 | Only lead II was working | (TC 03/13/2023 11:19:08 AM, Reid, Emmerson ) Replaced RA lead. Tested. Working now.
1054479 | 193631 | CM | CMPLT | CMPLT | 2023-03-13 11:31 | 2023-06-02 09:45 | Side rail won't go down | (TC 03/13/2023 11:30:57 AM, Canjimavadivel, Kamallesh ) Contacted vendor, technician came on site to assess and ordered part required. (TC 06/02/2023 09:44:19 AM, Canjimavadivel, Kamallesh ) See attached report
1059980 | 193632 | CM | CMPLT | CMPLT | 2023-03-13 11:33 | 2023-06-02 09:44 | Head section no working | (TC 03/13/2023 11:32:40 AM, Canjimavadivel, Kamallesh ) Cotnacted vendor, waiting technician to assess (TC 06/02/2023 09:42:29 AM, Canjimavadivel, Kamallesh ) See attached repoort from vendor
1055646 | 193633 | CM | CNP | ENL | 2023-03-13 11:39 |  | Side rail not working | (TC 03/13/2023 11:38:41 AM, Canjimavadivel, Kamallesh ) Contacted vendor, waiting for technician to assess (TC 06/02/2023 09:48:01 AM, Canjimavadivel, Kamallesh ) Closing WO as bed cannot be located and issue has not be reported again.
67632 | 193634 | PM | CMPLT | CMPLT | 2023-03-13 11:39 | 2023-03-29 11:08 | PM-MAJOR - Annual PM For: Air-oxygen Blender | (TC 03/29/2023 11:07:23 AM, Ferreira, Qwynn) Checkin and attached the report to the WO Functional Check - ok
65790 | 193635 | CM | CMPLT | CMPLT | 2023-03-13 13:56 | 2023-03-13 13:57 | FW and open interface upgrade | (TC 03/13/2023 01:56:27 PM, Samanian, Kamran) FW upgrade from 1310 to 1951. Open interface upgraded from A.1 to A.7
| 193636 | CM | CANCL | CANCL | 2023-03-13 14:44 |  | Ordering O2 cell | (TC 03/13/2023 02:46:01 PM, Baghirzade, Nara) SC 1002907649 (TC 03/15/2023 07:42:12 AM, Baghirzade, Nara) Incorrect O2 cell was provided by the vendor; Russ from Novus brought the correct O2 cell for the Hypair and traded it in at no cost.
501068 | 193637 | CM | CMPLT | CMPLT | 2023-03-13 16:40 | 2023-03-15 09:32 | MSH CEOU RM 2 Unit will not power up. Blinking blue light on power supply | (TC 03/13/2023 04:44:21 PM, Dolinar, John) MSH CEOU RM 2 Unit will not power up. Blinking blue light on power supply Called Philips Ref# 021358924 POWER SUPPLY ON ORDER (TC 03/15/2023 09:32:10 AM, Fan, Tak) PS replaced. | MSH CEOU RM 2 Unit will not power up. Blinking blue light on power supply Called Philips Ref# 021358924 POWER SUPPLY ON ORDER
| 193638 | CM | CMPLT | CMPLT | 2023-03-13 16:42 | 2023-03-13 16:42 | Site coverage.
| 193639 | CM | CMPLT | CMPLT | 2023-03-14 08:46 | 2023-04-17 10:28 | Hoses for RT NICU | (TC 03/14/2023 08:51:28 AM, Rendon Lopez , Elizabeth) ordering Air hoses requested by Rex (TC 04/14/2023 10:27:30 AM, Rendon Lopez , Elizabeth) Received hoses, tested, delivered to RT NICU
1056667 | 193640 | CM | CMPLT | CMPLT | 2023-03-14 09:26 | 2023-03-14 09:26 | Image on the main monitor is out of focus | (TC 03/14/2023 09:26:00 AM, Samanian, Kamran) verified the problem. camera focus knob ( located on right side of the microscope head ) must be adjsuted to get focused image on the monitor.
| 193641 | CM | CMPLT | CMPLT | 2023-03-14 09:33 | 2023-03-21 11:00 | HDMI to DVI cable for OR | (TC 03/14/2023 11:55:14 AM, Samanian, Kamran) placed the order, SC#1002908079 PO# 7202178496 (TC 03/21/2023 10:58:08 AM, Samanian, Kamran) received
| 193642 | CM | CMPLT | CMPLT | 2023-03-14 09:56 | 2023-03-14 09:57 | PillCam PC | (TC 03/14/2023 09:56:27 AM, Samanian, Kamran) Requested to move the patinet videos to V drive. tested yesterday but it did not work due to wirting persmission. tried today and it worked.
1056125 | 193643 | PM | CMPLT | CMPLT | 2023-03-14 10:43 | 2023-03-14 10:44 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/14/2023 10:43:31 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
501105 | 193644 | CM | CMPLT | CMPLT | 2023-03-14 11:57 | 2023-03-14 11:56 | MSH CEOU SPARE iU22 SETUP | (TC 03/14/2023 11:56:45 AM, Whittington, Scott) MSH CEOU SPARE iU22 SETUP MARCH 14, 2023 MOVE SYSTEM TO MSH OPG 3 CEOU RM 6 FROM MSH 5 SETUP AS MSH_US_CEOU-3 (FROM VOLUSON 4) SETUP PROTOCOLS AND PROCEDURES BY RESTORING WITH B/U FROM CEOU RM 1 SETUP NETWORK SETTINGS FOR BURNER AND SR | MSH CEOU SPARE iU22 SETUP MARCH 14, 2023 MOVE SYSTEM TO MSH OPG 3 CEOU RM 6 FROM MSH 5 SETUP AS MSH_US_CEOU-3 (FROM VOLUSON 4) SETUP PROTOCOLS AND PROCEDURES BY RESTORING WITH B/U FROM CEOU RM 1 SETUP NETWORK SETTINGS FOR BURNER AND SR
1059778 | 193645 | FC | CMPLT | CMPLT | 2023-03-14 12:38 | 2023-03-14 12:39 | QC checklist | (TC 03/14/2023 12:38:28 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1061333 | 193646 | FC | CMPLT | CMPLT | 2023-03-14 12:39 | 2023-03-14 12:39 | QC checklist | (TC 03/14/2023 12:39:19 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
| 193647 | TEDU | CMPLT | CMPLT | 2023-03-14 12:40 | 2023-03-13 15:45 | PFT Training | (TC 03/13/2023 03:45:00 PM, Baghirzade, Nara) PFT training with Medisoft and Novus
1051733 | 193648 | CM | CMPLT | CMPLT | 2023-03-14 15:07 | 2023-03-15 14:37 | Replaced Rail and IV Pole | (TC 03/14/2023 03:08:47 PM, Ferreira, Qwynn) replaced the guard rail and attached new IV pole mount
1053034 | 193649 | CM | CMPLT | CMPLT | 2023-03-14 15:10 | 2023-03-15 14:37 | Guardrail replacement and IV Pole | (TC 03/14/2023 03:09:51 PM, Ferreira, Qwynn) Replaced guardrail and IV Pole clamp
1053391 | 193650 | CM | CMPLT | CMPLT | 2023-03-15 07:27 | 2023-03-15 07:37 | Functional check before next patient use | (TC 03/15/2023 07:36:55 AM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 51 mmHg Set 125 mmHg - Measured 127 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 126 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
| 193651 | TEDU | CMPLT | CMPLT | 2023-03-15 07:38 | 2023-03-14 15:00 | PFT Training | (TC 03/14/2023 03:00:00 PM, Baghirzade, Nara) PFT training
1064382 | 193652 | CM | CMPLT | CMPLT | 2023-03-15 08:58 | 2023-03-15 08:59 | Fan keeps running | (TC 03/06/2023 08:58:25 AM, Rendon Lopez , Elizabeth) Replaced filter, defective activation cable, replaced. Tested-OK
67088 | 193653 | CM | CMPLT | CMPLT | 2023-03-15 09:05 | 2023-03-15 09:07 | BP not working | (TC 03/10/2023 09:05:49 AM, Rendon Lopez , Elizabeth) E36 - Thermometer failure. Replaced temp probe, tested all functions-OK
69573 | 193654 | CM | CMPLT | CMPLT | 2023-03-15 09:09 | 2023-03-15 09:10 | Channel error | (TC 03/10/2023 09:09:34 AM, Rendon Lopez , Elizabeth) Downloaded error log - 240.4150.0 - Bottle pressure sensor error. Replaced sensor and LCD under recall, calibrated-OK
65262 | 193655 | CM | CMPLT | CMPLT | 2023-03-15 09:14 | 2023-03-15 14:18 | Unable to switch to HFO | (TC 03/14/2023 09:14:14 AM, Rendon Lopez , Elizabeth) Downloaded error log - Continuous Faulty neonatal flow sensor. Run tests, and tested in normal mode (TC 03/15/2023 02:17:47 PM, Rendon Lopez , Elizabeth) Ran throughout the night, no issues found. Returned to RT NICU
67645 | 193656 | CM | CMPLT | CMPLT | 2023-03-15 09:16 | 2023-03-15 09:18 | Blue light | (TC 03/14/2023 09:17:08 AM, Rendon Lopez , Elizabeth) Replaced battery, charged. Tested-OK.
| 193657 | CON | CMPLT | CMPLT | 2023-03-15 09:37 | 2023-03-15 09:37 | Needs network cable | (TC 03/14/2023 09:37:01 AM, Fan, Tak) Needed network cable. Also chair needs WD40.
1055946 | 193658 | CM | CMPLT | CMPLT | 2023-03-15 10:35 | 2023-03-15 10:35 | does not work | Tested BP, SpO2 and temp. Functional.
1055947 | 193659 | CM | CMPLT | CMPLT | 2023-03-15 10:36 | 2023-03-15 10:36 | deoes not read BP | Tested BP, SpO2 and temp. Functional.
1055579 | 193660 | CM | CMPLT | CMPLT | 2023-03-15 10:38 | 2023-03-15 10:38 | oximeter and temperature are not working | Replaced SpO2 patient cable and temperature probe. Tested BP, SpO2 and temp. Functional.
| 193661 | PA | CMPLT | CMPLT | 2023-03-15 10:51 | 2023-03-24 12:29 | Ording SpO2 Cables for OPG | (TC 03/24/2023 12:28:02 PM, Reid, Emmerson ) Recieved full order
1057545 | 193662 | CM | CMPLT | CMPLT | 2023-03-15 11:35 | 2023-04-04 12:15 | SPO2 port not working | (TC 03/15/2023 11:35:01 AM, Scero, Tomlyn) Confirmed using a known good SPO2 cable and sensor. Device saying SPO2 Malfunction replace transducer, sensor or cable. Created a ticket using Philips repair portal. (TC 03/23/2023 09:32:47 AM, Scero, Tomlyn) quoted by philips for a new SPO2 board. $361.87. Ordered on ereq (TC 04/04/2023 12:14:33 PM, Scero, Tomlyn) Replaced SPO2 board. Working fine when tested on Mx800
66658 | 193663 | CM | CMPLT | CMPLT | 2023-03-15 11:37 | 2023-03-15 11:37 | BAT replacement | (TC 03/15/2023 11:37:27 AM, Samanian, Kamran) done! UPS is located in Biomed
1058049 | 193664 | CM | CMPLT | CMPLT | 2023-03-15 11:49 | 2023-03-15 11:55 | Temp not reading properly | (TC 03/15/2023 11:49:34 AM, Reid, Emmerson ) Used air to blow out any dust from the temp sensors. Ran and tested temperature with infared therm. Could be an issue with the temp probe cables but we cannot test. Working now.
| 193665 | TEDU | CMPLT | CMPLT | 2023-03-15 11:58 | 2023-03-15 11:59 | CPET Training | (TC 03/15/2023 11:59:13 AM, Baghirzade, Nara) CPET training
| 193666 | PA | CMPLT | CMPLT | 2023-03-15 12:01 | 2023-04-03 11:20 | O-rings for Smiths Fluid Warmers | (TC 03/24/2023 10:52:51 AM, Reid, Emmerson ) Recieved 5/64 small O-rings (TC 04/03/2023 11:19:37 AM, Reid, Emmerson ) Recieved all
1054725 | 193667 | CM | CMPLT | CMPLT | 2023-03-15 14:22 | 2023-03-15 14:23 | Vendor requesting authorization for remote support | (TC 03/15/2023 02:22:26 PM, Fan, Tak) Placed ticket but IT says port 443 is open. Mirza ask for MAC confirmation, confirmed. He suggests trying on wired. Cody will test in the future.
501066 | 193668 | CM | CMPLT | CMPLT | 2023-03-15 14:32 | 2023-03-22 08:16 | Monitor intermittently working. | (TC 03/15/2023 02:33:20 PM, Fan, Tak) Monitor was black but backlight on, system OK. Unplugged power cable and plugged it back in, the monitor started up with the image. (TC 03/21/2023 04:15:00 PM, Fan, Tak) Replaced monitor.
1048472 | 193669 | CM | CMPLT | CMPLT | 2023-03-15 17:02 | 2023-03-15 17:02 | Tube cover melted. |  | Replaced front tube cover and ordered a new top tube cover.
| 193670 | CM | CMPLT | CMPLT | 2023-03-15 17:03 | 2023-03-15 17:03 | Mandatory courses
1056552 | 193671 | CM | CMPLT | CMPLT | 2023-03-16 09:41 | 2023-03-16 09:43 | Battery replacement | (TC 03/16/2023 09:41:22 AM, Canjimavadivel, Kamallesh ) Replaced battery
1056535 | 193672 | CM | CMPLT | CMPLT | 2023-03-16 09:47 | 2023-03-16 09:52 | Battery replacement | (TC 03/16/2023 09:47:55 AM, Canjimavadivel, Kamallesh ) Replaced battery
1056538 | 193673 | CM | CMPLT | CMPLT | 2023-03-16 10:38 | 2023-03-16 10:39 | Battery replacement | (TC 03/16/2023 10:38:03 AM, Canjimavadivel, Kamallesh ) Replaced battery
1056554 | 193674 | CM | CMPLT | CMPLT | 2023-03-16 11:02 | 2023-03-16 11:03 | Battery replacement | (TC 03/16/2023 11:02:29 AM, Canjimavadivel, Kamallesh ) Replaced battery
1056583 | 193675 | CM | CMPLT | CMPLT | 2023-03-16 11:04 | 2023-03-16 11:05 | Battery replacement | (TC 03/16/2023 11:04:42 AM, Canjimavadivel, Kamallesh ) Replaced battery
66436 | 193676 | CM | CMPLT | CMPLT | 2023-03-16 11:10 | 2023-03-17 14:36 | no data tele | (TC 03/17/2023 02:34:38 PM, Scero, Tomlyn) Tested the unit with an ECG simulator. no issues showing data or connecting to the network. Returned the unit to CCU
1056573 | 193677 | CM | CMPLT | CMPLT | 2023-03-16 11:15 | 2023-03-16 11:16 | Battery replacement | (TC 03/16/2023 11:15:25 AM, Canjimavadivel, Kamallesh ) Battery replaced
65264 | 193678 | CM | CMPLT | CMPLT | 2023-03-16 11:51 | 2023-03-23 13:44 | ordering VN500 parts | received parts, PO#7202180651
501072 | 193679 | CM | CMPLT | CMPLT | 2023-03-16 11:58 | 2023-03-16 11:59 | MSH CEOU RM#5 - Error 511 | (TC 03/16/2023 11:57:51 AM, Chau, Ken) case #0121383112 created for FSE to clear data file in widnows -reloaded back up files and tested unit okay (TC 03/17/2023 11:11:20 AM, Chau, Ken) All verification passed on JDMIPACS, Downtime, Burner, Worklist. However unable to "send to" and shows blank. -reloaded backup and issue still not resolved, had to add and remove all, verify again and tested okay after
1055800 | 193680 | CM | CMPLT | CMPLT | 2023-03-16 11:59 | 2023-03-16 13:53 | BP cuff broken | (TC 03/16/2023 01:51:19 PM, Scero, Tomlyn) NIBP cuff replaced and tubing replaced. connector was broken
501069 | 193681 | CM | CMPLT | CMPLT | 2023-03-16 12:01 | 2023-03-16 13:14 | MSH CEOU RM # 11 - Track Ball Not Moving | (TC 03/16/2023 12:01:14 PM, Chau, Ken) trackball not working -disassembled, cleaned and tested okay
1062551 | 193682 | CM | CMPLT | CMPLT | 2023-03-16 12:44 | 2023-04-17 10:40 | Left hand rail not staying up | (TC 03/16/2023 01:03:20 PM, Rendon Lopez , Elizabeth) Confirmed problem. release button on hand rail is stuck. Contacted stryker support (TC 04/12/2023 10:39:44 AM, Rendon Lopez , Elizabeth) Repair completed by stryker
500846 | 193683 | CM | CMPLT | CMPLT | 2023-03-16 12:47 | 2023-07-11 07:32 | MSH 5 US APLIO XG RM 3 SYSTEM WOULD NOT SAVE CASE ON HDD; WORKLIST /PT BUTTON FICKLE | (TC 03/16/2023 02:19:35 PM, Whittington, Scott) MSH 5 US APLIO XG RM 3 SYSTEM WOULD NOT SAVE CASE ON HDD WORKLIST /PT BUTTON FICKLE (TC 03/16/2023 02:19:35 PM, Whittington, Scott) MSH 5 US APLIO XG RM 3 SYSTEM WOULD NOT SAVE CASE ON HDD WORKLIST /PT BUTTON FICKLE FULL HDD MNTCE | MSH 5 US APLIO XG RM 3 SYSTEM WOULD NOT SAVE CASE ON HDD WORKLIST /PT BUTTON FICKLE
1050631 | 193684 | CM | CMPLT | CMPLT | 2023-03-16 13:15 | 2023-04-24 11:28 | MSH - Labour and Delivery - TE7 - Not Working | (TC 03/16/2023 01:15:27 PM, Chau, Ken) To be sent out to vendor for replair -inquired about loaner until unit is repaired -error "warning! transmitting voltage is under protection! the image can not be displayed correctly! please shut down!" (TC 03/22/2023 10:17:15 AM, Chau, Ken) Loaner received, SN#7P-57000200 Sent out unit to be repaired Loaner given to LD (TC 03/22/2023 10:17:15 AM, Whittington, Scott) MSH 15 L&D TE7A REPLACE FAULTY WITH LOANER, BACKUP SYSTEM AND RESTORE TO LOANER SN# 70-57000200 DELIVER L&D R20 MEET WITH NICOLE , NE L&D MANAGER (TC 04/21/2023 11:13:36 AM, Whittington, Scott) MSH - Labour and Delivery - TE7 - Not Working PREVIOUSLY UNABLE TO LOCATE LOANER SYSTEM FOR SWAPPING IN REPAIRED AND RETURN OF LOANER TO TANYA @ MINDRAY (TC 04/21/2023 11:13:36 AM, Whittington, Scott) MSH - Labour and Delivery - TE7 - Not Working PREVIOUSLY UNABLE TO LOCATE LOANER SYSTEM FOR SWAPPING IN REPAIRED AND RETURN OF LOANER TO TANYA @ MINDRAY FEDEX TRK 7719 2790 9391 (TC 04/24/2023 11:16:32 AM, Dolinar, John) MSH TE7 REPAIRED LOANER SWAPPED OUT SENT TO MSH DOCK FOR P/U FEDEX TRK 7719 2790 9391 | case 191004
502089 | 193685 | CM | CMPLT | CMPLT | 2023-03-16 13:21 | 2023-03-16 13:22 | mounting screw for phantom was stripped and needed to be replaced, no washer was present | replaced the screw with a new one and added a washer
| 193686 | CM | CMPLT | CMPLT | 2023-03-16 13:44 | 2023-04-13 11:19 | Order for LVP pressure sensors, biomed stock | (TC 04/06/2023 11:18:28 AM, Cortez, John) Sensors ordered/arrived
1054974 | 193687 | CM | CMPLT | CMPLT | 2023-03-16 16:08 | 2023-03-16 16:09 | "O2 sat not working" | (TC 03/16/2023 04:08:29 PM, Cortez, John) Confirmed issue, replaced SPO2 sensor All parameters tested ok
65430 | 193688 | CM | CMPLT | CMPLT | 2023-03-16 16:11 | 2023-03-16 16:13 | "Blinking light" | (TC 03/16/2023 04:11:08 PM, Cortez, John) Unsure what user was referring to about a blinking light. Difficult to obtain BP reading. Replaced adult cuff and supplied small adult and large adult cuffs. Changed Temp unit to celsius Replaced temp probe with spare All parameters tested ok
67164 | 193689 | CM | CMPLT | CMPLT | 2023-03-16 16:14 | 2023-03-16 16:18 | "no reading/measurement" | (TC 03/16/2023 04:15:53 PM, Cortez, John) Adult BP cuff is faulty, swapped cuff and supplied small and large size cuffs All parameters tested ok
1049905 | 193690 | CM | CMPLT | CMPLT | 2023-03-16 16:35 | 2023-07-11 07:31 | MSH R65 DETECTOR NOT LINKING, DET BATTERIES NOT CHARGING | (TC 03/16/2023 04:34:54 PM, Whittington, Scott) MSH R65 DETECTOR NOT LINKING, DET BATTERIES NOT CHARGING LOWER DETECTOR BATTERY CHARGING SLOT HAS BENT PINS DETECTOR OK PROVIDED FRESHLY CHARGED BATTERY RETURNED LOW 27% BATTERY TO DR CONTROL FOR TOASTER CHARGE (TC 03/20/2023 09:34:00 AM, Whittington, Scott) 193690 MARCH 16 MSH R65 DETECTOR NOT LINKING, DET BATTERIES NOT CHARGING LOWER DETECTOR BATTERY CHARGING SLOT HAS BENT PINS DETECTOR OK PROVIDED FRESHLY CHARGED BATTERY RETURNED LOW 27% BATTERY TO DR CONTROL FOR TOASTER CHARGE | MSH R65 DETECTOR NOT LINKING, DET BATTERIES NOT CHARGING
65264 | 193691 | CM | CMPLT | CMPLT | 2023-03-17 09:51 | 2023-03-17 09:51 | not passing pre-use check. nebulizer test fails. | not passing pre-use check. nebulizer test fails. collected error logs, sent them to Draeger. Remedies: 1) Connect the device to the HIT software and start the test mode. In the HIT software, open the product page "PMS8 V500_Opt.". 2) Using the HIT software, perform the nebulizer/ejector functional test as per test instructions. 3) Switch V10 to AIR position. 4) V12 ejector valve in position "open". 5) S2 sensor O2 pressure lies between 800 and 2500 mbar. 6) Switch V10 to O2 position. 7) S2 sensor O2 pressure lies between 800 and 2500 mbar. 8) V12 ejector valve in position "closed". The nebulizer 02 or Air pressure regulators may need to be replaced if drifted. Replaced both AIR and O2 pressure regulators. Adjusted them. Tested unit. Passed pre-use check. Functional.
1059778 | 193692 | FC | CMPLT | CMPLT | 2023-03-20 07:23 | 2023-03-20 07:24 | QC Checklist | (TC 03/20/2023 07:23:26 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1055949 | 193693 | CM | CMPLT | CMPLT | 2023-03-20 08:02 | 2023-03-20 08:02 | BP not working | Replaced BP cuff. tested. BP. SpO2 and temp. Functional.
1066352 | 193694 | ACC | CMPLT | CMPLT | 2023-03-20 08:28 | 2023-03-20 08:43 | Incoming inspection and asset tagging | (TC 03/20/2023 08:40:57 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Leakage current = 8.1 uA Power cord resistance = 33 mohms Ready for use. Replaced surgitome on new cable as per Slavica's request. Dropped equipment off at the dental clinic's front desk.
67948 | 193695 | CM | CMPLT | CMPLT | 2023-03-20 08:54 | 2023-03-20 08:55 | Broken porthole door | (TC 03/20/2023 08:53:42 AM, Baghirzade, Nara) Porthole door was replaced. Spring was also missing in the hinge; replaced the spring. All functional checks passed. Set temp 35C and humidity to 75%; measured 35C and 75%.
1056451 | 193696 | CM | CMPLT | CMPLT | 2023-03-20 10:09 | 2023-03-20 10:14 | Replacement Root; incoming inspection | (TC 03/20/2023 10:11:17 AM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Loctite on all mount screws. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Probe well kit was replaced. SpO2 finger sensor and LNC cable was added. Functional tests passed of NIBP, temp, and SpO2.
68806 | 193697 | CM | CMPLT | CMPLT | 2023-03-20 10:30 | 2023-03-20 10:32 | Prong on power cord is bent | (TC 03/20/2023 10:30:39 AM, Baghirzade, Nara) Replaced power cord. Electrical Safety - Ground resistance: 26 mohms Leakage current: 78.1 uA IUIs and device cleaned.
| 193698 | CM | CMPLT | CMPLT | 2023-03-20 11:01 | 2023-03-20 11:01 | PFT report editing | (TC 03/20/2023 11:00:38 AM, Baghirzade, Nara) Gaby had requested that all parameters are displayed with two decimal points; changed the decimal points on all FGA reports. Asked Gaby to confirm.
65985 | 193699 | CM | CMPLT | CMPLT | 2023-03-20 11:02 | 2023-03-20 11:04 | Channel error; pump not engaged | (TC 03/20/2023 11:02:42 AM, Baghirzade, Nara) Downloaded error log; noted one error regarding pump not engaged. According to service manual, action to be taken if the error is repeatable. Error did not repeat more than once. No other major errors noted. Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. Ran a basic infusion at various rates; no errors or issues noted. Infusion completed successfully.
63210 | 193700 | CM | CMPLT | CMPLT | 2023-03-20 11:06 | 2023-03-20 11:15 | Power cord broken | (TC 03/20/2023 11:14:17 AM, Baghirzade, Nara) Replaced power cord. Ran a feed; no errors or issues noted while running the pump. Set loading area is good, no damage. Pump is fully functional.
32580 | 193701 | CM | CMPLT | CMPLT | 2023-03-20 13:50 | 2023-03-20 13:50 | Display console won't boot up | Display PC won't boot up -Ping is alive between Scan and Display PCs -E: (image) drive is present/accessible on Display PC -No error warning in Recon unit -X: temp file not accessible on Display PC -Replaced display PC system disk and issue was resolved. (TC 03/24/2023 02:47:29 PM, Debebe, Adam) Room was booked down to duplicate the D-con backup disk. -The system disk is 160 GB and all the spare disk available in our stock are 80 GB. -Copying not possible due to disk mismatch. -Will have to source 160 GB disk -Returned system for use.
1048472 | 193702 | CM | CMPLT | CMPLT | 2023-03-20 17:29 | 2023-03-20 17:29 | Ordered tube cover
| 193703 | CM | CMPLT | CMPLT | 2023-03-20 17:32 | 2023-03-20 17:32 | Site coverage
1046061 | 193704 | CM | ACTIV | ISSUE | 2023-03-21 08:13 |  | Not providing enough suction | (TC 03/20/2023 08:13:39 AM, Reid, Emmerson ) Unit came from OR 3 and was not providing much suction. Tried switching the filter but it did not help. Looking to replace the unit.
65640 | 193705 | CM | CMPLT | CMPLT | 2023-03-21 09:12 | 2023-03-21 09:12 | Alarming | Tested. Functional.
65989 | 193706 | CM | CMPLT | CMPLT | 2023-03-21 09:41 | 2023-03-21 09:42 | Requires calibration | (TC 03/21/2023 09:41:12 AM, Baghirzade, Nara) Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. Ran a basic infusion at various rates; no errors or issues noted. Infusion completed successfully.
65898 | 193707 | CM | CMPLT | CMPLT | 2023-03-21 09:47 | 2023-03-21 09:50 | Channel error, broken rear case and handle | (TC 03/21/2023 09:47:30 AM, Baghirzade, Nara) Replaced: -Left handle -Rear case -Front case -Pressure sensor Reprinted barcode sticker and asset sticker. Full calibration of barrel clamp, pressure sensor disc, plunger force and accuracy was completed. Check-in tests completed. IUI connectors and device cleaned.
| 193708 | CM | CMPLT | CMPLT | 2023-03-21 09:54 | 2023-03-20 09:54 | biomed huddle | (TC 03/20/2023 09:54:21 AM, Samanian, Kamran) done!
1059637 | 193709 | CM | CMPLT | NFF | 2023-03-21 10:16 | 2023-03-21 10:17 | User reported unable to turn off the device | (TC 03/21/2023 10:15:56 AM, Baghirzade, Nara) Checked power cord; no issue found. Checked that all buttons depress properly and not sticky; all good. Device powers on and off; no issues switching it off. Returned to NICU. No fault found.
| 193710 | CM | CMPLT | CMPLT | 2023-03-21 11:10 | 2023-03-21 11:10 | Hyland in Room# 2 not showing up the first patient | (TC 03/21/2023 11:09:42 AM, Samanian, Kamran) user added the patient demographic manually. teh rest of patients for the day shown up. NFF
| 193711 | CM | CMPLT | CMPLT | 2023-03-21 11:11 | 2023-03-21 11:12 | EEG computer, unhappy windows | (TC 03/21/2023 11:11:16 AM, Samanian, Kamran) asked Nara to take the computer in the safe mode, not able to open the application in safe mode. rebooted computer and resolved the issue.
65937 | 193712 | CM | CMPLT | CMPLT | 2023-03-21 11:57 | 2023-03-21 11:59 | Channel error | (TC 03/21/2023 11:57:37 AM, Baghirzade, Nara) Noted a repeated failure of the pressure sensor disc; replaced the pressure sensor. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUIs and device cleaned. Fully functional.
66463 | 193713 | CM | CMPLT | CMPLT | 2023-03-21 12:31 | 2023-03-21 12:31 | Alarming, no screen | Collected error logs: A Medibus communication error has occurred. There is no data communication via Medibus. Check internal and external Medibus connection. Restart device. Perform an error-free power-up test. Perform an error-free system self-test. Tested unit as per Draeger instructios. Functional.
1066409 | 193714 | ACC | CMPLT | CMPLT | 2023-03-21 16:35 | 2023-05-11 17:50 | Incoming Inspection | (TC 03/22/2023 05:46:53 PM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 130 mOhms 18 uAmps
1066410 | 193715 | ACC | CMPLT | CMPLT | 2023-03-21 16:40 | 2023-05-11 17:51 | Incoming Inspection | (TC 03/22/2023 05:50:31 PM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066411 | 193716 | ACC | CMPLT | PMFAIL | 2023-03-21 16:42 | 2023-04-05 12:45 | Incoming Inspection | (TC 03/21/2023 12:43:45 PM, Cortez, John) Display defective out of the box. Artifacts on screen. Will request exchange
1066412 | 193717 | ACC | CMPLT | CMPLT | 2023-03-21 16:44 | 2023-05-11 18:20 | Incoming Inspection | (TC 03/22/2023 06:18:31 PM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 110 mOhms 18 uAmps
1066413 | 193718 | ACC | CMPLT | CMPLT | 2023-03-21 16:45 | 2023-05-11 18:23 | Incoming Inspection | (TC 03/22/2023 06:21:58 PM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066414 | 193719 | ACC | CMPLT | CMPLT | 2023-03-21 16:46 | 2023-05-12 08:52 | Incoming Inspection | (TC 03/22/2023 08:50:16 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 110 mOhms 19 uAmps
1066415 | 193720 | ACC | CMPLT | CMPLT | 2023-03-21 16:48 | 2023-05-12 08:54 | Incoming Inspection | (TC 03/22/2023 08:52:52 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 110 mOhms 19 uAmps
1066416 | 193721 | ACC | CMPLT | CMPLT | 2023-03-21 16:49 | 2023-05-12 09:00 | Incoming Inspection | (TC 03/22/2023 08:59:28 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066417 | 193722 | ACC | CMPLT | CMPLT | 2023-03-21 16:51 | 2023-05-12 09:03 | Incoming Inspection | (TC 03/22/2023 09:02:12 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066418 | 193723 | ACC | CMPLT | CMPLT | 2023-03-21 16:53 | 2023-05-12 09:05 | Incoming Inspection | (TC 03/22/2023 09:04:33 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066419 | 193724 | ACC | CMPLT | CMPLT | 2023-03-21 16:55 | 2023-05-12 09:08 | Incoming Inspection | (TC 03/22/2023 09:06:46 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 110 mOhms 19 uAmps
1066420 | 193725 | ACC | CMPLT | CMPLT | 2023-03-21 16:57 | 2023-05-12 09:09 | Incoming Inspection | (TC 03/22/2023 09:08:52 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066421 | 193726 | ACC | CMPLT | CMPLT | 2023-03-21 16:59 | 2023-05-12 09:12 | Incoming Inspection | (TC 03/22/2023 09:10:46 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 110 mOhms 19 uAmps
1066422 | 193727 | ACC | CMPLT | CMPLT | 2023-03-21 17:00 | 2023-05-12 09:13 | Incoming Inspection | (TC 03/22/2023 09:12:45 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 110 mOhms 19 uAmps
1066423 | 193728 | ACC | CMPLT | CMPLT | 2023-03-21 17:02 | 2023-05-12 09:17 | Incoming Inspection | (TC 03/22/2023 09:16:34 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 110 mOhms 19 uAmps
1066424 | 193729 | ACC | CMPLT | CMPLT | 2023-03-21 17:03 | 2023-05-12 09:20 | Incoming Inspection | (TC 03/22/2023 09:18:28 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066425 | 193730 | ACC | CMPLT | CMPLT | 2023-03-21 17:05 | 2023-05-12 09:22 | Incoming Inspection | (TC 03/22/2023 09:21:12 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066426 | 193731 | ACC | CMPLT | CMPLT | 2023-03-21 17:06 | 2023-05-12 09:24 | Incoming Inspection | (TC 03/22/2023 09:22:43 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1066427 | 193732 | ACC | CMPLT | CMPLT | 2023-03-21 17:07 | 2023-05-12 09:26 | Incoming Inspection | (TC 03/22/2023 09:25:11 AM, Cortez, John) Unboxed Visual and functional test ok Electrical safety passed, 120 mOhms 19 uAmps
1062393 | 193733 | CM | CMPLT | CMPLT | 2023-03-22 10:10 | 2023-03-22 10:16 | Ground pin on ower cord damaged | (TC 03/22/2023 10:09:48 AM, Cortez, John) Was made aware of this issue by someone from building services Replaced with spare cord
1047865 | 193734 | CM | CMPLT | CMPLT | 2023-03-22 10:30 | 2023-03-22 10:31 | Pump shut off before completing full feed volume | (TC 03/22/2023 10:30:12 AM, Baghirzade, Nara) When switching on the pump, I noticed that the battery was at 2% level. Battery was depleted which is why the pump shut off prematurely. Checked battery status and capacity; battery is in good condition. Ran a feed and no errors or issues noted. Occlusion alarm is functional.
1047865 | 193735 | CM | CMPLT | CMPLT | 2023-03-22 10:32 | 2023-03-22 10:33 | Updated date/time and configuration | (TC 03/22/2023 10:32:32 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
69118 | 193736 | PM | CMPLT | CMPLT | 2023-03-22 14:11 | 2023-03-22 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2023 02:11:05 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1049905 | 193737 | CM | CMPLT | CMPLT | 2023-03-22 17:14 | 2023-03-22 17:14 | Battery contacts bent, Detector drop, Exposure cable streched |  | Replaced battery contacts, performed full detector cal, replaced exposure cable.
1049905 | 193738 | CM | CMPLT | CMPLT | 2023-03-22 17:15 | 2023-03-27 15:14 | MSH R65 Collimator exposure switch not working. |  | Ordered collimator switch Replaced collimator switch.
32560 | 193739 | CM | CMPLT | CMPLT | 2023-03-22 18:30 | 2023-03-22 18:49 | IDD battery error | IDD Batttey error -Replaced both IDD batterires from stock - bad batteries, IDD battery error came back -Ordered new batteries and replaced them. -Error cleared and returned system for use.
62704 | 193740 | CM | CMPLT | CMPLT | 2023-03-23 08:04 | 2023-03-23 08:06 | Lead cables were damaged and frayed | (TC 03/22/2023 03:30:00 PM, Baghirzade, Nara) Replaced leads V1-V3 and RL; cables were damaged. ECG is fully functional.
1061333 | 193741 | FC | CMPLT | CMPLT | 2023-03-23 08:07 | 2023-03-23 08:07 | QC checklist | (TC 03/23/2023 08:07:00 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1059778 | 193742 | FC | CMPLT | CMPLT | 2023-03-23 08:08 | 2023-03-23 08:08 | QC checklist | (TC 03/23/2023 08:07:50 AM, Baghirzade, Nara) QC checklist completed. Returned to 11South
502044 | 193743 | CM | CMPLT | CMPLT | 2023-03-23 10:04 | 2023-03-23 10:04 | Arm Brake Tight | (TC 03/23/2023 10:04:31 AM, Paz, Ramon (Tino)) Adjusted disk brake.
502088 | 193744 | CM | CMPLT | CMPLT | 2023-03-23 10:06 | 2023-05-16 08:09 | Motion Error | (TC 03/23/2023 10:05:50 AM, Paz, Ramon (Tino)) Overcurrent error on SLD1. Scheduling for next week.
67913 | 193745 | CM | CMPLT | CMPLT | 2023-03-23 10:25 | 2023-03-23 10:25 | Latch recall | (TC 03/23/2023 10:25:16 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
67913 | 193746 | CM | CMPLT | CMPLT | 2023-03-23 10:26 | 2023-03-23 10:27 | Broken hood latch | (TC 03/23/2023 10:25:55 AM, Baghirzade, Nara) Replaced hood latch. Functional tests passed. Set temp to 35C and humidity to 75%; measured 35C and 77% humidity. Incubator is fully functional.
| 193747 | PA | CMPLT | CMPLT | 2023-03-23 10:30 | 2023-03-24 14:16 | Batteries for Transport Incubators
65191 | 193748 | CM | CMPLT | CMPLT | 2023-03-23 10:33 | 2023-04-05 14:56 | Knob broke off | (TC 03/23/2023 10:33:38 AM, Reid, Emmerson ) Was approached to 3DP a new knob for the device. Designed a new knob and printing now. (TC 04/04/2023 02:56:00 PM, Reid, Emmerson ) Attached the knob and gave it back to Steve from RT. Working now.
1062779 | 193749 | PM | CMPLT | CMPLT | 2023-03-23 12:34 | 2023-03-23 12:35 | PM | (TC 03/15/2023 03:00:00 PM, Baghirzade, Nara) PM completed by vendor (Medisoft) Smain Amoura assisted by distributor (Novus Medical). (TC 03/29/2023 08:40:33 AM, Baghirzade, Nara) Attached service report
1054986 | 193750 | CM | CMPLT | CMPLT | 2023-03-23 12:37 | 2023-03-23 12:40 | Consumables for respiratory monitoring | (TC 03/23/2023 12:37:10 PM, Baghirzade, Nara) Sarabeth requested 2 boxes of RAS-45 sensors and 2 boxes of RAS-125C sensors; dropped it off with Iliana at the 11N central station.
1053034 | 193751 | CM | CMPLT | CMPLT | 2023-03-23 14:59 | 2023-03-24 14:17 | Transport stopped working when in use | (TC 03/24/2023 02:16:42 PM, Ferreira, Qwynn) battery replaced The battery was worn so unable to hold charge. New battery was fully charged and inserted. Minor cleaning done as well. Functional test performed - will grab the unit for PM next week when more available to grab
32520 | 193752 | CM | CMPLT | CMPLT | 2023-03-23 15:37 | 2023-08-23 12:13 | MSH DR2 TUBE HEAD BETA "ROLL" FAULTY; DIFFICULT TO REACH POSITION, SOUND OF MOTOR BUT NOT ENGAGING GEAR | (TC 03/23/2023 03:36:35 PM, Whittington, Scott) MSH DR2 TUBE HEAD BETA "ROLL" FAULTY; DIFFICULT TO REACH POSITION, SOUND OF MOTOR BUT NOT ENGAGING GEAR | MSH DR2 TUBE HEAD BETA "ROLL" FAULTY; DIFFICULT TO REACH POSITION, SOUND OF MOTOR BUT NOT ENGAGING GEAR
502489 | 193753 | CM | CMPLT | CMPLT | 2023-03-23 15:46 | 2023-08-23 12:13 | MSH OPG SPP VOLUSON 10 PROBE IC5-9-D SN# 1005974WX9 FAULTY | (TC 03/24/2023 01:45:56 PM, Whittington, Scott) MSH OPG SPP VOLUSON 10 PROBE IC5-9-D SN# 1005974WX9 FAULTY REVIEWED WITH JONNY & VIJI NO FAULT FOUND | MSH OPG SPP VOLUSON 10 PROBE IC5-9-D SN# 1005974WX9 FAULTY
1054294 | 193754 | ACC | CMPLT | CMPLT | 2023-03-23 15:48 | 2023-08-23 12:14 | MSH OPG SPP GE VOLUSON 10 PROBE RIC6-12-D SN# 253790KR0 FAULTY | (TC 03/24/2023 01:46:49 PM, Whittington, Scott) MSH OPG SPP GE VOLUSON 10 PROBE RIC6-12-D SN# 253790KR0 FAULTY REVIEW WITH JONNY & VIJI NO FAULT FOUND (TC 04/27/2023 02:12:14 PM, Whittington, Scott) MSH OPG SPP GE VOLUSON 10 PROBE RIC 6 12-D TV PROBE ACCEPTANCE APRIL 27 2023 SN# 273602KR3 REPLACED SN# 253790KR0 | MSH OPG SPP GE VOLUSON 10 PROBE RIC6-12-D SN# 253790KR0 FAULTY MECHANICAL 3D WILL NOT OSCILLATE MSH OPG SPP GE VOLUSON 10 PROBE RIC 6 12-D TV PROBE ACCEPTANCE APRIL 27 2023 SN# 273602KR3 REPLACED SN# 253790KR0
500846 | 193755 | CM | CMPLT | CMPLT | 2023-03-23 15:51 | 2023-03-24 09:27 | MSH 5 US APLIO XG WILL NOT BOOT UP ; GREEN LIGHT ON , FAN ON, NO VIDEO | (TC 03/24/2023 09:26:04 AM, Fan, Tak) Confirmed issue. Turning off PS and hard restart did not help. Disconnected HDD and started system up. Screen displays boot up, eventually prompts no operating system. Shutdown, reconnected HDD, started up OK. | MSH 5 US APLIO XG WILL NOT BOOT UP ; GREEN LIGHT ON , FAN ON, NO VIDEO
1065736 | 193756 | PM | CMPLT | CMPLT | 2023-03-24 07:39 | 2023-10-02 14:26 | Annual PM | (TC 03/24/2023 07:46:16 AM, Baghirzade, Nara) SC 1002917026 (TC 05/11/2023 02:59:15 PM, Baghirzade, Nara) PM completed by Clarion field service tech, Shawn. Waiting for service report to attach. (TC 10/02/2023 02:23:48 PM, Baghirzade, Nara) After many attempts throughout the months to get the service report, I had reached out to Kim directly to obtain it. Attached the service report.
1056555 | 193757 | CM | CMPLT | CMPLT | 2023-03-24 08:54 | 2023-03-24 08:57 | No checkmark and not reporting to dashboard | (TC 03/24/2023 08:54:30 AM, Canjimavadivel, Kamallesh ) Reinstalled config and certs Replaced battery
1065164 | 193758 | CM | CMPLT | CMPLT | 2023-03-24 10:28 | 2023-04-28 13:11 | Not powering up | (TC 03/24/2023 01:42:05 PM, Ferreira, Qwynn) Tested theunit with new battery - not turning on AC light is on so power is being supplied Performed a hard restart on the monitor and it was able to turn on for a short period of time. Device is under warranty, will be sending it out RMA: 302931472 (TC 04/28/2023 01:10:22 PM, Ferreira, Qwynn) Unboxed and assembled the unit Installed the correct config files for use in the Oncology space. Functional Check - OK Electrical Safety - OK
1066314 | 193759 | ACC | CMPLT | CMPLT | 2023-03-24 14:09 | 2023-03-24 14:09 | Incoming inspection | (TC 03/15/2023 02:08:42 PM, Rendon Lopez , Elizabeth) Incoming inspection
1066315 | 193760 | ACC | CMPLT | CMPLT | 2023-03-24 14:09 | 2023-03-24 14:09 | Incoming inspection | (TC 03/15/2023 02:09:14 PM, Rendon Lopez , Elizabeth) Incoming inspection
1066316 | 193761 | ACC | CMPLT | CMPLT | 2023-03-24 14:10 | 2023-03-24 14:10 | Incoming inspection | (TC 03/15/2023 02:09:46 PM, Rendon Lopez , Elizabeth) Incoming inspection
1066317 | 193762 | ACC | CMPLT | CMPLT | 2023-03-24 14:10 | 2023-03-24 14:13 | Incoming inspection | (TC 03/15/2023 02:12:21 PM, Rendon Lopez , Elizabeth) Incoming inspection
1066318 | 193763 | ACC | CMPLT | CMPLT | 2023-03-24 14:13 | 2023-03-24 14:13 | Incoming inspection | (TC 03/15/2023 02:13:00 PM, Rendon Lopez , Elizabeth) Incoming inspection
1066319 | 193764 | ACC | CMPLT | CMPLT | 2023-03-24 14:13 | 2023-03-24 14:14 | Incoming inspection | (TC 03/15/2023 02:13:33 PM, Rendon Lopez , Elizabeth) Incoming inspection
1066320 | 193765 | ACC | CMPLT | CMPLT | 2023-03-24 14:14 | 2023-03-24 14:15 | Incoming inspection | (TC 03/15/2023 02:14:14 PM, Rendon Lopez , Elizabeth) Incoming inspection
1051427 | 193766 | CM | CMPLT | CMPLT | 2023-03-24 14:15 | 2023-04-11 16:20 | not detected | (TC 03/15/2023 02:15:32 PM, Rendon Lopez , Elizabeth) Confirmed issue, to be sent out under warranty (TC 04/11/2023 04:19:36 PM, Rendon Lopez , Elizabeth) Received, repaired under warranty, incoming inspection-OK
64992 | 193767 | CM | CMPLT | CMPLT | 2023-03-24 14:23 | 2023-03-24 14:24 | not working | (TC 03/17/2023 02:23:05 PM, Rendon Lopez , Elizabeth) Sent down from 15M-Antenatal. Missing batteries. New batteries, tested all functions OK
1056440 | 193768 | CM | CMPLT | CMPLT | 2023-03-24 14:29 | 2023-03-24 14:30 | Temp not measuring | (TC 03/24/2023 02:29:05 PM, Baghirzade, Nara) Temp was not measuring however i noticed that there was no configuration loaded to the Root. Loaded the MSH VSM config, uploaded network settings, and completed clinical config of root and radical. Temp probe was not properly seated in the connector; re-seated the temp probe and its measuring accurately. NIBP and SpO2 both functional.
| 193769 | ACC | CMPLT | CMPLT | 2023-03-24 15:10 | 2023-03-27 12:07 | MSH OPG SPP DEMO CANON APLIO PRISM FOR SYMPOSIUM ACCEPTANCE | (TC 03/27/2023 12:06:26 PM, Whittington, Scott) MSH OPG SPP DEMO CANON APLIO PRISM FOR SYMPOSIUM ACCEPTANCE TUS-AI800 SN# SHF2222016 75 mOHM 14 uA ETL MARCH 24 PICKUP MARCH 27 | MSH OPG SPP DEMO CANON APLIO PRISM FOR SYMPOSIUM ACCEPTANCE TUS-AI800 SN# SHF2222016 75 mOHM 14 uA ETL
1054986 | 193770 | CM | CMPLT | NFF | 2023-03-27 08:34 | 2023-03-27 08:36 | Not holding charge | (TC 03/27/2023 08:34:52 AM, Baghirzade, Nara) Power cord was tightly wrapped around the stand and was not properly connected at the back of the device. Root and radical battery was completely depleted; charged over the weekend. Functional tests for NIBP, temp, and SpO2 passed. No fault found with device or battery.
1054981 | 193771 | PA | CMPLT | CMPLT | 2023-03-27 08:50 | 2023-05-09 07:07 | Ordering rollstand | (TC 03/27/2023 11:27:56 AM, Baghirzade, Nara) SC 1002918804 (TC 05/09/2023 07:06:42 AM, Baghirzade, Nara) PO 7202186408 | Basket and stand are damaged; basket broken.
1066454 | 193772 | ACC | CMPLT | CMPLT | 2023-03-27 09:22 | 2023-03-27 09:40 | Incoming inspection | (TC 03/27/2023 09:38:41 AM, Baghirzade, Nara) Inspection and asset tagging completed. Attached KCI quality inspection card. Sent VAC in its case to 14S Attention: Monica Frecea
1066460 | 193773 | ACC | CMPLT | CMPLT | 2023-03-27 09:44 | 2023-03-27 09:44 | Incoming inspection and asset tagging | (TC 03/27/2023 09:43:42 AM, Baghirzade, Nara) Inspection and asset tagging completed. Attached KCI quality inspection card. Sent VAC in its case to 14S Attention: Monica Frecea
1066462 | 193774 | ACC | CMPLT | CMPLT | 2023-03-27 09:53 | 2023-03-27 09:54 | Incoming inspection and asset tagging | (TC 03/27/2023 09:53:20 AM, Baghirzade, Nara) Inspection and asset tagging completed. Attached KCI quality inspection card. Sent VAC in its case to 14S Attention: Monica Frecea
1049902 | 193775 | CM | CMPLT | CMPLT | 2023-03-27 10:00 | 2023-03-27 10:01 | Touch screen cursor jumping all over | (TC 03/24/2023 03:59:00 PM, Fan, Tak) Removed touchscreen. Cleaned. Calibrated.
65931 | 193776 | FC | CMPLT | CMPLT | 2023-03-27 11:01 | 2023-03-27 11:02 | Not specified error; functional check | (TC 03/27/2023 11:00:53 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1055751 | 193777 | FC | CMPLT | CMPLT | 2023-03-27 11:02 | 2023-03-27 11:03 | No error specified; functional check | (TC 03/27/2023 11:02:14 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
65910 | 193778 | FC | CMPLT | CMPLT | 2023-03-27 11:03 | 2023-03-27 11:06 | No error specified; functional check | (TC 03/27/2023 11:05:26 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
32560 | 193779 | CM | CMPLT | CMPLT | 2023-03-27 11:13 | 2023-03-24 11:12 | Application crash reported | Application crash reported Service completed troubleshooting CT Apllication restored CT system tested for intermittant hardware failure CT phantom tested OK
| 193780 | CM | CMPLT | CMPLT | 2023-03-27 12:18 | 2023-03-27 12:19 | Biomed Huddle | (TC 03/27/2023 12:18:38 PM, Baghirzade, Nara) Biomed Huddle
67913 | 193781 | CM | CMPLT | CMPLT | 2023-03-27 12:48 | 2023-03-27 12:49 | Porthole door came off the hinges | (TC 03/27/2023 12:48:43 PM, Baghirzade, Nara) Repair the porthole door. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 73.5% humidity. Incubator is fully functional.
66043 | 193782 | CM | CMPLT | CMPLT | 2023-03-27 13:43 | 2023-03-27 13:45 | Does not turn on | (TC 03/27/2023 01:42:53 PM, Baghirzade, Nara) Doesn't switch on but the AC LED is lit. Replaced front case as part of recall. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 22 mohms Leakage current: 92.1 uA
| 193783 | CM | CMPLT | CMPLT | 2023-03-27 13:44 | 2023-08-23 12:14 | MSH 15 L&D OBU ELECTRICAL INTERFERENCE PICKED UP BY SEVERAL IU22 US | (TC 03/27/2023 01:45:35 PM, Whittington, Scott) MSH 15 L&D OBU ELECTRICAL INTERFERENCE PICKED UP BY SEVERAL IU22 US SPOKE TO PLANT ENG POSSIBILTIES CCU CATH LAB DIRECTLYOVERHEAD ON 16 MOTOR INSTALLED AND TUNED ON LOWER LEVEL -CLINICAL UNPLUGGED ALL OTHER EQUIPMENT IN SUITE INCLUDING HEATER? & ESU | MSH 15 L&D OBU ELECTRICAL INTERFERENCE PICKED UP BY SEVERAL IU22 US
| 193784 | CM | CMPLT | CMPLT | 2023-03-27 13:51 | 2023-03-27 13:51 | biomed hudle | (TC 03/27/2023 01:51:30 PM, Samanian, Kamran) done!
1047861 | 193785 | FC | CMPLT | CMPLT | 2023-03-27 14:13 | 2023-03-27 14:13 | Updated time and config | (TC 03/27/2023 02:12:50 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
68814 | 193786 | CM | CMPLT | CMPLT | 2023-03-27 14:25 | 2023-04-03 12:31 | Ordering logic board | (TC 03/27/2023 02:41:14 PM, Baghirzade, Nara) SC 1002919240 (TC 04/03/2023 12:30:04 PM, Baghirzade, Nara) PO 7202187007 | Connector on the logic board is broken; parts of the PCB board are damaged
| 193787 | CM | CMPLT | CMPLT | 2023-03-27 15:12 | 2023-03-27 15:12 | working o n capturing licenses in MSH PICiX | (TC 03/27/2023 03:11:48 PM, Samanian, Kamran) done!
1047866 | 193788 | FC | CMPLT | CMPLT | 2023-03-28 07:50 | 2023-03-28 07:51 | Updated time/date and config; functional check | (TC 03/28/2023 07:50:35 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1053031 | 193789 | PM | CMPLT | CMPLT | 2023-03-28 10:13 | 2023-03-28 10:14 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/28/2023 10:12:50 AM, Ferreira, Qwynn) PM Passed Reconditioned battery Functional Test - OK
| 193790 | CM | CMPLT | CMPLT | 2023-03-28 12:03 | 2023-03-31 10:29 | Order batteries for transport incubator suction and monitors | (TC 03/28/2023 12:15:30 PM, Chukri, Samir) replaced order for transport incubator monitor and suction . (TC 03/31/2023 10:27:36 AM, Chukri, Samir) Received all batteries
1066474 | 193791 | ACC | CMPLT | CMPLT | 2023-03-28 12:21 | 2023-03-28 12:23 | Incoming Inspection | (TC 03/28/2023 12:22:01 PM, Reid, Emmerson ) Electrical Safety Complete in proceedure. Loaner device due to ongoing issues. Incoming complete.
1066475 | 193792 | ACC | CMPLT | CMPLT | 2023-03-28 12:30 | 2023-03-28 12:32 | Incoming Inspection | (TC 03/28/2023 12:31:03 PM, Reid, Emmerson ) Electrical safety completed in proceedure. Incoming complete. Loaner unit as part of ongoing contamination issue.
1066476 | 193793 | ACC | CMPLT | CMPLT | 2023-03-28 12:36 | 2023-03-28 12:37 | Incoming Inspection | (TC 03/28/2023 12:36:43 PM, Reid, Emmerson ) Electrical safety completed in proceedure. Incoming complete. Loaner unit as part of ongoing contamination issue.
502048 | 193794 | CM | CMPLT | CMPLT | 2023-03-28 14:03 | 2023-03-29 15:53 | System fails to initialize | (TC 03/28/2023 02:02:55 PM, Fan, Tak) Restarted, but failed to boot. Contacted Philips. Remote support suggested trying Ctrl+Alt+R during progress bar to cancel into keybased login. Logged in and ran batch command "C:\system32\deletepatientdata". Did not resolve issue. Paged to FSE. Sara reloaded SW. REF 0121444128. (TC 03/28/2023 03:52:03 PM, Fan, Tak) Sara deleted a registry file and replaced it with a default file. Reloaded presets. | System fails to boot. Stops at progress bar on process 48 then gives error 511.
65612 | 193795 | CM | CMPLT | CMPLT | 2023-03-28 14:47 | 2023-03-28 14:58 | SpO2 not working | (TC 03/28/2023 02:47:06 PM, Canjimavadivel, Kamallesh ) Found signs of chemical corrosion in connector on extension. Replaced extension, tested ok
1066482 | 193796 | ACC | CMPLT | CMPLT | 2023-03-28 15:10 | 2023-03-29 10:44 | MSH SITE RITE 8 ACCEPTANCE | (TC 03/29/2023 10:44:18 AM, Whittington, Scott) MSH SITE RITE 8 ACCEPTANCE (TC 03/29/2023 10:44:42 AM, Dolinar, John) MSH SITE RITE 8 ACCEPTANCE | MSH SITE RITE 8 ACCEPTANCE
66765 | 193797 | CM | CMPLT | CMPLT | 2023-03-28 16:56 | 2023-06-30 11:15 | "PCA button not working" | (TC 03/28/2023 04:56:27 PM, Cortez, John) Female side of dose cord connector damaged. Pin stuck in port (TC 06/30/2023 11:13:49 AM, Cortez, John) Returned from vendor. Functional check ok (see procedure)
1053391 | 193798 | FC | CMPLT | CMPLT | 2023-03-29 07:52 | 2023-03-29 07:53 | Functional check before patient use | (TC 03/29/2023 07:52:23 AM, Baghirzade, Nara) Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 126 mmHg Set 200 mmHg - Measured 199 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
| 193799 | CM | CMPLT | CMPLT | 2023-03-29 07:54 | 2023-03-29 07:54 | ZOLL auto pulse meeting | (TC 03/28/2023 07:53:58 AM, Samanian, Kamran) done!
| 193800 | CM | CMPLT | CMPLT | 2023-03-29 07:55 | 2023-03-28 07:55 | meet with sean in OR for long cord | (TC 03/28/2023 07:54:59 AM, Samanian, Kamran) verified that IV pumps and blanket warmer needs long power cord,
1061220 | 193801 | CM | CMPLT | CMPLT | 2023-03-29 08:52 | 2023-03-29 08:53 | Olympus Console is on log on screen, touch does not work | (TC 03/29/2023 08:52:08 AM, Samanian, Kamran) verified the problem, VMC3 rebooted but not successful. The root cause is Cat6 cable connected to the silver box in back of console, Reseated the connector, powered off and on the silver box, it came back up in 10 sec and all worked fine
| 193802 | CM | CMPLT | CMPLT | 2023-03-29 09:00 | 2023-05-04 08:49 | Alaris pumps spare parts | Received power cord straps and latch modules All received
| 193803 | CM | CMPLT | CMPLT | 2023-03-29 09:18 | 2023-03-29 09:19 | Manual BP leaking and wrong cuff connectors | (TC 03/29/2023 09:18:03 AM, Baghirzade, Nara) Replaced coiled cable, cable, and cuffs. Provided large adult, adult, and long adult sizes. Checked BP; functional and no leaks. SN: 111115131154
1056420 | 193804 | CM | CMPLT | CMPLT | 2023-03-29 10:22 | 2023-03-29 10:25 | "BP not working" | (TC 03/28/2023 10:21:48 AM, Cortez, John) Replaced BP hose, SPO2 sensor straightened basket All parameters tested ok
1066487 | 193805 | CM | CMPLT | CMPLT | 2023-03-29 10:22 | 2023-03-29 10:23 | Initial Accept and config | (TC 03/29/2023 10:22:03 AM, Samanian, Kamran) this device is the replacement for the root , asset# 105548 which was sent for repair. this new one configured with wks file and connected to wifi. tested and passed
1056455 | 193806 | CM | CMPLT | CMPLT | 2023-03-29 10:42 | 2023-04-13 08:13 | ordering Masimo hoses | received power cords PO#7202188575
1053348 | 193807 | CM | CMPLT | CMPLT | 2023-03-29 10:47 | 2023-07-11 07:33 | MSH BI 1272 H3 12th Floor - North Side CRACKED DETECTOR HOUSING (BOTTOM), BLOOD FROM BIOPSY PROCEDURES | (TC 03/29/2023 10:47:05 AM, Whittington, Scott) MSH BI 1272 H3 12th Floor - North Side CRACKED DETECTOR HOUSING (BOTTOM), BLOOD FROM BIOPSY PROCEDURES (TC 03/29/2023 10:47:15 AM, Dolinar, John) MSH BI 1272 H3 12th Floor - North Side CRACKED DETECTOR HOUSING (BOTTOM), BLOOD FROM BIOPSY PROCEDURES | MSH BI 1272 H3 12th Floor - North Side CRACKED DETECTOR HOUSING (BOTTOM), BLOOD FROM BIOPSY PROCEDURES
66456 | 193808 | CM | CMPLT | CMPLT | 2023-03-29 11:18 | 2023-03-29 11:18 | air leak | removed cable tie and replaced air filter. Tested. functional.
| 193809 | PA | CMPLT | CMPLT | 2023-03-29 14:41 | 2023-05-05 08:46 | IV Pole Setup for Transport Incubators
1066459 | 193810 | CM | CMPLT | CMPLT | 2023-03-29 15:46 | 2023-03-29 15:46 | Missing button | (TC 03/27/2023 03:45:45 PM, Rendon Lopez , Elizabeth) Replaced with spare button, added to TMS, tested-OK
1049940 | 193811 | CM | CMPLT | CMPLT | 2023-03-29 15:54 | 2023-03-29 15:55 | Can't burn CD. | (TC 03/29/2023 03:53:50 PM, Fan, Tak) Burn write failed. Worked yesterday apparently. Burned on external DVD burner, OK. Checked CD, OK. Checked CD on original DVD drive, OK. Not sure what the issue is with burning.
68603 | 193812 | CM | CMPLT | CMPLT | 2023-03-29 16:19 | 2023-04-13 14:39 | sent down by OR, "Syr needs to be calibrated" | (TC 03/28/2023 04:19:27 PM, Cortez, John) Downloaded error/event logs, 351.666 plunger position error, Ran test infusion, unable to replicate error Rear/front case/ left handle have damage, Replaced (plastics recall) Calibration tasks completed in ASM (TC 03/31/2023 02:37:50 PM, Cortez, John) PM completed successfully Ran test infusions, no errors
66045 | 193813 | CM | CMPLT | CMPLT | 2023-03-29 16:22 | 2023-03-29 16:24 | Incoming inspection from Sick kids | (TC 03/27/2023 04:23:08 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67088 | 193814 | CM | CMPLT | CMPLT | 2023-03-29 16:25 | 2023-03-29 16:28 | Not working | (TC 03/27/2023 04:27:06 PM, Rendon Lopez , Elizabeth) BP set to neonatal. Changed setting, replaced worn cuff. Tested all functions-OK
65962 | 193815 | CM | CMPLT | CMPLT | 2023-03-29 16:34 | 2023-03-29 16:34 | Incoming inspection from sick kids | (TC 03/29/2023 04:33:44 PM, Rendon Lopez , Elizabeth) Check in tests done, all tests passed
65978 | 193816 | CM | CMPLT | CMPLT | 2023-03-29 16:41 | 2023-03-29 16:42 | incoming inspection from sick kids | (TC 03/27/2023 04:41:39 PM, Rendon Lopez , Elizabeth) Check in tests done, all passed.
67873 | 193817 | CM | CMPLT | CMPLT | 2023-03-29 16:44 | 2023-03-29 16:45 | PCEA button doesn't work | (TC 03/28/2023 04:44:29 PM, Rendon Lopez , Elizabeth) Bolus cord had pins bent, adjusted, tested-OK. Run on user mode, all functions working
1054419 | 193818 | CM | CMPLT | NFF | 2023-03-29 16:48 | 2023-03-29 16:51 | Not picking up contractions | (TC 03/28/2023 04:50:12 PM, Rendon Lopez , Elizabeth) Tested, no fault found
1059659 | 193819 | CM | CMPLT | CMPLT | 2023-03-29 16:51 | 2023-03-29 16:52 | Missing knob | (TC 03/28/2023 04:51:17 PM, Rendon Lopez , Elizabeth) Replaced button with spare, tested-OK
1060134 | 193820 | CM | CMPLT | CMPLT | 2023-03-29 16:52 | 2023-03-29 16:58 | Missing button | (TC 03/28/2023 04:56:10 PM, Rendon Lopez , Elizabeth) Replaced button with spare, tested-OK
64187 | 193821 | CM | CMPLT | CMPLT | 2023-03-29 16:59 | 2023-03-29 17:00 | Middle heater damaged | (TC 03/28/2023 04:58:50 PM, Rendon Lopez , Elizabeth) Bag sensor wires exposed. Replaced sensor with spare. Tested-OK
64186 | 193822 | CM | CMPLT | CMPLT | 2023-03-29 17:01 | 2023-03-29 17:04 | Heater not detecting bag properly | (TC 03/28/2023 05:03:18 PM, Rendon Lopez , Elizabeth) Replaced bag sensor with spare one. Tested-OK
1066485 | 193823 | CM | CMPLT | CMPLT | 2023-03-29 17:04 | 2023-03-29 17:05 | Adding to TMS | (TC 03/28/2023 05:04:17 PM, Rendon Lopez , Elizabeth) Added fluid warmer to TMS. L&D OR 1
68999 | 193824 | CM | CMPLT | CMPLT | 2023-03-29 17:07 | 2023-03-29 17:09 | Not charging | (TC 03/28/2023 05:07:39 PM, Rendon Lopez , Elizabeth) Damaged power cord, worn screw. Trying to remove screw (TC 03/28/2023 05:07:39 PM, Rendon Lopez , Elizabeth) Replaced power cord, strap screw. Charged. Check in tests, all passed
66049 | 193825 | CM | CMPLT | CMPLT | 2023-03-29 17:11 | 2023-03-29 17:11 | Sent down with broken LVP | (TC 03/29/2023 05:10:46 PM, Rendon Lopez , Elizabeth) Check in tests, all passed
69429 | 193826 | CM | CMPLT | CMPLT | 2023-03-29 17:11 | 2023-03-29 17:15 | Bottom connection broken off | (TC 03/29/2023 05:14:11 PM, Rendon Lopez , Elizabeth) Moduel latch and rear case damaged. Replaced module latch. Rear case and LCD under recall. PM completed, all tests passed
69630 | 193827 | CM | CMPLT | CMPLT | 2023-03-29 17:18 | 2023-03-29 17:22 | Damaged front case | (TC 03/29/2023 05:21:11 PM, Rendon Lopez , Elizabeth) Replaced front case/keypad. Replaced LCD under recall. PM completed all tests passed
1059778 | 193828 | FC | CMPLT | CMPLT | 2023-03-30 07:14 | 2023-03-30 07:14 | QC checklist | (TC 03/30/2023 07:13:51 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
| 193829 | CM | CMPLT | CMPLT | 2023-03-30 08:57 | 2023-04-21 10:36 | WelchAllyn SPO2 parts for ENDO | (TC 04/03/2023 09:35:23 AM, Samanian, Kamran) received the quote, S/C#1002924331 (TC 04/03/2023 09:35:23 AM, Samanian, Kamran) received the quote, S/C#1002924331 PO# 7202190096 (TC 04/21/2023 10:35:18 AM, Samanian, Kamran) delivered to Anita ( dropped off in the cabinet )
52170 | 193830 | CM | CMPLT | CMPLT | 2023-03-30 09:26 | 2023-03-30 09:33 | Latch recall | (TC 03/30/2023 09:32:53 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
1050729 | 193831 | CM | CMPLT | CMPLT | 2023-03-30 09:31 | 2023-05-08 08:50 | Not turning on | (TC 03/30/2023 09:31:57 AM, Rendon Lopez , Elizabeth) Not turning on, no suction at all. confirmed problem, tried new power cord and filter, not resolved. To be sent out for repair under consignment (TC 03/31/2023 09:46:15 AM, Rendon Lopez , Elizabeth) Shipped for repair under PO 7500186809. Conmed to ship in loaner (TC 05/04/2023 08:49:28 AM, Rendon Lopez , Elizabeth) Incoming inspection, tested-OK. Returned to L&D OR. See attached report
52170 | 193832 | PM | CMPLT | CMPLT | 2023-03-30 09:34 | 2023-03-30 09:40 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 03/30/2023 09:39:33 AM, Baghirzade, Nara) PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 77%. Functional tests passed. Electrical safety results - Leakage current = 105.0 uA Power cord resistance = 43.1 mohms
| 193833 | CM | CMPLT | CMPLT | 2023-03-30 09:45 | 2023-03-30 09:46 | Screw missing on slit lamp arm stand | (TC 03/30/2023 09:44:55 AM, Baghirzade, Nara) Henna reached out regarding a loose arm on the slit lamp stand; a screw was missing on the bottom part of the arm. Samir and I were able to find a screw of the same size and threading but it is only a temp fix. Let Henna know this and waiting for her to send over the vendor contact info so we can order the correct screw.
1066598 | 193834 | ACC | CMPLT | CMPLT | 2023-03-30 10:17 | 2023-03-30 10:19 | Incoming Inspection | (TC 03/30/2023 10:17:55 AM, Reid, Emmerson ) Certificate of Calibration attached to WO Incoming Complete.
1066599 | 193835 | ACC | CMPLT | CMPLT | 2023-03-30 10:32 | 2023-03-30 10:33 | Incoming Inspection | (TC 03/30/2023 10:32:57 AM, Reid, Emmerson ) Certificate of Calibration attached to WO Incoming Complete.
1066600 | 193836 | ACC | CMPLT | CMPLT | 2023-03-30 10:41 | 2023-03-30 10:48 | Incoming Inspection | (TC 03/30/2023 10:48:04 AM, Reid, Emmerson ) Certificate of Calibration attached to WO Incoming Complete.
1066604 | 193837 | ACC | CMPLT | CMPLT | 2023-03-30 10:56 | 2023-03-30 11:00 | Incoming Inspection | (TC 03/30/2023 10:59:50 AM, Reid, Emmerson ) Certificate of Calibration attached to WO Incoming Complete.
1066605 | 193838 | ACC | CMPLT | CMPLT | 2023-03-30 11:04 | 2023-03-30 11:09 | Incoming Inspection | (TC 03/30/2023 11:08:39 AM, Reid, Emmerson ) Certificate of Calibration attached to WO Incoming Complete.
| 193839 | OC | CMPLT | CMPLT | 2023-03-30 13:31 | 2023-04-17 11:11 | Adeel Alam On-Call; March 3 - March 9, 2023 TEST | (TC 03/11/2023 01:42:31 PM, Alam, Adeel) Got a call regarding an overview station in the NICU. Came into the hospital to restart the system. System was working after the restart.
| 193840 | PA | CMPLT | CMPLT | 2023-03-30 14:59 | 2023-04-24 10:41 | Ording parts for OPG | (TC 04/05/2023 03:09:33 PM, Reid, Emmerson ) Recieved the light bulb. (TC 04/06/2023 11:09:07 AM, Reid, Emmerson ) Replaced the light bulb, lamp is working now. Also fixed a BP monitor that needed it's battery changed. (TC 04/24/2023 10:40:09 AM, Reid, Emmerson ) Recieved everything.
1059763 | 193841 | CM | CMPLT | CMPLT | 2023-03-30 17:06 | 2023-07-11 07:33 | MSH 15 L&D TE7D SYSTEM ASKING FOR PW | (TC 03/30/2023 05:06:00 PM, Whittington, Scott) MSH 15 L&D TE7D SYSTEM ASKING FOR PW | MSH 15 L&D TE7D SYSTEM ASKING FOR PW
1047866 | 193842 | CM | CMPLT | CMPLT | 2023-03-31 08:40 | 2023-03-31 08:44 | Failed occlusion alarm during functional check | (TC 03/31/2023 08:40:17 AM, Baghirzade, Nara) When running an infusion/feed, the occlusion alarm failed and instead got alarm for "coarse travel, check plunger/clutch". Replaced the clutch as it was very worn out. Completed full calibration of: -syringe size and position -force sensor -plunger position sensor -syringe size sensor Ran feed with 60mL syringe; completed successfully and occlusion alarm was functional. Ran a second feed with 30mL syringe; completed successfully and occlusion alarm was functional.
| 193843 | PA | CMPLT | CMPLT | 2023-03-31 10:01 | 2023-04-06 13:13 | Sterile Water for Altrix
1054184 | 193844 | CM | CMPLT | CMPLT | 2023-03-31 11:46 | 2023-05-29 08:33 | not working | (TC 03/30/2023 11:45:41 AM, Rendon Lopez , Elizabeth) Alarm equipment malfunction. kept to send out for repair (TC 05/23/2023 08:32:41 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1056128 | 193845 | CM | CMPLT | CMPLT | 2023-03-31 11:48 | 2023-05-29 08:55 | Not working | (TC 03/30/2023 11:47:47 AM, Rendon Lopez , Elizabeth) Not FHR on monitor, kept to send out for repair (TC 05/23/2023 08:54:49 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1053263 | 193846 | CM | CMPLT | CMPLT | 2023-03-31 11:49 | 2023-03-31 11:49 | not working | (TC 03/30/2023 11:48:46 AM, Rendon Lopez , Elizabeth) No FHR on monitor. previously repaired, old device. Retired from service
502089 | 193847 | RQ | CMPLT | CMPLT | 2023-03-31 13:10 | 2023-05-16 08:09 | Symbia 2 unresponsive hand control | (TC 03/31/2023 08:12:14 AM, Green, Paul) Customer had persistant error. Worked with Siemens from TWH and logged onto system. Checked error log and found: MSH Symbia2 SLD1 persistent error Cleared error and will look into this further when time permits.
65893 | 193848 | CM | CMPLT | CMPLT | 2023-03-31 13:12 | 2023-03-31 13:16 | Not sensing the pressure disc | (TC 03/31/2023 01:12:42 PM, Baghirzade, Nara) Ran analog test and the pressure sensor failed. Pressure sensor was previously re-seated. Replaced the pressure sensor. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
68932 | 193849 | CM | CMPLT | CMPLT | 2023-03-31 13:16 | 2023-03-31 13:22 | Not charging | (TC 03/31/2023 01:21:04 PM, Baghirzade, Nara) Defective battery; replaced battery. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 20 mohms Leakage current: 65.5 uA
68804 | 193850 | CM | CMPLT | NFF | 2023-03-31 13:25 | 2023-03-31 13:26 | System error reported | (TC 03/31/2023 01:25:24 PM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 50 mohms Leakage current: 64.8 uA No fault found.
1046196 | 193851 | CM | CMPLT | CMPLT | 2023-03-31 15:48 | 2023-03-31 15:48 | Broken table side cover | Table side cover broken - most likely by being hit with strecher Part was ordered from Siemens -Replaced and removed all tapes and adhesives that was holding it together.
26106 | 193852 | PM | ACTIV | ISSUE | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26181 | 193853 | PM | CANCL | CANCL | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: LASERS, SURGICAL, CARBON DIOXIDE
50806 | 193854 | PM | CNP | ENL | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: Bladder Scanner
62974 | 193855 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-06-20 07:37 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/20/2023 07:36:46 AM, Rhuland, Courtney) PM complete Filters cleaned Light intensity check - OK Electrical safety - OK
62975 | 193856 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2024-05-06 10:45 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/06/2024 10:17:16 AM, Rhuland, Courtney) PM completed Replaced missing scews Replaced power cord Irradiance - 33uW/cm^2/nm Electrical safety - OK Leakage current - 46.6uA Ground leakage - 0.210 mOhms
62985 | 193857 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-05-12 12:20 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/12/2023 11:19:22 AM, Ferreira, Qwynn) PM Passed Removed cover and replaced the high.low intensity light switch Removed fans and filters to clean. Washed the filters and dried them Tested the light and is working, it was flickering previously. All tests passed Electrical Saftey - OK 84.1 mohms, 63.7 uA High intensity - 33.1 Low intensity - 13.3
63011 | 193858 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-05-12 15:14 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/12/2023 03:13:25 PM, Ferreira, Qwynn) PM Passed Hours: 5307 Avg Intensity: 47.6 Electrical Safety - OK
63023 | 193859 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-05-15 10:10 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/15/2023 10:01:37 AM, Rhuland, Courtney) Repaired bottom chasis screw housings with epoxy PM complete Average light intensity - 52 µW•cm-2•nm-1 Leakage current - 11.2 uA Ground resistance - 118 mOhms Hours - 7973
63072 | 193860 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2024-03-05 08:39 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 02/29/2024 01:58:29 PM, Chukri, Samir) work with Qwynn, Replaced battery and functional check . (TC 03/05/2024 08:38:30 AM, Ferreira, Qwynn) workedwith samir to perform the PM. Functional test - ok
63094 | 193861 | PM | ACTIV | ISSUE | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63301 | 193862 | PM | CANCL | DUP | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM changed to 2 year schedule
63874 | 193863 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-04-28 15:24 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/28/2023 03:22:47 PM, Baghirzade, Nara) Replaced humidity sensor. PM completed. VRef, Mains V, LV were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 74.9%. Functional tests passed. Electrical safety results - Leakage current = 29.1 uA Power cord resistance = 125 mohms
64495 | 193864 | PM | ACTIV | ISSUE | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
65877 | 193865 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-05-09 10:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/09/2023 10:30:36 AM, Rhuland, Courtney) Channel error present Pressure sensor replaced Calibration and check-in complete All tests passed Refer to ASM for details
65880 | 193866 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-10-20 11:43 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 10/20/2023 11:42:21 AM, Campbell, Connor) PM complete/passed, refer to ASM
65896 | 193867 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-09-07 09:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/06/2023 09:31:34 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
65899 | 193868 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-06-29 09:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/28/2023 09:20:41 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
65917 | 193869 | PM | CANCL | CANCL | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65918 | 193870 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-08-08 13:16 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/08/2023 01:15:40 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65933 | 193871 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-07-12 09:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/12/2023 09:11:16 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65947 | 193872 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-07-12 14:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/12/2023 02:14:25 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65970 | 193873 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-06-14 07:34 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/14/2023 07:29:27 AM, Rhuland, Courtney) Left side of handle replaced PM complete All tests passed Functional check - OK Refer to ASM for details
65977 | 193874 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-07-12 09:08 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/12/2023 09:07:21 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65979 | 193875 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-09-14 11:45 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/14/2023 11:44:42 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65987 | 193876 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-05-08 13:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/08/2023 01:42:19 PM, Rhuland, Courtney) Pressure sensor replaced Reseated spring holding disk sensor due to error Rear case replaced (recall) Calibration and check in completed All tests passed Refer to ASM for details
65989 | 193877 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-07-07 09:07 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/07/2023 09:06:12 AM, Baghirzade, Nara) Pressure sensor disc failed during PM; recalibrated the pressure sensor disc. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
66011 | 193878 | PM | CMPLT | PMFAIL | 2023-04-01 03:18 | 2024-03-14 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
66019 | 193879 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-05-16 15:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/16/2023 03:20:58 PM, Rhuland, Courtney) Cracked front and rear case replaced (recall) Replaced female IUI connector Battery replaced and conditioned PM completed All tests passed Refer to ASM for details
66031 | 193880 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-08-04 13:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/04/2023 01:28:02 PM, Canjimavadivel, Kamallesh ) Battery condition complete New capacity 4019 mamphr Cleaned PC PM tests pass
66032 | 193881 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-07-20 06:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/19/2023 08:12:46 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 93 mohms Leakage current: 64.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 07/20/2023 06:24:05 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4091 mamphr PM completed.
66048 | 193882 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-07-18 07:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/18/2023 07:42:40 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 38 mohms Leakage current: 89.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66052 | 193883 | PM | CANCL | CANCL | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66061 | 193884 | PM | CANCL | CANCL | 2023-04-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66456 | 193885 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-04-27 08:40 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 75uA and 83 ohm.
66469 | 193886 | PM | CMPLT | CMPLT | 2023-04-01 03:18 | 2023-04-28 08:44 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 81uA and 89 ohm.
66683 | 193887 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-21 09:44 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/21/2023 09:43:21 AM, Baghirzade, Nara) PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C. Functional tests passed. Electrical safety results - Leakage current = 89.1 uA Power cord resistance = 185 mohms
66693 | 193888 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-13 13:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/13/2023 10:05:21 AM, Rhuland, Courtney) Hood replaced with spare Recall parts replaced Screw cover replaced LV calibrated Main voltage calibrated Scale calibrated (TC 07/13/2023 01:02:27 PM, Rhuland, Courtney) PM complete RH calibrated Scale validated Functional check - OK Electrical safety - OK
66715 | 193889 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
66724 | 193890 | PM | CANCL | DUP | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM changed to 2 year schedule
66836 | 193891 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | done 2 years not every year
66836 | 193892 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | BP Check
66837 | 193893 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | BP Check
67373 | 193894 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67571 | 193895 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
67572 | 193896 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
67872 | 193897 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-19 09:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/12/2023 09:05:56 AM, Rendon Lopez , Elizabeth) Updated software and library. PM completed, all tests passed
67919 | 193898 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-28 14:22 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/28/2023 02:19:53 PM, Ferreira, Qwynn) Replaced Hood Replaced the recall parts Replaced the RH probe PM Passed Calibrated RH Calibrated Main voltage Calibrated line Voltage Calibrated scale Calibrated System Votlage Functional Check - OK Electrical Safety - OK (TC 04/28/2023 02:19:53 PM, Ferreira, Qwynn) Replaced Hood Replaced the recall parts Replaced the RH probe PM Passed Calibrated RH Calibrated Main voltage Calibrated line Voltage Calibrated scale Calibrated System Votlage Functional Check - OK Electrical Safety - OK
67920 | 193899 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-04 08:38 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/04/2023 08:35:16 AM, Ferreira, Qwynn) Came down with note stating door would not close and alarming when turning on. A recall door was used on a non-recall incubator which prohibited the door from closing Installed recall parts Re-seated fan as it was not seated correctly. Calibrated scale Calibrated LV Calibrated system voltage Calibrated mains voltage Electrical Safety - OK Functional Check - OK
67947 | 193900 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-09 12:49 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/09/2023 09:34:12 AM, Rhuland, Courtney) Recall parts replaced Chipped door replaced with spare Replaced 2 side wall hinges (TC 06/09/2023 12:07:18 PM, Rhuland, Courtney) During functional check, alarm sounds stating "Fan always in high speed" Heat sink temperature sensor replaced, does not clear alarm Control board replaced and error is cleared Functional check - OK Leakage current - 88.4 mA Ground resistance - 54mOhms
67951 | 193901 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-10 13:00 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/10/2023 12:58:41 PM, Ferreira, Qwynn) Went to the unit to obtain the device PM Passed Calibrated line voltage Calibrated main voltage Calibrated system vltage Calibrated scale Recall parts installed Electrical Safety - OK Functional Test - OK Returned to unit
68298 | 193902 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68368 | 193903 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Anesthesia Units | AGM not available
68372 | 193904 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2024-02-01 10:37 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 11/28/2023 02:36:21 PM, Scero, Tomlyn) 6 year PM started. Waiting for repair parts so I started replacing the PM parts while I wait. Batteries replaced. (TC 01/25/2024 09:03:10 AM, Scero, Tomlyn) All PM parts replaced during repair. Started PM procedures as per IPM-L (TC 01/26/2024 02:47:30 PM, Scero, Tomlyn) Continued with PM Procedures as per IPM-L (TC 01/30/2024 02:59:06 PM, Scero, Tomlyn) Working through PM procedures IPM-L. Everything going ok and results passing. Failed 4054 and 5054 of DIVA dosage module function tests. Now device failing self test. Test info says, Leak in breathing hoses. Fresh-gas valve defect. (TC 01/31/2024 10:35:16 AM, Scero, Tomlyn) replaced breathing system with the Biomed spare and self test passed with no leaks. Continued with the rest of the PM with no issues. PM tests and procedures complete and passed. Results sheet attached. Device returned to the OR hallway.
68389 | 193905 | PM | ACTIV | READY | 2023-04-01 03:19 |  | PM Major - Delivery Module SEVO
68404 | 193906 | PM | ACTIV | READY | 2023-04-01 03:19 |  | PM MAJOR - Delivery Module DESO
68406 | 193907 | PM | ACTIV | READY | 2023-04-01 03:19 |  | PM MAJOR - Delivery Module DESO
68410 | 193908 | PM | ACTIV | READY | 2023-04-01 03:19 |  | PM MAJOR - Delivery Module DESO
68525 | 193909 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
68526 | 193910 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
68527 | 193911 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
68538 | 193912 | PM | ACTIV | READY | 2023-04-01 03:19 |  | PM MAJOR - Delivery Module DESO
68583 | 193913 | PM | CMPLT | PMFAIL | 2023-04-01 03:19 | 2023-09-28 10:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/28/2023 10:09:49 AM, Campbell, Connor) Cracked rear case near pressure sensor cracked front case in the same spot barrel clamp not working
68588 | 193914 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-13 13:27 | PM-MAJOR-Annual PM for Servo-Ventilator | (TC 04/13/2023 01:26:06 PM, Canjimavadivel, Kamallesh ) PM performed by vendor, see attached service report
68614 | 193915 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-04 11:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/04/2023 11:27:53 AM, Canjimavadivel, Kamallesh ) Battery replaced Cleaned PC PM tests pass
68643 | 193916 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68650 | 193917 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-27 13:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/27/2023 01:13:01 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
68660 | 193918 | PM | CMPLT | PMFAIL | 2023-04-01 03:19 | 2023-09-18 12:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 12:39:18 PM, Campbell, Connor) bottle pressure sensor fatal error
68667 | 193919 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-25 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 02:24:07 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
68671 | 193920 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68677 | 193921 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-20 12:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/20/2023 12:19:35 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 120 mohms Leakage current: 81.9 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68712 | 193922 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68719 | 193923 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-20 12:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/20/2023 12:14:23 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 144 mohms Leakage current: 62.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68733 | 193924 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68740 | 193925 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-02 07:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 07:40:03 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 95 mohms Leakage current: 76.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68741 | 193926 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-10 11:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 11:30:37 AM, Cortez, John) Cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4102mAhs
68759 | 193927 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-10-23 08:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/13/2023 07:07:01 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 57 mohms Leakage current: 76.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 10/23/2023 08:31:14 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3993 mamphr PM completed.
68760 | 193928 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-04 11:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/04/2023 11:04:47 AM, Canjimavadivel, Kamallesh ) Battery replaced Cleaned PC PM tests pass
68761 | 193929 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68768 | 193930 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68789 | 193931 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-01 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/29/2023 10:38:56 AM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
68811 | 193932 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-25 14:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/25/2023 02:41:43 PM, Reid, Emmerson ) Replaced the battery. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68824 | 193933 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-10-12 09:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/12/2023 09:01:54 AM, Campbell, Connor) Replaced battery replaced power cord (came down saying not charging, it was the power cord connection that was faulty) PM complete/passed, refer to ASM
68830 | 193934 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-08 08:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 08:26:07 AM, Reid, Emmerson ) Replaced the battery. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68833 | 193935 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-09 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 11:23:49 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 66 mohms Leakage current: 72.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68838 | 193936 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68841 | 193937 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2024-04-03 13:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:48:04 PM, Rhuland, Courtney) PM completed by BD
68850 | 193938 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-27 15:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/27/2023 03:01:22 PM, Campbell, Connor) PCU wouldnt connect to ASM via ethernet replaced the network board, unit now connects Refer to procedure for electrical safety PM complete/passed, refer to ASM optimized the battery due to long time of disuse
68857 | 193939 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-01 10:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/29/2023 10:42:40 AM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
68888 | 193940 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-20 12:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/20/2023 12:11:25 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 131 mohms Leakage current: 75.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68890 | 193941 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68916 | 193942 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68953 | 193943 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68957 | 193944 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-02 13:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | See WO#196625
68960 | 193945 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-31 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/31/2023 09:55:31 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 30 mohms Leakage current: 62.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68974 | 193946 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-28 11:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 09:58:21 AM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioning battery overnight (TC 08/28/2023 11:25:27 AM, Cortez, John) 4033 mAhs
68981 | 193947 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-07 14:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/07/2023 02:06:09 PM, Reid, Emmerson ) Replaced the batteries. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69018 | 193948 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69021 | 193949 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69055 | 193950 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69070 | 193951 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2024-03-12 10:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69081 | 193952 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-08 13:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:01:06 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69090 | 193953 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69102 | 193954 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-10-05 10:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 10:43:11 AM, Campbell, Connor) PM complete/passed, refer to ASM
69131 | 193955 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69132 | 193956 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69139 | 193957 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-18 13:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/18/2023 01:24:56 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69150 | 193958 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69182 | 193959 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69222 | 193960 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-13 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/13/2023 02:23:24 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69228 | 193961 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-07 09:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 09:25:24 AM, Rendon Lopez , Elizabeth) Door latch not holding up, replaced front case. PM completed, all tests passed
69263 | 193962 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-08 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:41:52 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69314 | 193963 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69342 | 193964 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-21 09:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/21/2023 09:38:15 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69362 | 193965 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2024-03-22 13:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2024 01:19:50 PM, Reid, Emmerson ) ASM PM Complete.
69371 | 193966 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-23 13:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 01:34:57 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69373 | 193967 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-13 16:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/13/2023 03:30:56 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69375 | 193968 | PM | CMPLT | PMFAIL | 2023-04-01 03:19 | 2024-03-13 13:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:06:56 PM, Baghirzade, Nara) Failed rate calibration
69377 | 193969 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-26 13:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/26/2023 01:16:48 PM, Cortez, John) Cleaned, Downloaded error/event logs, Bottle-side pressure sensor errors were logged, Pressure sensor replaced, Calibrated pressure + infusion rate, PM completed successfully, results logged in ASM
69393 | 193970 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69397 | 193971 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-09 13:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/09/2023 01:12:15 PM, Rhuland, Courtney) PM completed All tests passed Refer to ASM for details
69427 | 193972 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-01 18:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 06:18:08 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69463 | 193973 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69477 | 193974 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-23 13:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 01:51:08 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69508 | 193975 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69541 | 193976 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-08 13:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/08/2023 01:55:36 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69563 | 193977 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-06 15:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/06/2023 03:16:22 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69564 | 193978 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-28 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/28/2023 09:38:24 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete
69577 | 193979 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-16 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 10:38:15 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69578 | 193980 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-16 15:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/16/2023 03:53:16 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69620 | 193981 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-25 13:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 01:27:36 PM, Canjimavadivel, Kamallesh ) PM tests pass
69633 | 193982 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-27 08:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/27/2023 08:41:40 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69638 | 193983 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-12-29 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 12/29/2023 10:38:47 AM, Campbell, Connor) PM passed/complete, refer to ASM
69664 | 193984 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-05 09:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2023 09:29:03 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69682 | 193985 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
32540 | 193986 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-08 15:31 | PM - Monthly | PM completed Mechanical PM done SC-595394 TOS phantom a(-993.9,9.8) b(331.9,15.3) c(133.7,11.8) d(95.1,14.0) e(-106.2,12.3) f(0.4,15.9) -passed all phantom tests
32560 | 193987 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-03 14:49 | PM - Monthly | Mechanical PM done sc-1179448 TOS phantom a(-991.0,10.4) b(333.0,14.6) c(133.5,13.5) d(96.2,13.1) e(-106.5,12.7) f(0.6,13.2) -passed all phantom tests
32580 | 193988 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-02 15:39 | PM - Monthly | PM Major done Display console system backup disk completed Slice count - 296209 TOS phantom a(-990.9, 10.8) b(333.9, 15.3) c(134.1, 14.2) d(96.5, 13.4) e(-107.2, 12.8) f(0.8, 13.8) -System passed all tests
502231 | 193989 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-05 18:31 | PM - Semi-Annual | PM was done.
502231 | 193990 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-19 16:08 | HARP - Semi Annual | HARP
1046670 | 193991 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-ELECTRICAL MECHANISM INSPECTION AND OPERATION TEST | Equipment in use | DEPARTMENT CONTACT THE VENDOR TO PROVIDE PM AND MAINTENANCE
1046671 | 193992 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-24 13:09 | PM-ELECTRICAL MECHANISM INSPECTION AND OPERATION TEST | (TC 07/24/2023 01:04:08 PM, Scero, Tomlyn) PM test complete. Electrical safety test complete. Bed is used with an extension cord. Cord should look into being replaced. Power cord doesn't have ground, transformer used so no ground resistance reading. LC 6uA without cord 14uA with cord | DEPARTMENT CONTACT THE VENDOR TO PROVIDE PM AND MAINTENANCE
1047048 | 193993 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047848 | 193994 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-10 11:07 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/10/2023 11:06:57 AM, Baghirzade, Nara) PM completed. See procedure for PM results.
1047852 | 193995 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-04 07:55 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 08/04/2023 07:53:34 AM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1047853 | 193996 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-17 10:41 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/17/2023 10:40:24 AM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1047856 | 193997 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-11 09:39 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/11/2023 09:38:13 AM, Baghirzade, Nara) PM completed. Battery replaced. See procedure for PM results.
1047857 | 193998 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-10 15:07 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/10/2023 03:05:27 PM, Baghirzade, Nara) PM completed. See procedure for PM results.
1047859 | 193999 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-12 10:44 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/12/2023 10:43:19 AM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1048734 | 194000 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Ventilators
1049056 | 194001 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-24 15:29 | Planned Maintenance | (TC 05/24/2023 03:28:01 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure (TC 05/25/2023 08:35:40 AM, Canjimavadivel, Kamallesh ) Repaired damaged hose
1049938 | 194002 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-18 11:21 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/18/2023 11:19:49 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
1050240 | 194003 | PM | CANCL | CANCL | 2023-04-01 03:19 | 2024-03-26 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1050244 | 194004 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-12 11:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/12/2023 11:19:20 AM, Reid, Emmerson ) Replaced the battery. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1050259 | 194005 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-01 14:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/01/2023 02:18:26 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1050266 | 194006 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-01 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050614 | 194007 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-13 09:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/13/2023 09:10:42 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1050620 | 194008 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-07-17 08:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/17/2023 08:05:58 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1051357 | 194009 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-07 11:16 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/07/2023 09:51:04 AM, Baghirzade, Nara) PM passed and completed. No physical damage. @12.0in distance between the surface and the light, the light intensity is: high = 30.8 uw/cm2/nm (30-35 uw/cm2/nm) low = 13.1 uw/cm2/nm (12-15uw/cm2/nm) All 6 rubber feet were damaged and 3 missing; replaced all rubber feet. Target light is good. Hours: 37.8 Cleaned the mesh filters and vacuumed the filters. Electrical safety: Power cord resistance = 102 mohms Current leakage = 51.3 uA
1051359 | 194010 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-07 13:33 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/07/2023 01:33:12 PM, Baghirzade, Nara) PM passed and completed. No physical damage. @12.0in distance between the surface and the light, the light intensity is: high = 31.1 uw/cm2/nm (30-35 uw/cm2/nm) low = 12.7 uw/cm2/nm (12-15uw/cm2/nm) Target light is good. Hours: 5033.9 Cleaned the mesh filters and vacuumed the filters. Electrical safety: Power cord resistance = 60.0 mohms Current leakage = 50.1 uA
1053010 | 194011 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-09 13:17 | PM - Annual | (TC 05/09/2023 01:15:16 PM, Fan, Tak) PM performed as per checklist. Gnd R = 181 ohms Gnd A = 158.2 / 310 uA Chassis A = 0.2 / 160.2 / 310 uA ECG = 0.4 / 2.9 uA ISO = 7.6 uA M5Sc = 0.5 / 1.8 uA 4Vc = 0.5 / 2.3 uA
1053347 | 194012 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-15 15:55 | MSH BI 1270 H1 HARP MAY 9 TRGT SEMI-ANNUAL |  | MSH BI 1270 H1 HARP MAY 9 TRGT SEMI-ANNUAL SYSTEM IS STILL NON-FUNCTIONAL , AS OF MAY 31 2023 WO# 194740 TUBE REQUIRES REPLACEMENT WILL PERFORM HARP AFTER TUBE REPLACEMENT
1053347 | 194013 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-06-19 12:40 | MSH BI 1270 -PM SEMI-ANNUAL | (TC 06/19/2023 12:40:04 PM, Whittington, Scott) MSH BI 1270 -PM SEMI-ANNUAL OERFORMED BY HOLOGIC ON THEIR SCHEDULE
1054039 | 194014 | PM | CANCL | CANCL | 2023-04-01 03:19 | 2023-10-18 14:05 | PM-MAJOR - ANNUAL PM For: Warmers
1054460 | 194015 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-17 08:32 | Lubrication cycle | (TC 05/17/2023 08:31:09 AM, Canjimavadivel, Kamallesh ) Lubricated O-rings
1054462 | 194016 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-17 08:33 | Lubrication cycle | (TC 05/17/2023 08:32:35 AM, Canjimavadivel, Kamallesh ) Lubricated O-rings
1054468 | 194017 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-17 08:35 | Lubrication cycle | (TC 05/17/2023 08:35:12 AM, Canjimavadivel, Kamallesh ) Lubricated O-rings
1055094 | 194018 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-01 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 11:23:49 AM, Rendon Lopez , Elizabeth) Battery conditioning running (TC 09/01/2023 11:23:49 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1055102 | 194019 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055108 | 194020 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-15 11:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/15/2023 11:50:27 AM, Rhuland, Courtney) Battery replaced and conditioned PM complete Refer to ASM for details
1055113 | 194021 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055119 | 194022 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055127 | 194023 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-24 09:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/22/2023 09:12:37 AM, Rendon Lopez , Elizabeth) Running battery conditioning (TC 08/23/2023 09:12:37 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1055156 | 194024 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055158 | 194025 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-01 11:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/31/2023 11:34:35 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
1055159 | 194026 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-11-06 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/01/2023 09:05:35 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1055166 | 194027 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-01 11:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/31/2023 11:36:08 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
1055167 | 194028 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-16 13:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2023 01:26:36 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1055169 | 194029 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055170 | 194030 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055173 | 194031 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2024-03-12 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055174 | 194032 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055175 | 194033 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055176 | 194034 | PM | CNP | ENL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055179 | 194035 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-01 11:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 11:01:12 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1055181 | 194036 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-05 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2023 09:05:08 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1055820 | 194037 | PM | ACTIV | READY | 2023-04-01 03:19 |  | Annual PM - Phototherapy Pads
1056055 | 194038 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-05 16:10 | HARP - Semi Annual | HARP
1056055 | 194039 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-04-27 17:33 | PM - Semi Annual | A complete PM was done.
1057838 | 194040 | PM | CANCL | CANCL | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057840 | 194041 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-11-03 10:20 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 11/03/2023 10:19:34 AM, Reid, Emmerson ) PM Complete, see attached report.
1057843 | 194042 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057844 | 194043 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-08-18 10:09 | PM-Major - ANNUAL PM for: Elecrosurgical Unit | (TC 08/18/2023 10:07:57 AM, Reid, Emmerson ) PM complete, see attached PM Checklist for details.
1058227 | 194044 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-01 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 11:16:52 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
1058269 | 194045 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-09-20 08:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 08:57:36 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed.
1059702 | 194046 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1059704 | 194047 | PM | ACTIV | ISSUE | 2023-04-01 03:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1060518 | 194048 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-26 16:12 | HARP - Semi Annual | HARP
1060518 | 194049 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-17 17:28 | PM - Semi Annual | PM was done.
1063026 | 194050 | PM | CMPLT | CMPLT | 2023-04-01 03:19 | 2023-05-09 13:19 | PM - Annual | (TC 05/09/2023 01:17:21 PM, Fan, Tak) PM performed as per checklist. Gnd R = 129 ohms Gnd A = 168.5 / 331 uA Chassis A = 0.2 / 168.6 / 335 uA ECG = 0.6 / 2.9 uA ISO = 7.7 uA M5Sc = 0.4 / 1.8 uA 4Vc = 0.4 / 2.2 uA
1064167 | 194051 | PM | CANCL | CANCL | 2023-04-01 03:19 | 2024-05-23 10:26 | PM MAJOR
| 194052 | CM | CMPLT | CMPLT | 2023-04-03 08:26 | 2023-04-03 08:27 | ENDO room#2 has problem with SPO2 | (TC 04/03/2023 08:26:32 AM, Samanian, Kamran) provided an used cable with a new sensor ( needs to be charged ENDO )
1066629 | 194053 | CM | CMPLT | CMPLT | 2023-04-03 09:25 | 2023-04-03 12:49 | MSH NICU ZONARE Z63 ACCEPTANCE Incoming Inspection | (TC 04/03/2023 12:48:28 PM, Ferreira, Qwynn) unboxed the unt set up according to instruction manual electrical safety - ok added necessary cabeling added SN of the transducers Handed over to scott (TC 04/03/2023 03:18:03 PM, Whittington, Scott) MSH NICU ZONARE Z63 ACCEPTANCE (TC 04/03/2023 03:18:16 PM, Dolinar, John) MSH NICU ZONARE Z63 ACCEPTANCE (TC 04/11/2023 09:10:03 PM, Whittington, Scott) MSH NICU ZONARE Z63 ACCEPTANCE Incoming Inspection NICU PREP FOR APPLICATIONS (TC 05/11/2023 02:43:26 PM, Whittington, Scott) RETURN 18 MHZ PROBE TO LISA & RICHARD ON THEIR SITE VISIT TODAY (TC 06/13/2023 02:05:26 PM, Whittington, Scott) GATHER INFO FOR CR, DUPLICATE CR 1122 DEC 2019 | MSH NICU ZONARE Z63 ACCEPTANCE Incoming Inspection
| 194054 | OC | CMPLT | CMPLT | 2023-04-03 09:37 | 2023-04-07 07:00 | Nara Baghirzade On-Call; March 31 - April 6, 2023 | (TC 03/31/2023 05:00:00 PM, Baghirzade, Nara) Got a call from L&D OR4 regarding warmer bed and when I asked for more information, the person put me on hold for about 10 mins. Hung up after 10 mins and called the L&D main desk, spoke to a person named Jessica (she tried to get more info from the staff in OR4 but was unable to) and then I guided her on how to send the biomed equipment for repair and where to find a spare warmer bed. (TC 04/05/2023 12:00:00 AM, Baghirzade, Nara) Got a call from CCU regarding X2 in room 10; when trying to connect to the MX monitor, it does not connect and shows error "MMS unplugged". When they swapped it with a different X2 from another room, it was working. The X2 in room 10 was defective. Since the MX monitors were not configured to "Ask User", it was a risk to have them connect the X2 as it could cause them to lose patient data on that monitor. I had to come in, take the spare CCU X2, go up to CCU, go to config mode on the monitor and change the MMS model use to "Ask User". The MCM wanted to be contacted through locating before I left CCU, so I had to wait about 10 minutes for her to call back to give the ok. Got home around midnight.
1063354 | 194055 | CM | CMPLT | CMPLT | 2023-04-03 10:21 | 2023-04-03 10:22 | Flow meter knob repair | (TC 04/03/2023 10:21:41 AM, Scero, Tomlyn) repaired the flow meter knob. tested with O2, ok
531067 | 194056 | CM | CMPLT | CMPLT | 2023-04-03 10:43 | 2023-04-26 11:18 | Ordering parts | (TC 04/03/2023 10:44:31 AM, Baghirzade, Nara) SC 1002924479 (TC 04/26/2023 11:17:04 AM, Baghirzade, Nara) PO 7202190224
67136 | 194057 | CM | CMPLT | CMPLT | 2023-04-03 11:10 | 2023-04-03 11:18 | "BP + temp not working" | (TC 04/03/2023 11:16:13 AM, Cortez, John) Confirmed issues, SPO2 also not working Replaced hose adult cuff with spare, replaced temp probe + well, replaced spo2 extension cable All parameters tested ok
| 194058 | CM | CMPLT | CMPLT | 2023-04-03 11:16 | 2023-03-31 11:17 | Experience with X3 when the ECG leads will be remved | (TC 03/31/2023 11:16:09 AM, Samanian, Kamran) tested ECG lead off when wires coming off. regardless of RESP lead, alwasy and always message with ECG leadsoff and RESP leads off will appear in the first two HL7 messages and then individual lead off message is appeared
66680 | 194059 | CM | CMPLT | CMPLT | 2023-04-03 11:20 | 2023-04-03 11:21 | Latch recall | (TC 04/03/2023 11:20:34 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
66680 | 194060 | CM | CMPLT | CMPLT | 2023-04-03 11:21 | 2023-04-03 11:22 | SW latch broken | (TC 04/03/2023 11:21:14 AM, Baghirzade, Nara) Replaced SW latch. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 76% humidity. Incubator is fully functional.
1066630 | 194061 | ACC | CMPLT | CMPLT | 2023-04-03 11:37 | 2023-04-03 11:38 | Incoming Inspection | (TC 04/03/2023 11:36:43 AM, Reid, Emmerson ) PO is attached. Double insulated. Incoming Complete.
502366 | 194062 | CM | CMPLT | CMPLT | 2023-04-03 15:06 | 2023-04-05 10:33 | "Failed to load patient, Root element is missing" | (TC 04/03/2023 03:06:34 PM, Fan, Tak) System displays error when in the patient screen or when changing dataflows. System not available, contacted GE, recommended SW upgrade. Inspected system. Issue occurs when changing dataflows as well, additional details say the source of the error is the "No archive" archive. Cannot change the settings, a SW reload would likely clear the issue. Ordered SW upgrade USB. 1-547588115914. (TC 04/05/2023 01:06:16 PM, Fan, Tak) Apparently a patient study was done on "no archive" and was thus lost in the limbo realm of no archive. Looked in D drive for study, somehow it was there. Copied it out and the other error went away. Imported exam and exported to viewpoint. Worked with Alexi to organize the study.
| 194063 | CM | CMPLT | CMPLT | 2023-04-03 15:07 | 2023-04-03 15:07 | PICiX with Tomi | (TC 04/03/2023 03:06:56 PM, Samanian, Kamran) almost done
| 194064 | CM | CMPLT | CMPLT | 2023-04-04 07:40 | 2023-04-04 07:42 | QC Checklist | (TC 04/04/2023 07:42:14 AM, Baghirzade, Nara) 14South Loaner VAC from KCI, SN: VFVR19454 QC checklist completed. Returned to 14S.
| 194065 | PA | CMPLT | CMPLT | 2023-04-04 07:46 | 2023-07-05 08:24 | Ordering stock for biomed | (TC 04/04/2023 07:52:56 AM, Baghirzade, Nara) SC 1002925437 (TC 06/26/2023 11:24:30 AM, Baghirzade, Nara) Received 3/4 items. MF4000 retainer not received. PO 7202190869 (TC 07/04/2023 09:05:09 AM, Baghirzade, Nara) Received full order. PO 7202190869
1046294 | 194066 | CM | CMPLT | CMPLT | 2023-04-04 09:49 | 2023-04-13 10:01 | spare unit is at L&D with the cart | spare unit is at L&D with the cart. We received it back.
1057869 | 194067 | PM | CMPLT | CMPLT | 2023-04-04 09:50 | 2023-04-04 14:09 | Annual PM | (TC 04/04/2023 02:08:35 PM, Reid, Emmerson ) PM Complete, see procedure.
1061451 | 194068 | CM | CMPLT | CMPLT | 2023-04-04 09:51 | 2023-04-17 07:53 | ordering power cord for Styrker beds | Received power cords PO#7202191287
1066411 | 194069 | CM | CMPLT | CMPLT | 2023-04-04 13:16 | 2023-04-26 09:13 | Defective display | (TC 03/28/2023 01:16:18 PM, Cortez, John) New unit has defective display out of the box Will request exchange (TC 04/26/2023 09:09:31 AM, Cortez, John) Received exchange unit (SN# V1616697SX)
67222 | 194070 | CM | CMPLT | CMPLT | 2023-04-04 13:28 | 2023-04-04 13:30 | Will not turn on properly, beeps when powering on | (TC 04/04/2023 01:28:23 PM, Ferreira, Qwynn) Troubleshoot the device and found the power disruption is coming from the fuse which is loosely fitted into its holder. Bent the spring a bit to make more contact with the fuse. Device is powering on correctly now. Sent back to unit
502089 | 194071 | RQ | CMPLT | CMPLT | 2023-04-04 14:13 | 2023-04-05 14:07 | fan needs replacemnent on S2 gantry | (TC 04/05/2023 02:06:36 PM, Green, Paul) Fan completly seized. Replaced lower CT blower fan.
1062169 | 194072 | CM | CMPLT | CMPLT | 2023-04-04 14:55 | 2023-04-04 14:57 | Unable to move the head board and no controls or touch screen | (TC 04/04/2023 02:55:39 PM, Ferreira, Qwynn) Received a call stating the head and foot boards wont raise, and the touch screen is no longer functional. Reset the bed, no power to the foot board. Noticed the matress was wedged, dislodging the connector. Aligned one of the pins and resecured the foot board. Functional Test - OK
66760 | 194073 | CM | CMPLT | CMPLT | 2023-04-04 16:42 | 2023-06-30 07:13 | “system fault 44,510 occurred” | (TC 03/28/2023 04:42:20 PM, Cortez, John) Confirmed issue in event logs Will send to smiths for repair (TC 06/29/2023 07:08:43 AM, Cortez, John) Returned from vendor. Functional check ok (see procedure)
62696 | 194074 | CM | CMPLT | CMPLT | 2023-04-04 16:55 | 2023-04-04 16:58 | Sent to biomed with no tag | (TC 04/04/2023 04:56:18 PM, Cortez, John) Keyboard cover ripped, replaced Functional test with simulator ok
1059778 | 194075 | FC | CMPLT | CMPLT | 2023-04-05 07:47 | 2023-04-05 07:56 | QC checklist | (TC 04/05/2023 07:55:55 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
65930 | 194076 | PA | CMPLT | CMPLT | 2023-04-05 08:11 | 2023-04-13 09:45 | Ordering linear sensor | (TC 04/11/2023 09:44:03 AM, Baghirzade, Nara) SC 1002929486 (TC 04/13/2023 09:43:57 AM, Baghirzade, Nara) PO 7202193955 | "Calibration error" displays when trying to use syringe; repeated linear position sensor errors.
67958 | 194077 | CM | CMPLT | CMPLT | 2023-04-05 12:04 | 2023-04-05 12:05 | Broken Hood lock and tilt | (TC 04/05/2023 12:04:08 PM, Ferreira, Qwynn) Replaced the hood lock Greased the slide assembly for the tilt Replaced recall parts and documented Functional Test - OK
1061387 | 194078 | CM | CMPLT | CMPLT | 2023-04-05 12:26 | 2023-04-05 12:26 | needs new batteries | (TC 04/05/2023 12:25:59 PM, Cortez, John) Replaced AAA batteries
62696 | 194079 | CM | CMPLT | CMPLT | 2023-04-05 12:27 | 2023-04-05 12:32 | Barcode scanner not working | (TC 04/05/2023 12:27:48 PM, Cortez, John) Replaced barcode scanner with spare, confirmed functionality
1066462 | 194080 | FC | CMPLT | CMPLT | 2023-04-05 12:59 | 2023-04-05 12:59 | QC checklist | (TC 04/05/2023 12:59:02 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
| 194081 | CM | CMPLT | CMPLT | 2023-04-05 13:09 | 2023-04-04 13:09 | Biomed huddle | (TC 04/04/2023 01:09:30 PM, Baghirzade, Nara) Biomed huddle
| 194082 | CM | CMPLT | CMPLT | 2023-04-05 13:10 | 2023-04-05 13:12 | Meeting with Lesley from NICU | (TC 04/05/2023 01:10:27 PM, Baghirzade, Nara) Met with Lesley to discuss the EEG/CNS setup for NICU AI Baby research; concerns about how the lancets are secured to the device and improvements we can make to help with their setup. Also needed to discuss the ECG from the moms and to figure out how to collect resp data from the mom.
1048334 | 194083 | CM | CMPLT | CMPLT | 2023-04-05 13:13 | 2023-04-05 13:14 | Reported not turning on | (TC 04/05/2023 01:12:48 PM, Baghirzade, Nara) Battery was very depleted however the device turned on when connected to AC; charged for a couple of hours and device still turns on. Ran the device for an hour on battery; all functional, no issues with power or turning on. Functional checks of NIBP and SpO2 passed.
| 194084 | PA | CMPLT | CMPLT | 2023-04-05 14:53 | 2023-04-05 14:54 | Power Bars for OR | (TC 04/05/2023 02:53:21 PM, Reid, Emmerson ) Gave two power bars to the OR for a temperary solution. One is to be charged back in Parts Cost
1061203 | 194085 | CM | CMPLT | CMPLT | 2023-04-05 14:55 | 2023-04-05 14:55 | Audio Cable Broken | (TC 04/05/2023 02:54:56 PM, Reid, Emmerson ) Replaced the audio cable in OR2. Working now.
501069 | 194086 | CM | CMPLT | CMPLT | 2023-04-06 07:52 | 2023-07-11 07:33 | MSH CEOU-3238(Rm11) SHUTS DOWN SPONTANEOUSLY ; ERROR LOGS GLEANED | (TC 04/06/2023 07:52:15 AM, Whittington, Scott) MSH CEOU-3238(Rm11) SHUTS DOWN SPONTANEOUSLY ; ERROR LOGS GLEANED SITE# 553548 REF# | MSH CEOU-3238(Rm11) SHUTS DOWN SPONTANEOUSLY ; ERROR LOGS GLEANED SITE# 553548 REF# 0121497202
501071 | 194087 | CM | CMPLT | CMPLT | 2023-04-06 07:55 | 2023-04-06 14:38 | MSH CEOU-3237(Rm12) MONITOR ARM DROOPS; TRACKBALL CLEANING | (TC 04/06/2023 07:55:35 AM, Whittington, Scott) MSH CEOU-3237(Rm12) MONITOR ARM DROOPS; TRACKBALL CLEANING CLEANED TRACKBALL; UNABLE TO ADJUST MONITOR ARM POSITION, REQUIRES REPACEMENT GAS STRUT SITE# 553550 | MSH CEOU-3237(Rm12) MONITOR ARM DROOPS; TRACKBALL CLEANING CLEANED TRACKBALL; UNABLE TO ADJUST MONITOR ARM POSITION, REQUIRES REPACEMENT GAS STRUT SITE# 553550 REF# 0121497175 Replaced shock
32520 | 194088 | CM | CMPLT | CMPLT | 2023-04-06 08:36 | 2023-04-17 12:04 | MSH DR2 EXPOSURE | (TC 03/21/2023 06:36:00 PM, Whittington, Scott) MSH DR2 EXPOSURE | MSH DR2 EXPOSURE
1061210 | 194089 | CM | CMPLT | CMPLT | 2023-04-06 08:38 | 2023-04-14 08:41 | OR9 Olympus wall monitor SN# 722017821472 , Dead , no power | (TC 03/31/2023 08:37:48 AM, Samanian, Kamran) verified the problem, needs to be taken out to investigate (TC 04/05/2023 08:39:01 AM, Samanian, Kamran) taken out of the wall and moved it to biomed, verified that power supply is defective, both fuses gone, Replaced fuses and same issue. swap the power suuply board with power supply of OR4 monitor ( SN# 722021835461 ) that was ready to ship out due to damaged LCD. Resolved the issue for OR9 monitor and ready to put the monitor back to the OR. Therefore OR4 monitor will be shipped to vendor for repair with two problems: damaged LCD and defective power supply. (TC 04/14/2023 08:40:30 AM, Samanian, Kamran) returned the monitor back to OR9 and mounted, tested and passed
| 194090 | CM | CMPLT | CMPLT | 2023-04-06 10:34 | 2023-04-06 10:34 | Equipment assembly | (TC 04/06/2023 10:34:02 AM, Baghirzade, Nara) Assembled equipment for procedure room and triage area for Family Medicine, requested by Mara.
| 194091 | CM | CMPLT | CMPLT | 2023-04-06 12:50 | 2023-04-06 12:50 | cyber security paper meeting | (TC 04/06/2023 12:50:26 PM, Samanian, Kamran) done!
| 194092 | CM | CMPLT | CMPLT | 2023-04-06 12:52 | 2023-04-06 12:51 | added 8 licenses to MSHICU3S | (TC 04/06/2023 12:51:35 PM, Samanian, Kamran) done!
67433 | 194093 | CM | CMPLT | CMPLT | 2023-04-06 13:08 | 2023-04-06 13:12 | Does not remain on when powered on | (TC 04/06/2023 01:11:05 PM, Ferreira, Qwynn) Tested the unit Measured all batteries, but they were reading good. Tested all connections to ensure they were snug and seated correctly After testing, found that one battery was not working regardless showing correct voltage. Replaced the battery and was working
1048472 | 194094 | CM | CMPLT | CMPLT | 2023-04-06 14:53 | 2023-04-06 14:53 | Tube head cover melted, exposure switch cable streched |  | replaced tube head cover, repaced exposure switch cable
32560 | 194095 | CM | CMPLT | CMPLT | 2023-04-06 22:02 | 2023-04-06 22:01 | Clinical requested STRYKER data copied from CT | Clinical requested STRYKER data copied from CT Service replaced DVD drive Service copied/transfer USB from Coral
| 194096 | OC | CMPLT | CMPLT | 2023-04-10 08:51 | 2023-04-14 09:17 | Emmerson Reid On-Call April 7 - April 14 2023 | (TC 04/10/2023 08:51:48 AM, Reid, Emmerson ) L&D called, I was put on hold then the call got disconnected but they called back again. A patient warmer screen was not working. They are going to leave a voicemail for biomed to look at it tomoroow.
1066754 | 194097 | CM | CMPLT | CMPLT | 2023-04-11 10:51 | 2023-04-11 10:52 | Incoming Inspection | (TC 04/11/2023 10:51:09 AM, Ferreira, Qwynn) Incoming inspection Assembly Electrical Safety - OK
1066755 | 194098 | CM | CMPLT | CMPLT | 2023-04-11 10:58 | 2023-04-11 10:59 | Incoming Inspection | (TC 04/11/2023 10:58:05 AM, Ferreira, Qwynn) Incoming Inspection Assembly Electrical Safety - OK
1065054 | 194099 | CM | CMPLT | CMPLT | 2023-04-11 12:51 | 2023-05-08 15:46 | Defective RF card; unit under warranty, sending for repair | (TC 04/11/2023 01:18:23 PM, Baghirzade, Nara) SC 1002930043 (TC 04/11/2023 01:22:47 PM, Baghirzade, Nara) RMA 302836574 (TC 05/08/2023 03:43:58 PM, Baghirzade, Nara) PO 7202194336 Received PCU back from repair; attached service report from BD. | Defective RF card; unable to load network configurations
62702 | 194100 | CM | CMPLT | CMPLT | 2023-04-11 14:05 | 2023-04-12 10:13 | Battery not holding a charge | (TC 04/12/2023 10:12:23 AM, Scero, Tomlyn) Replaced old batteries. Confirmed ECG OK with simulator.
1064341 | 194101 | CM | CMPLT | CMPLT | 2023-04-11 16:07 | 2023-04-17 10:42 | OR table from OR 3 removed- issues with feet part moving | (TC 04/11/2023 04:06:45 PM, Rendon Lopez , Elizabeth) Inspected table, confirmed issue, feet part not moving properly when table top is removed. Contacted McDavis for service, they no longer service tables. Contacted skytron directly, they will provide service, created estimated cost PO (TC 04/14/2023 10:40:21 AM, Rendon Lopez , Elizabeth) Medaris medical tech came to check the table. No isuse was found, it is normal for foot section to move separetely when top removed, it is slave and master relation. Table placed back into service.
67226 | 194102 | CM | CMPLT | CMPLT | 2023-04-11 16:16 | 2023-05-08 08:54 | Screen blank | (TC 04/11/2023 04:16:14 PM, Rendon Lopez , Elizabeth) Confirmed issue, removed warmer from room 1, brought to 18th (TC 04/12/2023 10:49:06 AM, Rendon Lopez , Elizabeth) Replaced control board. Running warmer, after 30 mins, screen turned blue and touch panel buttons were unresponsive. Replaced touch panel and LCD assembly, screen is still blue. Reached out to Getinge tech support (TC 04/13/2023 10:49:06 AM, Rendon Lopez , Elizabeth) Troubleshooting, issues is not resolved. (TC 04/14/2023 10:49:06 AM, Rendon Lopez , Elizabeth) Getinge will issue a replacement control board (TC 04/18/2023 12:56:33 PM, Rendon Lopez , Elizabeth) Removed control board to return to getinge. Replaced Blender, tested-OK (TC 05/05/2023 08:53:16 AM, Rendon Lopez , Elizabeth) Received new control board, replaced, tested warmer-OK. Testing all functions all working. Returned damaged control board to Getinge.
1054137 | 194103 | CM | CMPLT | CMPLT | 2023-04-11 16:22 | 2023-04-11 16:23 | Not detected by monitor | (TC 03/29/2023 04:21:47 PM, Rendon Lopez , Elizabeth) Confirmed, sent out for repair (TC 04/11/2023 04:21:47 PM, Rendon Lopez , Elizabeth) Received, incoming inspection-OK
1050676 | 194104 | CM | CMPLT | NFF | 2023-04-11 16:23 | 2023-04-11 16:24 | Not recording contractions | (TC 04/11/2023 04:23:20 PM, Rendon Lopez , Elizabeth) Tested >1hr, NFF
1053348 | 194105 | CM | CMPLT | CMPLT | 2023-04-11 21:14 | 2023-07-11 07:34 | MSH BI 1272 H3 12th Floor - North Side SYSTEM WILL NOT OPERATE | (TC 04/11/2023 09:14:32 PM, Whittington, Scott) MSH BI 1272 H3 12th Floor - North Side SYSTEM WILL NOT OPERATE PMC ERROR REF# 01263778 (TC 04/11/2023 09:14:32 PM, Whittington, Scott) MSH BI 1272 H3 12th Floor - North Side SYSTEM WILL NOT OPERATE PMC ERROR REF# 01263778 | MSH BI 1272 H3 12th Floor - North Side SYSTEM WILL NOT OPERATE PMC ERROR REF# 01263778
1049902 | 194106 | CM | CMPLT | CMPLT | 2023-04-11 21:35 | 2023-04-12 11:47 | MSH R62 POWER CORD HAS LOST GROUND PRONG |  | MSH R62 POWER CORD HAS LOST GROUND PRONG Replaced power cord reel assembly. Ordered new part.
| 194107 | PA | CMPLT | CMPLT | 2023-04-12 07:58 | 2023-05-05 11:03 | Ordering doppler | (TC 04/12/2023 08:01:57 AM, Baghirzade, Nara) SC 1002930672 (TC 05/05/2023 11:02:19 AM, Baghirzade, Nara) PO 7202194810
68814 | 194108 | CM | CMPLT | CMPLT | 2023-04-12 09:40 | 2023-04-11 14:00 | Flash/upgrade PCU using extender board following new logic board | (TC 04/11/2023 02:00:00 PM, Baghirzade, Nara) Main processor, boot software, and keyboard was pushed to the PCU using the extender board. Transferred network configuration using ASM. Flashed PCU.
| 194109 | CM | CMPLT | CMPLT | 2023-04-12 09:43 | 2023-04-11 15:00 | Capital/fleet management meeting | (TC 04/11/2023 03:00:00 PM, Baghirzade, Nara) Capital/fleet management meeting
| 194110 | CM | CMPLT | CMPLT | 2023-04-12 09:44 | 2023-04-11 15:43 | NICU Huddle | (TC 04/11/2023 03:43:00 PM, Baghirzade, Nara) NICU huddle
| 194111 | CM | CMPLT | CMPLT | 2023-04-12 09:45 | 2023-04-13 09:41 | Changing XDS DB access password on all MX800s in NICU | (TC 04/12/2023 09:44:51 AM, Baghirzade, Nara) Changing XDS DB password (TC 04/12/2023 01:28:56 PM, Baghirzade, Nara) Changing XDS DB password (TC 04/13/2023 09:40:11 AM, Baghirzade, Nara) Changing XDS DB password in all NICU rooms is completed. Updated network sheet for NICU Philips data drop info.
68966 | 194112 | CM | CMPLT | CMPLT | 2023-04-12 10:18 | 2023-04-12 10:21 | does not charge when plugged in | (TC 04/12/2023 10:19:48 AM, Baghirzade, Nara) Replaced front case; checked power cord, no damages or issues with it. After front case replacement, the AC indicator lights up. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 24 mohms Leakage current: 69.8 uA
1061206 | 194113 | CM | ACTIV | ISSUE | 2023-04-12 10:39 |  | Camera arm not able to adjust height | (TC 04/12/2023 10:39:12 AM, Reid, Emmerson ) The navigation camera seems to been too light for the arm as it will only rest at the maximum height. I tried to adjust the weight balance but that was not enough. I have reached out to Jaron for other solutions.
1061220 | 194114 | CM | CMPLT | CMPLT | 2023-04-12 10:41 | 2023-04-12 10:44 | Call bell not working | (TC 04/12/2023 10:41:07 AM, Reid, Emmerson ) Went up to OR 19 to check out the issue and contacted help desk and created a ticket. The call bell was continuously ringing, the team was using the silence option but every 5 minutes it would continue. Someone from IT reached out to me and I met with them at the room. They unplugged the call bell and are contacting the vendor to repair.
1061217 | 194115 | CM | CMPLT | CMPLT | 2023-04-12 10:52 | 2023-04-12 11:01 | Microscope Questions | (TC 04/12/2023 10:54:49 AM, Reid, Emmerson ) Was called up because they said the microscope brightness was not able to be decrease. They ended up figureing it out by the time I arrived but they had some other general microscope questions I was able to answer.
62696 | 194116 | CM | CMPLT | NFF | 2023-04-12 13:27 | 2023-04-12 13:33 | ECGs deleted in ED | (TC 04/12/2023 01:30:41 PM, Scero, Tomlyn) ECGs performed with simulator. All printed and saved, 1 transfered using biomed test ID. Unit turned on and off multiple times, left to monitor ECG. Not once did the archive ECGs get deleted. Settings confirmed that the only time an ECG is deleted is when it is transfered to iECG. Device returned to ED. Likely a user deleted the saved ECGs before they were transfered.
1059778 | 194117 | FC | CMPLT | CMPLT | 2023-04-12 13:30 | 2023-04-12 13:31 | QC checklist | (TC 04/12/2023 01:30:50 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1066756 | 194118 | ACC | CMPLT | CMPLT | 2023-04-12 16:26 | 2023-04-12 16:33 | Incoming inspection | (TC 04/11/2023 04:26:02 PM, Cortez, John) Unboxed/assembled unit performance/alarm/electrical tests passed: 60 mOhms 16 uAmps Disinfected and refilled reservoir Lubricated o-rings
1066757 | 194119 | ACC | CMPLT | CMPLT | 2023-04-12 16:34 | 2023-04-12 16:34 | Incoming inspection | (TC 04/11/2023 04:33:46 PM, Cortez, John) Unboxed/assembled unit performance/alarm/electrical tests passed: 60 mOhms 16 uAmps Disinfected and refilled reservoir Lubricated o-rings
1053348 | 194120 | CM | CMPLT | CMPLT | 2023-04-12 16:55 | 2023-04-12 16:55 | Erratic compression movement. |  | Help Jon from Hologic troubleshoot compression movement problem. Most likely problem was a loose communication jack.
1046035 | 194121 | CM | CMPLT | CMPLT | 2023-04-12 16:58 | 2023-04-12 16:58 | Chest bucky horizontal positioning not centred. |  | Reset horizontal position.
32520 | 194122 | CM | CMPLT | CMPLT | 2023-04-12 19:12 | 2023-07-11 07:35 | MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY | (TC 04/12/2023 07:13:05 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY (TC 04/13/2023 02:02:02 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY FOUND THAT CABLE FROM EXPOSE SWITCH HOLDER TO CABINET HAS DETERIORATED INSULATION , BOTH ENDS FOUND NW CABLE TERMINATION KIT (TC 04/13/2023 02:06:56 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY FOUND THAT CABLE FROM EXPOSE SWITCH HOLDER TO CABINET HAS DETERIORATED INSULATION , BOTH ENDS FOUND NW CABLE TERMINATION KIT (TC 04/14/2023 05:34:08 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY EXPOSURE CABLE ALL NSULATION DETERIORATED PLANNED AND AND INSTALLED NEW CABLE , TERMINATED WITH ARISTOS CONNECTORS FROM LOCAL STOCK GW TEST OK (TC 04/14/2023 05:36:08 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY EXPOSURE CABLE ALL NSULATION DETERIORATED PLANNED AND AND INSTALLED NEW CABLE , TERMINATED WITH ARISTOS CONNECTORS FROM LOCAL STOCK GW TEST OK (TC 04/13/2023 08:06:00 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY FOUND THAT CABLE FROM EXPOSE SWITCH HOLDER TO CABINET HAS DETERIORATED INSULATION , BOTH ENDS FOUND NW CABLE TERMINATION KIT (TC 04/12/2023 10:13:00 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY (TC 04/12/2023 10:13:00 PM, Whittington, Scott) MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY | MSH DR2 SYSTEM THROWING ERRORS RE EXPOSE READY
65262 | 194123 | CM | CMPLT | CMPLT | 2023-04-13 07:47 | 2023-05-25 08:31 | ordering PS500 batteries | received 17 x PS500 batteries
65960 | 194124 | CM | CMPLT | CMPLT | 2023-04-13 09:50 | 2023-04-13 09:51 | No barcode sticker | (TC 04/13/2023 09:50:10 AM, Baghirzade, Nara) Reprinted new barcode sticker. Check-in tests passed and completed. IUI connectors and device cleaned.
1062232 | 194125 | CM | CMPLT | CMPLT | 2023-04-13 10:46 | 2023-05-17 14:01 | Audio Feedback not working | (TC 04/13/2023 10:46:35 AM, Reid, Emmerson ) The audio feeback was quiet/non existant when using the cautery. I have reached out the the vendor for support.
| 194126 | PA | CMPLT | CMPLT | 2023-04-13 11:07 | 2023-04-19 09:42 | Ordering NiMH rechargeable batteries | (TC 04/14/2023 08:13:02 AM, Baghirzade, Nara) SC 1002932887 (TC 04/19/2023 09:41:15 AM, Baghirzade, Nara) PO 7202196595
1061968 | 194127 | CM | CMPLT | CMPLT | 2023-04-13 11:13 | 2023-04-13 11:17 | Movement controls not working (footboard/siderails) | (TC 04/13/2023 11:13:22 AM, Cortez, John) Confirmed poroblem, Power cycled bed. Issue seems to be resolved
| 194128 | CM | CMPLT | CMPLT | 2023-04-13 12:55 | 2023-07-27 11:13 | PICiX set up and training for techs | (TC 04/12/2023 12:55:26 PM, Samanian, Kamran) prepared MSHSPAREPICIX for the lab and tested (TC 04/13/2023 12:56:06 PM, Samanian, Kamran) worked with Tomi to practice (TC 04/13/2023 02:04:58 PM, Samanian, Kamran) met with Nara and reviewed steps (TC 04/14/2023 01:46:51 PM, Samanian, Kamran) Training with Qwynn (TC 04/17/2023 09:55:13 AM, Samanian, Kamran) put the spare biomed back to the lab mode (TC 04/20/2023 08:34:29 AM, Samanian, Kamran) john practiced. also changed the picix back to use for lab (TC 04/17/2023 08:34:29 AM, Samanian, Kamran) john practiced. also changed the picix back to use for lab (TC 04/20/2023 02:52:04 PM, Samanian, Kamran) Emmerson done! (TC 04/25/2023 08:24:33 AM, Samanian, Kamran) Done with Olcay (TC 05/01/2023 02:37:07 PM, Samanian, Kamran) Kam did the training (TC 05/04/2023 02:23:26 PM, Samanian, Kamran) done with Jessica
| 194129 | CM | CMPLT | CMPLT | 2023-04-13 12:57 | 2023-04-14 08:34 | looking to the LAB 11th floor, N2 gas source tank for incubators | (TC 04/13/2023 12:57:16 PM, Samanian, Kamran) verified the problem, N2 is zero that won't allow O2 to be decreased. sent out email to clarify who supplies and connect the tanks
| 194130 | CM | CMPLT | CMPLT | 2023-04-13 13:02 | 2023-04-13 13:02 | looking into RESP and ECG alarm for ICU | (TC 04/11/2023 01:01:56 PM, Samanian, Kamran) tested, checked HL7 messages and verified few things
1057007 | 194131 | CM | CMPLT | CMPLT | 2023-04-13 13:28 | 2023-04-13 13:29 | SpO2 not working | (TC 04/13/2023 01:28:05 PM, Canjimavadivel, Kamallesh ) Faulty extension trunk cable, replaced cable working okay
1058131 | 194132 | CM | CMPLT | CMPLT | 2023-04-13 13:33 | 2023-04-13 13:46 | Needs batteries | (TC 04/13/2023 01:45:40 PM, Canjimavadivel, Kamallesh ) Replaced batteries
1061211 | 194133 | CM | ACTIV | ISSUE | 2023-04-13 13:36 |  | Wall monitor not showing HD Scope | (TC 04/13/2023 01:46:10 PM, Reid, Emmerson ) HD scope was not working on the monitor, but other inputs were working. I also noticed that the user profile was not saving properly. It kept reurning back to primary SDI with no fallover even after I set it to have HDMI primary and SDI fallover. Contacted Jaron to see what we can do to resolve this. (TC 05/17/2023 02:00:04 PM, Reid, Emmerson ) This is still not working.
1054286 | 194134 | CM | CMPLT | CMPLT | 2023-04-13 13:47 | 2023-04-13 13:52 | Low batteries | (TC 04/13/2023 01:51:21 PM, Canjimavadivel, Kamallesh ) Replaced batteries
1046035 | 194135 | CM | CMPLT | CMPLT | 2023-04-13 13:48 | 2023-04-13 13:49 | MSH Evo - Not Working | (TC 04/13/2023 01:48:15 PM, Chau, Ken) Room calibration was off by a lot causing issues -pin cal and room cal -recal positions as off -tested okay
1066116 | 194136 | CM | CMPLT | CMPLT | 2023-04-13 14:29 | 2023-04-13 14:30 | returned to ZOLL after AED ZOLL repaired and received | (TC 04/13/2023 02:29:29 PM, Samanian, Kamran) done! UPS tracking# 1Z0AT6169141268627
1054522 | 194137 | CM | CMPLT | CMPLT | 2023-04-13 14:35 | 2023-04-13 14:36 | Not picking up any signals | (TC 04/13/2023 02:35:10 PM, Canjimavadivel, Kamallesh ) RA lead completely severed. Replaced lead, tested okay
69121 | 194138 | CM | CMPLT | CMPLT | 2023-04-13 14:51 | 2023-05-26 10:34 | Wire harness to front display board damaged | (TC 04/13/2023 10:10:43 AM, Cortez, John) Replaced bezel and cover for front display board (TC 05/26/2023 10:33:24 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
| 194139 | CM | CMPLT | CMPLT | 2023-04-13 16:54 | 2023-04-13 16:54 | Replaced Xenon bulb in Room 4 in ENDO | (TC 04/13/2023 04:53:48 PM, Samanian, Kamran) replaced the bulb, reset the time, tested and passed
| 194140 | CM | CMPLT | CMPLT | 2023-04-13 16:55 | 2023-04-13 16:55 | ARGON tank regulator in room ENDO 1 | (TC 04/13/2023 04:55:13 PM, Samanian, Kamran) regulator connected to the tank in room ENDO 1 and ended to APC device,
| 194141 | CM | CMPLT | CMPLT | 2023-04-13 17:37 | 2023-05-09 15:35 | Spare parts for panda Warmers | Received all parts
| 194142 | PA | CMPLT | CMPLT | 2023-04-14 08:10 | 2023-06-13 10:29 | Incubator Hoods | waiting for one more
1061210 | 194143 | CM | CMPLT | CMPLT | 2023-04-14 08:38 | 2023-04-14 08:40 | Pharmacy cart alarming | (TC 04/14/2023 08:38:12 AM, Reid, Emmerson ) Pharmacy cart was alarming for an open drawer, but the drawer was not open. I called tech support and they instructed me to restart the cart. Alarm did not come on again but I have ticket 957644 for SN 210686
| 194144 | CM | CMPLT | CMPLT | 2023-04-14 12:13 | 2023-04-14 12:13 | meeting with Adil ( UHN ) cath lab | (TC 04/14/2023 12:12:41 PM, Samanian, Kamran) verifeid that we can have 12 Lead ECG from GE Tram on the cosnole outside
65930 | 194145 | CM | CMPLT | CMPLT | 2023-04-14 13:09 | 2023-04-14 13:12 | Parts replacement/repair | (TC 04/14/2023 01:09:31 PM, Baghirzade, Nara) Replaced the linear sensor. Rear case was broken; replaced the rear case and transferred latch and handle to new case. Performed full calibraton of: -barrel clamp -plunger force -plunger position -pressure disc Check-in tests passed and completed. Cleaned the IUIs. Ran basic infusion at various rates; infusion completed successfully with no errors/issues.
67958 | 194146 | CM | CMPLT | CMPLT | 2023-04-14 13:54 | 2023-04-14 14:01 | System Failure: 25 | (TC 04/14/2023 01:59:59 PM, Ferreira, Qwynn) System failure 25 was present Fan was not correctly seated Functional Test - OK
1050314 | 194147 | CM | CMPLT | CMPLT | 2023-04-14 14:13 | 2023-04-14 14:15 | Thermometer not working | (TC 04/14/2023 02:14:03 PM, Canjimavadivel, Kamallesh ) Thermometer faulty, replaced and tested ok
| 194148 | CM | CMPLT | CMPLT | 2023-04-14 14:40 | 2023-04-14 14:41 | Meeting with Donna about RESP/ECG lead off | (TC 04/14/2023 02:40:29 PM, Samanian, Kamran) sent requested SNs to Donna when she opens the ticket to BIU
1046196 | 194149 | CM | CMPLT | CMPLT | 2023-04-17 06:44 | 2023-04-14 06:43 | CT computer applications crashing/mouse/keyboard | CT computer applications crashing/mouse/keyboard Service ordered replacement computer from SIEMENS
| 194150 | OC | CMPLT | CMPLT | 2023-04-17 07:46 | 2023-04-26 10:42 | Tomi Scero On call; April 14-April 21, 2023 | (TC 04/16/2023 01:57:00 PM, Scero, Tomlyn) Donna from CCU called. A patient lost their telepak and they wanted help finding it. I explained that it didn't have a GPS and the only way to "track" it would be to check it's last AP location that it connected to. She said they'd keep looking for it and call me back if they wanted me to come in to give them a spare unit. Just before the call ended they found the telepak in the kitchen.
63807 | 194151 | CM | CMPLT | CMPLT | 2023-04-17 08:36 | 2023-04-20 08:42 | Does not turn on when not plugged | (TC 04/17/2023 08:36:57 AM, Canjimavadivel, Kamallesh ) Battery needed replacement, replaced with new battery. Installed cuff and BP hose
| 194152 | PA | CMPLT | CMPLT | 2023-04-17 09:22 | 2023-04-27 11:29 | Cuffs for stock | (TC 04/17/2023 09:27:32 AM, Canjimavadivel, Kamallesh ) Received quote via phone from vendor. Placed order for cuffs (TC 04/27/2023 11:29:04 AM, Canjimavadivel, Kamallesh ) Received
62644 | 194153 | CM | CMPLT | CMPLT | 2023-04-17 10:42 | 2023-04-17 10:43 | Alarming | (TC 04/14/2023 10:42:30 AM, Rendon Lopez , Elizabeth) Confirmed alarms. Retire from service, not needed by Rt NICU
1066770 | 194154 | ACC | CMPLT | CMPLT | 2023-04-17 10:45 | 2023-04-17 10:46 | Incoming inspection | (TC 04/12/2023 10:45:26 AM, Rendon Lopez , Elizabeth) Unpacking, entere in TMS (TC 04/14/2023 10:45:26 AM, Rendon Lopez , Elizabeth) Check-in tests done, all tests passed
1066771 | 194155 | ACC | CMPLT | CMPLT | 2023-04-17 10:47 | 2023-04-17 10:48 | Incoming inspection | (TC 04/12/2023 10:47:04 AM, Rendon Lopez , Elizabeth) Unpacked, entered in TMS (TC 04/14/2023 10:47:04 AM, Rendon Lopez , Elizabeth) Check-in tests done, all tests passed
1066772 | 194156 | ACC | CMPLT | CMPLT | 2023-04-17 10:48 | 2023-04-17 10:49 | Incoming inspection | (TC 04/12/2023 10:48:12 AM, Rendon Lopez , Elizabeth) Unpacking, entered in TMS (TC 04/14/2023 10:48:12 AM, Rendon Lopez , Elizabeth) Check-in tests done, all tests passed
1050726 | 194157 | CM | CMPLT | CMPLT | 2023-04-17 10:53 | 2023-04-17 10:54 | Not working | (TC 04/13/2023 10:53:34 AM, Rendon Lopez , Elizabeth) Activation cable malfunction, filter low. Replaced activation cable and filter, tested-OK. Returned to OR 2
1066840 | 194158 | CM | CMPLT | CMPLT | 2023-04-17 12:56 | 2023-04-24 14:47 | 6 Year PM | (TC 04/24/2023 02:46:47 PM, Scero, Tomlyn) PM parts replaced and tests complete. passed.
68546 | 194159 | CM | CMPLT | CMPLT | 2023-04-17 12:58 | 2023-04-17 13:45 | Inspiratory flow sensor defective | (TC 04/17/2023 01:43:33 PM, Scero, Tomlyn) Device in OR 5 failing selftest. Flow sensor error. swapped flow sensor with expiratory and issue now seen there. replaced flow sensor to fix the issue.
1057863 | 194160 | PM | CMPLT | CMPLT | 2023-04-17 14:27 | 2023-10-24 11:48 | Annual PM - For Fluid Warmer | (TC 04/18/2023 09:11:11 AM, Reid, Emmerson ) Changed circulating fluid, O-rings, and reattached the drain spout holder with doublesided tape. The unit speaker is not sounding for alarms. Looking into purchasing needed part to fix. (TC 10/19/2023 01:19:42 PM, Campbell, Connor) Tried to perform PM before looking up asset# Came across same problem as before, no audible alarms Probably need to replace board, will update when ordered parts (TC 10/24/2023 10:12:36 AM, Campbell, Connor) replaced speaker from a donor board audible alarm now functional researched parts to replace the ones from the donor board will complete PM when water arrives (TC 10/24/2023 11:46:02 AM, Campbell, Connor) Had to adjust all 3 pots as per the manual to acheive the proper LCD reading, over temp alarm and operating temp PM complete/passed, refer to procedure
68045 | 194161 | CM | CMPLT | CMPLT | 2023-04-17 15:12 | 2023-04-17 15:15 | BMI not working | (TC 04/17/2023 03:13:18 PM, Reid, Emmerson ) I looked at the scale this afternoon and everything was working properly for me. The only thing I could see happening is that the height has to be measured by the scale and cannot be entered in manually. If you do not use the height measurement arm the BMI will not be correct.
502231 | 194162 | CM | CMPLT | CMPLT | 2023-04-17 16:40 | 2023-04-17 16:39 | Boot problems in the morning. | Fixed.
1053432 | 194163 | CM | CMPLT | CMPLT | 2023-04-18 07:42 | 2023-05-29 08:32 | Not detected | (TC 04/17/2023 07:41:53 AM, Rendon Lopez , Elizabeth) under warranty, to be sent out (TC 05/23/2023 08:31:37 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1063185 | 194164 | CM | CMPLT | CMPLT | 2023-04-18 07:43 | 2023-05-19 15:02 | Not reading FHR | (TC 04/17/2023 07:43:48 AM, Rendon Lopez , Elizabeth) Confirmed issue, to be sent out for repair (TC 05/19/2023 03:01:11 PM, Rendon Lopez , Elizabeth) Returned, broken beyond repair. Retired
52181 | 194165 | CM | CMPLT | CMPLT | 2023-04-18 07:48 | 2023-04-18 07:53 | Wheel is squeaking | (TC 04/18/2023 07:52:23 AM, Ferreira, Qwynn) Lifted the incubator and oiled the bearins in the wheels Functional test - ok
67995 | 194166 | CM | CMPLT | CMPLT | 2023-04-18 07:55 | 2023-04-18 07:56 | Malfunction | (TC 04/17/2023 07:54:51 AM, Rendon Lopez , Elizabeth) Tele malfunction message. L&D doesn't need a replacement, retire from service.
1047227 | 194167 | CM | CMPLT | CMPLT | 2023-04-18 07:56 | 2023-04-18 07:57 | Incorrect temperature | (TC 04/17/2023 07:56:40 AM, Rendon Lopez , Elizabeth) Cleaned contacts, tested, NFF. Temps consistent and accurate
501173 | 194168 | CM | CMPLT | CMPLT | 2023-04-18 09:59 | 2023-04-26 09:30 | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 IMAGE QUALITY ISSUES WHEN USING C5-1, DOUBLE SWAP INDICATES SYSTEM ISSUE | (TC 04/18/2023 09:58:53 AM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 IMAGE QUALITY ISSUES WHEN USING C5-1, DOUBLE SWAP INDICATES SYSTEM ISSUE (TC 04/19/2023 04:11:33 PM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 IMAGE QUALITY ISSUES WHEN USING C5-1, DOUBLE SWAP INDICATES SYSTEM ISSUE SITE# 556348 REF# #0121548612 C5-1 PROBE SN# B3LH9T Narayan Mahalingam SYSTEM TEST FAILED 3 TIMES AFTER 3RD FAILURE, SENT IN ERROR LOGS (TC 04/20/2023 02:47:15 PM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 IMAGE QUALITY ISSUES WHEN USING C5-1, DOUBLE SWAP INDICATES SYSTEM ISSUE SITE# 556348 REF# #0121548612 C5-1 PROBE SN# B3LH9T Narayan Mahalingam CODY ON-SITE , DIAGNOSED FAULTY CONGA PROGRAMMMABLE CHIP ON DSC BD IN UMB HE WILL ORDER AND INSTALL TOMORROW (TC 04/26/2023 09:29:54 AM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 IMAGE QUALITY ISSUES WHEN USING C5-1, DOUBLE SWAP INDICATES SYSTEM ISSUE SITE# 556348 REF# #0121548612 C5-1 PROBE SN# B3LH9T Narayan Mahalingam CODY REPLACED DSC BD, OK | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 IMAGE QUALITY ISSUES WHEN USING C5-1, DOUBLE SWAP INDICATES SYSTEM ISSUE SITE# 556348 REF# #0121548612 C5-1 PROBE SN# B3LH9T Narayan Mahalingam
1057570 | 194169 | CM | CMPLT | CMPLT | 2023-04-18 11:08 | 2023-04-18 11:10 | Put on Cart for OR 1 | (TC 04/18/2023 11:08:00 AM, Reid, Emmerson ) Fixed the cart in biomed with Samir. Cleaned the cart and equipment. Connected the FT 10, smoke evac (1058598), and foot pedals together and tested to make sure everything was working.
69652 | 194170 | CM | CMPLT | CMPLT | 2023-04-18 13:50 | 2023-04-18 13:50 | "Channel error" | (TC 04/17/2023 01:49:36 PM, Rendon Lopez , Elizabeth) Downloaded error log, no errors, running on regular mode no issues (TC 04/18/2023 01:49:36 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1058280 | 194171 | CM | CMPLT | CMPLT | 2023-04-18 13:51 | 2023-04-18 13:52 | Will not turn on or charge at all | (TC 04/18/2023 01:50:46 PM, Rendon Lopez , Elizabeth) Keypad issue under recall. Network communication hardware error upon turned on. Replaced keypad and network assembly.
67355 | 194172 | CM | CMPLT | CMPLT | 2023-04-18 15:12 | 2023-10-24 15:15 | Leaking/Cracked enclosure cover, green LED is broken | (TC 09/25/2023 10:46:44 AM, Campbell, Connor) Spoke to and emailed with smiths medical representative about the process of getting a replacement board and cover will update with information once the order is placed for the parts (TC 09/26/2023 09:07:12 AM, Reid, Emmerson ) Placed an order for the man board, front cover, and screws. (TC 10/04/2023 09:07:33 AM, Reid, Emmerson ) Recieved all items except missing the last enclosure cover on B.O. (TC 10/19/2023 07:45:52 AM, Reid, Emmerson ) Recieved last enclosure cover (TC 10/23/2023 01:50:55 PM, Campbell, Connor) replaced the main board waiting for water to test functionality and PM (TC 10/24/2023 10:13:51 AM, Campbell, Connor) cleaned the water resovoir as there was lots of mould Repairs complete, will PM in next WO # 198605
1059778 | 194173 | FC | CMPLT | CMPLT | 2023-04-19 08:08 | 2023-04-19 08:09 | QC checklist | (TC 04/19/2023 08:08:33 AM, Baghirzade, Nara) QC checklist completed. Returned to 11S
1053391 | 194174 | CM | CMPLT | CMPLT | 2023-04-19 08:09 | 2023-04-19 08:10 | Functional check before patient use | (TC 04/19/2023 08:09:46 AM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 124 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
1055921 | 194175 | CM | CMPLT | CMPLT | 2023-04-19 09:07 | 2023-04-19 09:08 | Temperature module error | (TC 04/19/2023 09:07:09 AM, Baghirzade, Nara) Replaced temperature probe; temperature module is fully functional. Functional checks for NIBP and SpO2 also passed.
1056447 | 194176 | CM | CMPLT | NFF | 2023-04-19 09:08 | 2023-04-19 09:09 | Blood pressure cuff leaking | (TC 04/19/2023 09:08:18 AM, Baghirzade, Nara) Took several NIBP measurement, did not find any leaks in the hose or cuff. No physical damage to hose or cuff. Functional checks of temp and SpO2 also passed; no fault found with the device.
1064817 | 194177 | CM | CMPLT | CMPLT | 2023-04-19 09:12 | 2023-04-19 09:14 | SpO2 Malfunction | (TC 04/19/2023 09:11:53 AM, Reid, Emmerson ) ICU called down about SpO2 not working in Room 1801. I went to check it out and the sensor was not lighting up. I brought the cable down to see if the sensor or trunk had an issue and it ended up being the trunk that was the problem. Replaced the trunk cable, working now.
65974 | 194178 | CM | CMPLT | NFF | 2023-04-19 09:35 | 2023-04-19 09:37 | Channel error | (TC 04/19/2023 09:36:14 AM, Baghirzade, Nara) Re-seated pressure sensor. (TC 04/19/2023 09:36:34 AM, Baghirzade, Nara) Re-seated pressure sensor. Check-in tests passed and completed. IUI connectors and device cleaned. No fault found
1049219 | 194179 | CM | CMPLT | CMPLT | 2023-04-19 09:36 | 2023-04-21 13:35 | Alarm volume too low | (TC 04/18/2023 09:38:48 AM, Scero, Tomlyn) Met with Dianne in CCU regarding red alarm volume limits. There was an incident where a red alarm wasn't heard on a tele patient. Incident report is being created. Someone had turned the central station volume down to 2 overnight (lowest limit) and it can't be heard during the day on the other side of the room. Showed the clinicians where to adjust the alarm volume. (TC 04/19/2023 09:41:19 AM, Scero, Tomlyn) received an email from Dianne asking me to set the lowest alarm volume limit to 4. I sent the Philips monitoring request form for her to fill out and have approved for our documentation. Alarm low limit volume was set to 4. (TC 04/21/2023 01:33:55 PM, Scero, Tomlyn) Change request form completed and attached to work order.
| 194180 | CM | CMPLT | CMPLT | 2023-04-19 09:43 | 2023-04-19 09:46 | Replaced NiMH batteries on otoscope handles and drop off stock | (TC 04/19/2023 09:45:52 AM, Baghirzade, Nara) Replaced batteries on otoscope handles; dropped off battery stock for ENT with Ana and Elena.
1062235 | 194181 | CM | CMPLT | CMPLT | 2023-04-19 09:58 | 2023-06-14 14:00 | Scale not functioning - 3009-1-29-802 | (TC 04/19/2023 09:59:20 AM, Ferreira, Qwynn) Tried to weigh the patient but error message was displayed. Lifted the patient to zero the bed but still not working Reset the bed and checked under the matress for any damage or loose cabling running to the sensor Nothing was noticed Tried to trouble shoot deeper and turned off eye watch to see however nothing was solving the problem. Moved the bed out of the patient room and helped to locate a spare. Notified stryker of the issue (TC 06/14/2023 01:58:53 PM, Ferreira, Qwynn) Mark reset the bed and scale was working fine, seems to be a hiccup in the system. Returned to unit
1063746 | 194182 | CM | CMPLT | CMPLT | 2023-04-19 10:02 | 2023-04-19 10:03 | Cleaning | (TC 04/19/2023 10:02:32 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
| 194183 | CM | CMPLT | CMPLT | 2023-04-19 12:11 | 2023-04-19 12:15 | 11N Decant Meeting | (TC 04/19/2023 12:14:59 PM, Baghirzade, Nara) 11N Decant Meeting with Naz
65898 | 194184 | FC | CMPLT | CMPLT | 2023-04-19 12:17 | 2023-04-19 12:17 | Functional check | (TC 04/19/2023 12:16:54 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
65365 | 194185 | CM | CMPLT | CMPLT | 2023-04-19 12:20 | 2023-04-19 12:21 | Functional check and transfer to 60M Rheumatology | (TC 04/19/2023 12:20:11 PM, Baghirzade, Nara) Spare VSM from the endo move; confirmed with Kamran that I can transfer this VSM over to 60M Rheumatology dept. All functional tests for NIBP and SpO2 passed. Provided the SpO2 sensor; NIBP hose + cuff and power cord with the device. Dropped it off with Dieter.
| 194186 | CM | CMPLT | CMPLT | 2023-04-19 14:08 | 2023-04-25 08:24 | relabeling 17 Philips APs on ICU 5th floor due to mix with AP7 label | (TC 04/19/2023 02:08:33 PM, Samanian, Kamran) attached label inside and outside - configured PICiX as well. (TC 04/25/2023 08:23:35 AM, Samanian, Kamran) done on AP5ICU-S01
65997 | 194187 | CM | CMPLT | CMPLT | 2023-04-19 14:20 | 2023-04-19 14:23 | Functional check/dataset upload confirmation | (TC 04/19/2023 02:23:11 PM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 22 mohms Leakage current: 91.1 uA Confirmed that the latest dataset was uploaded via network: April 18-2023
1066843 | 194188 | CM | CMPLT | CMPLT | 2023-04-19 14:21 | 2023-06-27 13:41 | CLV-190 intermitent problem- bad contacts- broken front panel | (TC 04/17/2023 02:21:09 PM, Samanian, Kamran) verified the problem. Replaced it with the good one which was brough up from BRONCH. Requested a rush loaner device so that defective one will be sent to Olympus for repair. attached is paper works. PO#7202197685 by Anita (TC 04/19/2023 02:25:21 PM, Samanian, Kamran) received the loaner (TC 04/20/2023 10:14:15 AM, Samanian, Kamran) shipped to KJV, order# 364087 (TC 06/28/2023 01:38:20 PM, Samanian, Kamran) received the device after repair. installed in the room 1 by Rob, tested and passed
68907 | 194189 | CM | CMPLT | CMPLT | 2023-04-19 14:24 | 2023-04-19 14:25 | Functional check/dataset upload confirmation | (TC 04/19/2023 02:24:09 PM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 30 mohms Leakage current: 73.3 uA Confirmed that the latest dataset was uploaded via network: April 18-2023
68881 | 194190 | CM | ACTIV | ISSUE | 2023-04-19 14:57 |  | Wrong Serial Number!?
| 194191 | CM | CMPLT | CMPLT | 2023-04-20 08:36 | 2023-04-17 08:36 | huddle meeting | (TC 04/17/2023 08:35:57 AM, Samanian, Kamran) done!
1066851 | 194192 | CM | CMPLT | CMPLT | 2023-04-20 08:58 | 2023-04-20 10:13 | Initial Accept | (TC 04/20/2023 08:58:37 AM, Samanian, Kamran) Performed electrical safety test, GR=80mOHM, LC=72uA. unit passed the tests. this unit is loaner (TC 04/20/2023 10:13:19 AM, Samanian, Kamran) installed in Bronch cart
64715 | 194193 | CM | CMPLT | CMPLT | 2023-04-20 09:51 | 2023-04-20 09:52 | BP cuff leak | (TC 04/20/2023 09:51:31 AM, Scero, Tomlyn) Slow BP leak. replaced cuff and tubing. Leaking resolved.
| 194194 | CM | CMPLT | CMPLT | 2023-04-20 10:17 | 2023-04-20 10:17 | meeting with Katrina about Hyland | (TC 04/20/2023 10:17:35 AM, Samanian, Kamran) done!
68710 | 194195 | PM | CMPLT | CMPLT | 2023-04-20 10:45 | 2023-04-20 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/20/2023 10:47:11 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 111 mohms Leakage current: 67.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
67960 | 194196 | CM | CMPLT | CMPLT | 2023-04-20 10:58 | 2023-04-20 10:58 | Broken cover for door hinge | (TC 04/20/2023 10:58:00 AM, Baghirzade, Nara) Replaced door hinge cover. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75%. Incubator is fully functional.
68786 | 194197 | PM | CMPLT | CMPLT | 2023-04-20 11:19 | 2023-04-20 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/20/2023 11:19:05 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Power cord was replaced. Electrical Safety - Ground resistance: 45 mohms Leakage current: 75.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
67110 | 194198 | FC | CMPLT | CMPLT | 2023-04-20 11:23 | 2023-04-20 11:25 | BP and temperature module functional check | (TC 04/20/2023 11:22:50 AM, Baghirzade, Nara) Device is in good physical condition. Set 120/80; measured 120/82 Set 80/40; measured 83/41 Set 190/120 ; measured 190/121 Temperature module is functional. Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
| 194199 | CM | CMPLT | CMPLT | 2023-04-20 11:30 | 2023-04-20 14:51 | CAB meeting for BD server upgrade | (TC 04/20/2023 11:30:05 AM, Samanian, Kamran) done!
64078 | 194200 | PM | CMPLT | CMPLT | 2023-04-20 14:23 | 2023-04-20 14:29 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/20/2023 02:25:38 PM, Ferreira, Qwynn) Received coupler Installedthe new coupler and re assembled the motor assembly and transformer Assembled the canopy and attached the wiring and covers Testing of the coupler - OK Motor Function - Good PM Passed Calibrated scale Calibrated line voltage Calibrated Main voltage Calibrated System votlage Calibrated RH Electrical Safety - OK Functional Test - OK
1064458 | 194201 | CM | CMPLT | CMPLT | 2023-04-20 14:25 | 2023-07-04 13:32 | MSH ER GE VENUE GO POCUS FAULTY BAR CODE READER, CAUSING ISSUES WITH OTHER FUNCTIONS | (TC 04/20/2023 02:40:56 PM, Whittington, Scott) MSH ER GE VENUE GO POCUS FAULTY BAR CODE READER, CAUSING ISSUES WITH OTHER FUNCTIONS msh er venue go barcode scan SITE 5153965UL12 REF# 1-547769965565 18006680732 (TC 06/30/2023 01:30:09 PM, Whittington, Scott) MSH ER GE VENUE GO POCUS FAULTY BAR CODE READER, CAUSING ISSUES WITH OTHER FUNCTIONS msh er venue go barcode scan SITE 5153965UL12 REF# 1-547769965565 18006680732 BAR CODE READER DRECEIVED AND INSTALLED | MSH ER GE VENUE GO POCUS FAULTY BAR CODE READER, CAUSING ISSUES WITH OTHER FUNCTIONS msh er venue go barcode scan SITE 5153965UL12 REF# 1-547769965565 18006680732
1064457 | 194202 | CM | CMPLT | CMPLT | 2023-04-20 14:43 | 2023-07-11 07:38 | MSH ER GE POCUSC VENUE 40 IAN CHERNOFF RELATED SCAN/OPEARTIONAL CONCERNS TO MIKE MERCER & KEVIN BOU KHERS | (TC 04/20/2023 02:43:46 PM, Whittington, Scott) MSH ER GE POCUS VENUE 40 IAN CHERNOFF RELATED SCAN/OPEARTIONAL CONCERNS TO MIKE MERCER & KEVIN BOU KHERS msh er venue go scan issues from ian chernoff | MSH ER GE POCUS VENUE 40 IAN CHERNOFF RELATED SCAN/OPEARTIONAL CONCERNS TO MIKE MERCER & KEVIN BOU KHERS msh er venue go scan issues from ian chernoff
1066107 | 194203 | CM | CMPLT | CMPLT | 2023-04-20 14:45 | 2023-04-20 14:46 | MSH GIM 10TH FLR GE POCUS VENUE 40 ECG MODULE INSTALL | (TC 04/20/2023 02:45:14 PM, Whittington, Scott) MSH GIM 10TH FLR GE POCUS VENUE 40 ECG MODULE INSTALL | MSH GIM 10TH FLR GE POCUS VENUE 40 ECG MODULE INSTALL
| 194204 | CM | CMPLT | CMPLT | 2023-04-20 14:53 | 2023-04-20 14:53 | PM biomed meeting | (TC 04/20/2023 02:52:59 PM, Samanian, Kamran) done!
32510 | 194205 | CM | CMPLT | CMPLT | 2023-04-20 16:06 | 2023-04-03 16:07 | MSH DR1 DETECTOR CONTROLLER NOT FUNCTIONING | (TC 04/01/2023 04:06:16 PM, Whittington, Scott) MSH DR1 DETECTOR CONTROLLER NOT FUNCTIONING, REPAIRED | MSH DR1 DETECTOR CONTROLLER NOT FUNCTIONING
1059828 | 194206 | CM | CMPLT | CMPLT | 2023-04-21 09:26 | 2023-04-21 09:26 | Cleaning | (TC 04/21/2023 09:26:19 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1066861 | 194207 | CM | CMPLT | CMPLT | 2023-04-21 10:14 | 2023-04-21 10:15 | Initial Accept | (TC 04/17/2023 10:13:47 AM, Samanian, Kamran) Unboxed the cart with John (TC 04/21/2023 10:14:09 AM, Samanian, Kamran) performed elecytrical safety test, GR=160mhom, LC=160uA. Moved the cart to ENDO 6th floor
1050479 | 194208 | CM | CMPLT | CMPLT | 2023-04-21 10:58 | 2023-04-21 10:59 | Expiratory cassette membrane replacement | (TC 04/21/2023 10:58:27 AM, Canjimavadivel, Kamallesh ) Replaced diaphragm, reset membrane capacity counter
67944 | 194209 | CM | CMPLT | CMPLT | 2023-04-21 11:39 | 2023-04-21 11:43 | Tilt not working and scale issue | (TC 04/21/2023 11:40:40 AM, Ferreira, Qwynn) Fixed tilt assembly Tightened the tilt latch assembly Calibrated scale and it was working correctly Functional test - OK
63809 | 194210 | CM | CMPLT | NFF | 2023-04-21 12:55 | 2023-04-21 12:58 | No description | (TC 04/21/2023 12:57:01 PM, Canjimavadivel, Kamallesh ) Tested all functions, working ok Device charging and battery capacity good No fault found | Sent down with a blank repair tag
67152 | 194211 | CM | CMPLT | NFF | 2023-04-21 14:11 | 2023-04-21 14:11 | Checking BP | (TC 04/21/2023 02:10:54 PM, Reid, Emmerson ) Was asked to come and check the BP, it was working fine.
66867 | 194212 | CM | CMPLT | NFF | 2023-04-21 14:13 | 2023-04-21 14:13 | Checking BP | (TC 04/21/2023 02:12:56 PM, Reid, Emmerson ) Was asked to come and check the BP, it was working fine.
1066872 | 194213 | ACC | CMPLT | CMPLT | 2023-04-21 14:35 | 2023-04-21 14:46 | Incoming inspection | (TC 04/21/2023 02:43:06 PM, Canjimavadivel, Kamallesh ) Assembled device and mobile stand Added to database Functional tests pass
1055005 | 194214 | CM | CMPLT | CMPLT | 2023-04-21 14:47 | 2023-04-21 14:48 | Broken Handle and bin | (TC 04/21/2023 02:47:34 PM, Ferreira, Qwynn) Fixed Handle and fixed bin Sent back to unit
65843 | 194215 | CM | CMPLT | CMPLT | 2023-04-24 08:35 | 2023-04-24 08:38 | Hose ripped | (TC 04/24/2023 08:34:54 AM, Reid, Emmerson ) Cut off the ripped section near the end and reattached. Working now.
1062232 | 194216 | CM | CMPLT | CMPLT | 2023-04-24 08:35 | 2023-11-16 15:14 | not working
66809 | 194217 | CM | CMPLT | CMPLT | 2023-04-24 09:32 | 2023-04-24 09:33 | No yellow tag | (TC 04/24/2023 09:32:22 AM, Baghirzade, Nara) Replaced handle. No yellow tag so did a full functional test of all modules; all functional tests passed. | Broken handle
67915 | 194218 | CM | CMPLT | CMPLT | 2023-04-24 09:34 | 2023-04-24 09:35 | Latch recall | (TC 04/24/2023 09:34:33 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
65911 | 194219 | CM | CMPLT | CMPLT | 2023-04-24 09:57 | 2023-04-24 10:10 | Channel error | (TC 04/24/2023 10:06:30 AM, Baghirzade, Nara) Replaced pressure sensor. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65957 | 194220 | CM | CMPLT | CMPLT | 2023-04-24 10:16 | 2023-04-24 10:18 | Pressure disc not working | (TC 04/24/2023 10:17:00 AM, Baghirzade, Nara) Replaced pressure sensor. Replaced rear case. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
| 194221 | CM | CMPLT | CMPLT | 2023-04-24 10:44 | 2023-05-11 09:49 | Lamp Halogen for OPG | (TC 05/11/2023 09:48:20 AM, Reid, Emmerson ) Recieved the bulb. Went over to replace and it was working. NFF
1062325 | 194222 | CM | CMPLT | CMPLT | 2023-04-24 11:06 | 2023-04-24 11:07 | Broken secure connect power cord/adapter | (TC 04/24/2023 11:05:47 AM, Baghirzade, Nara) Replaced secure connect power cord/adapter. Removed broken power cord and the connector part that was stuck inside the secure connect device.
1059778 | 194223 | FC | CMPLT | CMPLT | 2023-04-24 11:11 | 2023-04-24 11:11 | QC checklist | (TC 04/24/2023 11:10:48 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
32510 | 194224 | CM | CMPLT | CMPLT | 2023-04-24 12:01 | 2023-04-24 12:02 | MSH DR1 DETECTOR Z NOT MOVING | (TC 04/21/2023 10:00:00 PM, Whittington, Scott) MSH DR1 DETECTOR Z NOT MOVING ABLE TOIMPROVE DETECTOR TELESCOPE ISSUE | MSH DR1 DETECTOR Z NOT MOVING
502452 | 194225 | CM | CMPLT | CMPLT | 2023-04-24 12:10 | 2023-08-08 12:04 | MSH OPG SPP VOLUSON A PROBE FAULT RIC5-9-D RE# H48651MS SN# 239785KR9 MFG FEB 2019 ARTIFACT | (TC 04/24/2023 12:15:55 PM, Whittington, Scott) MSH OPG SPP VOLUSON A PROBE FAULT RIC5-9-D RE# H48651MS SN# 239785KR9 MFG FEB 2019 ARTIFACT (TC 08/08/2023 12:02:59 PM, Whittington, Scott) MSH OPG SPP VOLUSON A PROBE FAULT RIC5-9-D RE# H48651MS SN# 239785KR9 MFG FEB 2019 ARTIFACT UPDATED GE CALL CENTRE THAT PROBE HAD BEEN REPLACED | MSH OPG SPP VOLUSON A PROBE FAULT RIC5-9-D RE# H48651MS SN# 239785KR9 MFG FEB 2019 ARTIFACT site# 416596VE10A REF# 1-547808702432
| 194226 | PA | CMPLT | CMPLT | 2023-04-24 14:15 | 2023-04-28 08:17 | Ordering syringe pressure sensors | (TC 04/24/2023 02:17:35 PM, Baghirzade, Nara) SC 1002940327 (TC 04/28/2023 08:16:14 AM, Baghirzade, Nara) PO 7202202201
501260 | 194227 | CM | CMPLT | CMPLT | 2023-04-24 15:06 | 2023-05-02 11:00 | MSH ICU SONOSITE MTURBO BLUE SCREEN FAILURE CODE CTRL : 0x41e4 (16868) D: Ox30 | (TC 04/24/2023 03:09:22 PM, Whittington, Scott) MSH ICU SONOSITE MTURBO BLUE SCREEN FAILURE CODE CTRL : 0x41e4 (16868) D: Ox30 (TC 04/24/2023 03:09:22 PM, Dolinar, John) MSH ICU SONOSITE MTURBO BLUE SCREEN FAILURE CODE CTRL : 0x41e4 (16868) D: Ox30 (TC 04/25/2023 02:40:04 PM, Whittington, Scott) MSH ICU SONOSITE MTURBO BLUE SCREEN FAILURE CODE CTRL : 0x41e4 (16868) D: Ox30 SONOSITE QUOTED $13K-$14K FOR MAIN BOARD, AND ALL THREE PROBES NEED TO BE REPLACED ESTIMATE $30K TOTAL (TC 05/02/2023 10:58:27 AM, Whittington, Scott) MSH ICU SONOSITE MTURBO BLUE SCREEN FAILURE CODE CTRL : 0x41e4 (16868) D: Ox30 SONOSITE QUOTED $13K-$14K FOR MAIN BOARD, AND ALL THREE PROBES NEED TO BE REPLACED ESTIMATE $30K TOTAL SYSTEM REPLACED BY NEW SPARQ SYSTEM RETIRED AND STORED IN TGH GW | MSH ICU SONOSITE MTURBO BLUE SCREEN FAILURE CODE CTRL : 0x41e4 (16868) D: Ox30 SONOSITE QUOTED $13K-$14K FOR MAIN BOARD, AND ALL THREE PROBES NEED TO BE REPLACED ESTIMATE $30K TOTAL
| 194228 | OC | CMPLT | CMPLT | 2023-04-24 15:10 | 2023-05-03 10:54 | John Cortez On-Call; Apr 21-27, 2023 | (TC 04/21/2023 11:30:00 PM, Cortez, John) ICU called regarding an SPO2 malfunction error on monitor in room 1814. Caller mentioned that swapping cables/sensor did not resolve issue, and other parameters were reading ok. They also said that they would use defib to read SPO2 in the short term, but asked for issue to be resolved for weekend Came on-site next morning to troubleshoot. Swapping X3 also didn't resolve issue, but restarting MX800 did. They asked for an alternative spare SPO2 device, so I supplied them with a spare Masimo Radical+dock. (TC 04/24/2023 10:30:00 PM, Cortez, John) ER called saying they were only seeing lead III, and weren't getting resp rate in room 43. They asked first if I was on site, when I replied no, they said that it wasn't urgent, and that they would try swapping leads/electrodes. Did not need to come on site
1046196 | 194229 | CM | CMPLT | CMPLT | 2023-04-24 15:35 | 2023-04-24 15:35 | Computer mosue and keyboard freezes intermittently | Computer mosue and keyboard freezes intermittently -Hard shut down has to be done to get the system working again Replacement PC was ordered by Siemens (Different Model was delivered, which required full software reload) -Upon troubleshooting it was found the front USB ports work ok even when the rear ports of the tower stops working. -Mouse and keyboard were plugged it at the front port. -In the meantime compatible PC tower was ordred that will require less time to install.
67106 | 194230 | CM | CMPLT | CMPLT | 2023-04-24 15:48 | 2023-04-24 15:51 | "BP will not register" | (TC 04/24/2023 03:48:38 PM, Cortez, John) Confirmed problem, error C04 (leak) It's a cuff issue, swapped All parameters tested ok
67099 | 194231 | CM | CMPLT | CMPLT | 2023-04-24 15:52 | 2023-04-24 15:53 | "bp will not register" | (TC 04/24/2023 03:52:32 PM, Cortez, John) Confirmed problem, error C04 (leak) It's a cuff issue, swapped All parameters tested ok
1058273 | 194232 | CM | CMPLT | CMPLT | 2023-04-24 16:01 | 2023-05-04 15:56 | unable to update drug library | (TC 04/21/2023 08:53:11 AM, Cortez, John) error 600.6875, communication hardware fault Swapped network card and issue resolved, dataset downloaded Will order replacement (TC 05/04/2023 03:53:48 PM, Cortez, John) Removed spare network card and replaced with ordered replacement (TC 05/04/2023 03:53:48 PM, Cortez, John) Removed spare network card and replaced with ordered replacement Completed PM successfully
502171 | 194233 | CM | CMPLT | CMPLT | 2023-04-24 23:57 | 2023-04-26 15:58 | Studies not transferring | (TC 04/24/2023 11:57:11 PM, Schreier, Ryan) Studies not showing up No DICOM errors Manual transfer seems to work then studies disappear (TC 04/26/2023 03:56:57 PM, Schreier, Ryan) Test and verify operation OK
| 194234 | CM | CMPLT | CMPLT | 2023-04-25 08:08 | 2023-04-25 08:31 | Defib accessory charge back | (TC 04/25/2023 08:30:44 AM, Canjimavadivel, Kamallesh ) Provided missing adapter for defib
1066851 | 194235 | CM | CMPLT | CMPLT | 2023-04-25 08:31 | 2023-04-21 08:33 | moving the loaner | (TC 04/21/2023 08:31:29 AM, Samanian, Kamran) Initially this loaner was installed on the BRONCH cart but because light source in ENDO EMERG cart has damaged front panel then the loaner was moved to ENDO emerg cart. light source with the damaged front panel was moved to BRONCH cart and reuqested to get a loaner for that so we will send it for repair.
1054470 | 194236 | CM | CMPLT | CMPLT | 2023-04-25 08:45 | 2023-06-02 10:32 | Patient right side head rail does not go down | (TC 04/25/2023 08:45:11 AM, Canjimavadivel, Kamallesh ) Confirmed issue reported by unit. Contacted Baxter tech support to create a repair ticket (TC 05/23/2023 10:29:47 AM, Cortez, John) Bed repaired by Vince from Hillrom Checked functionality before turning to service
1061895 | 194237 | CM | CMPLT | CMPLT | 2023-04-25 08:49 | 2023-05-25 16:00 | Head section actuator damaged | (TC 04/24/2023 08:49:12 AM, Cortez, John) Notified by 14N about broken bed. Confirmed issue with head section movement. Actuator damaged. Requested service call with stryker. (TC 05/25/2023 03:59:19 PM, Cortez, John) Bed repaired, Mark from stryker sent me the service report
| 194238 | CM | CMPLT | CMPLT | 2023-04-25 08:55 | 2023-04-25 09:06 | No O2 sat reading on monitor in room 13 | (TC 04/25/2023 09:04:35 AM, Scero, Tomlyn) Gave nurse a replacement extention cable to replace the defective one in room 13. patient was in isolation so I couldn't get in to get the asset number of the X2.
65416 | 194239 | CM | CMPLT | CMPLT | 2023-04-25 09:06 | 2023-04-21 15:00 | Latch recall | (TC 04/21/2023 03:06:00 PM, Ferreira, Qwynn) Latch recall completed.
67146 | 194240 | CM | CMPLT | CMPLT | 2023-04-25 09:11 | 2023-04-26 08:11 | not charging | (TC 04/25/2023 09:11:23 AM, Samanian, Kamran) replaced the battery, added a cuff. left it to charge the battery, tested all functions and passed
| 194241 | CON | CMPLT | CMPLT | 2023-04-25 11:08 | 2023-05-08 15:50 | CPET report for study | (TC 04/25/2023 11:08:10 AM, Baghirzade, Nara) Working on editing the CPET report as per the study specifications. Upcoming study for CPET lab. Waiting to confirm report adjustments with Gaby. (TC 05/08/2023 03:48:34 PM, Baghirzade, Nara) Met with Gaby to discuss the feedback from research team; added PetO2 to the graph. Tried to change the layout of the table as per researcher's request but expair software is very limited on layout editing and was unable to do a layout as specified in the study report. E-mailed this to Gaby and she had sent it to the research team. As for now, the report is adequate for the study.
67945 | 194242 | CM | CMPLT | CMPLT | 2023-04-25 11:29 | 2023-04-25 11:30 | Latch recall | (TC 04/25/2023 11:29:30 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
1065042 | 194243 | CM | CMPLT | CMPLT | 2023-04-25 12:26 | 2023-04-25 12:27 | labeling and configuration | (TC 04/25/2023 12:26:32 PM, Samanian, Kamran) labeled the fiber optics, got a copy of config file . made remote access to switch as well.
1065043 | 194244 | CM | CMPLT | CMPLT | 2023-04-25 12:28 | 2023-04-25 12:28 | labeling and configuration | (TC 04/25/2023 12:27:51 PM, Samanian, Kamran) labeled the fiber optics, got a copy of config file . made remote access to switch as well.
| 194245 | CM | CMPLT | CMPLT | 2023-04-25 12:51 | 2023-05-11 09:48 | Ordering AA batteries for Biomed | (TC 04/25/2023 01:07:44 PM, Reid, Emmerson ) placed the order
1062175 | 194246 | CM | CMPLT | CMPLT | 2023-04-25 12:55 | 2023-04-25 12:56 | Patient helper mount and arm installation | (TC 04/25/2023 12:55:29 PM, Baghirzade, Nara) Installed the patient helper mount and patient helper for the bed.
67945 | 194247 | CM | CMPLT | CMPLT | 2023-04-25 13:34 | 2023-04-25 13:35 | Broken hinge | (TC 04/25/2023 01:34:06 PM, Baghirzade, Nara) Replaced porthole door. Repaired door hinge. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35.2C and 74% humidity. Incubator is fully functional.
501583 | 194248 | CM | CMPLT | CMPLT | 2023-04-25 14:39 | 2023-04-25 14:39 | SPP 15-217 Rose L&D Hallway (IU22) 501583 PT DB SEND ALL UNSENT STUDIES CLEAR PT STUDIES EXCEPT SAME WEEK | (TC 04/25/2023 02:39:09 PM, Whittington, Scott) SPP 15-217 Rose L&D Hallway (IU22) 501583 PT DB SEND ALL UNSENT STUDIES CLEAR PT STUDIES EXCEPT SAME WEEK | SPP 15-217 Rose L&D Hallway (IU22) 501583 PT DB SEND ALL UNSENT STUDIES CLEAR PT STUDIES EXCEPT SAME WEEK
1064418 | 194249 | CM | CMPLT | CMPLT | 2023-04-25 15:19 | 2023-05-05 11:16 | Port Door Fell off | (TC 04/25/2023 03:18:50 PM, Ferreira, Qwynn) Met with John to discuss will be sending emails to the clinical team as well as Michael from AME to disucss a solution for the repairs as well as future steps to reduce the occurances of broken parts Took photos of the broken parts and currentl issues with the MRI Tried to find a solution for a fix which included the use of a plastic clear hinge (TC 05/05/2023 10:45:55 AM, Chukri, Samir) Replaced the port door and spoke with Michael about the warranty issues being faced with AME/LMT Noticed the main door seal was also not in place but rather fell off. (TC 05/05/2023 11:15:33 AM, Ferreira, Qwynn) replaced seal for the port door. Discussed with Michael about the warranty issues and he mentioned that LMT will be in Toronto June 4-9 so we may discuss our problems.
| 194250 | CM | CMPLT | CMPLT | 2023-04-26 07:48 | 2023-04-26 07:49 | Blue BNC cable in ENDO room#1 | (TC 04/26/2023 07:48:18 AM, Samanian, Kamran) BNC connector on the Blue cable ( used with salve monitor behind the processor was loose. Arranged with Olympus tech ( Steve ). He came to the site and fixed the connector. Tested and passed
68027 | 194251 | CM | CMPLT | CMPLT | 2023-04-26 08:38 | 2023-04-26 08:40 | No reading | (TC 04/26/2023 08:38:38 AM, Canjimavadivel, Kamallesh ) Cleaned tip Calibration verification pass
65159 | 194252 | PM | CMPLT | CMPLT | 2023-04-26 08:55 | 2023-04-26 09:16 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/26/2023 09:15:17 AM, Reid, Emmerson ) PM Complete, see proceedure for details.
1047866 | 194253 | CM | CMPLT | CMPLT | 2023-04-26 09:25 | 2023-04-26 09:29 | Error "check syringe plunger not in place" | (TC 04/26/2023 09:24:46 AM, Baghirzade, Nara) Replaced plungerhead assembly. Calibrated force sensor. All functional tests passed; refer to procedure document for results.
1066411 | 194254 | ACC | CMPLT | CMPLT | 2023-04-26 09:41 | 2023-04-26 09:45 | Incoming inspection | (TC 04/26/2023 09:41:45 AM, Cortez, John) Visual inspection, function test passed 120 mOhms, 33 uA SFC
1066882 | 194255 | ACC | CMPLT | CMPLT | 2023-04-26 09:43 | 2023-04-27 11:27 | Incoming inspection | (TC 04/27/2023 11:26:37 AM, Canjimavadivel, Kamallesh ) Functional test pass
1066883 | 194256 | ACC | CMPLT | CMPLT | 2023-04-26 09:46 | 2023-04-27 11:26 | Incoming inspection | (TC 04/27/2023 11:21:35 AM, Canjimavadivel, Kamallesh ) Functional test pass
1066884 | 194257 | ACC | CMPLT | CMPLT | 2023-04-26 09:48 | 2023-04-27 11:21 | Incoming inspection | (TC 04/27/2023 11:20:21 AM, Canjimavadivel, Kamallesh ) Functional test pass
1066885 | 194258 | ACC | CMPLT | CMPLT | 2023-04-26 09:49 | 2023-04-27 11:20 | Incoming inspection | (TC 04/27/2023 11:19:50 AM, Canjimavadivel, Kamallesh ) Functional test pass
1056055 | 194259 | CM | CMPLT | CMPLT | 2023-04-26 09:49 | 2023-04-26 09:51 | MSH-5-506 ANGIO AZURION 7 M20 ASSIST MIKE FROM BIRNIE ELECTRIC IN LOCATING UPS, AND NON-STANDARD KEY IN JOANNE NOSEWORTHYS DESK | (TC 04/26/2023 09:50:46 AM, Whittington, Scott) MSH-5-506 ANGIO AZURION 7 M20 ASSIST MIKE FROM BIRNIE ELECTRIC IN LOCATING UPS, AND NON-STANDARD KEY IN JOANNE NOSEWORTHYS DESK (TC 04/26/2023 09:50:46 AM, Whittington, Scott) MSH-5-506 ANGIO AZURION 7 M20 ASSIST MIKE FROM BIRNIE ELECTRIC IN LOCATING UPS, AND NON-STANDARD KEY IN JOANNE NOSEWORTHYS DESK | MSH-5-506 ANGIO AZURION 7 M20 ASSIST MIKE FROM BIRNIE ELECTRIC IN LOCATING UPS, AND NON-STANDARD KEY IN JOANNE NOSEWORTHYS DESK
1066886 | 194260 | ACC | CMPLT | CMPLT | 2023-04-26 09:50 | 2023-04-27 11:19 | Incoming inspection | (TC 04/27/2023 11:18:37 AM, Canjimavadivel, Kamallesh ) Functional test pass
1066887 | 194261 | ACC | CMPLT | CMPLT | 2023-04-26 09:50 | 2023-04-27 11:18 | Incoming inspection | (TC 04/27/2023 11:18:00 AM, Canjimavadivel, Kamallesh ) Functional test pass
| 194262 | CM | CMPLT | CMPLT | 2023-04-26 10:07 | 2023-08-30 10:44 | ordered thermometers for MBU
1056125 | 194263 | PM | CMPLT | CMPLT | 2023-04-26 10:18 | 2023-04-26 10:21 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/26/2023 10:19:38 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
| 194264 | CM | CMPLT | CMPLT | 2023-04-26 10:38 | 2023-04-26 10:38 | sent two maismo extension cable to PNAC
69213 | 194265 | PM | CMPLT | CMPLT | 2023-04-26 11:36 | 2023-04-26 11:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/26/2023 11:37:24 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68689 | 194266 | CM | CMPLT | CMPLT | 2023-04-26 11:39 | 2023-04-26 11:42 | unable to update dataset | (TC 04/24/2023 11:39:27 AM, Cortez, John) Network profile not configured Configured in ASM, and updated dataset
| 194267 | CM | CMPLT | CMPLT | 2023-04-26 12:38 | 2023-04-26 12:39 | BD server upgrade meeting | (TC 04/26/2023 12:38:33 PM, Baghirzade, Nara) BD server upgrade meeting
69440 | 194268 | PM | CMPLT | CMPLT | 2023-04-26 12:41 | 2023-04-26 13:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/26/2023 01:08:01 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69660 | 194269 | PM | CMPLT | CMPLT | 2023-04-26 13:15 | 2023-04-26 13:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/26/2023 01:15:11 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 194270 | PA | CMPLT | CMPLT | 2023-04-26 13:54 | 2023-05-12 22:08 | SpO2 cables | (TC 04/26/2023 02:23:26 PM, Canjimavadivel, Kamallesh ) Contacted sales rep for quote (TC 04/27/2023 07:36:36 AM, Canjimavadivel, Kamallesh ) Placed order for trunk cables
1053661 | 194271 | CM | CMPLT | CMPLT | 2023-04-26 13:55 | 2023-06-15 13:43 | CLV-190 damaged front panel | (TC 04/26/2023 01:55:25 PM, Samanian, Kamran) received the loaner ( PO#7400208051 ) shiped this device to company for repair. (TC 06/15/2023 07:54:44 AM, Samanian, Kamran) It was returned by OLympus and installed in the Emerg Cart
| 194272 | CM | CMPLT | CMPLT | 2023-04-26 14:08 | 2023-04-26 14:08 | capturing ICU philips network info from closets and rooms | (TC 04/26/2023 02:08:15 PM, Samanian, Kamran) done!
| 194273 | CM | CMPLT | CMPLT | 2023-04-26 14:09 | 2023-05-24 14:39 | BD server upgrade | (TC 04/26/2023 02:09:13 PM, Samanian, Kamran) discussed about cutover. MAY 24 and plans (TC 05/24/2023 02:38:46 PM, Samanian, Kamran) BD server upgarde done. tested and passed
1056577 | 194274 | CM | CMPLT | CMPLT | 2023-04-26 14:56 | 2023-04-26 15:00 | Low battery | (TC 04/26/2023 02:59:55 PM, Canjimavadivel, Kamallesh ) Replaced battery
1056562 | 194275 | CM | CMPLT | CMPLT | 2023-04-26 15:01 | 2023-04-26 15:02 | Low battery | (TC 04/26/2023 03:01:31 PM, Canjimavadivel, Kamallesh ) Replaced battery
| 194276 | CM | CMPLT | CMPLT | 2023-04-26 15:13 | 2023-04-26 15:12 | STRYKER DICOM support | STRYKER DICOM support
| 194277 | CM | CMPLT | CMPLT | 2023-04-27 07:58 | 2023-04-27 08:01 | Moving Force Triads to Mobile Carts | (TC 04/27/2023 07:58:52 AM, Reid, Emmerson ) Moved the FT10 from OR3 to OR14. Moved the 2450 from OR12 to OR3 Moved the Force Triade from OR 14 to a mobile cart. Synced the smoke evacs in every case.
1061333 | 194278 | FC | CMPLT | CMPLT | 2023-04-27 08:27 | 2023-04-27 08:33 | QC checklist | (TC 04/27/2023 08:30:19 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1054019 | 194279 | CM | CMPLT | NFF | 2023-04-27 09:01 | 2023-04-27 09:03 | "changed batteries and not turning on" | (TC 04/26/2023 09:01:44 AM, Cortez, John) Charged battery, unit turned on normally
| 194280 | CM | CMPLT | CMPLT | 2023-04-27 10:12 | 2023-04-27 10:13 | Ritter 354 bulb replacement in Procedure 1 | (TC 04/27/2023 10:12:21 AM, Samanian, Kamran) received the call from IVF ( wendy ) that bulb on exam light in procedure room 1 is gone. arrived in the room and replaced the bulb, tested and passed. noticed the socket contacts are burnt out too. advised that socket needs to be replaced soon as it has to be ordered.
| 194281 | CM | CMPLT | CMPLT | 2023-04-27 12:10 | 2023-05-01 12:32 | MSH SURGICAL SKILLS WORKSHOP SONOSITE PX LOANER | (TC 05/01/2023 12:32:09 PM, Whittington, Scott) MSH SURGICAL SKILLS WORKSHOP SONOSITE PX LOANER | MSH SURGICAL SKILLS WORKSHOP SONOSITE PX LOANER
| 194282 | CM | CMPLT | CMPLT | 2023-04-27 14:01 | 2023-04-27 14:01 | Patient Committe meeting | (TC 04/27/2023 02:01:12 PM, Samanian, Kamran) done!
1054294 | 194283 | CM | CMPLT | CMPLT | 2023-04-27 14:05 | 2023-04-27 14:09 | MSH OPG SPP GE VOLUSON 10 PROBE RIC 6 12-D TV PROBE ACCEPTANCE APRIL 27 2023 SN# 273602KR3 REPLACED SN# 253790KR0 | (TC 04/27/2023 02:08:33 PM, Whittington, Scott) MSH OPG SPP GE VOLUSON 10 PROBE RIC 6 12-D TV PROBE ACCEPTANCE APRIL 27 2023 SN# 273602KR3 REPLACED SN# 253790KR0 | MSH OPG SPP GE VOLUSON 10 PROBE RIC 6 12-D TV PROBE ACCEPTANCE APRIL 27 2023 SN# 273602KR3 REPLACED SN# 253790KR0
63874 | 194284 | CM | CMPLT | CMPLT | 2023-04-27 14:11 | 2023-04-27 14:12 | Latch recall | (TC 04/27/2023 02:11:33 PM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
63874 | 194285 | CM | CMPLT | CMPLT | 2023-04-27 14:13 | 2023-04-27 14:17 | Hood replacement | (TC 04/27/2023 02:13:43 PM, Baghirzade, Nara) Replaced hood. Replaced porthole door. Replaced corner grommets and big grommet. Replaced hinge pin on door.
1066917 | 194286 | ACC | CMPLT | CMPLT | 2023-04-27 14:47 | 2023-04-28 15:05 | Incoming inspection | (TC 04/28/2023 03:04:47 PM, Canjimavadivel, Kamallesh ) Asset tagged Set clock Changed nurse call setting
1066918 | 194287 | ACC | CMPLT | CMPLT | 2023-04-27 14:50 | 2023-04-28 15:07 | Incoming inspection | (TC 04/28/2023 03:06:40 PM, Canjimavadivel, Kamallesh ) Asset tagged Set clock Changed nurse call setting
1066919 | 194288 | ACC | CMPLT | CMPLT | 2023-04-27 14:51 | 2023-04-28 15:08 | Incoming inspection | (TC 04/28/2023 03:07:58 PM, Canjimavadivel, Kamallesh ) Asset tagged Set clock Changed nurse call setting
1066920 | 194289 | ACC | CMPLT | CMPLT | 2023-04-27 14:52 | 2023-04-28 15:11 | Incoming inspection | (TC 04/28/2023 03:10:22 PM, Canjimavadivel, Kamallesh ) Asset tagged Set clock Changed nurse call setting
1066921 | 194290 | ACC | CMPLT | CMPLT | 2023-04-27 14:54 | 2023-04-28 15:12 | Incoming inspection | (TC 04/28/2023 03:11:26 PM, Canjimavadivel, Kamallesh ) Asset tagged Set clock Changed nurse call setting
1066922 | 194291 | ACC | CMPLT | CMPLT | 2023-04-27 14:55 | 2023-04-28 15:13 | Incoming inspection | (TC 04/28/2023 03:12:29 PM, Canjimavadivel, Kamallesh ) Asset tagged Set clock Changed nurse call setting
| 194292 | CM | CMPLT | CMPLT | 2023-04-27 17:33 | 2023-04-27 17:33 | Biomed Support Tool upgrade | (TC 04/27/2023 05:33:20 PM, Samanian, Kamran) done, upgraded to 16.0
68542 | 194293 | CM | CMPLT | CMPLT | 2023-04-28 09:33 | 2023-04-28 10:57 | Not passing self test - MAINS failure | (TC 04/28/2023 10:53:10 AM, Scero, Tomlyn) Self test wouldn't pass after weekly system reset and the error that was displayed in red was MAINS on the test 2 page. I noticed the battery light wasn't lit at all. We unplugged the device and it completely shut off. I plugged it back in and checked that the battery connection wasn't loose. After restarting the device the battery light came on and it passed the self check. We should keep an eye on this device just to be sure. Possibly a loose connection or some sort of disconnect in the software that was resolved when the device was abbruptly shut off and forced to restart.
| 194294 | CM | CMPLT | CMPLT | 2023-04-28 10:35 | 2023-04-28 10:35 | OR bed training by emmerson | (TC 04/28/2023 10:35:22 AM, Samanian, Kamran) Emmerson showed how to work on OR table in OR15
1059709 | 194295 | CM | CMPLT | CMPLT | 2023-04-28 10:49 | 2023-04-28 10:51 | Only Lead II working | (TC 04/28/2023 10:49:22 AM, Reid, Emmerson ) Come down saying only lead II was working, meaning maybe an issue with LA I believe. Before turning on I made sure all of the leads were securely connected. Connected leads to my simulator and everything was working. I probably pushed in the cable and fixed it. There is no visible damage either. Working now.
1056441 | 194296 | FC | CMPLT | CMPLT | 2023-04-28 10:53 | 2023-04-28 10:54 | Functional check | (TC 04/28/2023 10:53:42 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1054980 | 194297 | FC | CMPLT | CMPLT | 2023-04-28 10:54 | 2023-04-28 10:55 | Functional check | (TC 04/28/2023 10:54:54 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055916 | 194298 | FC | CMPLT | CMPLT | 2023-04-28 10:57 | 2023-04-28 10:57 | Functional check | (TC 04/28/2023 10:56:49 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1054978 | 194299 | FC | CMPLT | CMPLT | 2023-04-28 10:57 | 2023-04-28 10:58 | Functional check | (TC 04/28/2023 10:57:39 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056434 | 194300 | CM | CMPLT | CMPLT | 2023-04-28 10:58 | 2023-04-28 10:59 | Functional check | (TC 04/28/2023 10:58:17 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055919 | 194301 | FC | CMPLT | CMPLT | 2023-04-28 10:59 | 2023-04-28 10:59 | Functional check | (TC 04/28/2023 10:59:07 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056437 | 194302 | FC | CMPLT | CMPLT | 2023-04-28 11:00 | 2023-04-28 11:01 | Functional check | (TC 04/28/2023 11:00:17 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056438 | 194303 | CM | CMPLT | CMPLT | 2023-04-28 11:21 | 2023-04-28 11:22 | Functional check + MSH config upload | (TC 04/28/2023 11:21:44 AM, Baghirzade, Nara) Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Removed the radical battery. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055918 | 194304 | FC | CMPLT | CMPLT | 2023-04-28 11:37 | 2023-04-28 11:38 | Functional check | (TC 04/28/2023 11:36:54 AM, Baghirzade, Nara) Changed the lock on the radical. Radical was not being detected by Root; 1st pin on the radical was depressed. Popped the pin up and it was being detected by Root. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056436 | 194305 | FC | CMPLT | CMPLT | 2023-04-28 11:48 | 2023-04-28 11:49 | Functional check | (TC 04/28/2023 11:48:10 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional. Replaced lock on radical. (TC 04/28/2023 11:48:10 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional. Replaced lock on radical. Replaced SpO2 sensor.
1056435 | 194306 | FC | CMPLT | CMPLT | 2023-04-28 12:01 | 2023-04-28 12:02 | Functional check | (TC 04/28/2023 12:01:17 PM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 194307 | CM | CMPLT | CMPLT | 2023-04-28 12:29 | 2023-06-07 10:11 | Shaker problem. order new motor | (TC 06/01/2023 07:41:10 AM, Chukri, Samir) contact and follow up with MI Motor comany reg, the motor availability . contact EBERBACH to get another price (TC 06/06/2023 02:36:31 PM, Chukri, Samir) Picked up the motor from Motion Scarbrough location, replaced the motor and tested working fine waiting for the right 5A Fuse, to complete the job. (TC 06/07/2023 10:09:45 AM, Chukri, Samir) completed repair and replaced fuse 5a . tested working fine, emailed Yasmin for pick up. | EBERBACH Model # E6010.00 SN# 042015
1066938 | 194308 | ACC | CMPLT | CMPLT | 2023-04-28 13:37 | 2023-04-28 13:40 | Incoming Inspection | (TC 03/01/2023 01:37:25 PM, Cortez, John) Unboxed/assembled bladder scanner and rollstand Visual/functional test passed
| 194309 | CM | CMPLT | CMPLT | 2023-04-28 14:42 | 2023-04-28 14:42 | Meeting with Sarabeth 11N | (TC 04/28/2023 02:41:46 PM, Baghirzade, Nara) Met with Sarabeth to discuss VSM, resp accessories, and management of their VSMs.
51580 | 194310 | PM | ACTIV | ISSUE | 2023-05-01 03:25 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52244 | 194311 | PM | CANCL | CANCL | 2023-05-01 03:25 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
60094 | 194312 | PM | ACTIV | ISSUE | 2023-05-01 03:25 |  | PM-MAJOR - Annual PM For: Scales, Infant
62125 | 194313 | PM | CANCL | CANCL | 2023-05-01 03:25 |  | PM-MAJOR - Annual PM For: LASERS, SURGICAL, CARBON DIOXIDE
62192 | 194314 | PM | CMPLT | UNL | 2023-05-01 03:25 | 2023-08-16 13:26 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 08/16/2023 01:25:30 PM, Rendon Lopez , Elizabeth) Unable to locate +3yrs. Retired from service
62226 | 194315 | PM | CANCL | CANCL | 2023-05-01 03:25 |  | PM-MAJOR - Annual PM For: Simulators, Multiparameter
62267 | 194316 | PM | CMPLT | CMPLT | 2023-05-01 03:25 | 2023-07-19 10:49 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 10:48:05 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
62288 | 194317 | PM | CNP | ENL | 2023-05-01 03:25 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency | No longer located on CCU cart
62369 | 194318 | PM | CMPLT | CMPLT | 2023-05-01 03:25 | 2023-06-02 11:03 | PM-MAJOR - Annual PM For: Simulators, Transducer Output, Invasive Blood Pressure | (TC 06/02/2023 11:00:44 AM, Chukri, Samir) Completed PM performed by Getinge Certificate of Calibraion attached.
62639 | 194319 | PM | ACTIV | ISSUE | 2023-05-01 03:25 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
62971 | 194320 | PM | CMPLT | CMPLT | 2023-05-01 03:25 | 2023-09-01 10:09 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 09/01/2023 10:05:39 AM, Baghirzade, Nara) PM passed and completed. Replaced timer. Replaced main switch. No physical damage. @30.5cm distance between the surface and the light, the light intensity is: high = 30.7 uw/cm2/nm (30-35 uw/cm2/nm) low = 12.5 uw/cm2/nm (12-15uw/cm2/nm) Target light is good. Hours: 0.0 Cleaned the mesh filters and vacuumed the filters. Electrical safety: Power cord resistance = 48 mohms Current leakage = 50.4 uA
62980 | 194321 | PM | CANCL | CANCL | 2023-05-01 03:26 | 2024-07-25 11:40 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62981 | 194322 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-13 08:51 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 09/13/2023 08:29:47 AM, Campbell, Connor) PM complete/passed Performance verified and within ranges for intensity Earth Leakage: 56.7 uA Ground resistance: 53 mili Ohm replaced timer as old battery had run out replaced rubber feet
62983 | 194323 | PM | CANCL | CANCL | 2023-05-01 03:26 | 2024-09-13 10:07 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62987 | 194324 | PM | ACTIV | ISSUE | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62990 | 194325 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-07 13:02 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/07/2023 01:01:44 PM, Baghirzade, Nara) PM passed and completed. No physical damage. @12.0in distance between the surface and the light, the light intensity is: high = 30.4 uw/cm2/nm (30-35 uw/cm2/nm) low = 13..3 uw/cm2/nm (12-15uw/cm2/nm) Target light is good. Hours: 16860.4 Cleaned the mesh filters and vacuumed the filters. Electrical safety: Power cord resistance = 20.0 mohms Current leakage = 50.2 uA
63010 | 194326 | PM | CANCL | CANCL | 2023-05-01 03:26 | 2024-06-27 09:58 | PM-MAJOR - Annual PM For: Radiometer | It is for the old Phototherapy lights.
63026 | 194327 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-10-19 15:27 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/11/2023 08:37:47 AM, Rhuland, Courtney) Standby switch light not illuminating Bottom plate screw housing broken off and top plate bent Contacted Getinge to get quote for parts (TC 10/19/2023 03:24:28 PM, Rhuland, Courtney) Replaced main housing Replaced standby switch Repaired bottom plate screw housing Replaced power cord retainer PM complete Radiance - 48.9 uW/cm^2/nm Electrical safety - OK
63036 | 194328 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-12 15:22 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/12/2023 03:18:08 PM, Ferreira, Qwynn) PM Passed Hours: 6764 Avg Intensity: 51uW Electrical Safety - OK
63054 | 194329 | PM | ACTIV | ISSUE | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63175 | 194330 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-02 14:05 | PM-MAJOR - Annual PM For: Simulators, Cardiac, Electrocardiography | (TC 06/02/2023 12:05:26 PM, Chukri, Samir) Completed PM performed by Getinge . Certificate of Calibration attached
63275 | 194331 | PM | ACTIV | ISSUE | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
63447 | 194332 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-02 14:13 | PM-MAJOR-Annual PM for: Oscilloscope | (TC 06/02/2023 02:11:11 PM, Chukri, Samir) Completed PM Performed by Getinge . Certificate of Calibration attached
63726 | 194333 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-02 14:27 | PM-MAJOR - Annual PM For: Testers | (TC 06/02/2023 02:25:17 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate Of Calibration attached
63876 | 194334 | PM | CMPLT | PMFAIL | 2023-05-01 03:26 | 2023-08-03 11:50 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/03/2023 11:48:04 AM, Baghirzade, Nara) RH sensor failed during PM. Calibrated the RH sensor and measurement was not accurate. Re-calibrated the sensor and letting incubator rest and stabilize before I start the functional check again. If RH sensor fails again, will be replacing the RH sensor.
63931 | 194335 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-02 14:37 | PM-MAJOR-Annual -PM for : EEG simulator | (TC 06/02/2023 02:33:28 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached.
64920 | 194336 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-02 14:40 | PM-MAJOR - Annual PM For: multimeter | (TC 06/02/2023 02:38:52 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached
65102 | 194337 | PM | ACTIV | ISSUE | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
65148 | 194338 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-01-02 14:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 10/27/2023 02:09:04 PM, Reid, Emmerson ) Occlusion sensor seal is damaged and the unit needs to be sent out. (TC 10/31/2023 02:57:47 PM, Reid, Emmerson ) emailed ICU medical for RMA and created a PO for the repair. (TC 01/02/2024 02:12:30 PM, Reid, Emmerson ) Attached the pole mount and added the APS IV PCA Mar 2022. Attached is the repair report. PM Complete.
65151 | 194339 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65156 | 194340 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-30 09:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/30/2023 09:34:21 AM, Reid, Emmerson ) PM complete, needed to run the 20ml test 3 times before passing as it was over infusing. See proceedure for more details.
65162 | 194341 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-30 13:25 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/30/2023 01:24:51 PM, Reid, Emmerson ) PM complete, see proceedure for details.
65171 | 194342 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-29 09:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/29/2023 09:39:16 AM, Reid, Emmerson ) PM Complete, see proceedure for details.
65907 | 194343 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-08 07:59 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/04/2023 01:39:30 PM, Rhuland, Courtney) Front case replaced Mylar gasket replaced (TC 08/08/2023 07:58:24 AM, Rhuland, Courtney) Calibration complete Check-in PM complete All tests passed Refer to ASM for details
65921 | 194344 | PM | CANCL | CANCL | 2023-05-01 03:26 | 2023-06-09 14:49 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65963 | 194345 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-08 14:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/08/2023 02:24:04 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65967 | 194346 | PM | CMPLT | PMFAIL | 2023-05-01 03:26 | 2023-08-14 09:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/09/2023 09:23:26 AM, Rendon Lopez , Elizabeth) Rear case and handle cracked, replaced under recall. Fails clamp close position test
65973 | 194347 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-08 08:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/07/2023 02:13:39 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, pressure disk failed. ran pressure disk calibration and also failed. (TC 06/07/2023 03:17:53 PM, Reid, Emmerson ) Opened unit and replaced the pressure sensor cable. Ran ASM Calibration, all tests passed. Still need to complete PM. (TC 06/08/2023 08:09:58 AM, Reid, Emmerson ) Ran ASM PM, all tests passed. PM Complete.
65974 | 194348 | PM | CMPLT | PMFAIL | 2023-05-01 03:26 | 2023-09-26 14:17 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/26/2023 02:16:08 PM, Campbell, Connor) Channel error Pressure sensor not working at all front case standoff is cracked on the inside near the pressure sensor
65976 | 194349 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-19 14:30 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/19/2023 02:29:40 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
65983 | 194350 | PM | CMPLT | PMFAIL | 2023-05-01 03:26 | 2023-08-21 09:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/21/2023 09:45:35 AM, Baghirzade, Nara) Pressure disc accuracy verification failed during PM; checked pressure disc with analog sensor test and pressure is not being detected. Will replace the pressure sensor disc.
65996 | 194351 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-16 11:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 11:41:50 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
65999 | 194352 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66004 | 194353 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66005 | 194354 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-12 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/12/2023 02:25:24 PM, Rhuland, Courtney) Front and rear cases replaced (recall) Battery changed and conditioned PM complete All tests passed Refer to ASM for details
66010 | 194355 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-08 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/08/2024 02:25:01 PM, Reid, Emmerson ) Replaced the rear case. Downloaed and cleared log history. Ran ASM PM, keyboard test failed Replaced the front case Ran ASM PM, all tests passed. PM Complete
66027 | 194356 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-17 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/17/2023 10:36:47 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 48 mohms Leakage current: 78.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66030 | 194357 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-08 13:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/08/2023 01:18:40 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 81 mohms Leakage current: 44.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66033 | 194358 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-17 13:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/17/2023 01:19:14 PM, Rhuland, Courtney) Cracked front and back cases replaced (recall) Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
66039 | 194359 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-05 09:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/20/2023 11:46:10 AM, Rhuland, Courtney) Replaced broken rear case Waiting on power cord to complete PM (TC 07/05/2023 09:31:46 AM, Rhuland, Courtney) Power cord replaced Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
66050 | 194360 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66345 | 194361 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-02 14:51 | PM-MAJOR - Annual PM For: Testers | (TC 06/02/2023 02:49:24 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached
66453 | 194362 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-03 14:24 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 71 uA and 87ohm.
66682 | 194363 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-08-29 09:50 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 08/29/2024 09:50:11 AM, Rhuland, Courtney) PM completed Replaced battery Swapped with device that was faulty in NICU
66684 | 194364 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-16 13:11 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 06/16/2023 01:09:51 PM, Ferreira, Qwynn) Replaced battery Cleaned the area and cleaned the calbes up - unorganized One screw was stripped - replaced the screw once it was out
66705 | 194365 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-29 12:18 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/29/2024 12:10:28 PM, Rhuland, Courtney) PM completed Replaced porthole seal Replaced corner grommet Ref and Main voltage calibrated Scale calibrated 9V rechargeable battery replaced Electrical safety - OK Set 35C & 75% humidity Read 35.1C & 75.2% humidity
66709 | 194366 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-02 15:00 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/02/2023 02:52:24 PM, Rhuland, Courtney) Recall porthole and side door latches replaced PM completed Vref - 1.2316v Line voltage - 117.6v Scale calibrated All tests passed
66716 | 194367 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-04 09:50 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 04/04/2024 09:40:06 AM, Rhuland, Courtney) Reported device was turning on/off Replaced standby rocker switch average irradiance - 40.2 uW/cm^2nm Electrical safety - 136mOhms 0.9uA
66764 | 194368 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66839 | 194369 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | BP Check
66840 | 194370 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66848 | 194371 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | BP Check
66850 | 194372 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | BP Check
66862 | 194373 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | BP Check
66876 | 194374 | PM | ACTIV | ISSUE | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67631 | 194375 | PM | CANCL | DUP | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM every two years
67885 | 194376 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-26 08:29 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/20/2023 08:28:20 AM, Rendon Lopez , Elizabeth) Upgraded software, downloaded library. PM completed, all tests passed
67914 | 194377 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-31 09:09 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/31/2023 09:05:33 AM, Rhuland, Courtney) Hood replaced Side wall and porthole latches replaced (recall) Ran functional check, fan was loud Replaced fan and issue was resolved PM complete RH calibrated Scale calibrated LV and main voltage calibrated 9V battery replaced All tests passed Electrical safety - OK
67928 | 194378 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-25 10:57 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/25/2023 10:53:59 AM, Rhuland, Courtney) Side door and porthole latches replaced (recall) Missing grommet replaced PM complete Main voltage calibrated LV calibrated Scale calibrated 9V battery replaced Electrical safety - OK
67935 | 194379 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-29 14:51 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/29/2023 02:49:19 PM, Rhuland, Courtney) Hood replaced Porthole door replaced PM complete Scale calibrated LV and main voltage calibrated Humidity and scale check passed Electrical safety - OK
67936 | 194380 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-23 09:38 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/16/2023 09:03:14 AM, Rhuland, Courtney) Bed was tilting downwards, fixed loose screws on elevating platform with Samir Hood cracked, waiting for new hood (TC 05/23/2023 09:33:06 AM, Rhuland, Courtney) Hood replaced Porthole latches and side wall latches replaced (recall) Ventilator slot grommet replaced PM complete Humidifier probe calibrated All tests passed 9V battery replaced Electrical Safety - OK
67954 | 194381 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-29 09:14 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/28/2023 02:29:11 PM, Rhuland, Courtney) Hood replaced 3 porthole doors replaced Ventilator tube grommet replaced Porthole seal replaced Both doors were missing, replaced with spares Recall parts replaced LV and main voltage calibrated Scale calibrated and verified (TC 06/29/2023 09:12:46 AM, Ferreira, Qwynn) Ran functional test - OK Electrical Safety - OK Sent back to unit
67955 | 194382 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-12-13 12:55 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 12/13/2023 12:54:24 PM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced in 2022. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 81.2% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 85.1 uA Power cord resistance = 140 mohms
67958 | 194383 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-25 13:06 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/25/2023 01:04:11 PM, Baghirzade, Nara) Originally sent down for system failure 25, fan was not seated properly. Re-seated the fan. Replaced air probe cover. PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Greased the tilt as it was very dry and hard to move. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. RH sensor was calibrated. Set RH to 75%, measured on the analyzer 75%. Functional tests passed. Electrical safety results - Leakage current = 31.4 uA Power cord resistance = 50 mohms
67961 | 194384 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
68302 | 194385 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68371 | 194386 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-11 10:19 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 08/08/2023 10:17:03 AM, Scero, Tomlyn) All PM parts replaced. Oring seals replaced on the 2 gas outlets not used. (TC 08/09/2023 10:17:59 AM, Scero, Tomlyn) Continued with PM tests as per IPM-L (TC 08/10/2023 10:18:30 AM, Scero, Tomlyn) All PM test complete and passed. Results sheet attached for reference.
68390 | 194387 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-11 10:38 | PM Major - Delivery Module SEVO | (TC 08/11/2023 10:36:52 AM, Scero, Tomlyn) 3 year PM parts replaced and tests complete and passed. Shown on PM results sheet with 6 year PM for 68371.
68560 | 194388 | PM | ACTIV | READY | 2023-05-01 03:26 |  | PM Major - Delivery Module SEVO
68624 | 194389 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-25 13:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 01:22:14 PM, Canjimavadivel, Kamallesh ) Replaced battery PM tests pass
68634 | 194390 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-03-21 18:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68672 | 194391 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-07 09:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 08:58:59 AM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
68680 | 194392 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-10 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 10:08:32 AM, Cortez, John) Cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4082mAhs
68686 | 194393 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-03 13:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:19:18 PM, Rhuland, Courtney) PM completed by BD
68696 | 194394 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68702 | 194395 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-10-26 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/25/2023 01:51:36 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 94 mohms Leakage current: 74.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 10/26/2023 09:49:30 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3423 mamphr New battery capacity: 3963 mamphr PM completed.
68705 | 194396 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-14 09:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/08/2023 09:00:41 AM, Rendon Lopez , Elizabeth) Battery replaced, charging (TC 08/09/2023 09:00:41 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
68709 | 194397 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-18 13:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/18/2023 01:04:20 PM, Cortez, John) Kam cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4174 mAhs
68721 | 194398 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-12 09:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 09:22:12 AM, Rhuland, Courtney) Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
68747 | 194399 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-01 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/01/2023 02:46:56 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 38 mohms Leakage current: 71.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68771 | 194400 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-05 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:44:12 PM, Rhuland, Courtney) PM completed by BD
68782 | 194401 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68793 | 194402 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68810 | 194403 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-02 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 10:43:49 AM, Rhuland, Courtney) Battery replaced and conditoned PM complete All tests passed Refer to ASM for details
68823 | 194404 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68851 | 194405 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-12 14:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/12/2023 02:25:42 PM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioning battery overnight (TC 09/12/2023 09:09:20 AM, Cortez, John) 3911 mAhs
68867 | 194406 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-10 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 11:16:20 AM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4006 mAhs
68870 | 194407 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-10-23 08:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/13/2023 01:39:58 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 128 mohms Leakage current: 73.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 10/23/2023 08:28:30 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3515 mamphr New battery capacity: 3991 mamphr PM completed.
68874 | 194408 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68881 | 194409 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-08 08:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 11:14:41 AM, Campbell, Connor) 2023/09/07 PM complete, refer to ASM Battery conditioning started 2023/09/08 Battery conditioning finished - 3923 mAh old capacity - 4078 mAh new capacity
68898 | 194410 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68914 | 194411 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68936 | 194412 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-19 15:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 03:12:13 PM, Campbell, Connor) replaced battery plug prongs were a bit bent, bent them back straight for leakage test refer to procedure Pm complete/passed, refer to ASM
68937 | 194413 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-03 13:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:20:48 PM, Rhuland, Courtney) PM completed by BD
68940 | 194414 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-15 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 11:29:02 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 63 mohms Leakage current: 74.3 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/15/2023 02:23:25 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4053 mamphr PM completed.
68949 | 194415 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-12 09:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/12/2023 08:50:03 AM, Reid, Emmerson ) Ran Optimal battery conditioning, new battery capacity 4054 mamphr. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68962 | 194416 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-14 08:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/14/2023 08:13:46 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 174 mohms Leakage current: 77.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4072 mamphr PM completed.
68969 | 194417 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-03-25 10:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68992 | 194418 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69004 | 194419 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-21 09:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/21/2023 08:51:25 AM, Reid, Emmerson ) Ran optimal battery conditioning test, new battery capacity 4046 mamphr, Pass. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
69010 | 194420 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-05 10:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 03:10:26 PM, Reid, Emmerson ) Replaced the battery and left to charge (TC 09/05/2023 10:21:02 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69013 | 194421 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69024 | 194422 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-19 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/19/2023 10:15:11 AM, Reid, Emmerson ) Replaced the battery. Ran ASM MSH PM Pcu, all tests passed. PM complete.
69027 | 194423 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-15 09:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/15/2023 09:22:03 AM, Rhuland, Courtney) Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
69033 | 194424 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-03-25 11:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69038 | 194425 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-25 12:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/25/2023 12:14:03 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69040 | 194426 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-15 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/15/2023 01:11:01 PM, Rhuland, Courtney) Rate calibrated PM complete All tests passed post calibration Refer to ASM for details
69054 | 194427 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69060 | 194428 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-28 13:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 01:04:12 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69085 | 194429 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-27 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/27/2023 08:50:07 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69091 | 194430 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-04 08:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/04/2023 08:23:47 AM, Rhuland, Courtney) Device cleaned PM complete All tests passed Refer to ASM for details
69093 | 194431 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-23 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 11:23:47 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69142 | 194432 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-01 14:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/01/2023 02:16:04 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69166 | 194433 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-26 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69170 | 194434 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-08 13:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:47:16 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69172 | 194435 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-08 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/08/2023 02:58:09 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69173 | 194436 | PM | CMPLT | PMFAIL | 2023-05-01 03:26 | 2023-09-19 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 10:43:39 AM, Campbell, Connor) IUI male connector severly corroded bottle pressure sensor broken
69181 | 194437 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-06 13:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/06/2023 01:00:14 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69188 | 194438 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-13 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/13/2023 02:40:48 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69194 | 194439 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-19 15:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/19/2023 03:24:56 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69212 | 194440 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-20 11:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/20/2023 10:59:57 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
69214 | 194441 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69218 | 194442 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-01 08:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/01/2023 08:58:50 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69269 | 194443 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-26 10:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/26/2023 10:58:56 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69273 | 194444 | PM | CANCL | CANCL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69282 | 194445 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-01 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/29/2023 11:14:43 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
69290 | 194446 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-29 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/29/2023 02:14:05 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69293 | 194447 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-18 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/18/2023 09:48:47 AM, Cortez, John) Cleaned by Kam, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69313 | 194448 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-30 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 01:02:06 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
69315 | 194449 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-11-14 11:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/14/2023 11:28:59 AM, Campbell, Connor) Pm complete/passed, refer to ASM
69327 | 194450 | PM | CMPLT | PMFAIL | 2023-05-01 03:26 | 2023-10-20 12:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/20/2023 12:17:58 PM, Campbell, Connor) PM failed pressure calibration
69341 | 194451 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-30 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69358 | 194452 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-06-08 13:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:57:21 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69399 | 194453 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-12 16:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 04:48:29 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69419 | 194454 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69442 | 194455 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69447 | 194456 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-06 09:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/06/2023 08:27:38 AM, Reid, Emmerson ) Computer restarted in the middle of the PM. Downloaded and cleared log history Pump needed calibration, passed. ASM PM, pressure test failed again. Ran Pressure calibration, passed. Ran ASM PM, rate accuracy failed. Ran Rate calibration, passed. Ran ASM PM, all tests passed. PM Complete.
69451 | 194457 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-30 13:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69469 | 194458 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2024-04-23 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69475 | 194459 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69476 | 194460 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-01 08:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 08:56:21 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69501 | 194461 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-20 08:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/20/2023 08:06:33 AM, Campbell, Connor) PM complete/passed, refer to ASM
69507 | 194462 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-09-01 09:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 09:32:20 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69514 | 194463 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69519 | 194464 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69522 | 194465 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-10-24 09:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/24/2023 09:29:30 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69524 | 194466 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-28 14:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 01:59:56 PM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details (TC 07/28/2023 02:40:03 PM, Rhuland, Courtney) During cleaning, noticed crack on side light pannel Replaced front case Functional check - OK
69534 | 194467 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-13 17:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/13/2023 05:13:37 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69538 | 194468 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-08-17 09:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/17/2023 09:18:13 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69555 | 194469 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-05-26 10:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/26/2023 10:47:35 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
69557 | 194470 | PM | CNP | ENL | 2023-05-01 03:26 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69572 | 194471 | PM | CMPLT | CMPLT | 2023-05-01 03:26 | 2023-07-05 10:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/05/2023 10:11:35 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69574 | 194472 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69596 | 194473 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-08-08 09:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/08/2023 09:26:12 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69598 | 194474 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69611 | 194475 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 09:43:54 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
69619 | 194476 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69628 | 194477 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2024-04-04 12:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 12:48:30 PM, Rhuland, Courtney) PM completed by BD
69641 | 194478 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69659 | 194479 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
32520 | 194480 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 11:48 | MSH DR2 PM - Semi-Annual | (TC 06/08/2023 11:45:52 AM, Whittington, Scott) MSH DR2 PM - Semi-Annual | MSH DR2 PM - Semi-Annual
32520 | 194481 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-09 08:26 | MSH DR2 HARP MAY 15 TRGT - Semi Annual | (TC 05/09/2023 08:23:49 AM, Whittington, Scott) MSH DR2 HARP MAY 15 TRGT - Semi Annual HARP FOLLOWUP (TC 05/09/2023 08:24:18 AM, Whittington, Scott) MSH DR2 HARP MAY 15 TRGT - Semi Annual HARP (TC 05/09/2023 08:24:18 AM, Whittington, Scott) MSH DR2 HARP MAY 15 TRGT - Semi Annual HARP | MSH DR2 HARP MAY 15 TRGT - Semi Annual
32540 | 194482 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-12 19:35 | PM - Monthly | PM Major done Slice count - 602530 TOS phantom a(-997.7, 10.3) b(335.4, 14.1) c(134.2, 14.1) d(98.1, 14.7) e(-105.1, 13.4) f(1.7, 14.3) -System passed all tests
32560 | 194483 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-06 15:01 | PM - Monthly | PM completed Mechanical PM done SC-1203963 TOS phantom a(-992.0,10.0) b(332.2,14.3) c(133.0,13.8) d(100.5,13.2) e(-106.1,11.8) f(1.1,13.7) -passed all phantom testing
32580 | 194484 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 22:22 | PM - Monthly | PM completed -Mechanical PM done -Scan computer replaced/boot errors -SC-55355 TOS phantom a(-991.0,10.1) b(332.1,15.6) c(133.1,13.5) d(94.6,12.3) e(-107.4,12.0) f(0.6,13.9) -passed all phantom tests
501640 | 194485 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-07 11:11 | PM - Annual | (TC 07/07/2023 11:10:34 AM, Green, Paul) Not used and only used for backup.
1046457 | 194486 | PM | CMPLT | UNL | 2023-05-01 03:27 | 2023-06-09 15:29 | PM-MAJOR-Annual-PM for : Pressure meter | (TC 06/09/2023 03:26:25 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached.
1046612 | 194487 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1046625 | 194488 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-10-19 13:54 | Annual PM | (TC 10/19/2023 01:51:54 PM, Scero, Tomlyn) Annual PM test complete with ESU tester. All tests passed. LC 17 uA GWR 140mOhms
1046781 | 194489 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-03-20 16:13 | PM - Semi-Annual | PM
1046781 | 194490 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-03-17 16:15 | HARP - Semi Annual | HARP
1046789 | 194491 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | Annual PM | Laser end of support, no PM available. Laser will be replaced through Capital
1047038 | 194492 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047041 | 194493 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 15:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 03:11:42 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete. NOTE one segment on the display is fairly dim and should be replaced next PM.
1047042 | 194494 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-08-30 13:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 12:59:27 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete
1047043 | 194495 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-18 09:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/18/2023 09:26:30 AM, Cortez, John) Cleaned by Kam, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1047850 | 194496 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-11 15:08 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/11/2023 03:04:19 PM, Rhuland, Courtney) Error with barrel sizer when completing PM Opened device, potentiometer sizer was disengaged Reseated barrel sizer potentiometer Replaced left and right plunger case (spare) replaced plunger seal replaced top housing (spare) Sizer calibrated PM complete All tests passed Refer to procedure for details
1047860 | 194497 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe
1048377 | 194498 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM
1048522 | 194499 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2024-03-28 13:10 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/05/2023 02:05:42 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, pressure disk test failed. Ran pressure disk PM, also failed. Opened the pump to check the pressure sensorand the back case broken by the bottom screw. Waiting for back cases to come in.
1049204 | 194500 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-04 11:31 | PM-MAJOR - Annual PM For: Central Monitor | Cleaned during go live to L5
1049215 | 194501 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-04 11:31 | PM-MAJOR - Annual PM For: Central Monitor | Cleaned during go-live to L5
1049216 | 194502 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor | SSDU merged with ICU and no longer requires this PICIX. Currently PICIX is on storage cart on level 6 and not used. Cancelled PM as it is not being used.
1049217 | 194503 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-27 13:52 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/27/2023 01:51:33 PM, Scero, Tomlyn) PM complete during patch reboot
1049218 | 194504 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-25 13:49 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/25/2023 01:48:30 PM, Scero, Tomlyn) PM complete. Patch upgrade down, rebooted machine.
1049221 | 194505 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-27 13:51 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/27/2023 01:51:06 PM, Scero, Tomlyn) PM complete during patch reboot
1049222 | 194506 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-25 13:48 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/25/2023 01:47:02 PM, Scero, Tomlyn) PM complete during patch upgrade. Central rebooted.
1049410 | 194507 | PM | CANCL | CANCL | 2023-05-01 03:27 | 2024-07-25 07:32 | PM-MAJOR - Annual PM For: Radiometer
1049901 | 194508 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 15:38 | MSH R61 HARP JUNE 12 TRGT - Semi Annual |  | MSH R61 HARP JUNE 12 TRGT - Semi Annual HARP completed.
1049901 | 194509 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 12:33 | PM | (TC 06/13/2023 12:32:37 PM, Chau, Ken) PM okay -left handle cover requires replacement
1049902 | 194510 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 11:57 | HARP - Semi Annual | (TC 06/07/2023 11:56:40 AM, Chau, Ken) HARP Completed
1049902 | 194511 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 11:59 | PM | (TC 06/07/2023 11:57:30 AM, Chau, Ken) PM completed -replaced battery charger prongs -system battery was at 50%, charged and tested okay -if system battery does not hold charge, requires replacement
1049903 | 194512 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-05 13:18 | HARP - Semi Annual | (TC 06/05/2023 01:17:21 PM, Fan, Tak) HARP Completed.
1049903 | 194513 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-05 13:17 | PM | (TC 06/05/2023 11:13:00 AM, Fan, Tak) PM completed. Replaced boom dock cover and wear pads.
1049904 | 194514 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 15:36 | HARP - Semi Annual |  | HARP completed.
1049904 | 194515 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 12:34 | PM | (TC 06/13/2023 12:33:38 PM, Chau, Ken) PM okay
1049905 | 194516 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 14:49 | HARP - Semi Annual |  | Harp completed.
1049905 | 194517 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 14:50 | PM |  | PM Minor
1050258 | 194518 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-15 09:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/15/2023 09:17:38 AM, Rhuland, Courtney) Sensor broken high failure error present Confirmed with analog sensor test and bottle side pressure sensor stayed high Bottle side pressure sensor replaced Pressure and rate calibrated PM complete All tests passed Refer to ASM for details
1050263 | 194519 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-19 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/19/2023 11:23:19 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1050271 | 194520 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2024-04-23 09:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:34:44 AM, Huang, Jessica) replaced the female iui connector completed PM
1050617 | 194521 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-09-22 10:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/22/2023 10:46:21 AM, Campbell, Connor) PM complete/passed, refer to ASM
1051356 | 194522 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
1051613 | 194523 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-02 15:13 | PM-MAJOR - Annual PM For: Testers | (TC 06/02/2023 03:05:36 PM, Chukri, Samir) Completed PM Performed by Getinge. Certificate of Calibration attached
1051830 | 194524 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PLANNED MAINTENANCE-VACUUM PC
1051831 | 194525 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor
1051840 | 194526 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-08-10 09:39 | PM MAJOR - PM For Central Monitor PC | (TC 08/10/2023 09:38:17 AM, Reid, Emmerson ) Updated the system
1053213 | 194527 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-03 08:50 | PM-MAJOR - Annual PM For: Testers | (TC 06/03/2023 08:48:46 AM, Chukri, Samir) Completed PM Performed by Getinge. Certifiate of Calibration attached .
1053439 | 194528 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 13:35 | PM-MAJOR - Annual PM For: Simulators, Multiparameter | (TC 06/09/2023 01:32:41 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached
1053440 | 194529 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 13:37 | PM-MAJOR - Annual PM For: Testers, Pulse Oximetry Monitor | (TC 06/09/2023 01:35:47 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached
1053441 | 194530 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 13:40 | PM-MAJOR - Annual PM For: Testers, Pulse Oximetry Monitor | (TC 06/09/2023 01:38:13 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached.
1053475 | 194531 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 14:23 | Annual PM | (TC 06/09/2023 02:19:09 PM, Chukri, Samir) Completed PM Performed by Getinge. Certificate of Calibration attached.
1054591 | 194532 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-15 09:34 | PM MAJOR-Annual Philips Monitoring, Central Station | (TC 06/15/2023 09:34:23 AM, Samanian, Kamran) Serviced
1054599 | 194533 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:45 | PM MAJOR-Annual Philips Monitoring, Central Station | (TC 06/13/2023 09:44:04 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1055110 | 194534 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-08-14 09:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/08/2023 09:02:56 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed
1055117 | 194535 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055124 | 194536 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2024-03-20 15:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055155 | 194537 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2024-04-26 10:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055171 | 194538 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055711 | 194539 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
1055712 | 194540 | PM | CANCL | CANCL | 2023-05-01 03:27 | 2024-08-08 10:24 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
1055721 | 194541 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055730 | 194542 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055731 | 194543 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055732 | 194544 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055733 | 194545 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055734 | 194546 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055735 | 194547 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055736 | 194548 | PM | CNP | ENL | 2023-05-01 03:27 |  | PM ACTIVE-MICROSTREAM EXTENSION
1055743 | 194549 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-10-27 13:13 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 10/27/2023 01:11:49 PM, Rhuland, Courtney) Replaced rear case (recall) PM complete All tests passed Refer to ASM for details
1055746 | 194550 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-14 12:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/14/2023 12:31:56 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1055774 | 194551 | PM | CMPLT | PMFAIL | 2023-05-01 03:27 | 2023-06-14 15:28 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/14/2023 03:27:03 PM, Cortez, John) Discovered faulty speaker during PM Will call tech support
1055793 | 194552 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-02 16:04 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/02/2023 04:03:47 PM, Cortez, John) PM passed, see procedure for results
1055794 | 194553 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 15:39 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/07/2023 03:34:10 PM, Cortez, John) PM passed, see procedure for results
1055795 | 194554 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-05 09:42 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/01/2023 09:36:31 AM, Cortez, John) PM passed, see procedure for results
1055796 | 194555 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 16:49 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/06/2023 04:46:07 PM, Cortez, John) PM passed, see procedure for results
1055797 | 194556 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-02 15:57 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/02/2023 03:56:05 PM, Cortez, John) PM passed, see procedure for results
1055798 | 194557 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-14 15:08 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/14/2023 03:07:25 PM, Cortez, John) PM passed, see procedure for results battery swapped with conditioned spare
1055800 | 194558 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-09-27 10:34 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 09/27/2023 10:31:07 AM, Scero, Tomlyn) All PM tests complete and passed.
1055803 | 194559 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 15:20 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/13/2023 03:19:08 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure for results
1055804 | 194560 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 14:02 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/08/2023 02:01:51 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure for results
1055805 | 194561 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 15:00 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/09/2023 02:59:14 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure for results
1055806 | 194562 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 15:28 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/07/2023 03:27:28 PM, Cortez, John) PM passed, see procedure for results
1055807 | 194563 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 13:06 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/13/2023 01:04:35 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure for results
1055808 | 194564 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:37 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/13/2023 09:36:54 AM, Canjimavadivel, Kamallesh ) PM complete, see procedure for results
1055810 | 194565 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 09:09 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/05/2023 09:06:32 AM, Cortez, John) Started PM (TC 06/06/2023 09:08:31 AM, Cortez, John) PM passed, see procedure for results
1055811 | 194566 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-01 09:26 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 05/31/2023 09:24:52 AM, Cortez, John) PM passed, see procedure for results
1055812 | 194567 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-01 09:16 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 05/31/2023 09:14:49 AM, Cortez, John) PM passed, see procedure for results
1056037 | 194568 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 14:29 | PM-MAJOR-Annual-PM for : Tester | (TC 06/09/2023 02:25:03 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached
1056190 | 194569 | PM | CMPLT | PMFAIL | 2023-05-01 03:27 | 2023-06-06 15:08 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 06/06/2023 03:06:31 PM, Canjimavadivel, Kamallesh ) Patches installion failed. PC auto shutdown prior to completion.
1057302 | 194570 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-14 08:58 | PM - Annual | (TC 06/14/2023 08:55:40 AM, Fan, Tak) PM performed. Gnd R = 110 mohms Max A = 2.1 / 79.3 / 158.8 uA C2-9-D (359931YP3) = 0.5 / 2.5 uA RM6C (238232KR3) = 0.6 / 4.3 uA C1-6-D (363058YP9) = 0.6 / 2.8 uA
1057699 | 194571 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-15 10:04 | Battery Cycling | (TC 06/15/2023 10:03:02 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
1057700 | 194572 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-30 07:34 | Battery Cycling | (TC 05/30/2023 07:32:58 AM, Canjimavadivel, Kamallesh ) Recondition cycle complete
1057710 | 194573 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-31 10:40 | Battery Cycling | (TC 05/31/2023 10:40:04 AM, Canjimavadivel, Kamallesh ) Battery cycle complete
1057711 | 194574 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 08:48 | Battery Cycling | (TC 06/09/2023 08:45:59 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
1057713 | 194575 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 08:29 | Battery Cycling | (TC 06/08/2023 08:28:17 AM, Canjimavadivel, Kamallesh ) Recondition complete, battery capacity test pass
1057714 | 194576 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 10:15 | Battery Cycling | (TC 06/05/2023 10:14:50 AM, Cortez, John) Reconditioned and swapped battery
1057715 | 194577 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 10:13 | Battery Cycling | (TC 06/05/2023 10:12:34 AM, Cortez, John) Reconditioned and swapped battery
1057716 | 194578 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-14 08:40 | Battery Cycling | (TC 06/14/2023 08:39:35 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
1057717 | 194579 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 08:29 | Battery Cycling | (TC 06/08/2023 08:29:09 AM, Canjimavadivel, Kamallesh ) Recondition complete, battery capacity test pass
1057718 | 194580 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 08:28 | Battery Cycling | (TC 06/08/2023 08:27:30 AM, Canjimavadivel, Kamallesh ) Recondition complete, battery capacity test pass
1057719 | 194581 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-02 12:33 | Battery Cycling | (TC 06/02/2023 12:31:24 PM, Canjimavadivel, Kamallesh ) Reconditioning cycle complete, battery test pass
1057720 | 194582 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-31 10:32 | Battery Cycling | (TC 05/31/2023 10:30:57 AM, Canjimavadivel, Kamallesh ) Battery cycle complete
1057721 | 194583 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 08:22 | Battery Cycling | (TC 06/13/2023 08:21:21 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
1057722 | 194584 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-31 10:39 | Battery Cycling | (TC 05/31/2023 10:38:30 AM, Canjimavadivel, Kamallesh ) Battery cycle complete
1057723 | 194585 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-31 10:44 | Battery Cycling | (TC 05/31/2023 10:43:29 AM, Canjimavadivel, Kamallesh ) Battery cycle complete
1057724 | 194586 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 08:27 | Battery Cycling | (TC 06/08/2023 08:26:25 AM, Canjimavadivel, Kamallesh ) Recondition complete, battery capacity test pass
1057726 | 194587 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 08:40 | Battery Cycling | (TC 06/09/2023 08:39:31 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
1057727 | 194588 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-14 08:39 | Battery Cycling | (TC 06/14/2023 08:37:33 AM, Canjimavadivel, Kamallesh ) Recondition complete, capactiy 3/5 LEDs
1057730 | 194589 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 08:22 | Battery Cycling | (TC 06/09/2023 08:20:52 AM, Reid, Emmerson ) Battery Cycling complete. Batery capacity shows 3/5 LED lights.
1057732 | 194590 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 14:00 | Battery Cycling | (TC 06/08/2023 01:59:36 PM, Reid, Emmerson ) Battery cycle done, 3/5 lights shown for capacity.
1057733 | 194591 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 08:44 | Battery Cycling | (TC 06/09/2023 08:29:43 AM, Reid, Emmerson ) Battery Cycling complete. Batery capacity shows 4/5 LED lights.
1057734 | 194592 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 13:57 | Battery Cycling | (TC 06/08/2023 01:54:45 PM, Reid, Emmerson ) Battery cycle done, 4/5 lights shown for capacity.
1057735 | 194593 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 08:24 | Battery Cycling | (TC 06/09/2023 08:22:58 AM, Reid, Emmerson ) Battery Cycling complete. Batery capacity shows 3/5 LED lights.
1057736 | 194594 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-12 10:19 | Battery Cycling | (TC 06/12/2023 10:18:13 AM, Reid, Emmerson ) Batery cycleing complete. Full Capacity shows 4/5 LEDs.
1057737 | 194595 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | Battery Cycling
1057738 | 194596 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-08 13:58 | Battery Cycling | (TC 06/08/2023 01:58:01 PM, Reid, Emmerson ) Battery cycle done, 4/5 lights shown for capacity.
1057739 | 194597 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-09-18 09:14 | Battery Cycling | (TC 09/11/2023 09:13:08 AM, Rendon Lopez , Elizabeth) battery cycling - passed. Battery with defib 1055776
1057740 | 194598 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-12 10:17 | Battery Cycling | (TC 06/12/2023 10:16:32 AM, Reid, Emmerson ) Batery cycleing complete. Full Capacity shows 3/5 LEDs.
1057758 | 194599 | PM | CMPLT | PMFAIL | 2023-05-01 03:27 | 2023-06-06 15:03 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 06/06/2023 03:02:14 PM, Canjimavadivel, Kamallesh ) Installed patches failed. PC auto shutdown prior to complete
1057760 | 194600 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-06 15:01 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 06/06/2023 02:59:46 PM, Canjimavadivel, Kamallesh ) Patches installed successfully and vacuumed PC
1057761 | 194601 | PM | CMPLT | PMFAIL | 2023-05-01 03:27 | 2023-06-26 14:28 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 06/26/2023 02:25:04 PM, Canjimavadivel, Kamallesh ) Cleaned PC and restarted PC to complete OS patch install. Host keeps shutting down on its own
1057837 | 194602 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-16 14:39 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 05/16/2023 02:34:00 PM, Reid, Emmerson ) Electrical Safety Ground resistance 123 mohms Leakage current 16.4 uA PM Complete, see attched PM Report.
1057839 | 194603 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-10-30 09:15 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 10/30/2023 09:11:55 AM, Reid, Emmerson ) Completed PM with Connor, see attached PM Checklist
1058188 | 194604 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-21 14:38 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/21/2023 02:37:29 PM, Reid, Emmerson ) PM Complete, see proceedure for details.
1058224 | 194605 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-09-07 09:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 09:28:38 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
1058239 | 194606 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2024-03-15 15:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058250 | 194607 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-26 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/26/2023 10:38:48 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
1058278 | 194608 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059330 | 194609 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:45 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:45:04 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059530 | 194610 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-21 10:35 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/21/2023 10:33:41 AM, Reid, Emmerson ) Updated PiC, removed internal USB and rebuilt the PC. Configured back to default. PM Complete.
1059531 | 194611 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-20 11:22 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/20/2023 11:20:24 AM, Reid, Emmerson ) Update was pushed. Shut down PC and it started installing. Vacuumed the PC while it was off. Turned PC back on and it updated successfully. Ran comand prompts and checked updates were successful. PM Complete.
1059532 | 194612 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059533 | 194613 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-08-10 09:38 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 08/10/2023 09:37:12 AM, Reid, Emmerson ) Updated the system
1059534 | 194614 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059541 | 194615 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:49 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:48:44 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059542 | 194616 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:49 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:49:14 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059543 | 194617 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor | Cancelling - used previous workorder
1059544 | 194618 | PM | CNP | EIU | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor
1059545 | 194619 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:51 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:50:34 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059546 | 194620 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:51 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:51:03 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059547 | 194621 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:52 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:51:32 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059548 | 194622 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:52 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:52:01 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059549 | 194623 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor
1059550 | 194624 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor
1059551 | 194625 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-07-26 09:23 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/25/2023 09:22:21 AM, Rendon Lopez , Elizabeth) Patch installation completed. Vacuum centra, PM completed
1059552 | 194626 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor | Cancelling - used previous workorder
1059553 | 194627 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor | Canclling - used previous workorder
1059554 | 194628 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:56 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:55:20 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059555 | 194629 | PM | CANCL | CANCL | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor | cancelling - used previous workorder
1059556 | 194630 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 09:57 | PM-MAJOR - Annual PM For: Central Monitor | (TC 06/13/2023 09:56:42 AM, Ferreira, Qwynn) Cleaned the grills and opened side panel and vacuumed inside.
1059557 | 194631 | PM | CNP | EIU | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor
1059877 | 194632 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-08-14 10:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/14/2023 10:26:53 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
1060155 | 194633 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-02 12:34 | Battery Cycling | (TC 06/02/2023 12:33:57 PM, Canjimavadivel, Kamallesh ) Reconditioning test complete, battery test pass
1060156 | 194634 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-14 08:37 | Battery Cycling | (TC 06/14/2023 08:34:35 AM, Canjimavadivel, Kamallesh ) Reconditioning complete, capacity 3/5 LEDs
1060157 | 194635 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-13 08:19 | Battery Cycling | (TC 06/13/2023 08:18:41 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
1060195 | 194636 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-30 07:35 | Battery Cycling | (TC 05/30/2023 07:34:50 AM, Canjimavadivel, Kamallesh ) Recondition cycle complete
1060198 | 194637 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-30 07:34 | Battery Cycling | (TC 05/30/2023 07:34:06 AM, Canjimavadivel, Kamallesh ) Recondition cycle complete
1060336 | 194638 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 08:42 | Battery Cycling | (TC 06/09/2023 08:41:04 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 4/5 LEDs
1060337 | 194639 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-07 10:07 | Battery Cycling | (TC 06/05/2023 10:05:29 AM, Cortez, John) Reconditioned and swapped battery
1060338 | 194640 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-15 10:03 | Battery Cycling | (TC 06/15/2023 10:01:44 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 4/5 LEDs
1060579 | 194641 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-02 10:37 | PM-MAJOR - Annual PM For: multimeter | (TC 06/02/2023 10:37:51 AM, Chukri, Samir) PM completed performed by Getinge Certificate for Calibration attached
1060609 | 194642 | PM | CANCL | CANCL | 2023-05-01 03:27 | 2024-07-05 11:02 | PM-MAJOR-Annual PM for ESU Tester
1060901 | 194643 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-05-30 07:36 | Battery Cycling | (TC 05/30/2023 07:35:49 AM, Canjimavadivel, Kamallesh ) Recondition cycle complete
1060902 | 194644 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-09 08:38 | Battery Cycling | (TC 06/09/2023 08:37:34 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 4/5 LEDs
1061254 | 194645 | PM | CMPLT | CMPLT | 2023-05-01 03:27 | 2023-06-14 08:02 | Battery Cycling | (TC 06/14/2023 08:00:36 AM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
1063358 | 194646 | PM | ACTIV | ISSUE | 2023-05-01 03:27 |  | PM-MAJOR - Annual PM For: Central Monitor
1064418 | 194647 | PM | CMPLT | CMPLT | 2023-05-01 03:28 | 2024-01-30 14:30 | PM-MAJOR-Annual PM for:MR-Diagnostics Incubator | (TC 01/23/2024 11:20:09 AM, Chukri, Samir) working with Michael from AME to prepair the incubator for the PM service (TC 01/23/2024 12:40:06 PM, Chukri, Samir) Picked up the incubator after cleaned and delivered to 18th floor Biomed work shop. (TC 01/24/2024 08:48:38 AM, Chukri, Samir) went to check the repaired MRI incubator was in use, check the parts left by Michael. whaiting the repair report to close the PM wo. (TC 01/30/2024 02:25:08 PM, Chukri, Samir) completed PM by Vendor. please see attached.
1064496 | 194648 | PM | CMPLT | CMPLT | 2023-05-01 03:28 | 2023-09-18 09:21 | Battery Cycling | (TC 09/14/2023 09:20:16 AM, Rendon Lopez , Elizabeth) Battery cycled - OK
1064497 | 194649 | PM | CMPLT | CMPLT | 2023-05-01 03:28 | 2023-06-12 10:20 | Battery Cycling | (TC 06/12/2023 10:19:29 AM, Reid, Emmerson ) Batery cycleing complete. Full Capacity shows 5/5 LEDs.
1066939 | 194650 | PM | CMPLT | CMPLT | 2023-05-01 03:28 | 2023-07-28 09:53 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/28/2023 09:51:27 AM, Canjimavadivel, Kamallesh ) Rebooted and successfully installed patches.
| 194651 | PA | CMPLT | CMPLT | 2023-05-01 08:30 | 2023-05-09 07:09 | Ordering for biomed stock | (TC 05/01/2023 08:32:38 AM, Baghirzade, Nara) SC 1002944929 (TC 05/09/2023 07:08:58 AM, Baghirzade, Nara) PO 7202206133
| 194652 | PA | CNP | ENL | 2023-05-01 08:41 |  | Ordering sterile water for altrix cleaning NICU | (TC 05/01/2023 08:45:22 AM, Baghirzade, Nara) SC 1002944955 (TC 05/08/2023 03:50:23 PM, Baghirzade, Nara) Order was cancelled as 1000mL sterile water is out of stock; Iqbal provided 48x500mL sterile water for NICU altrix cleaning. Received 48x500mL water
| 194653 | PA | CMPLT | CMPLT | 2023-05-01 08:47 | 2023-05-17 13:09 | X3 bezels for stock | (TC 05/01/2023 08:51:11 AM, Canjimavadivel, Kamallesh ) Placed order for stock parts (TC 05/17/2023 01:08:27 PM, Canjimavadivel, Kamallesh ) Received parts
1066940 | 194654 | ACC | CMPLT | CMPLT | 2023-05-01 09:03 | 2023-05-11 17:12 | Incoming Inspection | (TC 05/08/2023 05:10:23 PM, Cortez, John) Electrical safety passed, 100 mohms, 44 uAmps (TC 05/08/2023 05:10:23 PM, Cortez, John) see WO#194659
67122 | 194655 | FC | CMPLT | CMPLT | 2023-05-01 09:06 | 2023-05-01 09:16 | Functional check | (TC 05/01/2023 09:14:09 AM, Baghirzade, Nara) Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed. Replaced SpO2 trunk cable and temperature probe with spare parts. No charge.
67131 | 194656 | FC | CMPLT | CMPLT | 2023-05-01 09:16 | 2023-05-01 09:17 | Functional check | (TC 05/01/2023 09:16:41 AM, Baghirzade, Nara) Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
| 194657 | CM | CMPLT | CMPLT | 2023-05-01 09:45 | 2023-05-01 09:45 | upgraded 4 x X2s from Rev J to Rev M | (TC 05/01/2023 09:45:04 AM, Samanian, Kamran) done!
67123 | 194658 | FC | CMPLT | CMPLT | 2023-05-01 10:23 | 2023-05-01 10:23 | Functional check/ replacement of accessories | (TC 05/01/2023 10:22:51 AM, Baghirzade, Nara) Battery ok. All functional tests passed. Replaced SpO2 trunk cable, SpO2 finger sensor and temperature probe with spare parts. No charge.
1066941 | 194659 | ACC | CMPLT | CMPLT | 2023-05-01 10:28 | 2023-05-11 17:05 | Incoming Inspection | (TC 05/08/2023 05:03:21 PM, Cortez, John) Electrical safety passed, 100 mohms, 44 uAmps
| 194660 | ACC | CMPLT | CMPLT | 2023-05-01 12:28 | 2023-05-01 12:29 | MSH SURGICAL SKILLS SONOSITE PX ACCEPTANCE | (TC 04/28/2023 05:28:00 PM, Whittington, Scott) MSH SURGICAL SKILLS SONOSITE PX ACCEPTANCE | MSH SURGICAL SKILLS SONOSITE PX ACCEPTANCE
1059938 | 194661 | CM | CMPLT | CMPLT | 2023-05-01 12:44 | 2023-05-01 12:45 | RA lead pin broken | (TC 05/01/2023 12:44:21 PM, Canjimavadivel, Kamallesh ) Replaced RA lead, tested okay with simulator
501585 | 194662 | CM | CMPLT | CMPLT | 2023-05-01 12:49 | 2023-05-01 12:50 | MSH OPG SPP C2 WON'T START | (TC 05/01/2023 12:49:53 PM, Whittington, Scott) MSH OPG SPP C2 WON'T START started (TC 05/01/2023 12:49:53 PM, Whittington, Scott) MSH OPG SPP C2 WON'T START started | MSH OPG SPP C2 WON'T START
| 194663 | CM | CMPLT | CMPLT | 2023-05-01 14:09 | 2023-05-01 14:10 | AED cabinet 20th floor hallway chirping | (TC 05/01/2023 02:08:58 PM, Scero, Tomlyn) changed the 9V battery in the AED cabinet. 20th floor hallway
1054777 | 194664 | CM | CMPLT | CMPLT | 2023-05-01 14:46 | 2023-05-01 14:48 | Repurposed for PFT | (TC 05/01/2023 02:46:22 PM, Baghirzade, Nara) Repurposed VSM for PFT lab. Confirmed with Kamran. All functional tests passed. Dropped it off with Gaby. Informed her that any parts and accessories is their responsibility to pay for.
65362 | 194665 | FC | CMPLT | CMPLT | 2023-05-01 14:48 | 2023-05-01 14:49 | Functional check | (TC 05/01/2023 02:48:14 PM, Baghirzade, Nara) Battery ok. All functional tests passed. Replaced SpO2 trunk cable and temperature probe with spare parts. No charge. Replaced BP cuff - charged.
1059778 | 194666 | FC | CMPLT | CMPLT | 2023-05-01 14:53 | 2023-05-01 14:56 | QC checklist | (TC 05/01/2023 02:53:25 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1061364 | 194667 | CM | CMPLT | CMPLT | 2023-05-01 16:53 | 2023-05-01 16:52 | One of teh wheel is broken needs replacement. | Replaced the wheel. System fully functional.
1059795 | 194668 | CM | CMPLT | CMPLT | 2023-05-01 16:56 | 2023-05-01 16:55 | Preparation for upgrading the system to a newer software version. | Collected information from both the Zenitions for the upgrade.
| 194669 | CM | CMPLT | CMPLT | 2023-05-01 17:16 | 2023-05-01 17:16 | Site Coverage.
| 194670 | PA | CMPLT | CMPLT | 2023-05-02 07:36 | 2023-05-12 22:08 | Brackets for humidifier on servo | (TC 05/01/2023 03:00:00 PM, Canjimavadivel, Kamallesh ) Requested quote from vendor (TC 05/02/2023 07:43:07 AM, Canjimavadivel, Kamallesh ) Placed order for parts
1066945 | 194671 | ACC | CMPLT | CMPLT | 2023-05-02 09:42 | 2023-05-02 09:46 | Incoming Inspection | (TC 05/02/2023 09:41:43 AM, Cortez, John) Unboxed and assembled, Functional test passed- test load 30kg, reading 30kg
| 194672 | CM | CMPLT | CMPLT | 2023-05-02 09:52 | 2023-05-11 19:40 | capturing HL7 message for alarming in Biomed lab | (TC 05/02/2023 09:52:24 AM, Samanian, Kamran) captured and emailed to ASCOM team
| 194673 | CM | CMPLT | CMPLT | 2023-05-02 10:30 | 2023-05-02 10:31 | Chest lead long needed replacing | (TC 05/02/2023 10:30:16 AM, Scero, Tomlyn) ECG tech came down and needed a replacement chest lead long.
1063808 | 194674 | CM | CMPLT | CMPLT | 2023-05-02 11:11 | 2023-07-11 07:39 | MSH 15 SPP GE VOLUSON SETUP IN OBU OR1 TO SEE IF IT PICKED UP ELECTRICAL/ENVIRONMENTA EMF THAT MAY HAVE BEEN AFFECTING IU22 SYSTEMS | (TC 05/02/2023 11:10:52 AM, Whittington, Scott) MSH 15 SPP GE VOLUSON SETUP IN OBU OR1 TO SEE IF IT PICKED UP ELECTRICAL/ENVIRONMENTA EMF THAT MAY HAVE BEEN AFFECTING IU22 SYSTEMS AT THE TIME OF TEST MORNING OF MAY 1, NO INTERERENCE DETECTED ON FREE AIR SCAN | MSH 15 SPP GE VOLUSON SETUP IN OBU OR1 TO SEE IF IT PICKED UP ELECTRICAL/ENVIRONMENTA EMF THAT MAY HAVE BEEN AFFECTING IU22 SYSTEMS AT THE TIME OF TEST MORNING OF MAY 1, NO INTERERENCE DETECTED ON FREE AIR SCAN
1049218 | 194675 | CM | CMPLT | CMPLT | 2023-05-02 12:09 | 2023-05-03 09:54 | slow to load patient review | (TC 05/02/2023 09:53:04 AM, Scero, Tomlyn) nurse complaining that the CCU overview was very slow to load patient review alarms. I restarted the application and it speeded things up a bit.
1055801 | 194676 | CM | CMPLT | CMPLT | 2023-05-02 12:09 | 2023-05-03 09:53 | printer jammed | (TC 05/02/2023 09:52:12 AM, Scero, Tomlyn) printer paper jammed in printer. replaced paper and then confirmed that printer working ok.
1055800 | 194677 | CM | CMPLT | CMPLT | 2023-05-02 12:09 | 2023-05-03 09:52 | printing but no trace | (TC 05/02/2023 09:51:33 AM, Scero, Tomlyn) Printer paper in upside down
| 194678 | CM | ACTIV | ISSUE | 2023-05-02 12:35 |  | adeel test for weekly procedures
| 194679 | CM | CMPLT | CMPLT | 2023-05-02 16:19 | 2023-05-02 16:17 | MSH Emerg CT computer Syngo client application fault | MSH Emerg CT computer Syngo client application fault Service contacted PACS application support/reloaded/tested software Syngo tested OK
| 194680 | CM | CMPLT | CMPLT | 2023-05-03 09:00 | 2023-05-23 13:03 | Ordering exam lamp socket for IVF | (TC 05/23/2023 01:03:20 PM, Samanian, Kamran) received
1062926 | 194681 | CM | CMPLT | CMPLT | 2023-05-03 09:33 | 2023-11-16 15:14 | not working
| 194682 | CM | CMPLT | CMPLT | 2023-05-03 10:15 | 2023-05-03 10:15 | online focal point training | (TC 05/03/2023 10:14:50 AM, Samanian, Kamran) taken and completed
67970 | 194683 | CM | CMPLT | CMPLT | 2023-05-03 10:26 | 2023-05-03 10:28 | Broken plate under mattress; functional check | (TC 05/03/2023 10:27:03 AM, Baghirzade, Nara) Replaced plate under mattress. Functional tests passed and completed. Scale is functional. Tilt function is good. Set temp to 35C and humidity to 75%; measured 35.1C and humidity 77%. Incubator is fully functional.
1059709 | 194684 | CM | CMPLT | NFF | 2023-05-03 10:47 | 2023-05-03 10:48 | "not working" | (TC 05/02/2023 10:47:15 AM, Cortez, John) Visual/functional inspection passed. Unable to find issue with cart
1054975 | 194685 | CM | CMPLT | NFF | 2023-05-03 10:58 | 2023-05-03 11:00 | "Not working" | (TC 05/03/2023 10:58:12 AM, Rhuland, Courtney) Came down stating it was not working No fault found Functional check ran, all tests passed
1054995 | 194686 | CM | CMPLT | CMPLT | 2023-05-03 11:00 | 2023-05-03 11:01 | Not working | (TC 05/03/2023 11:00:19 AM, Rhuland, Courtney) Came down stating not working While doing functional check, noticed probe well was missing sensor Thermometer probe well replaced with spare Functional check complete, all tests passed
501585 | 194687 | CM | CMPLT | CMPLT | 2023-05-03 13:10 | 2023-07-11 07:39 | MSH OPG SPP C2 TCS ERASE CALIPERS VIRTUAL KEY NOT OPERATING: TCS FAILING? | (TC 05/03/2023 01:10:27 PM, Whittington, Scott) MSH OPG SPP C2 TCS ERASE CALIPERS VIRTUAL KEY NOT OPERATING: TCS FAILING? (TC 05/03/2023 01:11:53 PM, Whittington, Scott) MSH OPG SPP C2 TCS ERASE CALIPERS VIRTUAL KEY NOT OPERATING: TCS FAILING? (TC 05/04/2023 12:41:50 PM, Whittington, Scott) MSH OPG SPP C2 TCS ERASE CALIPERS VIRTUAL KEY NOT OPERATING: TCS FAILING? peating correctly at present (TC 05/03/2023 05:11:00 PM, Whittington, Scott) MSH OPG SPP C2 TCS ERASE CALIPERS VIRTUAL KEY NOT OPERATING: TCS FAILING? | MSH OPG SPP C2 TCS ERASE CALIPERS VIRTUAL KEY NOT OPERATING: TCS FAILING?
| 194688 | CM | CMPLT | CMPLT | 2023-05-03 14:46 | 2023-05-03 14:46 | STRYKER DICOM research images transferred to USB | STRYKER DICOM research images transferred to USB
500851 | 194689 | CM | CMPLT | CMPLT | 2023-05-03 16:38 | 2023-05-03 16:38 | MSH Aplio XG - System Disassembly | (TC 05/03/2023 04:37:46 PM, Chau, Ken) System taken apart, throwing out the "shell"
67948 | 194690 | CM | CMPLT | CMPLT | 2023-05-04 11:07 | 2023-05-04 11:13 | Chasis tilted downwards | (TC 05/04/2023 11:06:50 AM, Rhuland, Courtney) Chasis of incubator was not level Worked with Samir to disassemble Tightened screws on the elevating column that was causing the issue Reassembled incubator and performed functional check All tests passed
1058057 | 194691 | CM | CMPLT | CMPLT | 2023-05-04 11:57 | 2023-05-04 11:58 | data intermittent to PC | (TC 05/04/2023 11:57:11 AM, Scero, Tomlyn) Data cable to EC10 was damaged and not always working. I replaced the cable and confirmed that data was sending
| 194692 | CM | CMPLT | CMPLT | 2023-05-04 12:27 | 2023-05-04 12:28 | N2 hose and Y connector ( from Andrew Egonu ) | (TC 05/04/2023 12:26:47 PM, Samanian, Kamran) connected the hose to the Y connector through the adapters
| 194693 | CM | CMPLT | CMPLT | 2023-05-04 12:28 | 2023-05-04 12:28 | ASCOM in ICU with Nima | (TC 05/04/2023 12:28:12 PM, Samanian, Kamran) done!
| 194694 | CM | CMPLT | CMPLT | 2023-05-04 12:29 | 2023-05-04 12:29 | CAB meeting for BD Alaris Pump | (TC 05/04/2023 12:29:05 PM, Samanian, Kamran) done!
32510 | 194695 | CM | CMPLT | CMPLT | 2023-05-04 12:48 | 2023-05-08 12:33 | MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN | (TC 05/04/2023 12:48:13 PM, Whittington, Scott) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN (TC 05/04/2023 12:48:52 PM, Dolinar, John) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN (TC 05/04/2023 05:35:08 PM, Whittington, Scott) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN (TC 05/04/2023 12:48:52 PM, Dolinar, John) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN (TC 05/05/2023 12:05:20 PM, Whittington, Scott) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN REPLACED HT TANK, NO CHANGE ORIGINAL 2005 SIEMENS TUBE HAD HIGH (100 OHM)RESISTANCE ON SF, LF (AT 1.5 OHM) OK) (TC 05/05/2023 12:05:20 PM, Whittington, Scott) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN REPLACED HT TANK, NO CHANGE ORIGINAL 2005 SIEMENS TUBE HAD HIGH (100 OHM)RESISTANCE ON SF, LF (AT 1.5 OHM) OK) REPACED TUBE WITH 2004 VARIAN FROM LOCAL STOCK CALIBRATE (TC 05/05/2023 12:26:50 PM, Dolinar, John) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN REPLACED HT TANK, NO CHANGE ORIGINAL 2005 SIEMENS TUBE HAD HIGH (100 OHM)RESISTANCE ON SF, LF (AT 1.5 OHM) OK) REPACED TUBE WITH 2004 VARIAN FROM LOCAL STOCK CALIBRATE (TC 05/04/2023 05:35:08 PM, Whittington, Scott) MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN TUBE REPLACEMENT | MSH DR1 ERROR 402 POSSIBLE TUBE FILAMENT OPEN
67945 | 194696 | CM | CMPLT | NFF | 2023-05-04 12:57 | 2023-05-04 12:59 | Air probe error | (TC 05/04/2023 12:57:14 PM, Rhuland, Courtney) Came down stating "air probe failure" alarm was appearing No error appeared when device turned on No error found in the error log Completed functional check and checked air probe connections Ran for 45 minutes and no error presented No fault found
| 194697 | CM | CMPLT | CMPLT | 2023-05-04 13:00 | 2023-05-15 15:17 | Panda warmer blender for overhaul | Received
67390 | 194698 | CM | CMPLT | CMPLT | 2023-05-04 13:08 | 2023-05-10 15:45 | Blank screen | (TC 05/04/2023 08:44:24 AM, Rendon Lopez , Elizabeth) Confirmed issue, screen is blank, no alarms. Backlight board issue, ordered new control board kit (TC 05/09/2023 03:42:42 PM, Rendon Lopez , Elizabeth) Replaced control and alarm board. Tested all functions, working properly
| 194699 | CM | CMPLT | CMPLT | 2023-05-04 15:30 | 2023-05-04 15:30 | MSH CT images copied from Coral to USB/STRYKER analysis | MSH CT images copied from Coral to USB/STRYKER analysis
1058273 | 194700 | CM | CMPLT | CMPLT | 2023-05-04 16:12 | 2023-05-16 12:59 | Battery due for replacement | (TC 05/04/2023 04:12:37 PM, Cortez, John) Replaced battery, conditioning overnight (TC 05/16/2023 12:57:56 PM, Cortez, John) PM completed successfully, see ASM for results Battery capacity: 4132 mAhrs
67118 | 194701 | CM | CMPLT | CMPLT | 2023-05-05 10:27 | 2023-05-05 10:29 | Retiring device, disassembly, and disposal | (TC 05/05/2023 10:27:14 AM, Baghirzade, Nara) Retired device; the SpO2 module is inoperational. It never reads higher than 80% saturation; tried with different cables and sensors but no change. Device is end of support, no parts or repairs available. Disposed of device in e-waste.
1067011 | 194702 | ACC | CMPLT | CMPLT | 2023-05-05 11:05 | 2023-05-05 11:05 | Incoming inspection | (TC 05/05/2023 11:05:00 AM, Baghirzade, Nara) Unit is functional. Asset tagged. Dropped off at 11South
| 194703 | CM | CMPLT | CMPLT | 2023-05-05 11:26 | 2023-05-05 11:26 | Installing expair software on Dr. Blazer's PC | (TC 05/05/2023 11:26:01 AM, Baghirzade, Nara) Installing expair software on Dr. Blazer's PC. Quick overview of the software features
1067012 | 194704 | ACC | CMPLT | CMPLT | 2023-05-05 13:40 | 2023-05-16 08:42 | Incoming inspection | (TC 05/05/2023 01:49:12 PM, Canjimavadivel, Kamallesh ) Assembled cart and added to database (TC 05/15/2023 08:40:26 AM, Scero, Tomlyn) Uploaded ECG department configuration and network config. Checked settings, adjusted clock, and changed the asset label. Ready for use.
| 194705 | CM | CMPLT | CMPLT | 2023-05-05 14:49 | 2023-05-05 14:49 | labeling fibers in MCR for CORE switches & ICU data sheet | (TC 05/05/2023 02:48:57 PM, Samanian, Kamran) done!
1050100 | 194706 | CM | CMPLT | CMPLT | 2023-05-05 15:33 | 2023-10-10 10:58 | Defective DIVA module | (TC 05/05/2023 03:33:33 PM, Rendon Lopez , Elizabeth) Diva module is failing with AGM 68545 during self test. DIVA has physical damage and confirmed error when being used. Sending out for repair estimate (TC 09/20/2023 02:01:22 PM, Rendon Lopez , Elizabeth) Received device back, see attached report. (TC 10/05/2023 10:57:30 AM, Rendon Lopez , Elizabeth) Inspection completed with AGM 68543, all tests passed.
1063797 | 194707 | PM | ACTIV | ISSUE | 2023-05-05 15:54 |  | Annual PM | (TC 05/05/2023 03:54:25 PM, Green, Paul) Performed PM. Returned repaired SPO2 and tested ok. Found IBP POD has no readings. Will look further inti this as warranty until Aug 2023, Informed Jason.
| 194708 | OC | CMPLT | CMPLT | 2023-05-07 10:17 | 2023-05-12 13:11 | Qwynn Ferreira On-Call May 5 - May 12, 2023 | (TC 05/07/2023 10:17:46 AM, Ferreira, Qwynn) Received a call from OR regarding patient information not showing up in Cerner from the Philips monitor. Troubelshoot over the phone to just reboot the monitor and was able to populate patient info back into cerner.
66455 | 194709 | CM | CMPLT | CMPLT | 2023-05-08 07:47 | 2023-05-16 07:57 | ordering VN500 parts for RT | ordering flow sensor cables for RT. ordering Connector cable flow sensor for VN500 vents. We received cables. PO#7202210145
62696 | 194710 | CM | CMPLT | CMPLT | 2023-05-08 08:04 | 2023-05-08 09:06 | Not charging | (TC 05/05/2023 08:08:51 AM, Canjimavadivel, Kamallesh ) Batteries needed to be repalced. Reconditioned batteries that were on the shelf. (TC 05/08/2023 09:04:59 AM, Canjimavadivel, Kamallesh ) Installed new batteries and tested ok. Removed debris from castor that was stuck
1054987 | 194711 | CM | CMPLT | CMPLT | 2023-05-08 08:25 | 2023-05-08 08:30 | Respiratory monitoring sensors | (TC 05/08/2023 08:24:42 AM, Baghirzade, Nara) Dropped off 2 boxes of RAS-125C sensors, 2 boxes of RAS-45 sensors, 2 boxes of neo adhesive sensors, and 1 box of replacement tapes with Sarabeth for stock.
1054986 | 194712 | CM | CMPLT | CMPLT | 2023-05-08 08:31 | 2023-05-08 08:33 | Respiratory monitoring sensors | (TC 05/08/2023 08:30:57 AM, Baghirzade, Nara) Dropped off 1 box of RAS-125C sensors, 1 box of RAS-45 sensors, 1 boxes of neo adhesive sensors, and 1 box of replacement tapes with Sarabeth for stock.
1053379 | 194713 | CM | CMPLT | CMPLT | 2023-05-08 08:46 | 2023-05-08 08:47 | Expiratory cassette membrane low capacity | (TC 05/04/2023 08:46:47 AM, Rendon Lopez , Elizabeth) Replaced membrane on cassette SN: 103086. Performed tests, all passed
1055655 | 194714 | CM | CMPLT | CMPLT | 2023-05-08 08:50 | 2023-05-08 08:52 | FHR not printing | (TC 05/04/2023 08:50:57 AM, Rendon Lopez , Elizabeth) Cleaned printerhead, calibrated. testing in demo mode, working properly. Returned to L&D room 4
1066987 | 194715 | ACC | CMPLT | CMPLT | 2023-05-08 08:55 | 2023-05-08 08:56 | Incoming inspection | (TC 05/05/2023 08:55:16 AM, Rendon Lopez , Elizabeth) Incoming inspection, added to TMS. Stored in 18th Fl Biomed
68545 | 194716 | CM | CMPLT | CMPLT | 2023-05-08 08:59 | 2023-05-08 09:01 | AGM failing self test | (TC 05/05/2023 08:59:04 AM, Rendon Lopez , Elizabeth) AA removed AGM from OR 3, brought downstairs, downloaded error logs - Faulty DIVA. Removed faulty DIVA #1050100, testing with new DIVA, working no issues.
1056116 | 194717 | CM | CMPLT | CMPLT | 2023-05-08 09:34 | 2023-05-08 09:35 | Not receiving any patient list in Hyland | (TC 05/08/2023 09:33:55 AM, Samanian, Kamran) verified the problem, noticed that network cable connector is broken and rest is left on the wall jack. Repalced the cable, repair the wall jack and reconnected the cable, tested and passed. all patients are shown in the list.
1062393 | 194718 | CM | CMPLT | CMPLT | 2023-05-08 10:42 | 2023-05-08 13:10 | Display not working | (TC 05/08/2023 01:08:59 PM, Baghirzade, Nara) Headboard was not seated properly because the extender was not fully locked in. Locked the extender in properly and reseated headboard; display is working. Bed boots up properly and all functions/controls are working. Let Sorina know that the issue was resolved. (TC 05/08/2023 01:09:59 PM, Rhuland, Courtney) Headboard was not seated properly because the extender was not fully locked in. Locked the extender in properly and reseated headboard; display is working. Bed boots up properly and all functions/controls are working. Let Sorina know that the issue was resolved.
| 194719 | CM | CMPLT | CMPLT | 2023-05-08 12:28 | 2023-05-08 12:28 | huddel meeting | (TC 05/08/2023 12:28:14 PM, Samanian, Kamran) done!
| 194720 | CM | CMPLT | CMPLT | 2023-05-08 12:29 | 2023-05-08 12:30 | Olympus Server Vault - D drive had 40GB used from MAY 2 | (TC 05/08/2023 12:29:33 PM, Samanian, Kamran) verified the email alert, connected to the server and verified the issue. Engaged with Chris Ramon in TAC team in Olympus and he moved the studies to F from safezone on D and reindexed the patient list. all good now.
| 194721 | CM | CMPLT | CMPLT | 2023-05-08 13:32 | 2023-05-08 13:32 | Biomed huddle | (TC 05/08/2023 01:31:58 PM, Baghirzade, Nara) Biomed huddle
1063746 | 194722 | PM | CMPLT | CMPLT | 2023-05-08 14:05 | 2023-05-08 14:07 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/08/2023 02:04:59 PM, Rhuland, Courtney) Device came down with no water Label added to water reservoir "do not empty water" Disinfected, rinsed and refilled Buzzing noise noted from behind outlets but no alarms sounded.
1049727 | 194723 | CM | CMPLT | CMPLT | 2023-05-08 14:13 | 2023-05-08 14:16 | Repurposed for SDU | (TC 05/08/2023 02:14:29 PM, Baghirzade, Nara) Original thermometer (SN: A1570425X) sent down from SDU had a broken display and broken probe. Repurposed spare BME thermometer for SDU, updated SN, and reprinted asset tag. Thermometer is fully functional.
66429 | 194724 | CM | CMPLT | CMPLT | 2023-05-08 14:14 | 2023-05-08 14:15 | CCU was not able to view | (TC 05/08/2023 02:14:01 PM, Reid, Emmerson ) Remoted in and the central was showing that the cell was created and no data was showing. I went up and got them to turn the tele pack on. I then confirmed that TRxBio1 was meant for Tele 3 and it was working.
1050074 | 194725 | CM | CMPLT | CMPLT | 2023-05-08 14:34 | 2023-05-26 15:13 | BP cuff not working | (TC 05/17/2023 10:16:01 AM, Canjimavadivel, Kamallesh ) Cuff connector damaged, needs to be replaced. Rear housing also cracked and needs to be replaced. Reached out to vendor for parts. (TC 05/18/2023 08:52:46 AM, Canjimavadivel, Kamallesh ) Received price from vendor, placed order for rear housing (TC 05/24/2023 02:10:00 PM, Canjimavadivel, Kamallesh ) Received back housing. While prepping to reassemble discovered front housing was also damaged and needs to be replaced. Contacted vendor for quote, part is on back order for 2 weeks. (TC 05/25/2023 02:47:38 PM, Canjimavadivel, Kamallesh ) Swapped the components for the back housing. Waiting for front housing. (TC 05/26/2023 03:11:55 PM, Canjimavadivel, Kamallesh ) Swapped front housing from old retired unit for part.
| 194726 | MEP | ACTIV | ISSUE | 2023-05-08 14:41 |  | 3D Printing Work | (TC 05/08/2023 02:40:58 PM, Reid, Emmerson ) Completed several different test prints to determine the best method for creating clear prints (TC 05/08/2023 02:41:52 PM, Reid, Emmerson ) Made a file with 7 different prints to try differet methods for optimal print clarity. Also worked on running two of them. (TC 05/09/2023 07:55:18 AM, Reid, Emmerson ) Set up second batch of clearness testing (TC 05/10/2023 10:07:36 AM, Reid, Emmerson ) Met with Qwynn to discuss and design a new MRI depressor (TC 05/11/2023 02:15:15 PM, Reid, Emmerson ) Finished the 3D model of V2 of the MRI depressor and assembled it into 2 GCODE files to be printed next week (TC 05/12/2023 09:04:14 AM, Reid, Emmerson ) Discussed plans for 3D printing and test setup for clairity (TC 05/17/2023 01:40:50 PM, Reid, Emmerson ) Deisgned a filter cover to fit on a ResMed VPAP asset 1046727. Made the first version on the Dremel to test fit. With sanding on the sides to round them it fit well. I rounded the sides on the file to print with biocompatable material on the 3B+ (TC 05/18/2023 02:32:19 PM, Reid, Emmerson ) Printed the filter cover for steve with the biomed clear resin and it appears to have warped and does not fit. Will need to try again.. (TC 05/18/2023 02:33:20 PM, Reid, Emmerson ) Worked with Samir to adjust the sizes of the calibration fixture and print a new set. (TC 05/19/2023 09:47:51 AM, Reid, Emmerson ) Removed supports and gave prints to samir (TC 05/26/2023 02:38:24 PM, Reid, Emmerson ) Increased the thickness on the Filter cover print for Steve. This was done to avoid warping durring the curing process. Printed two of these, washed, cured, and removed supports. Gave a gentle sand to remove the attachment points. Both fit well. I gave them to Steve with the device. (TC 05/30/2023 12:02:33 PM, Reid, Emmerson ) Worked on V2 of the PVL model (TC 05/31/2023 02:35:43 PM, Reid, Emmerson ) Complete the right side of the PVL model and exporting it as a .STL (TC 06/01/2023 09:45:13 AM, Reid, Emmerson ) Setup the new clear resin and started a test print. (TC 06/06/2023 01:26:13 PM, Reid, Emmerson ) Setup the first of two prints for the V2 of the abdominal depressor on the dremel. (TC 06/06/2023 02:56:55 PM, Reid, Emmerson ) 3D printing meeting (TC 06/06/2023 03:28:47 PM, Reid, Emmerson ) Setup VR PC at my desk and started the abdominal depressor print again. (TC 06/09/2023 10:42:29 AM, Reid, Emmerson ) Cleaned and fixed clog in the Dremel printer. (TC 06/12/2023 11:04:33 AM, Reid, Emmerson ) Moved the dremel into the lunch room and started the depressor print again. (TC 06/12/2023 03:08:42 PM, Reid, Emmerson ) Troubleshooting the dremel printer. The nozzel kept dragging the filament and getting caught up. Ran ABC test and raised temperature, looks to be working better. Will wait to run the full print tomorrow. (TC 06/13/2023 10:01:11 AM, Reid, Emmerson ) Adjusted the print settings to 230C and 70C nozzel and bed temp. Printing the first build plate of 2 for the depressor. (TC 06/13/2023 01:16:13 PM, Reid, Emmerson ) Ran ABC test print again and adjusted the distance to +0.1. Prints look better now. (TC 06/19/2023 09:46:55 AM, Reid, Emmerson ) Setup the GCODE for Black PLA for the Depressor and started printing on the dremel. (TC 06/20/2023 03:13:36 PM, Reid, Emmerson ) Finished 3D printing the first set of depressor parts and fit them together. Just need to print the base tomorrow. (TC 06/21/2023 08:47:15 AM, Reid, Emmerson ) Started printing the base component of V2 of the MRI depressor. (TC 06/22/2023 08:54:58 AM, Reid, Emmerson ) Removed supports from base component and assembled the full V2 MRI depressor. Tested on myself and it seemed to work well. (TC 06/22/2023 02:29:00 PM, Reid, Emmerson ) Meeting with Elucis VR (TC 06/28/2023 01:55:38 PM, Reid, Emmerson ) Presented the MRI depressor to Dr. Haider. He liked the idea and suggested some minor edit to the model. He is also starting a REB application. (TC 06/28/2023 01:55:38 PM, Reid, Emmerson ) Presented the MRI depressor to Dr. Haider. He liked the idea and suggested some minor edit to the model. He is also starting a REB application. (TC 06/29/2023 10:57:06 AM, Reid, Emmerson ) Worked on updating the MRI depressor designs as per our meeting with Dr. Haider yesterday. (TC 06/29/2023 10:57:06 AM, Reid, Emmerson ) Worked on updating the MRI depressor designs as per our meeting with Dr. Haider yesterday. (TC 06/29/2023 03:08:31 PM, Reid, Emmerson ) Changed MRI depressor from Hexagon-based to octagon based after consulting with Qwynn. Also created a bent column at 15 deg. (TC 07/04/2023 03:14:36 PM, Reid, Emmerson ) 3D printing meeting (TC 07/06/2023 02:55:36 PM, Reid, Emmerson ) Modified the depressor and started printing V3 on dremel. Setup print of the Brain V2 for Dr. Pourva (TC 07/06/2023 02:57:46 PM, Reid, Emmerson ) emailed OR to discuss seting up a biomed booth to demostraite our 3DP and VR technology (TC 07/07/2023 10:52:39 AM, Reid, Emmerson ) Checked both of my prints. The depressor came out well but the hole for the column was not deep enough to securly hold the octagon shape and I will need to improve this. I will increase the height of the walls to secure this better. The PVL model turned out well. The clear resin seems to always have a yellow tint to it, however it is clear. I washed it and cleared out the cavities with alcohol. I did need to break a layer of cured resin that appeared for some reason to expose the empty cavities. I did this with a small screwdriver and a hammer then used a syringe to wash it internally. Letting it dry now. Also spoke with Qwynn and Adeel about making a poster for the VR booth and possibly working with the communication comittee. (TC 07/07/2023 03:02:04 PM, Reid, Emmerson ) CLeaned, removed supports, and cured. Ready for apoxy to be added. (TC 07/10/2023 09:24:50 AM, Reid, Emmerson ) Mixed and added resin to the PVL model. Letting it set now. I used a mixture of blue and red to make purple. The blue was much stronger than ther red in the mixture but with more red it became purple. (TC 07/11/2023 09:59:21 AM, Reid, Emmerson ) Coloured apoxy cured nicely. Used the clear resin to fill the holes. Dripped it in with a syringe and a small screwdriver. Letting it sit now to help remove bubbles. (TC 07/11/2023 07:42:52 AM, Reid, Emmerson ) Finished adding clear resin to the holes and cured. (TC 07/12/2023 09:31:40 AM, Reid, Emmerson ) Sanded the PVL model and added magnets (TC 07/12/2023 01:04:41 PM, Reid, Emmerson ) Elucis Training (TC 07/12/2023 01:04:41 PM, Reid, Emmerson ) Elucis Training (TC 07/13/2023 10:42:33 AM, Reid, Emmerson ) Reviewed and signed the Elucis trail license agreement (TC 07/13/2023 10:46:49 AM, Reid, Emmerson ) Worked on making a poster for out 3DP booth (TC 07/14/2023 10:59:56 AM, Reid, Emmerson ) Worked and finished the first version of the poster with Qwynn (TC 07/14/2023 11:24:33 AM, Reid, Emmerson ) Sent Reeba an email with a draft email for her to forward to her team regarding our 3DP booth in the OR (TC 07/18/2023 12:35:26 PM, Reid, Emmerson ) Meeting to discuss 3DP booth (TC 07/18/2023 12:35:26 PM, Reid, Emmerson ) Meeting to discuss 3DP booth (TC 07/18/2023 01:22:59 PM, Reid, Emmerson ) Went to NICU to see a 3DP project for an EEG camera mount to IV pole (TC 07/18/2023 01:22:59 PM, Reid, Emmerson ) Went to NICU to see a 3DP project for an EEG camera mount to IV pole (TC 07/18/2023 02:26:29 PM, Reid, Emmerson ) Made final adjustments to the poster and made 2 versions, letter and ledger sizes. reached out to the printing house to get an estimate for have 6 ledger and 75 letter sizes printed. (TC 07/19/2023 12:01:49 PM, Reid, Emmerson ) Went to staples to print pasters and handouts for 3DP booth. Hung up 3 posters on the 3rd floor and gave a stack of handout to Tania for the huddle Modified sinai logo and sent one to test print (TC 07/19/2023 01:51:41 PM, Reid, Emmerson ) Created a GCODE file with 19 sinai logos with a raft and sent to printer (TC 07/19/2023 02:55:33 PM, Reid, Emmerson ) Modified the MRI depressor as per Dr. Haider Included raising the coil cover and reducing the size of the strap rings (TC 07/19/2023 03:13:39 PM, Reid, Emmerson ) put up another poster on the 7th floor by the murray elevators (TC 07/20/2023 11:25:52 AM, Reid, Emmerson ) Started another sinai logo print with white PLA. Added clear coat to the Neo PVL case print (TC 07/20/2023 02:34:51 PM, Reid, Emmerson ) added several coatings of clear coat for PVL model (TC 07/21/2023 01:16:31 PM, Reid, Emmerson ) Made a 3D model of a patient lift clip for Samir (TC 07/21/2023 01:22:32 PM, Reid, Emmerson ) Started an other batch of sinailogos in white (TC 07/24/2023 10:11:26 AM, Reid, Emmerson ) Cleared off and started another print of Sinai Logos (TC 07/24/2023 10:43:15 AM, Reid, Emmerson ) Started a print for the patient lift clip in tough 2000 (TC 07/24/2023 08:31:56 AM, Reid, Emmerson ) Seting up, taking down and running the 3D printing booth in the lobby (TC 07/25/2023 10:25:44 AM, Reid, Emmerson ) 3DP booth (TC 07/26/2023 11:19:10 AM, Reid, Emmerson ) Made some edits to the patient lift clip. Shrunk the pin holes and lengthened the center part. Sent to printer. (TC 07/26/2023 09:33:40 AM, Reid, Emmerson ) 3DP Booth (TC 07/27/2023 10:24:17 AM, Reid, Emmerson ) Tried the V2 out and it was better but still not enough spring to hold shut. I compaired to the original part and it was much more flexible compaired. I started editing the V3 and thickened the tough to increase springyness. I also further reduced the diameter of the pins and increased the size of the arms closer to the base by about 50%. Then sent to the printer (TC 07/28/2023 08:44:28 AM, Reid, Emmerson ) 3DP Booth (TC 07/27/2023 08:44:28 AM, Reid, Emmerson ) 3DP Booth (TC 07/28/2023 08:55:00 AM, Reid, Emmerson ) Washed and cured the new patient lift clip (TC 07/28/2023 09:57:36 AM, Reid, Emmerson ) Test fit the patient lift clips. The slot was getting caught so I carved it out a bit to make it larger, and it worked. I made changes to the drawing to reflect this change so future prints will work without modification. (TC 07/28/2023 03:08:20 PM, Reid, Emmerson ) 3DP Booth (TC 07/31/2023 10:45:28 AM, Reid, Emmerson ) Showed a student our 3D printing lab (TC 07/31/2023 08:55:38 AM, Reid, Emmerson ) 3DP booth in OR (TC 08/01/2023 09:07:20 AM, Reid, Emmerson ) 3DP booth in the OR (TC 08/02/2023 08:52:10 AM, Reid, Emmerson ) 3DP Booth in the OR (TC 08/03/2023 10:58:38 AM, Reid, Emmerson ) 3D scanner demo (TC 08/03/2023 07:54:39 AM, Reid, Emmerson ) 3DP Booth (TC 08/04/2023 01:03:18 PM, Reid, Emmerson ) 3DP booth (TC 08/09/2023 12:29:15 PM, Reid, Emmerson ) Met at TMU to discuss VR for medical training (TC 08/09/2023 03:03:12 PM, Reid, Emmerson ) Tried printing MRI depressor parts but failed twice due to it getting tangled. Also started modeling a part for cath lab. (TC 08/10/2023 09:42:37 AM, Reid, Emmerson ) Setup the MRI depressor print again and made the spool much more robust to prevent it from getting caught (TC 08/10/2023 10:32:51 AM, Reid, Emmerson ) Finished modeling and started printed the cath lab screw handle on the SLA printer (TC 08/10/2023 10:32:51 AM, Reid, Emmerson ) Finished modeling and started printed the cath lab screw handle on the SLA printer (TC 08/11/2023 10:32:00 AM, Reid, Emmerson ) Glued and sanded the cath lab handle. Letting it set now (TC 08/11/2023 01:02:21 PM, Reid, Emmerson ) Sanded the print for cath lab more. also added a layer of tough 2000 around the crease to seal it and let it cure. I also inspected the print for Qwynn. (TC 08/11/2023 03:00:37 PM, Reid, Emmerson ) Added clear coat to the cath lab handle and letting dry (TC 08/14/2023 02:20:53 PM, Reid, Emmerson ) Went with Qwynn to drop off some prints for a patient (TC 08/15/2023 08:47:24 AM, Reid, Emmerson ) Helped Qwynn change the tank and mix the resin of the 3B+ (TC 08/16/2023 12:10:20 PM, Reid, Emmerson ) 3DP meeting (TC 08/17/2023 10:07:29 AM, Reid, Emmerson ) Setup depressor print base comp (TC 08/23/2023 02:41:04 PM, Reid, Emmerson ) Downloaded and viewed Endometrial Cancer case (TC 08/24/2023 08:40:11 AM, Reid, Emmerson ) Research on Endometrial Carcinoma for case (TC 08/25/2023 01:55:25 PM, Reid, Emmerson ) did post processes for the MRI depressor part. also had another part fail in the dremel (TC 08/25/2023 02:11:10 PM, Reid, Emmerson ) More research for EC case (TC 08/29/2023 09:46:12 AM, Reid, Emmerson ) Nara brought down the clamp with a screw that fit into the camera for the NICU. I was able to recycle a failed 3D print to make a platorm for the camera to attach to. I glued two pieced together to make it more thick and attached it to the pole clamp with 3 long wood screws. I also put the screw for mounting the camera in the middle of the two pieces with glue to hold it in place. Nara will do a final check with the camera. (TC 08/29/2023 03:11:35 PM, Reid, Emmerson ) Made changes to the NICU camera pole clamp (TC 08/30/2023 02:43:15 PM, Reid, Emmerson ) 3DP meeting (TC 08/31/2023 01:26:01 PM, Reid, Emmerson ) Reviewed different VR programs. Also working on making the EC model and did research (TC 09/05/2023 03:14:17 PM, Reid, Emmerson ) Finished making the EC model on Meshmixer (TC 09/06/2023 02:34:21 PM, Reid, Emmerson ) Downloaded ortho case files and started segmentations (TC 09/08/2023 03:11:45 PM, Reid, Emmerson ) Worked on 3Dslicer for hip case. Also helped qwynn load the PC into his car and showed Adeel meshmixer (TC 09/21/2023 08:32:35 AM, Reid, Emmerson ) Setup, ran, and took down 3D printing "VR Technology" booth for Sinai BBQ (TC 09/20/2023 08:32:35 AM, Reid, Emmerson ) Setup, ran, and took down 3D printing "VR Technology" booth for Sinai BBQ (TC 09/21/2023 10:11:30 AM, Reid, Emmerson ) Meeting with Dr. Kim (TC 09/22/2023 02:09:21 PM, Reid, Emmerson ) Meeting with Dr. Kim and team to discuss the uterine models (TC 09/22/2023 02:09:58 PM, Reid, Emmerson ) Worked on Ortho hip segmentation (TC 09/25/2023 02:43:44 PM, Reid, Emmerson ) Worked on creating uterus model. Also filled in the SoC update form (TC 09/26/2023 03:15:03 PM, Reid, Emmerson ) immersiveTouch meeting for VR (TC 09/26/2023 03:15:31 PM, Reid, Emmerson ) Meeting with Dr.Haider for the MRI depressor. also took pictures to send to him and for documentation. Also discussed the placenta device with Qwynn (TC 09/28/2023 03:15:28 PM, Reid, Emmerson ) Finished writing the update for SoC and save to T-Drive (TC 09/28/2023 03:16:04 PM, Reid, Emmerson ) Worked on Pelvis Segmentation (TC 09/29/2023 10:59:06 AM, Reid, Emmerson ) Met with Qwynn and Adeel to discuss the placenta stand design (TC 10/01/2023 08:18:36 AM, Reid, Emmerson ) Worked on Segmenting the Pelvic case. Worked at home as well, (TC 10/02/2023 08:19:51 AM, Reid, Emmerson ) Worked on Segmenting the Pelvic case. Worked at home as well, (TC 10/03/2023 09:24:09 AM, Reid, Emmerson ) Worked on finishing segmentation for pelvis and started creating the molds (TC 10/03/2023 09:24:57 AM, Reid, Emmerson ) Worked on Uterus EC model at home (TC 10/11/2023 03:20:52 PM, Reid, Emmerson ) Printed the Uterus model and fionished making the negative for the Hip case. Also started printing the hip mold (TC 10/10/2023 03:21:49 PM, Reid, Emmerson ) Worked on the hip mold. started printing the uterus model (TC 10/19/2023 03:01:15 PM, Reid, Emmerson ) CoP meeting (TC 10/18/2023 03:01:32 PM, Reid, Emmerson ) SoC presentations (TC 10/16/2023 03:02:19 PM, Reid, Emmerson ) Worked on SoC presentation (TC 10/23/2023 03:09:45 PM, Reid, Emmerson ) Tried fixing the dremel printer (TC 10/23/2023 03:11:13 PM, Reid, Emmerson ) Went to try and drop of the Uterus model to Dr. Kim (TC 10/24/2023 03:10:19 PM, Reid, Emmerson ) Discussed 3D printing plans and Ortho case with Qwynn (TC 10/27/2023 11:07:01 AM, Reid, Emmerson ) Went with Qwynn to drop of the Ovarian Mass model to OPG (TC 11/01/2023 08:05:33 AM, Reid, Emmerson ) Met with Dr. Paul Kuzyk to discuss making models for him. He mentioned several applications and we are going to try making a model for a case on the 15th (TC 11/06/2023 03:16:21 PM, Reid, Emmerson ) Printing the full sized right hip for display (TC 11/07/2023 03:00:31 PM, Reid, Emmerson ) 3D printing meeting (TC 11/07/2023 03:00:49 PM, Reid, Emmerson ) washed cured and sand 3D pelvis model for display. Also started printing 1/2 sized full pelvis model for display (TC 11/08/2023 03:13:31 PM, Reid, Emmerson ) Work on printing hip case for display. (TC 11/10/2023 01:06:58 PM, Reid, Emmerson ) Coordinating meeting with Dr. Greenberg and collecting images for case next week (TC 11/10/2023 01:07:41 PM, Reid, Emmerson ) Setup 3B+ printer to be ready to print tonight. Changed the tank and resin for biomed white (TC 11/10/2023 08:54:03 AM, Reid, Emmerson ) Met with Dr Greenberg and Kuzyk to discuss printing for their upcoming case. Also got images and generated a threshold model. (TC 11/10/2023 08:56:11 AM, Reid, Emmerson ) Made the segmentation for for the pelvic case, made it printable and sent the first print to the printer (TC 11/11/2023 09:22:11 AM, Reid, Emmerson ) Segmented the femur, Sent a second print and washed and cured the fist print. (TC 11/12/2023 09:23:33 AM, Reid, Emmerson ) Second print failed. Sent it again to print. (TC 11/13/2023 09:24:10 AM, Reid, Emmerson ) Washed and cured the second print. Removed supports and dremeled all prints. Setup printer with surgical guide resin (TC 11/13/2023 01:58:06 PM, Reid, Emmerson ) Picked up a milk crate for Qwynn (TC 11/14/2023 07:46:20 AM, Reid, Emmerson ) Showed Dr. Greenberg the final prints, cleaned, and handed over to MDRD. (TC 11/15/2023 08:34:10 AM, Reid, Emmerson ) Attended the case with Kuzyk and Greenberg while they used out models (TC 11/17/2023 02:20:07 PM, Reid, Emmerson ) Went with Nara to look at EEG that needs a new wand holder and made measurements. Created a model and printed. (TC 11/20/2023 09:43:57 AM, Reid, Emmerson ) Went with Nara to check the fit of the EEG wand holder and it fit perfectly. I press fit the part into place and started pritning the second one. (TC 11/30/2023 08:50:49 AM, Reid, Emmerson ) Testing a new filament, needed to clear the nozel multiple times because it kept jamming. (TC 11/29/2023 08:50:49 AM, Reid, Emmerson ) Testing a new filament, needed to clear the nozel multiple times because it kept jamming. (TC 12/01/2023 11:20:14 AM, Reid, Emmerson ) Finished exporting the larger ovarian cancer model (1.5x) and sent to the printer (TC 12/04/2023 02:54:23 PM, Reid, Emmerson ) unclogged the nozzel and setup a test print with new filament. Sanded and cleaned the print. worked out well (TC 12/05/2023 10:58:35 AM, Reid, Emmerson ) Filled the larger ovarian cancer model with coloured resin and letting it cure. (TC 12/06/2023 03:00:30 PM, Reid, Emmerson ) Read through a research paper outlining how to create a spinal epidural model for thoraxic insertion via 3DP. I downloaded .stl files from the paper and aligned them in meshmixer. I then also created a back overtop to help landmark the insertion point, This will be used to create a silicone mold. (TC 12/07/2023 02:29:04 PM, Reid, Emmerson ) 3DP meeting (TC 12/07/2023 02:29:25 PM, Reid, Emmerson ) Mixed more resin and topped up the ovarian cancer epoxy filling (TC 12/07/2023 02:30:00 PM, Reid, Emmerson ) Started printing the vertibrae for epidural model. First attempt failed so I adjusted settings and sending again. (TC 12/08/2023 09:37:52 AM, Reid, Emmerson ) Covered the hole on the ovarian cancer model and used resin to attach the two ovaries to the model. Letting cure now. (TC 12/08/2023 02:40:17 PM, Reid, Emmerson ) Dremeled and clear coated the ovarian cancer model (TC 12/11/2023 08:14:47 AM, Reid, Emmerson ) Worked on spinal epidural model (TC 12/15/2023 03:45:42 PM, Reid, Emmerson ) Meeting with Ortho to discuss case next week (TC 12/15/2023 03:46:05 PM, Reid, Emmerson ) segmented pelvis model (TC 12/19/2023 10:40:08 AM, Reid, Emmerson ) Met with Dr. Greenberg to discuss the creation of a surgical guide for upcoming case (TC 12/19/2023 10:43:28 AM, Reid, Emmerson ) Worked on cleaning prints and designing the surgical guide (TC 12/20/2023 10:44:05 AM, Reid, Emmerson ) Finished deisgning 3 versions of the guide and sent to printer. (TC 12/21/2023 03:17:28 PM, Reid, Emmerson ) Went over to Mars building to see Autodesk office. (TC 12/21/2023 03:19:53 PM, Reid, Emmerson ) worked on case for Greenberg. Cleaned and dropped off models to MDRD (TC 12/22/2023 01:38:38 PM, Reid, Emmerson ) Washed and cured my second R_pelvis print for cutting. Learned some basic operations in Blender for guide creation. (TC 12/22/2023 01:38:38 PM, Reid, Emmerson ) Washed and cured my second R_pelvis print for cutting. Learned some basic operations in Blender for guide creation. (TC 12/27/2023 08:36:41 AM, Reid, Emmerson ) Redesigned the guide (TC 01/04/2024 09:33:14 AM, Reid, Emmerson ) Worked on creating ortho PAO guides, also practiced making the actual PAO cuts on a model with guide. Created 2 versions with 2 sizes each. (TC 01/05/2024 02:14:29 PM, Reid, Emmerson ) Printed, washed, cured, and removed supports for medfusion pump clamp print. Handed off to Samir. (TC 01/08/2024 02:16:09 PM, Reid, Emmerson ) Sent the guide print on saunday morning. Washed, cured, and sanded the guides (TC 01/11/2024 02:36:10 PM, Reid, Emmerson ) Used two solvents to try and disolve ABS. Both eneded up working. Added a layer of silicone to the ABS print and letting it set (TC 01/12/2024 03:25:03 PM, Reid, Emmerson ) Worked on various 3D printing things. (TC 01/15/2024 11:16:08 AM, Reid, Emmerson ) Worked on Modeling the FMR Allograft (TC 01/13/2024 11:16:35 AM, Reid, Emmerson ) Modelling for upcoming ortho case (TC 01/17/2024 10:09:10 AM, Reid, Emmerson ) Modeling (TC 01/18/2024 10:14:13 AM, Reid, Emmerson ) Printed several femor models. Discussed paper topics and structure. Collaborated to make a surgical guide for the femor case. (TC 01/19/2024 10:16:12 AM, Reid, Emmerson ) Washed, stained, and cured the two femor models Also resent the femor for cutting model on the dremel with changes in print settings (TC 01/23/2024 10:32:03 AM, Reid, Emmerson ) Washed the guides for the upcoming PAO case and handed off to Qwynn (TC 01/22/2024 12:32:42 PM, Reid, Emmerson ) Worked on 3DP for Biomed paper submission (TC 02/20/2024 02:07:11 PM, Reid, Emmerson ) Worked with Kam to setup the mold with urothane for the AVM model (TC 02/20/2024 09:01:51 AM, Reid, Emmerson ) 3DP Meeting and had Qwynn teach me about the new bambu printer (TC 02/22/2024 08:31:08 AM, Reid, Emmerson ) Worked on creating the AVM model with epoxy resin and trying to disolve the internal structure in acetone, this ended up breaking the epoxy after being left overnight. Also completed segmentation for ENT case. Also attended a ortho case to view the insertion of the cutting guide. (TC 02/23/2024 03:12:46 PM, Reid, Emmerson ) Worked on AVM model, which broke. Tried to figure out next steps. Worked on an aneurysm model on 3D slicer and blender. (TC 02/26/2024 02:12:42 PM, Reid, Emmerson ) setup bambu printer file (TC 02/26/2024 02:14:06 PM, Reid, Emmerson ) Removed AVM model from mold. Setup ultrasonic cleaner. Ran ultrasonic cleaner with the AVM mold to test it out (TC 02/26/2024 03:22:15 PM, Reid, Emmerson ) More work on the AVM model and added tape to the bambu print to hold down the ABS (TC 02/27/2024 10:59:35 AM, Reid, Emmerson ) Tried to clean Samir's tool with the ultrasonic cleaner and alcohol, this damaged the tool. Also tried to setup the bambo to print the AVM model but the nozzel showed it was clogged. After some troubleshooting it looks like the clog detaction sensor is having issues and it looks like the connector has been melted. (TC 02/27/2024 03:14:00 PM, Reid, Emmerson ) Working to learn blender, boleen features (TC 02/28/2024 11:19:54 AM, Reid, Emmerson ) Created asset tags for the Bambu printer and AMS (TC 02/28/2024 11:21:12 AM, Reid, Emmerson ) Cleaned up the 3D printing space in the workshop (TC 02/28/2024 02:15:24 PM, Reid, Emmerson ) Made aneurysm model with 0.7mm thickness and set up on the 3DP PC. (TC 02/28/2024 02:52:52 PM, Reid, Emmerson ) Discussing the format for 3D printing safety presentation for biomed huddle (TC 02/29/2024 03:17:25 PM, Reid, Emmerson ) Worked on making the blender aneurysm fit the base, was able to complete version 1. (TC 03/01/2024 11:05:00 AM, Reid, Emmerson ) Added supports for the arms on the aneurysm model. Made it printable and sent to printer. Had issue with connecting and sending due to network problems. (TC 03/01/2024 02:38:09 PM, Reid, Emmerson ) Made edtis and added images to the safety and considerations ppt (TC 03/04/2024 02:23:39 PM, Reid, Emmerson ) Aneurysm model printed, I washed out thuroughly, and cured in water. I used distilled water and pre-cured it for 5 minutes before the water cure for 12 minutes. This left it still very clear and not sticky. However, it ripped when I tried to remove the supports. (TC 03/04/2024 02:25:22 PM, Reid, Emmerson ) Working on a thicker aneurysm model (1mm instead of 0.7mm) also will try removing supports before curing. (TC 03/05/2024 03:09:40 PM, Reid, Emmerson ) 3DP meeting and oc health inspection (TC 03/05/2024 03:15:38 PM, Reid, Emmerson ) Tried to remove supports from aneurysm but I broke the model, printing another with adjustments to supports and thicker shell. Also worked on a FMR resection case for practice. (TC 03/06/2024 11:01:14 AM, Reid, Emmerson ) Cleaned the aneurysm print and cured it. Had to repair a small rip. Looks good and ready to be tested. (TC 03/06/2024 03:22:37 PM, Reid, Emmerson ) Working on a second Aneurysm model, also met with Emily to discuss the first model. (TC 03/12/2024 11:27:03 AM, Reid, Emmerson ) Cleaned and cured the aneurysm model (TC 03/13/2024 03:20:51 PM, Reid, Emmerson ) Worked on the 3rd aneurysm model (TC 03/14/2024 11:03:58 AM, Reid, Emmerson ) Sent Aneurysm3 print in 50A resin (TC 03/14/2024 03:15:06 PM, Reid, Emmerson ) booked RSNA 3DP course. Worked on blender resection (TC 03/15/2024 09:55:25 AM, Reid, Emmerson ) Booking RSNA conference for April 2024 (TC 03/15/2024 09:55:48 AM, Reid, Emmerson ) cleaned the aneurysm print and removed from the plate, however it broken while removing the supports. Also setup the printer for surgical resin. (TC 03/18/2024 02:43:35 PM, Reid, Emmerson ) Sent another aneurysm print in 50A resin (TC 03/19/2024 09:35:45 AM, Reid, Emmerson ) Broke another aneurysm model, will need to make it thicker for 50A. Trying it now in 80A (TC 03/21/2024 10:02:33 AM, Reid, Emmerson ) Washed, removed supports and cured the aneurysm3 model. (TC 03/21/2024 11:03:22 AM, Reid, Emmerson ) 3DP OHS meeting (TC 03/21/2024 02:59:34 PM, Reid, Emmerson ) Organized aneurysm models by proper names in my files (TC 03/25/2024 03:10:09 PM, Reid, Emmerson ) Worked on Aneurysm models (TC 03/26/2024 11:14:51 AM, Reid, Emmerson ) Working on LICA Aneurysm model (TC 03/26/2024 02:03:49 PM, Reid, Emmerson ) Discussed with the Bone bank about a new project to create a test tube fitted drying rack for anmiotic placental tissue. (TC 03/26/2024 03:27:09 PM, Reid, Emmerson ) More aneurysm modeling and trying to remove cups (TC 03/27/2024 02:50:05 PM, Reid, Emmerson ) Cleaned cured and removed supports for LICA_0.7mm 50A print. Modeled the LICA_0.5mm for 80A. Modeled a coupon to test the biomed white print for extreme tempureatures for the amniotic platenta project. Printed this coupon, washed and cured. Printed the LICA_0.5mm 80A print (TC 03/28/2024 12:53:58 PM, Reid, Emmerson ) Cleaned and cured LICA model, however it had a printing error at one of the connectors likely due to a lack of support. I cut the connectors off and will use this print as a demo model for the stand. (TC 04/02/2024 03:04:45 PM, Reid, Emmerson ) 3DP meeting (TC 04/02/2024 03:08:12 PM, Reid, Emmerson ) Sent LICA_0.5mm_80A. Also sanded, cleaned, and handed over prints for a surgical case (TC 04/03/2024 09:38:33 AM, Reid, Emmerson ) cleaned, removed supprts and cured the LICA_80A print (TC 04/04/2024 11:14:03 AM, Reid, Emmerson ) Sent the coupon print again but at 45 deg (TC 04/04/2024 11:15:09 AM, Reid, Emmerson ) Progress report for SoC (TC 04/04/2024 02:53:53 PM, Reid, Emmerson ) cleaned and cured the coupon print. setup the printer for surgical guide resin. Finished up and sent in the progress report for SoC (TC 04/10/2024 09:03:49 AM, Reid, Emmerson ) Washed and cured a print for Qwynn (TC 04/10/2024 10:47:16 AM, Reid, Emmerson ) Removed supports and sanded the two models for Joel case. (TC 04/10/2024 11:13:16 AM, Reid, Emmerson ) Sent a temp coupon print (TC 04/10/2024 01:16:03 PM, Reid, Emmerson ) Dropped off model for Joel (TC 04/11/2024 08:27:17 AM, Reid, Emmerson ) washed and cured coupon (TC 04/12/2024 09:41:00 AM, Reid, Emmerson ) Coupon was less curled this time. Trying another variation. (TC 04/12/2024 03:28:58 PM, Reid, Emmerson ) printing and quality discussions (TC 04/15/2024 03:13:17 PM, Reid, Emmerson ) Designed aneurysm connectors in fusion (TC 04/16/2024 02:31:05 PM, Reid, Emmerson ) Meeting (TC 04/16/2024 03:13:58 PM, Reid, Emmerson ) Looked for matierial to make aneurysm sim stand. Reached out to Emily to borrow their to remake. Also downloaded mimics for RSNA conference. (TC 04/17/2024 03:22:04 PM, Reid, Emmerson ) Modeling and sourcing parts for Aneurysm stand (TC 04/22/2024 08:34:25 AM, Reid, Emmerson ) Got the patient images from eImages and uploaded them to Slicer, was having issues so instralled slicer 5.2 and reuploaded images there. Completed pelvic segmentation and sent email to schedule a meeting with Ortho. Found out later that the case has been moved up and we will not be able to print for them. (TC 04/23/2024 01:02:35 PM, Reid, Emmerson ) Made final edits to the Aneurysm connetors and printed 1 large and 1 small to test the fit. (TC 04/23/2024 03:25:03 PM, Reid, Emmerson ) Worked with VR and washed and cured my connector print (TC 04/24/2024 08:55:38 AM, Reid, Emmerson ) Removed supports from my connector print and test fit. The nut was not fitting on the connector and on the large one the nut was not fitting over the rivots. Made changes and sent another print. Modeling took extra long because my computer was really struggling running the programs. (TC 04/24/2024 03:00:12 PM, Reid, Emmerson ) Modeled a new LICA model to fit my connectors at a smaller length (14cm instead of 17cm) for better printability. (TC 04/25/2024 02:15:54 PM, Reid, Emmerson ) Created 3 variations in size for both the large and small nut to test threading. washed and cured the print (TC 04/25/2024 03:17:06 PM, Reid, Emmerson ) Removed supports and determined the best fit for the nut prints. Updated Fusion file so the correct fit is reflected in the model. (TC 04/29/2024 09:58:06 AM, Reid, Emmerson ) Created a 3D model for a smalled stand, and sent it to the printer. (TC 04/30/2024 03:23:42 PM, Reid, Emmerson ) Worked on the post processing for the aneurysm stand(lost of supports and sanding), created a Basilar tip aneurysm model. setup a build plate and started a print for a basilar tip and LICA model. (TC 05/01/2024 10:02:17 AM, Reid, Emmerson ) 3DP meeting (TC 05/01/2024 10:21:30 AM, Reid, Emmerson ) Sanding and clear coated the aneurysm stand (TC 05/01/2024 03:13:21 PM, Reid, Emmerson ) DId some organizing of the 3DP shop space. Downlaoded and setup the ACOM aneurysm model in blender. (TC 05/02/2024 08:41:48 AM, Reid, Emmerson ) Gather tubing and a syringe for the auerysm stand setup (TC 05/02/2024 11:07:46 AM, Reid, Emmerson ) Went to Dollarama for a glass container for curing Aneurysm models (TC 05/02/2024 11:09:42 AM, Reid, Emmerson ) Washed the models. Attempted to remove supports on the first one and it ripped. Curing the second with supports. Tried cutting a glass jar with metal wire but it didn't work. (TC 05/02/2024 02:38:59 PM, Reid, Emmerson ) Lots of work post processing the aneurysm models. Lost one of them in the process due to a rip. Found the connectors fit very well but model should be made a few cms smaller (TC 05/03/2024 12:23:48 PM, Reid,
1061333 | 194727 | FC | CMPLT | CMPLT | 2023-05-08 14:59 | 2023-05-08 14:59 | QC checklist | (TC 05/08/2023 02:59:00 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1055177 | 194728 | CM | CMPLT | CMPLT | 2023-05-08 15:00 | 2023-05-08 15:01 | No yellow tag or problem description | (TC 05/08/2023 02:59:51 PM, Baghirzade, Nara) Ran a check-in and occlusion pressure was high; Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
1059938 | 194729 | CM | CMPLT | CMPLT | 2023-05-08 15:27 | 2023-05-08 15:27 | V3 not working | (TC 05/08/2023 03:27:05 PM, Canjimavadivel, Kamallesh ) Replaced chest lead set, tested okay
1065054 | 194730 | FC | CMPLT | CMPLT | 2023-05-08 15:55 | 2023-05-10 15:50 | Check-in/functional check | (TC 05/08/2023 03:55:12 PM, Baghirzade, Nara) Received back from BD repair, wireless card was replaced; Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 218 mohms Leakage current: 56.0 uA
| 194731 | CM | CMPLT | CMPLT | 2023-05-08 17:18 | 2023-05-08 17:18 | Performance Appraisal paperwork and latch repair for Civco door, site coverage
66465 | 194732 | PM | CMPLT | CMPLT | 2023-05-09 07:50 | 2023-05-09 07:52 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. NO PS500 and NO GS500. Tested. Functional. SW.2.42 Leakage: 73 uA and 85ohm.
67955 | 194733 | CM | CMPLT | NFF | 2023-05-09 08:28 | 2023-05-09 08:30 | Temperature sensor issue | (TC 05/09/2023 08:28:21 AM, Rhuland, Courtney) Came down stating "temperature sensor?" ran functionality test for 30 minutes Temp held steady No issue found
65964 | 194734 | CM | CMPLT | CMPLT | 2023-05-09 08:37 | 2023-05-09 08:38 | Channel error | (TC 05/09/2023 08:36:46 AM, Rhuland, Courtney) Pressure sensor error Pressure sensor replaced Calibration and check in completed Refer to ASM for details
1054981 | 194735 | CM | CMPLT | CMPLT | 2023-05-09 08:46 | 2023-05-09 08:48 | Assembled new stand, config load, and functional check | (TC 05/09/2023 08:46:03 AM, Baghirzade, Nara) Assembled new stand and transferred Root on to it. Stand was already charged in previous WO to this asset. When booting the Root, noted that there is no MSH configuration loaded. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Removed the radical battery. Battery was swollen. Printed asset tag with 15M as the location. Functional tests of NIBP, temp, and SpO2 completed and passed.
| 194736 | RDS | ACTIV | ISSUE | 2023-05-09 09:27 |  | Weekly check for NICU - Test
1061217 | 194737 | CM | CMPLT | CMPLT | 2023-05-09 09:42 | 2023-05-09 09:45 | Tabletop micrscope issue | (TC 05/09/2023 09:41:40 AM, Reid, Emmerson ) Went to OR16 because the monitors were cutting out one AUX 1 with the tabletop microscope. I watched as this happened, it would momentarily cut out every minute or so. It didn't seem to be caused by anything. They treid reinstering the DVI end of the cable. I reinsterted the HDMI end at the micropscope and rested the cable up so it wasn't pulling down as hard. This seemed to help. It is working now but they will let me know if they have more issues.
1067048 | 194738 | ACC | CMPLT | CMPLT | 2023-05-09 10:43 | 2023-05-11 16:49 | Incoming Inspection | (TC 05/09/2023 04:47:46 PM, Cortez, John) Visual/functional test ok electrical safety passed, 50 mOhms, 19 uAmps
1067049 | 194739 | ACC | CMPLT | CMPLT | 2023-05-09 10:48 | 2023-05-11 16:47 | Incoming Inspection | (TC 05/09/2023 04:42:32 PM, Cortez, John) Visual/functional test ok electrical safety passed, 60 mOhms, 24 uAmps
1053347 | 194740 | CM | CMPLT | CMPLT | 2023-05-09 10:50 | 2023-07-11 07:40 | MSH BI 1270 H1 UNABLE TO EXPOSE GENERATOR ERRORS 25:8, 25:11 | (TC 05/09/2023 10:50:09 AM, Whittington, Scott) MSH BI 1270 H1 UNABLE TO EXPOSE GENERATOR ERRORS 25:8, 25:11 REQUIRES ELECTRICAL SUPPLY UPGRADE (TC 05/10/2023 11:43:23 AM, Whittington, Scott) MSH BI 1270 H1 UNABLE TO EXPOSE GENERATOR ERRORS 25:8, 25:11 REQUIRES ELECTRICAL SUPPLY UPGRADE CONSULT WITH STEVE FLANAGAN, BIRNIE ELECTRIC, CHECKING ON-SITE ELECTRICAL SUPPLY (TC 05/31/2023 01:34:27 PM, Whittington, Scott) From: Steven Flanagan <sflanagan@birnie.com> Sent: May 26, 2023 9:59 PM To: Francis, Kyle - Sinai Health System <Kyle.Francis@sinaihealth.ca> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Noseworthy, Joanne - Sinai Health System <Joanne.Noseworthy@sinaihealth.ca>; Veeneman, Jon <Jon.Veeneman@hologic.com>; Moe - Dynamic Imaging <moe@dynamicimaging.ca>; Whittington, Scott <Scott.Whittington@uhn.ca>; Quinn, Frank - Sinai Health System <Frank.Quinn@sinaihealth.ca>; Dolinar, John <John.Dolinar@uhn.ca>; Fan, Tak <Tak.Fan@uhn.ca>; Manzo, Peter <Peter.Manzo@hologic.com>; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca>; Kyle Brunt <kyleb3333@gmail.com> Subject: Re: [External] Room 1270 Electrical Issue Thanks Kyle, For our records, room 1270’s equipment is now fed from distribution panel PP-12L Recorded readings at the disconnect and at the unit are as follows, Phase A to isolated ground -121V Phase B to isolated ground - 121V Phase A-B - 209V (TC 06/05/2023 02:44:41 PM, Whittington, Scott) JUNE 5 2023 CPU REPLACED 3C:EC:EF:C5:3D:9C replaces AC-1F-6B-B2-AE-64 HOLOGIC REPLACED TUBE | MSH BI 1270 H1 UNABLE TO EXPOSE GENERATOR ERRORS 25:8, 25:11 REQUIRES ELECTRICAL SUPPLY UPGRADE Work with Hologic and Electrical. Found x-ray tube faulty. JUNE 5 2023 CPU REPLACED 3C:EC:EF:C5:3D:9C replaces AC-1F-6B-B2-AE-64
1059828 | 194741 | PM | CMPLT | CMPLT | 2023-05-09 10:50 | 2023-05-09 10:51 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/09/2023 10:49:53 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water Label placed on water reservoir "DO NOT EMPTY WATER"
1067050 | 194742 | ACC | CMPLT | CMPLT | 2023-05-09 10:50 | 2023-05-11 16:42 | Incoming Inspection | (TC 05/09/2023 04:38:35 PM, Cortez, John) Visual/functional/electrical safety test ok
1061107 | 194743 | CM | ACTIV | ISSUE | 2023-05-09 10:52 |  | Not able to setup bed exit alarm | (TC 05/09/2023 10:51:48 AM, Reid, Emmerson ) Went to look at the bed and I could not set up the bed exit alarm. Sent an email to Mark from stryker to come in.
1061203 | 194744 | CM | CMPLT | CMPLT | 2023-05-09 10:53 | 2023-05-12 09:04 | Needed light camera for case | (TC 05/09/2023 10:53:36 AM, Reid, Emmerson ) Maureen called about setting up the light camera for a case today. I set it up on light three and will need to take it down later today.
1061443 | 194745 | CM | CMPLT | NFF | 2023-05-09 10:56 | 2023-05-09 10:57 | Reported that buttons on left side control not working | (TC 05/09/2023 10:55:46 AM, Baghirzade, Nara) Power-cycled the bed; bed booted up properly. Checked all functions on footboard display, side controls, and patient controls; all functions are working. No fault found with bed. Let Sorina know that the bed is functional.
68607 | 194746 | CM | CMPLT | CMPLT | 2023-05-09 11:04 | 2023-05-09 11:05 | Missing barcode | (TC 05/09/2023 11:04:37 AM, Rhuland, Courtney) Barcode printed Check-in completed Refer to ASM for details
1067051 | 194747 | ACC | CMPLT | CMPLT | 2023-05-09 12:34 | 2023-05-17 13:08 | Incoming inspection | Asset tagged
1067052 | 194748 | ACC | CMPLT | CMPLT | 2023-05-09 12:37 | 2023-05-17 13:08 | Incoming inspection | Asset tagged
1067053 | 194749 | ACC | CMPLT | CMPLT | 2023-05-09 14:28 | 2023-05-10 07:42 | Incoming inspection and asset tagging | (TC 05/10/2023 07:40:22 AM, Baghirzade, Nara) Assembled stand and asset tagged the pump. Electrical safety: Power cord resistance = 250 mohms Leakage current = 17.4 uA Incoming inspection completed. Coordinated between Sarabeth, Wendy, and Brittany for storage of the device in PACU for the 11N decant.
| 194750 | CM | CMPLT | CMPLT | 2023-05-10 08:06 | 2023-05-11 13:46 | Crash cart from MRI ( extension power bar loss ground.) | (TC 05/11/2023 01:45:05 PM, Chukri, Samir) Replaced power cable. tested working fine.
1062557 | 194751 | CM | CMPLT | CMPLT | 2023-05-10 08:20 | 2023-06-06 12:07 | Headboard size not being detected and scale requires calibration | (TC 05/10/2023 08:19:53 AM, Baghirzade, Nara) Tried to reset the size and headboard multiple times; headboard size is not being detected by the bed. This is causing the bed to continously beep when plugged in to AC. Scale requires calibration. Placed a service request with Stryker. (TC 06/06/2023 12:05:49 PM, Baghirzade, Nara) Attached service report.
1054522 | 194752 | CM | CMPLT | CMPLT | 2023-05-10 08:35 | 2023-05-10 08:47 | RA lead broken | (TC 05/10/2023 08:46:21 AM, Reid, Emmerson ) Replaced the RA lead and tested. Working now.
1047858 | 194753 | CM | CMPLT | CMPLT | 2023-05-10 09:47 | 2023-05-10 14:03 | Uploaded MSH config/date and time adjustment | (TC 05/10/2023 02:03:18 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
65984 | 194754 | CM | CMPLT | CMPLT | 2023-05-10 10:22 | 2023-05-10 10:26 | Channel error | (TC 05/10/2023 10:22:17 AM, Rhuland, Courtney) Pressure sensor error 354.6770 Reseated pressure sensor calibration and check-in completed All tests passed Refer to ASM for details
32510 | 194755 | CM | CMPLT | CMPLT | 2023-05-10 11:46 | 2023-07-11 07:40 | MSH DR1 GRIDS NOT DETECTED | (TC 05/10/2023 11:47:08 AM, Whittington, Scott) MSH DR1 GRIDS NOT DETECTED CHECKED | MSH DR1 GRIDS NOT DETECTED
1047863 | 194756 | CM | CMPLT | CMPLT | 2023-05-10 13:59 | 2023-05-10 13:59 | Uploaded MSH config/adjust time and date | (TC 05/10/2023 01:59:01 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1047864 | 194757 | CM | CMPLT | CMPLT | 2023-05-10 14:00 | 2023-05-10 14:00 | Updated config, adjusted date and time | (TC 05/10/2023 01:59:55 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1047848 | 194758 | CM | CMPLT | CMPLT | 2023-05-10 14:02 | 2023-05-10 14:03 | Uploaded MSH config. Adjusted date and time | (TC 05/10/2023 02:02:33 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1047851 | 194759 | CM | CMPLT | CMPLT | 2023-05-10 14:39 | 2023-05-10 14:39 | Uploaded MSH config. Adjusted time and date | (TC 05/10/2023 02:38:45 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1047857 | 194760 | CM | CMPLT | CMPLT | 2023-05-10 14:39 | 2023-05-10 14:40 | Uploaded MSH config. Adjusted date and time | (TC 05/10/2023 02:39:37 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1047851 | 194761 | PM | CMPLT | CMPLT | 2023-05-10 14:55 | 2023-05-10 14:57 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/10/2023 02:56:16 PM, Baghirzade, Nara) PM completed. See procedure for PM results.
| 194762 | RDS | CMPLT | CMPLT | 2023-05-10 15:42 | 2023-05-10 15:42 | Rounds | (TC 05/10/2023 03:42:11 PM, Baghirzade, Nara) Rounds in NICU.
| 194763 | RDS | CMPLT | CMPLT | 2023-05-10 15:43 | 2023-05-10 15:43 | Rounds | (TC 05/10/2023 03:42:57 PM, Baghirzade, Nara) Rounds in 11N and 11S
1063133 | 194764 | CM | CMPLT | CMPLT | 2023-05-10 15:46 | 2023-05-17 08:29 | Not picking up heart rate | (TC 05/09/2023 03:49:58 PM, Rendon Lopez , Elizabeth) Damaged probe crystals. needs replacement probe (TC 05/17/2023 08:27:42 AM, Rendon Lopez , Elizabeth) New probe tested-OK
67088 | 194765 | CM | CMPLT | NFF | 2023-05-10 15:51 | 2023-05-10 15:52 | No measure of BP | (TC 05/10/2023 03:51:33 PM, Rendon Lopez , Elizabeth) Tested, NFF Returned to unit
1046541 | 194766 | CM | CMPLT | CMPLT | 2023-05-10 15:53 | 2023-05-10 15:54 | Will not turn off | (TC 05/08/2023 03:53:05 PM, Rendon Lopez , Elizabeth) Malfunction. Replaced thermometer with a spare one, changed SN in TMS, returned to 15M
| 194767 | CM | CMPLT | CMPLT | 2023-05-10 17:19 | 2023-05-10 17:19 | Conversion SDI to HDMI in SPP procedure room |  | Needs special converter to scale down to 30Hz
1059795 | 194768 | CM | CMPLT | CMPLT | 2023-05-10 17:24 | 2023-05-10 17:23 | An upgrade is needed, FCO is ready. | Upgrade was completed.
1055030 | 194769 | CM | CMPLT | CMPLT | 2023-05-11 07:29 | 2023-05-11 07:30 | Battery replacement + upgrade | (TC 05/11/2023 07:29:04 AM, Baghirzade, Nara) Replaced battery in Radical. Used the battery upgrade SDcard to upgrade battery; returned the original SDcard after upgrade completed.
1054980 | 194770 | CM | CMPLT | NFF | 2023-05-11 07:30 | 2023-05-11 07:32 | Reported temperature probe is not working | (TC 05/11/2023 07:30:21 AM, Baghirzade, Nara) Temperature probe was not seated properly; re-seated probe. Temperature module is functional. No fault found. Functional check for NIBP and SpO2 passed.
1054537 | 194771 | CM | CMPLT | CMPLT | 2023-05-11 07:52 | 2023-05-11 07:53 | Cleaning cycle | (TC 05/09/2023 07:51:58 AM, Cortez, John) Completed cleaning cycle
1047855 | 194772 | PM | CMPLT | CMPLT | 2023-05-11 08:41 | 2023-05-11 09:29 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/11/2023 09:27:59 AM, Rhuland, Courtney) PM complete All tests passed Refer to procedure for details
1047855 | 194773 | CM | CMPLT | CMPLT | 2023-05-11 08:48 | 2023-05-11 08:48 | Uploaded config. Adjusted date and time. | (TC 05/11/2023 08:47:43 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1047854 | 194774 | CM | CMPLT | CMPLT | 2023-05-11 08:48 | 2023-05-11 08:49 | Uploaded MSH config. Adjusted date and time | (TC 05/11/2023 08:48:33 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1050547 | 194775 | CM | CMPLT | NA | 2023-05-11 08:50 | 2023-05-11 08:51 | Low readings | (TC 05/11/2023 08:49:53 AM, Baghirzade, Nara) Tried to calibrate the thermometer; calibration failed. Hardware error 8. Device is over 5 years old and no longer covered by warranty. Device will be retired.
1047856 | 194776 | CM | CMPLT | CMPLT | 2023-05-11 08:53 | 2023-05-11 08:53 | Uploaded config and adjusted date and time | (TC 05/11/2023 08:53:13 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1066114 | 194777 | FC | CMPLT | CMPLT | 2023-05-11 09:03 | 2023-05-11 09:03 | QC check | (TC 05/11/2023 09:02:49 AM, Canjimavadivel, Kamallesh ) QC check pass
| 194778 | PA | CMPLT | CMPLT | 2023-05-11 09:47 | 2023-06-05 08:26 | 3D Printer Resin Order
1049903 | 194779 | CM | CMPLT | CMPLT | 2023-05-11 09:52 | 2023-05-11 09:55 | No network | (TC 05/11/2023 09:52:57 AM, Fan, Tak) Checked network connection. Connected to MSH network, no connection to JDMIPACS, WL, or downtime. Checked settings, OK. Restarted system, connection OK. Sent pending images, transferred OK. Not sure if related to Coral network issues. Updated time. Detector calibrated. | No connection to PACS and downtime and HIS/RIS. Detector needs calibration as well.
62908 | 194780 | PM | CMPLT | CMPLT | 2023-05-11 11:11 | 2023-06-08 10:55 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 06/08/2023 10:54:25 AM, Ferreira, Qwynn) Tested and working properly. PM sheet attached
67431 | 194781 | PM | CMPLT | CMPLT | 2023-05-11 11:15 | 2023-06-08 11:09 | PM-MAJOR - Annual PM For: Air-oxygen Blender | (TC 06/08/2023 11:08:23 AM, Ferreira, Qwynn) Tested and working properly. PM sheet attached from company
67634 | 194782 | PM | CMPLT | CMPLT | 2023-05-11 13:16 | 2023-06-08 11:07 | PM-MAJOR - Annual PM For: Air-oxygen Blender | (TC 06/08/2023 11:06:03 AM, Ferreira, Qwynn) Tested and working properly. PM sheet attached from company
501583 | 194783 | CM | CMPLT | CMPLT | 2023-05-11 14:48 | 2023-07-11 07:41 | MSH SPP 15-217 Rose L&D OVER TEMP WARNING | (TC 05/11/2023 02:48:27 PM, Whittington, Scott) SPP 15-217 Rose L&D OVER TEMP WARNING | SPP 15-217 Rose L&D OVER TEMP WARNING
501442 | 194784 | CM | CMPLT | CMPLT | 2023-05-11 14:49 | 2023-08-23 12:15 | MSH US PROBE PVT 375 ARTIFACT | (TC 05/11/2023 02:49:23 PM, Whittington, Scott) MSH US PROBE PVT 375 ARTIFACT | MSH US PROBE PVT 375 ARTIFACT
1058648 | 194785 | CM | CMPLT | CMPLT | 2023-05-11 14:50 | 2023-08-23 12:15 | MSH 5 US PROBE PVT 661 ARTIFACT | (TC 05/11/2023 02:50:22 PM, Whittington, Scott) MSH 5 US PROBE PVT 661 ARTIFACT | MSH 5 US PROBE PVT 661 ARTIFACT
| 194786 | CM | CMPLT | CMPLT | 2023-05-11 15:01 | 2023-05-11 15:04 | Procedure recall notice meeting | (TC 05/11/2023 03:00:45 PM, Baghirzade, Nara) Shawn from Clarion dropped off a procedure recall notice letter and USB with updated RFU for the Ellex Tangp. Samir copied the RFU and notice to T Drive. Met with Henna to discuss the notice and informed that she needs to pass it on to the physicians for their review.
| 194787 | PA | CMPLT | CMPLT | 2023-05-11 15:27 | 2023-06-19 12:44 | NICU BP Cuffs Sizes 2-4 | Handed these over to materials management
66155 | 194788 | CM | CMPLT | CMPLT | 2023-05-11 15:41 | 2023-05-11 15:49 | Temp not reading correctly | (TC 05/11/2023 03:48:44 PM, Ferreira, Qwynn) Tested the probe and it was significantly off. Replaced the probe and tested 37 = 37deg
65719 | 194789 | CM | CMPLT | CMPLT | 2023-05-11 16:25 | 2023-05-12 14:14 | Touchscreen not working | (TC 05/10/2023 04:25:15 PM, Cortez, John) Noticed touchscreen not working during installation in new ICU. Difficult to navigate UI even with remote/mouse, screen may be defective. Will try to calibrate via support tool (TC 05/12/2023 02:02:40 PM, Reid, Emmerson ) Worked wit Courtney to replace the screen and recallibrate. Working now.
65871 | 194790 | ACC | CMPLT | CMPLT | 2023-05-11 17:24 | 2023-05-11 17:26 | Installed on new tower | (TC 05/08/2023 05:24:36 PM, Cortez, John) Installed on new tower by Olympus tech Electrical safety passed, 50 mohms 41 uAmps
65870 | 194791 | ACC | CMPLT | CMPLT | 2023-05-11 17:28 | 2023-05-11 17:29 | Installed on new tower | (TC 05/08/2023 05:28:20 PM, Cortez, John) see WO# 194790
1067040 | 194792 | ACC | CMPLT | CMPLT | 2023-05-11 17:36 | 2023-05-11 17:37 | Incoming Inspection | (TC 05/08/2023 05:36:15 PM, Cortez, John) Electrical safety passed, 50 mohms, 42 uAmps
1067039 | 194793 | ACC | CMPLT | CMPLT | 2023-05-11 17:39 | 2023-05-11 17:40 | Incoming Inspection | (TC 05/08/2023 05:39:09 PM, Cortez, John) see WO# 194792
| 194794 | TEDU | CMPLT | CMPLT | 2023-05-11 19:29 | 2023-05-12 08:46 | HillRom TS7000 training- Batevillw Indiana | (TC 05/09/2023 07:29:00 PM, Samanian, Kamran) fligt (TC 05/12/2023 08:45:37 AM, Samanian, Kamran) return
| 194795 | CM | CMPLT | CMPLT | 2023-05-11 19:32 | 2023-05-08 19:32 | meeting with Anita-katrina-Stephanie and Adeel for Hyland and wrong patients | (TC 05/08/2023 07:31:57 PM, Samanian, Kamran) discussed
| 194796 | CM | CMPLT | CMPLT | 2023-05-11 19:54 | 2023-05-23 08:49 | EC5#104 for thrid partyy device | (TC 05/12/2023 08:48:41 AM, Samanian, Kamran) PO# 7202213273
66453 | 194797 | CM | CMPLT | CMPLT | 2023-05-12 08:48 | 2023-05-18 08:25 | ordering VN500 parts | ordering VN500 parts. received parts. PO#7202213320
1066632 | 194798 | ACC | CMPLT | CMPLT | 2023-05-12 10:10 | 2023-05-12 10:11 | Incoming inspection | (TC 04/05/2023 10:09:47 AM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066639 | 194799 | ACC | CMPLT | CMPLT | 2023-05-12 10:12 | 2023-05-12 10:12 | Incoming inspection | (TC 04/05/2023 10:11:48 AM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066643 | 194800 | ACC | CMPLT | CMPLT | 2023-05-12 10:13 | 2023-05-12 10:14 | Incoming inspection | (TC 04/05/2023 10:13:40 AM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066645 | 194801 | ACC | CMPLT | CMPLT | 2023-05-12 10:16 | 2023-05-12 10:17 | Incoming inspection | (TC 04/05/2023 10:16:06 AM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066670 | 194802 | ACC | CMPLT | CMPLT | 2023-05-12 10:18 | 2023-05-12 10:19 | Incoming inspection | (TC 04/05/2023 10:18:38 AM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1053349 | 194803 | CM | CMPLT | CMPLT | 2023-05-12 10:48 | 2023-05-19 17:40 | MSH BI 1274 H2 ARTIFACT, UNABLE TO PASS DAILY FLAT FIELD QC | (TC 05/12/2023 10:48:53 AM, Whittington, Scott) MSH BI 1274 H2 ARTIFACT, UNABLE TO PASS DAILY FLAT FIELD QC PETER MANZOS TUNNELD IN AND FOUND FAULT CASUSED BY DEBRIS ON TUBE SIDE REQ CLEAN | MSH BI 1274 H2 ARTIFACT, UNABLE TO PASS DAILY FLAT FIELD QC New Tube installed HARP completed.
| 194804 | CM | CMPLT | CMPLT | 2023-05-12 12:04 | 2023-05-12 12:01 | Bi-Weekly Department Meetings
| 194805 | CM | CMPLT | CMPLT | 2023-05-12 12:06 | 2023-05-12 12:01 | Bi-Weekly Department Meetings
| 194806 | CM | CMPLT | CMPLT | 2023-05-12 12:13 | 2023-05-12 12:11 | Bi-Weekly Department Meetings
| 194807 | CM | CMPLT | CMPLT | 2023-05-12 12:16 | 2023-05-12 12:16 | Bi- Weekly DepartmentMeetings
1066678 | 194808 | ACC | CMPLT | CMPLT | 2023-05-12 13:51 | 2023-05-12 13:52 | Incoming inspection | (TC 04/05/2023 01:51:34 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066679 | 194809 | ACC | CMPLT | CMPLT | 2023-05-12 13:53 | 2023-05-12 13:55 | Incoming inspection | (TC 04/05/2023 01:53:44 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066680 | 194810 | ACC | CMPLT | CMPLT | 2023-05-12 13:57 | 2023-05-12 13:58 | Incoming inspection | (TC 04/05/2023 01:57:28 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066682 | 194811 | ACC | CMPLT | CMPLT | 2023-05-12 14:00 | 2023-05-12 14:03 | Incoming inspection | (TC 04/05/2023 02:00:52 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
66474 | 194812 | PM | CMPLT | CMPLT | 2023-05-12 14:04 | 2023-05-12 14:04 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. NO PS500 and NO GS500. Tested. Functional. SW.2.42 Leakage: 77 uA and 83ohm.
1066890 | 194813 | ACC | CMPLT | CMPLT | 2023-05-12 14:28 | 2023-05-12 14:29 | Incoming inspection | (TC 05/03/2023 02:27:49 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066891 | 194814 | ACC | CMPLT | CMPLT | 2023-05-12 14:33 | 2023-05-12 14:35 | Incoming inspection | (TC 05/03/2023 02:34:44 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066892 | 194815 | ACC | CMPLT | CMPLT | 2023-05-12 14:38 | 2023-05-12 14:39 | Incoming inspection | (TC 05/03/2023 02:37:58 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066893 | 194816 | ACC | CMPLT | CMPLT | 2023-05-12 14:39 | 2023-05-12 14:40 | Incoming inspection | (TC 05/03/2023 02:39:38 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066894 | 194817 | ACC | CMPLT | CMPLT | 2023-05-12 14:41 | 2023-05-12 14:41 | Incoming inspection | (TC 05/03/2023 02:40:47 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066895 | 194818 | ACC | CMPLT | CMPLT | 2023-05-12 14:42 | 2023-05-12 14:43 | Incoming inspection | (TC 05/03/2023 02:42:30 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066896 | 194819 | ACC | CMPLT | CMPLT | 2023-05-12 14:44 | 2023-05-12 14:44 | Incoming inspection | (TC 05/03/2023 02:44:00 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066897 | 194820 | ACC | CMPLT | CMPLT | 2023-05-12 14:45 | 2023-05-12 14:47 | Incoming inspection | (TC 05/03/2023 02:46:22 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066898 | 194821 | ACC | CMPLT | CMPLT | 2023-05-12 14:48 | 2023-05-12 14:48 | Incoming inspection | (TC 05/03/2023 02:47:43 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066899 | 194822 | ACC | CMPLT | CMPLT | 2023-05-12 14:49 | 2023-05-12 14:50 | Incoming inspection | (TC 05/03/2023 02:49:26 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066900 | 194823 | ACC | CMPLT | CMPLT | 2023-05-12 14:51 | 2023-05-12 14:51 | Incoming inspection | (TC 05/03/2023 02:51:01 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066901 | 194824 | ACC | CMPLT | CMPLT | 2023-05-12 14:55 | 2023-05-12 14:56 | Incoming inspection | (TC 05/03/2023 02:55:19 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066902 | 194825 | ACC | CMPLT | CMPLT | 2023-05-12 14:57 | 2023-05-12 14:58 | Incoming inspection | (TC 05/03/2023 02:57:15 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066903 | 194826 | ACC | CMPLT | CMPLT | 2023-05-12 14:58 | 2023-05-12 15:00 | Incoming inspection | (TC 05/03/2023 02:59:08 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066904 | 194827 | ACC | CMPLT | CMPLT | 2023-05-12 15:01 | 2023-05-12 15:01 | Incoming inspection | (TC 05/03/2023 03:00:44 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066905 | 194828 | ACC | CMPLT | CMPLT | 2023-05-12 15:02 | 2023-05-12 15:07 | Incoming inspection | (TC 05/03/2023 03:02:48 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066906 | 194829 | ACC | CMPLT | CMPLT | 2023-05-12 15:08 | 2023-05-12 15:09 | Incoming inspection | (TC 05/03/2023 03:08:40 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066907 | 194830 | ACC | CMPLT | CMPLT | 2023-05-12 15:11 | 2023-05-12 15:11 | Incoming inspection | (TC 05/03/2023 03:10:49 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
63027 | 194831 | PM | CMPLT | CMPLT | 2023-05-12 15:11 | 2023-05-12 15:13 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/12/2023 03:11:34 PM, Ferreira, Qwynn) PM Passed Hours: 5128 Avg Intensity: 40.3 Electrical Safety - OK
1066908 | 194832 | ACC | CMPLT | CMPLT | 2023-05-12 15:12 | 2023-05-12 15:13 | Incoming inspection | (TC 05/03/2023 03:12:35 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
1066909 | 194833 | ACC | CMPLT | CMPLT | 2023-05-12 15:14 | 2023-05-12 15:14 | Incoming inspection | (TC 05/03/2023 03:13:47 PM, Cortez, John) Unboxed Visual/functional test ok Alarm check and performance verification passed (Ran for 10 mins at 150 mL/hr)
63035 | 194834 | PM | CMPLT | CMPLT | 2023-05-12 15:16 | 2023-05-12 15:17 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/12/2023 03:16:44 PM, Ferreira, Qwynn) PM Passed Hours: 3403 Avg Intensity: 48.4uW Electrical Safety - OK
66714 | 194835 | PM | CMPLT | CMPLT | 2023-05-12 15:26 | 2023-05-12 15:30 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/12/2023 03:26:59 PM, Ferreira, Qwynn) PM Passed Hours: 6735 avg Intensity: 46.9 Electrical Safety - OK
68106 | 194836 | CM | CANCL | CANCL | 2023-05-12 16:20 |  | Created by Failed PM WO # 183545 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Device set to inactive as it is not currently used | (TC 02/17/2022 04:18:28 PM, Cortez, John) Set up PM software and prepared/gathered testing equipment for first time Will PM unit in case it will be used in the near future Started PM procedures (TC 02/22/2022 11:30:09 AM, Cortez, John) PM results: -strain guage pressure measurement stage 1 - 12.5N stage 3 - 29N stage 8 - 65N -motor power limitation stage 1 - 14.2N stage 3 - 29.5N stage 6 - 52.2N (fail, acceptable range 58-80N) All other checks passed. Will try to calibrate pump due to failure
1062999 | 194837 | CM | CMPLT | CMPLT | 2023-05-12 17:07 | 2023-05-12 17:08 | Cleaning | (TC 04/27/2023 05:07:09 PM, Cortez, John) Completed cleaning cycle
1046015 | 194838 | RQ | CANCL | CANCL | 2023-05-15 01:11 |  | Test |  | This is a Test WO.
67970 | 194839 | CM | CMPLT | CMPLT | 2023-05-15 07:54 | 2023-05-15 08:08 | Bed tilt stuck | (TC 05/15/2023 08:01:28 AM, Rhuland, Courtney) Bed tilt latch not moving Reseated tilt mechanism on tilt platform Functional check - OK
53076 | 194840 | CM | CMPLT | CMPLT | 2023-05-15 08:00 | 2023-05-15 12:00 | battery not charge. | (TC 05/15/2023 12:00:13 PM, Chukri, Samir) repaired and replaced AC power cable .tested working fine retun to department.
63874 | 194841 | CM | CMPLT | CMPLT | 2023-05-15 08:10 | 2023-05-15 08:11 | Missing door | (TC 05/15/2023 08:10:21 AM, Rhuland, Courtney) Came down with two right doors Replaced door and door hinges functional check - OK
1054544 | 194842 | CM | CMPLT | CMPLT | 2023-05-15 08:47 | 2023-05-16 10:03 | GlideScope Overheated | (TC 05/15/2023 08:47:25 AM, Reid, Emmerson ) Monitor had a message: "Monitor has exceeded the safe operating temperature" and the screen was frozen. Left to sit unplugged over the weekend. Reached out to the vendor. (TC 05/16/2023 09:12:09 AM, Reid, Emmerson ) Attached email from CHS and saved to t-drive. It looks like the unit had connected the HDMI for the scope into the HDMI video output on the monitor. Tech support advised that this can cause the overheat error. Ran the monitor with the scope for 1 hour to make sure it is working. Working now.
1046952 | 194843 | CM | CMPLT | CMPLT | 2023-05-15 10:26 | 2023-05-15 17:14 | WIFI not connected. |  | Ref # 021676534 Was never connected.
1061208 | 194844 | CM | CMPLT | CMPLT | 2023-05-15 10:32 | 2023-05-15 10:33 | PACS PC not connecting to monitors | (TC 05/15/2023 10:32:30 AM, Reid, Emmerson ) Calledhelp desk as it was the video cable that was not working.
1056125 | 194845 | PM | CMPLT | CMPLT | 2023-05-15 12:49 | 2023-05-15 12:50 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/15/2023 12:49:11 PM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
1057384 | 194846 | CM | CMPLT | CMPLT | 2023-05-15 14:19 | 2023-05-15 17:13 | MSH CR READER MOVE & NETWORK | (TC 05/12/2023 02:20:09 PM, Dolinar, John) MSH CR READER MOVE & NETWORK (TC 05/15/2023 02:20:48 PM, Whittington, Scott) MSH CR READER MOVE & NETWORK | MSH CR READER MOVE & NETWORK Called MSH IT Ref# N-SD034764125 Jack # 5SE-492-2B
1056665 | 194847 | CM | CMPLT | CMPLT | 2023-05-15 15:12 | 2023-07-20 09:12 | Provido needs the bulbs to be replaced | (TC 05/15/2023 03:12:26 PM, Reid, Emmerson ) Was called up because a message appeared about the life of the bulb. I showed them how to switch between bulbs if needed. Looking to purchase new bulbs. (TC 06/19/2023 10:23:21 AM, Reid, Emmerson ) Placed an order for new bulbs for all the leica microscopes (TC 07/19/2023 08:34:13 AM, Reid, Emmerson ) Recieved 2/3 of the 300W bulbs (TC 07/20/2023 09:10:56 AM, Reid, Emmerson ) Replaced Lamp 1 with Jessica and reset the hours. Lamp 2 is at 270 hours. Tested and working. Also tightened the right handle.
32510 | 194848 | CM | CMPLT | CMPLT | 2023-05-15 17:20 | 2023-05-15 17:20 | Grid would not park properly |  | Silcon oiled the guide tracts.
1064810 | 194849 | CM | CMPLT | CMPLT | 2023-05-16 08:10 | 2023-05-16 09:45 | Not connecting on monitor | (TC 05/15/2023 08:10:09 AM, Canjimavadivel, Kamallesh ) Pins on MSL connector damaged, need to replace MSL board (TC 05/16/2023 09:25:32 AM, Canjimavadivel, Kamallesh ) Replaced MSL board
502044 | 194850 | CM | CMPLT | CMPLT | 2023-05-16 08:11 | 2023-05-16 08:11 | High Humidity | (TC 05/15/2023 08:10:42 AM, Paz, Ramon (Tino)) Replaced head dessicant
1054542 | 194851 | PM | CMPLT | CMPLT | 2023-05-16 08:41 | 2023-05-16 11:16 | Annual PM - For: Lift | Functional check working fine (2023-05-16 10:44 AM Samir Chukri) Functional check and battery (TC 05/16/2023 02:22:16 PM, Rhuland, Courtney) Battery swapped with spare Waiting for internal battery order
| 194852 | CM | CMPLT | CMPLT | 2023-05-16 08:52 | 2023-05-12 12:11 | Bi-Weekly Department Meetings
| 194853 | CM | ACTIV | ISSUE | 2023-05-16 08:52 |  | Working on TMS mobile
| 194854 | CM | ACTIV | ISSUE | 2023-05-16 08:52 | 2023-05-16 08:50 | Alarming not work and the scale not Zero | Zeroing the bed scale and testing the alarm working fine
| 194855 | CM | ACTIV | ISSUE | 2023-05-16 08:53 |  | Working on TMS mobile | (TC 05/30/2023 12:44:45 PM, Chukri, Samir) working on Active unassigned scheduled WO. (TC 06/08/2023 10:13:35 AM, Chukri, Samir) Retired one blender and MRI spo2 monitor and cancel the PM wo.
| 194856 | CM | CMPLT | CMPLT | 2023-05-16 08:53 | 2023-05-16 08:50 | Alarming not work and the scale not Zero | Zeroing the bed scale and testing the alarm working fine
1059587 | 194857 | CM | CMPLT | CMPLT | 2023-05-16 09:48 | 2023-05-16 09:49 | Needs batteries | (TC 05/16/2023 09:48:24 AM, Reid, Emmerson ) Replaced the batteries and ran calibration. Working
1059590 | 194858 | CM | CMPLT | CMPLT | 2023-05-16 10:09 | 2023-05-16 10:10 | Not reading properly | (TC 05/16/2023 10:09:08 AM, Reid, Emmerson ) Therm was reported to be reading wrong temp. I tried calibrating it many times over serveral days as well as cleaning the lense multiple times. Today it passed caliobration. Working now.
1065114 | 194859 | CM | CMPLT | CMPLT | 2023-05-16 11:13 | 2023-05-16 11:14 | Funnel too long | (TC 05/16/2023 11:13:20 AM, Reid, Emmerson ) Went to OPG to cut down a plastic funnel to size for the aquarius system.
1061204 | 194860 | CM | CMPLT | NFF | 2023-05-16 12:45 | 2023-05-16 12:46 | Camera not connecting to monitors | (TC 05/16/2023 12:44:49 PM, Reid, Emmerson ) Went up to OR3, found that the camera SDI cable was not connected to the boom. I connected it to AUX 2 and set the monitors up. Working now, NFF.
| 194861 | CM | CMPLT | CMPLT | 2023-05-16 13:23 | 2023-05-18 08:21 | order batteries for Biomed stock. | (TC 05/18/2023 08:19:06 AM, Chukri, Samir) received PO# 7202214755
| 194862 | CM | CMPLT | CMPLT | 2023-05-16 13:44 | 2023-07-05 15:31 | Order for Alaris PCU power cords | (TC 05/16/2023 01:50:26 PM, Cortez, John) Order for biomed stock
62701 | 194863 | CM | CMPLT | CMPLT | 2023-05-16 14:12 | 2023-05-17 13:23 | check batteries and leads | (TC 05/16/2023 02:21:59 PM, Scero, Tomlyn) Tech said batteries wouldn't charge. Powercord wasn't connected in the back. Plugged in a let charge to see if that's the only issue. replaced USB cable, RL lead, and V4-V6 long chest lead. ECG simulation ok. (TC 05/17/2023 01:21:12 PM, Scero, Tomlyn) reconditioned batteries. they wouldn't fully charge past 75% by plugging them in. replaced batteries with reconditioned ones from the shelf.
68744 | 194864 | CM | CMPLT | CMPLT | 2023-05-16 14:14 | 2023-05-16 14:16 | "Not charging" | (TC 05/16/2023 02:14:44 PM, Cortez, John) PCU able to charge normally, but power cord ground pin bent. Replaced cord Electrical safety check passed
| 194865 | CON | CMPLT | CMPLT | 2023-05-16 14:17 | 2023-08-23 12:16 | MSH DR CR CASSETTE REPLACEMENT SUGGESTION WITH CARESTREAM 3' DIGITAL DETECTOR | (TC 05/17/2023 12:49:53 PM, Whittington, Scott) Hi Scott, Thanks for contacting us. I will send you some details on the DRX-LC tomorrow. Talk to you soon. Jonathan Neville Director of Sales - Canada - Directeur des ventes Medical Imaging Systems - Systèmes d'imagerie médicale jonathan.neville@carestream.com Cell: 613.290.2822 Follow us on EverythingRad Blog Instagram LinkedIn Facebook Twitter YouTube From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Tuesday, May 16, 2023 12:36 PM To: Jordan Berry <jordan.berry@carestream.com>; Mason Nordstrom <mason.nordstrom@carestream.com>; Jonathan Neville <jonathan.neville@carestream.com> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Chan-Dinevski, Ivy - Sinai Health System <Ivy.Chan-Dinevski@sinaihealth.ca>; Noseworthy, Joanne - Mount Sinai Hospital <joanne.noseworthy@sinaihealthsystem.ca>; Sean Young <sean.young@carestream.com>; Pejcic, Sonja <Sonja.Pejcic@uhn.ca> Subject: [EXTERNAL] MSH DR, Long Legs CAUTION: This Email is from an EXTERNAL source. Ensure you trust this sender before clicking on any links or attachments. Hi Could you please provide feasibility, info & pricing on using the new CARESTREAM 3' detector to replace our aging CR system for LLI ? I understand the detector will interface with our Revolution mobiles Thanks Scott 416-402-2462 (TC 05/17/2023 12:49:53 PM, Whittington, Scott) MSH DR CR CASSETTE REPLACEMENT SUGGESTION WITH CARESTREAM 3' DIGITAL DETECTOR Hi Scott, Thanks for contacting us. I will send you some details on the DRX-LC tomorrow. Talk to you soon. Jonathan Neville Director of Sales - Canada - Directeur des ventes Medical Imaging Systems - Systèmes d'imagerie médicale jonathan.neville@carestream.com Cell: 613.290.2822 Follow us on EverythingRad Blog Instagram LinkedIn Facebook Twitter YouTube From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Tuesday, May 16, 2023 12:36 PM To: Jordan Berry <jordan.berry@carestream.com>; Mason Nordstrom <mason.nordstrom@carestream.com>; Jonathan Neville <jonathan.neville@carestream.com> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Chan-Dinevski, Ivy - Sinai Health System <Ivy.Chan-Dinevski@sinaihealth.ca>; Noseworthy, Joanne - Mount Sinai Hospital <joanne.noseworthy@sinaihealthsystem.ca>; Sean Young <sean.young@carestream.com>; Pejcic, Sonja <Sonja.Pejcic@uhn.ca> Subject: [EXTERNAL] MSH DR, Long Legs CAUTION: This Email is from an EXTERNAL source. Ensure you trust this sender before clicking on any links or attachments. Hi Could you please provide feasibility, info & pricing on using the new CARESTREAM 3' detector to replace our aging CR system for LLI ? I understand the detector will interface with our Revolution mobiles Thanks Scott 416-402-2462 | MSH DR CR CASSETTE REPLACEMENT SUGGESTION WITH CARESTREAM 3' DIGITAL DETECTOR
1056055 | 194866 | CM | CMPLT | CMPLT | 2023-05-16 14:53 | 2023-05-16 14:52 | Adjust time with standard time on both the Angio machines and Patient monitoring system. | Fixed.
502231 | 194867 | CM | CMPLT | CMPLT | 2023-05-16 14:55 | 2023-05-16 14:54 | Compression cone problems, C-Arm not working. | Fixed.
1046196 | 194868 | CM | CMPLT | CMPLT | 2023-05-17 11:41 | 2023-05-18 15:59 | CT system upgrade | CT system upgrade ICS tower replaced and software reload/upgrade to Win10 -Download all necessary service packs -Backup/resotore all required system backups (TC 05/17/2023 08:04:16 PM, Debebe, Adam) -Day two of system upgrade -Service pack 6 installaiton and setup completion -All data erased from the old ICS tower (TC 05/18/2023 03:56:53 PM, Debebe, Adam) -Finishid up the past part of the upgrade -Completed all necessary tests -Packed and made ready all parts and components ready for shipment. -System returned for use.
502088 | 194869 | RQ | CMPLT | CMPLT | 2023-05-17 11:56 | 2023-05-18 19:05 | drawer error head #1 | (TC 05/18/2023 07:03:37 PM, Paz, Ramon (Tino)) Lubricated drawer 1. Tested OK. Adviced customer to notify us if the problem comes back. | able to clear and proceed, but recurrent
1061211 | 194870 | CM | CMPLT | CMPLT | 2023-05-17 13:01 | 2023-05-17 13:55 | VMC froze | (TC 05/17/2023 01:54:06 PM, Reid, Emmerson ) VMC froze, and the touch screen was not responding. I restarted the VMC and it is working now. Note the wall monitor was not able to show the 3D scope.
| 194871 | CM | ACTIV | ISSUE | 2023-05-17 13:34 | 2023-05-17 13:14 | Room 1 imaging issue | Checked the room no problem found monitor working fine
| 194872 | CM | ACTIV | ISSUE | 2023-05-17 13:34 | 2023-05-17 13:32 | Alarming off ( adjusted the width) | Re adjusted the width and rest the alarm working fine
| 194873 | CM | CMPLT | CMPLT | 2023-05-17 13:43 | 2023-05-17 13:14 | Room 1 imaging issue | Checked the room no problem found monitor working fine (2023-05-17 1:13 PM Samir Chukri) No problem found
1062548 | 194874 | CM | CMPLT | CMPLT | 2023-05-17 13:43 | 2023-05-17 13:32 | Alarming off ( adjusted the width) | Re adjusted the width and rest the alarm working fine
| 194875 | CM | ACTIV | ISSUE | 2023-05-17 13:53 | 2023-05-17 13:51 | Alarm off adjust width | Re-adjusted and rest the alarm working fine
| 194876 | CM | CANCL | CANCL | 2023-05-17 13:55 | 2023-05-17 13:51 | Alarm off adjust width
1061206 | 194877 | CM | CMPLT | CMPLT | 2023-05-17 13:56 | 2023-05-17 13:56 | Vitals not on monitor | (TC 05/17/2023 01:55:53 PM, Reid, Emmerson ) AGM vital monitor was not connected to the boom. Working now.
1061214 | 194878 | CM | ACTIV | ISSUE | 2023-05-17 14:02 |  | Interference from ForceTriad | (TC 05/17/2023 02:02:11 PM, Reid, Emmerson ) Went with Maisha to look at OR13, also showed her the forcetriade, olympus cabinates, lim monitor and explained how the image is displayed and ways we tried to troubleshoot (TC 05/17/2023 02:04:02 PM, Reid, Emmerson ) Brought down the ForceTriad to show Samir, we rans some tests but nothing conclusive. Went up to OR to test more things in the room.
63870 | 194879 | CM | CMPLT | CMPLT | 2023-05-17 14:28 | 2023-05-25 15:47 | Noisy resuscitation system | (TC 05/17/2023 12:21:33 PM, Rendon Lopez , Elizabeth) Received blender from overhaul, replaced. system is still noisy. Troubleshooting with Getinge, issue is with check valve close to cylinder input (staying open). Ordered parts needed (TC 05/24/2023 03:41:24 PM, Rendon Lopez , Elizabeth) Replaced cylinder check valves, issue not resolved, replaced air hose inlet check valve, working. Running for some time, noise started to happen again, blender O2 reading is off value. (TC 05/25/2023 03:41:24 PM, Rendon Lopez , Elizabeth) Replaced air and O2 hose inlet check valves, adjusted blender knob. All check valves are brand new, system is no longer leaking or noisy. Tested all functions-OK
| 194880 | CM | CMPLT | CMPLT | 2023-05-17 14:30 | 2023-05-23 11:25 | Panda warmer spare parts | Received
| 194881 | CM | CMPLT | CMPLT | 2023-05-17 15:05 | 2023-05-17 14:57 | Interference with OR monitors | (2023-05-17 2:10 PM Samir Chukri) Working With Emmerson try to replicate the error We are unable to replicate the problem
1062233 | 194882 | CM | CMPLT | CMPLT | 2023-05-17 15:13 | 2023-05-17 14:57 | Interference with OR monitors | (2023-05-17 2:10 PM Samir Chukri) Working With Emmerson try to replicate the error We are unable to replicate the problem
| 194883 | CM | CMPLT | CMPLT | 2023-05-17 17:04 | 2023-05-17 17:04 | Site Coverage
67936 | 194884 | CM | CMPLT | CMPLT | 2023-05-18 08:11 | 2023-06-06 08:23 | ordering hoods for Giraffe incubators | ordering hoods for Giraffe incubators. Received hoods. PO# 7202216154
| 194885 | CM | CMPLT | CMPLT | 2023-05-18 09:54 | 2023-05-18 09:53 | Replaced battery | (2023-05-18 9:49 AM Samir Chukri) Completed and replaced two batteries
| 194886 | CM | CMPLT | CMPLT | 2023-05-18 10:01 | 2023-05-18 09:53 | Replaced battery | (2023-05-18 9:49 AM Samir Chukri) Completed and replaced two batteries
1054806 | 194887 | CM | CMPLT | CMPLT | 2023-05-18 10:05 | 2023-05-18 11:40 | Replaced Batteries. | (TC 05/18/2023 10:05:25 AM, Chukri, Samir) Repalced two batteries.
1056665 | 194888 | CM | CMPLT | CMPLT | 2023-05-18 10:15 | 2023-05-18 10:17 | Microscope handle is loose | (TC 05/18/2023 10:15:19 AM, Reid, Emmerson ) Went to tighten the handle, also sent an email regarding how to properly adjust in the future to avoid this issue.
| 194889 | CM | CMPLT | CMPLT | 2023-05-18 12:14 | 2023-05-24 13:26 | Fetal monitors replacement recorder and spare | Received
1066115 | 194890 | FC | CMPLT | CMPLT | 2023-05-18 12:57 | 2023-05-18 12:57 | QC check | (TC 05/18/2023 12:56:41 PM, Canjimavadivel, Kamallesh ) QC check complete
1054381 | 194891 | CM | CMPLT | CMPLT | 2023-05-18 13:21 | 2023-05-18 13:27 | Blood pressure not reading | (TC 05/18/2023 01:20:58 PM, Rhuland, Courtney) BP cuff had a leak Replaced cuff Functional - OK
1065164 | 194892 | CM | CMPLT | CMPLT | 2023-05-18 13:42 | 2023-08-08 13:49 | Will not power on | (TC 07/06/2023 01:45:22 PM, Ferreira, Qwynn) Tested the device and it is able to power up occasionally but also powers up on its own It will also shut down randomly, and the power button does not always work. Will need to send the device out. (TC 08/08/2023 01:48:12 PM, Ferreira, Qwynn) received device Uploaded certificate to TMS Assembled the device Imported configuration Tested the device for functionality and connection to the server
| 194893 | PA | CMPLT | CMPLT | 2023-05-18 13:59 | 2023-08-18 10:43 | Drain pan for Altrix | (TC 06/02/2023 08:03:12 AM, Canjimavadivel, Kamallesh ) Placed order for drain pan (TC 08/18/2023 10:30:31 AM, Canjimavadivel, Kamallesh ) Received drain pan
1051031 | 194894 | CM | CMPLT | CMPLT | 2023-05-18 14:05 | 2023-05-29 10:32 | System crashing | (TC 05/18/2023 10:05:00 AM, Fan, Tak) System unable to store and crashes when looking at spooler. Could not access spooler to clear studies. Logged into windows and renamed Spooler file to Spooler2. Restarted application. OK. Sent latest study, which was also stuck on the spooler OK. Ashraf came by. Other images are research from an external USB. She will check with Faith to see how the other images were transferred and see what can be done in the dataflow. Leaving images on Spooler2 for now. (TC 05/25/2023 03:25:24 PM, Fan, Tak) Deleted saved spooler data from last time. Troubleshot sending with Faith and Ashraf. Had two dataflows setup with "DICOM Storage", not sure if that confuses the system, so one was removed from Query Retrieve dataflow. Changing the compression and quality only reduced the study by 25%. For some reason sending from USB to local was 1.26 gb but USB to Viewpoint was 10.1 gb. USB to local to Viewpoint was also 10.1 gb. Not sure why. Contacted GE, they suggested sw upgrade. (TC 05/29/2023 03:27:23 PM, Fan, Tak) Followup with GE. Tony will have a look. Tony requested system info but system is not on floor.
| 194895 | CON | CMPLT | CMPLT | 2023-05-18 14:14 | 2023-05-18 14:14 | Checking DR 1 for cassette issue
1056665 | 194896 | CM | CMPLT | CMPLT | 2023-05-18 14:26 | 2023-05-18 14:32 | Issue with the Focusing motor | (TC 05/18/2023 02:26:18 PM, Reid, Emmerson ) Was called to OR 13 because the hand controlls were not working to adjust the focusing (working) distance. They switched microscopes for the case and I was able to look at it and speak with steve. I found that the controlls got stick at either end of the max focusing distance, however the on screen measurement of the distance did not line up with the position. With the lense all the way up (furthest distance) it read 571 when it should have read 600. Then it would decrease to 225 (the minimum) before reaching all the way back down. With another half turn down on the manual knob the screen would read 600 when it should have read 225. All the values seem shifted up and this was messing with the controlls. Even after a few restarts it did not resolve this issue. I moved the microscope and it was working again. Ran all motor calibration from the service menue and all passed. Working now but will have a service call if this happens again.
1057835 | 194897 | PM | CMPLT | CMPLT | 2023-05-19 05:22 | 2023-05-19 08:56 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 05/19/2023 05:28:48 AM, Reid, Emmerson ) Ran PM proceedure, all tests passed. See results attached to WO. Set ESU back to 45, 45, 20 as it was, connected foot pedals and made sure they were working. PM complete.
1046455 | 194898 | PM | CMPLT | CMPLT | 2023-05-19 06:32 | 2023-05-19 06:45 | PM MAJOR - ANNUAL PM for: electrosurgical units | (TC 05/19/2023 06:32:29 AM, Reid, Emmerson ) Moved from cart to OR 1 boom, synced smoke evac. See attached PM report for details. PM Complete.
1049050 | 194899 | PM | CMPLT | CMPLT | 2023-05-19 07:00 | 2023-05-19 07:20 | Planned Maintenance | (TC 05/19/2023 07:19:11 AM, Reid, Emmerson ) PM complete, see proceedure for details.
66464 | 194900 | CM | CMPLT | CMPLT | 2023-05-19 07:37 | 2023-05-31 07:38 | Ordering VN500 parts | Ordering VN500 parts. Received valves. PO#702217501
65387 | 194901 | PM | CMPLT | CMPLT | 2023-05-19 07:55 | 2023-05-18 07:55 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. NO PS500 and NO GS500. Year 2010. replaced exp.vave cover. Tested. Functional. SW.2.42 Leakage: 79 uA and 81ohm.
1061202 | 194902 | CM | CMPLT | CMPLT | 2023-05-19 09:00 | 2023-05-19 09:01 | Maintenance Check | (TC 05/19/2023 09:00:27 AM, Reid, Emmerson ) Checked all the booms, monitors, lights and the tablet. Had to tighten monitor 1 a little bit.
1061205 | 194903 | CM | CMPLT | CMPLT | 2023-05-19 09:02 | 2023-05-19 09:03 | Maintenance check | (TC 05/19/2023 09:02:08 AM, Reid, Emmerson ) Checked all the booms, monitors, lights and the tablet.
1061208 | 194904 | CM | CMPLT | CMPLT | 2023-05-19 09:03 | 2023-05-19 09:05 | Maintenance check | (TC 05/19/2023 09:03:27 AM, Reid, Emmerson ) Checked all the booms, monitors, lights and the tablet. The foot boom has the middle brake (orange) not working. I opened it up and reseated the cable, working now. The anesthesia boom had the purple brake very loose so I made it more firm.
1064207 | 194905 | CM | CMPLT | CMPLT | 2023-05-19 09:04 | 2023-05-19 09:05 | Installed power bar | (TC 05/19/2023 09:04:08 AM, Canjimavadivel, Kamallesh ) Installed power bar on ventilator
| 194906 | CM | CMPLT | CMPLT | 2023-05-19 09:25 | 2023-05-19 09:22 | Installed power bar | Helping Kam to install power bar
1058534 | 194907 | CM | CMPLT | CMPLT | 2023-05-19 09:27 | 2023-05-19 09:29 | Bladder scanner Not scanning | (TC 05/19/2023 09:27:37 AM, Reid, Emmerson ) Was emailed by Eden saying the device was reading inaccuratly. The tag on the device said it was not scanning. When I ran the self-test, the probe SN was not appearing. I reconnected the probe and it registered and passed the self test.
| 194908 | CM | CMPLT | CMPLT | 2023-05-19 09:28 | 2023-05-19 09:22 | Installed power bar | Helping Kam to install power bar
1063372 | 194909 | CM | CMPLT | CMPLT | 2023-05-19 10:19 | 2023-06-20 09:49 | Surface fault error code | (TC 05/19/2023 10:19:29 AM, Canjimavadivel, Kamallesh ) Found multiple error codes on the bed. Also found wire near left side foot that was severed. Need to bring in vendor for repair. (TC 06/20/2023 09:47:19 AM, Cortez, John) Vince from Hillrom completed repair I confirmed functionality, checked if error codes cleared | Multiple error codes: 1041 A0C5 A0C9
| 194910 | PA | CMPLT | CMPLT | 2023-05-19 11:02 | 2023-08-23 08:03 | Parts for ICU beds | (TC 05/19/2023 11:25:34 AM, Canjimavadivel, Kamallesh ) Placed order for ICU bed PM parts (TC 08/09/2023 01:52:50 PM, Canjimavadivel, Kamallesh ) Partially received the order. Unboxed the batteries and cords. (TC 08/10/2023 01:14:51 PM, Canjimavadivel, Kamallesh ) Received batteries. Unloaded skid and brought boxes into biomed (TC 08/23/2023 08:01:56 AM, Canjimavadivel, Kamallesh ) Received all items from order
1061385 | 194911 | CM | CMPLT | CMPLT | 2023-05-19 11:34 | 2023-05-19 11:35 | Navigation When viewed on the monitor image was very distrupted | (TC 05/19/2023 11:34:03 AM, Reid, Emmerson ) Image was not viewable on monitor. Tried different AUX ports, looked the same. Needed to be restarted. working now. (TC 05/23/2023 01:50:36 PM, Whittington, Scott) JADE DOMINICIS FROM MEDTRONIC INQUIRING RE PACS CONNECTION? OF SYSTEMS?
1061214 | 194912 | CM | CMPLT | CMPLT | 2023-05-19 11:36 | 2023-05-19 11:36 | Power Not wokring on the boom | (TC 05/19/2023 11:36:13 AM, Reid, Emmerson ) Power on working for Smoke evac. The boom power plug LP-7SW-NO-18 had no power.
1055181 | 194913 | CM | CMPLT | CMPLT | 2023-05-19 14:38 | 2023-05-19 14:39 | Channel erros | (TC 05/18/2023 02:38:29 PM, Rendon Lopez , Elizabeth) Downloaded error log, Bottle pressure sensor fatal error. Replaced presure sensor, calibrated (TC 05/19/2023 02:39:14 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69461 | 194914 | CM | CMPLT | CMPLT | 2023-05-19 14:40 | 2023-05-19 14:46 | Very loud red alarm | (TC 05/18/2023 02:43:35 PM, Rendon Lopez , Elizabeth) Downloaded error, bottle pressure sensor error. Replaced pressure sensor, calibrated-OK
66500 | 194915 | CM | CMPLT | CMPLT | 2023-05-19 14:50 | 2023-05-19 14:52 | Unspecified problem, not working | (TC 05/18/2023 02:50:35 PM, Rendon Lopez , Elizabeth) Constant check paper alarm on monitor. Checked recorder, defective paper sensor. Replaced recorder with spare, ordered more spare. Tested all functions-OK
52470 | 194916 | CM | CMPLT | CMPLT | 2023-05-19 14:55 | 2023-05-19 14:56 | Returned to service | (TC 05/18/2023 02:54:55 PM, Rendon Lopez , Elizabeth) Returned to service in L&D OR to replaced defective H1000 unit there that was defective and retired.
500857 | 194917 | CM | CMPLT | CMPLT | 2023-05-19 17:43 | 2023-05-19 17:43 | Memory Parity Error |  | Rebooted System. Tested Ok
| 194918 | CM | CMPLT | CMPLT | 2023-05-19 17:44 | 2023-05-19 17:44 | Site Coverage
502231 | 194919 | CM | CMPLT | CMPLT | 2023-05-19 18:02 | 2023-05-19 18:01 | System needs re-boot every morning. | Problem was due to an air bubble in the tube cooling unit, removed the bubble to fix the problem.
| 194920 | TEDU | CMPLT | CMPLT | 2023-05-23 07:30 | 2023-05-19 07:32 | CMBES45 - VANCOUVER | (TC 05/15/2023 07:30:17 AM, Samanian, Kamran) TO VANCOIVER FOR CMBES45 (TC 05/16/2023 07:30:54 AM, Samanian, Kamran) CMBES45 (TC 05/17/2023 07:31:22 AM, Samanian, Kamran) CMBES45 (TC 05/18/2023 07:31:44 AM, Samanian, Kamran) CMBES45 (TC 05/19/2023 07:32:00 AM, Samanian, Kamran) BACK FROM CMBES45
| 194921 | OC | CMPLT | CMPLT | 2023-05-23 07:40 | 2023-05-23 07:44 | Kamran Samanian On-Call, 22 MAY 2023 | (TC 05/22/2023 09:15:00 AM, Samanian, Kamran) received a call from ICU that the computer in alcove ( POD B room 14 ) is black screen however staff checked all power cords and cables and they seem to be fine. Guided staff ( Lena ) to contact helpdesk to open a ticket and ask to resolve this issue. Also mentioned that when it's fixed and if there is any problem with Philips MPV she can call me back.
1056125 | 194922 | PM | CMPLT | CMPLT | 2023-05-23 08:41 | 2023-06-07 11:48 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 06/07/2023 11:47:09 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1061214 | 194923 | CM | CMPLT | CMPLT | 2023-05-23 09:54 | 2023-05-24 08:07 | Light iPad not charging | (TC 05/23/2023 09:57:16 AM, Reid, Emmerson ) Went to OR 13 to check the iPad, it was dead and not charging. I opened up the wall compartment and tried reseating the charging cable on both ends, it did not work. The cable was damaged from the way the iPad was put into the wall. This crimped the cable and broke it. Not sure when this happened or how. I have attached a picture to the WO. I took the cable from OR 12 and installed it into OR13 so it is charging now. Need to order a new cable for OR 12. (TC 05/23/2023 10:52:45 AM, Chukri, Samir) Ordered from Amazon.ca (TC 05/24/2023 08:06:45 AM, Reid, Emmerson ) Recieved both cables.
1059618 | 194924 | CM | CMPLT | CMPLT | 2023-05-23 10:10 | 2023-05-23 11:09 | Microscope light not working | (TC 05/23/2023 11:06:19 AM, Reid, Emmerson ) Complaining that the light was too dim. I saw the LED was very dim and flickered a bit. I thought i would need a new bulb but when I looked at the age of the microscope I thought otherwise and brought it down to biomed to look at it. I tested the voltage and it looked good. Then I adjusted one of the knobs and turned the brightness up. It was just that the knob was turned all the way down. NFF
62618 | 194925 | CM | CMPLT | CMPLT | 2023-05-23 10:22 | 2023-05-23 11:44 | Heater door does not close | (TC 05/23/2023 11:43:40 AM, Rhuland, Courtney) Completed in PM work order
1061205 | 194926 | CM | CMPLT | CMPLT | 2023-05-23 11:13 | 2023-05-23 11:15 | Light iPad was not charging | (TC 05/23/2023 11:13:35 AM, Reid, Emmerson ) iPad was not charging. I reseated the USB end of the connection and it started charging. Checked 20 minutes later and it was still charging. Working now.
500879 | 194927 | ACC | CMPLT | CMPLT | 2023-05-23 13:24 | 2023-05-23 13:32 | MSH-5-US PVT 375 ARTIFACTS SN# FMA1722130 | (TC 05/23/2023 01:26:17 PM, Whittington, Scott) MSH-5-US PVT 375 ARTIFACTS SN# FMA1722130 MFG FEB 2017 REPLACED BY TWH 502253 SN# FMC1893334 MFG SEP 2018 | MSH-5-US PVT 375 ARTIFACTS SN# FMA1722130
1058649 | 194928 | CM | CMPLT | CMPLT | 2023-05-23 13:34 | 2023-08-23 12:12 | MSH 5 US PROBE PVT661 HAS ARTIFACTS, MAY BE MEMBRANE | (TC 05/23/2023 01:33:50 PM, Whittington, Scott) MSH 5 US PROBE PVT661 HAS ARTIFACTS, MAY BE MEMBRANE | MSH 5 US PROBE PVT661 HAS ARTIFACTS, MAY BE MEMBRANE
1046196 | 194929 | CM | CMPLT | CMPLT | 2023-05-23 13:44 | 2023-05-23 13:43 | Remote connectivity issue/slow Application reported | Remote connectivity issue/slow Application reported SIEMENS will follow up with VPN testing on Remote
66069 | 194930 | PM | CMPLT | CMPLT | 2023-05-23 13:56 | 2023-05-24 10:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/24/2023 10:18:32 AM, Reid, Emmerson ) Replaced the battery. Ran ASM MSH PM Pcu, all tests passed. PM Complete.
| 194931 | CM | CMPLT | CMPLT | 2023-05-23 13:57 | 2023-05-25 12:09 | order three power bar from digikey.ca | (TC 05/23/2023 01:57:31 PM, Chukri, Samir) Ordered three power bar for servo-u vent. (TC 05/25/2023 12:09:16 PM, Chukri, Samir) received
1059795 | 194932 | CM | CMPLT | CMPLT | 2023-05-23 14:43 | 2023-05-23 14:42 | Misceleneous
| 194933 | CM | CMPLT | PMFAIL | 2023-05-23 15:31 | 2023-05-23 15:31 | Initial meeting for Equipment Replacement Planning Project Team
1054963 | 194934 | CM | CMPLT | NFF | 2023-05-23 16:48 | 2023-05-23 16:54 | sent down with yellow tag, no problem description | (TC 05/23/2023 04:48:36 PM, Cortez, John) All parameters tested ok, no power cord, provided one wrong probe covers, removed No fault found
1061991 | 194935 | CM | CMPLT | CMPLT | 2023-05-23 16:55 | 2023-05-23 16:56 | bed alarm not triggering | (TC 05/23/2023 04:55:36 PM, Cortez, John) Power cycled bed, Functinional check passed (scale, motions, bed alarms)
67936 | 194936 | CM | CMPLT | CMPLT | 2023-05-24 08:37 | 2023-05-24 08:39 | Broken porthole door | (TC 05/24/2023 08:37:28 AM, Rhuland, Courtney) Broken porthole door replaced Functional check - OK
32580 | 194937 | CM | CMPLT | CMPLT | 2023-05-24 09:15 | 2023-05-23 18:00 | XC Error | XC error Filament timeout Checked all error logs Comleted IF cal Did mutiple test scans - also using patient protocol that was used during when the error occured.
1054543 | 194938 | CM | CMPLT | CMPLT | 2023-05-24 09:30 | 2023-05-24 09:55 | "Not charging" | (TC 05/15/2023 09:30:05 AM, Cortez, John) Cord loose on power supply, charged battery overnight, Next morning, battery still at 25% Verified power supply output, and monitored charging behaviour in system info Unsure why it didnt charge overnight, will continue to monitor Secured power cord with tyrap
1055906 | 194939 | CM | CMPLT | CMPLT | 2023-05-24 09:58 | 2023-06-09 13:40 | Broken mount | (TC 05/24/2023 09:42:17 AM, Cortez, John) Ordered Root quick connect kit (TC 06/09/2023 01:39:13 PM, Cortez, John) Installed stand, functional check ok
67918 | 194940 | CM | CMPLT | CMPLT | 2023-05-24 10:43 | 2023-05-24 11:36 | Wrong door attached | (TC 05/24/2023 10:43:30 AM, Rhuland, Courtney) Door with new latches attached on the wrong side of incubator Replaced door (spare) and replaced side door latches and porthole latches (recall) Power cord replaced Functional check - OK
| 194941 | CM | CMPLT | CMPLT | 2023-05-24 11:02 | 2023-07-05 15:31 | Alaris parts order- PCU batteries, LVP bezels | (TC 05/24/2023 11:07:24 AM, Cortez, John) Ordered LVP bezels and PCU batteries for stock (TC 06/05/2023 02:06:26 PM, Cortez, John) batteries arrived
| 194942 | CM | CMPLT | CMPLT | 2023-05-24 11:25 | 2023-07-18 14:08 | Order for X3 batteries, biomed stock | (TC 05/24/2023 11:29:13 AM, Cortez, John) X3 battery order for biomed stock (TC 07/18/2023 02:07:25 PM, Cortez, John) received items
501106 | 194943 | CM | CMPLT | CMPLT | 2023-05-24 11:32 | 2023-05-29 16:44 | MSH 5 US IU22 POWER CORD REPLACEMENT | (TC 05/24/2023 11:32:15 AM, Dolinar, John) MSH 5 US IU22 POWER CORD REPLACEMENT | MSH 5 US IU22 POWER CORD REPLACEMENT Replaced power cord
| 194944 | CM | CMPLT | CMPLT | 2023-05-24 12:40 | 2023-05-24 12:40 | Working with Brandon to set up APs on 18th floor | (TC 05/24/2023 12:40:06 PM, Samanian, Kamran) done!
| 194945 | CM | CMPLT | CMPLT | 2023-05-24 13:02 | 2023-05-23 13:02 | replacing the lamp socket in OR exam light | (TC 05/23/2023 01:02:24 PM, Samanian, Kamran) replaced the socket, replaced the lamp, tested and passed
1061220 | 194946 | CM | CMPLT | CMPLT | 2023-05-24 13:30 | 2023-05-24 13:36 | Stryker scope tower image was looking green | (TC 05/23/2023 01:30:26 PM, Reid, Emmerson ) Stryker tower light source was visibly green and this was causing the camera image to look very red. Discussed with Jessica and she said that the tower is very old and not supported, as well as Reeba does not want to spend money to repair. Left for now.
68369 | 194947 | CM | CMPLT | CMPLT | 2023-05-24 14:12 | 2023-05-24 14:13 | monitor reconfig | (TC 05/24/2023 02:12:20 PM, Scero, Tomlyn) moved philips monitor location for 3 month trial
68377 | 194948 | CM | CMPLT | CMPLT | 2023-05-24 14:13 | 2023-05-24 14:14 | monitor reconfiguration | (TC 05/24/2023 02:13:33 PM, Scero, Tomlyn) moved philips monitor location for 3 month trial
68554 | 194949 | CM | CMPLT | CMPLT | 2023-05-24 14:14 | 2023-05-24 14:15 | monitor reconfiguration | (TC 05/24/2023 02:14:18 PM, Scero, Tomlyn) moved philips monitor location for 3 month trial
| 194950 | CM | CMPLT | CMPLT | 2023-05-24 17:27 | 2023-05-24 17:27 | Site Coverage
1064964 | 194951 | CM | CMPLT | CMPLT | 2023-05-24 17:58 | 2023-05-24 17:57 | SoftwARE UPGRADE NEEDED. | Complete
67928 | 194952 | CM | CMPLT | CMPLT | 2023-05-25 07:56 | 2023-08-10 07:57 | ordering Giraffe incubator parts | ordering Giraffe incubator parts. Received parts. PO#7202219643
66454 | 194953 | CM | CMPLT | CMPLT | 2023-05-25 08:29 | 2023-05-25 08:29 | internal air leak | replaced air filter without cable tie. Tested. Functional. No leak.
| 194954 | CM | CMPLT | CMPLT | 2023-05-25 09:19 | 2023-05-25 09:20 | EEG windows 7 problem, stuck on the page windows cleaning up | (TC 05/25/2023 09:19:23 AM, Samanian, Kamran) verified the problem, taken the windows to the safe mode, and then restarted normaly , issue was resolved.
1062031 | 194955 | CM | CMPLT | CMPLT | 2023-05-25 09:21 | 2023-05-25 09:22 | Head section not moving | (TC 05/25/2023 09:21:01 AM, Cortez, John) Confirmed error code 3009-1-9-1331 (limited motion functionality) Power cycled bed, cleared error log Tested ok
1061205 | 194956 | CM | CMPLT | CMPLT | 2023-05-25 09:31 | 2023-05-25 09:32 | Nurse PC and VMC monitor were not working | (TC 05/25/2023 09:30:50 AM, Reid, Emmerson ) Unplugged VMC monitor and plugged in again fixed the touchscreen issue. Checked the cabling for the nurse PC and tried turning on and it worked.
1061209 | 194957 | CM | ACTIV | ISSUE | 2023-05-25 09:33 |  | OR 8 Wall Monitor not working | (TC 05/25/2023 09:39:33 AM, Reid, Emmerson ) Attached an image of the issue. Wall monitor image has a lot of noise showing up while the other monitors are fine. I tried reconnecting the SDI cable but did not working. Reached out to Derrick from Olympus to help. (TC 06/08/2023 11:31:28 AM, Reid, Emmerson ) Was up with Olympus to look at this room. I showed them the issue and went through some troubleshooting. Nothing uncovered yet.
| 194958 | CM | CMPLT | CMPLT | 2023-05-25 09:57 | 2023-05-25 09:55 | Battery chargers not connect | (2023-05-25 9:55 AM Samir Chukri) Completed and connected the battery chargers
| 194959 | CM | CMPLT | CMPLT | 2023-05-25 09:57 | 2023-05-25 08:49 | Wall receptacle need to be check no power zone C NICU | Informed NICU administrator to request service building repair
| 194960 | CM | CMPLT | CMPLT | 2023-05-25 10:02 | 2023-05-25 10:02 | Neptune - quick Reconnect | (TC 05/25/2023 10:01:58 AM, Chukri, Samir) Meeting.
| 194961 | CM | CMPLT | CMPLT | 2023-05-25 10:18 | 2023-05-25 09:55 | Battery chargers not connect | (2023-05-25 9:55 AM Samir Chukri) Completed and connected the battery chargers
| 194962 | CM | CMPLT | CMPLT | 2023-05-25 10:18 | 2023-05-25 08:49 | Wall receptacle need to be check no power zone C NICU | Informed NICU administrator to request service building repair (2023-05-25 8:46 AM Samir Chukri) Completed
32520 | 194963 | CM | CMPLT | CMPLT | 2023-05-25 10:21 | 2023-05-29 11:56 | MSH DR2 DETECTOR Z NOT MOVING | (TC 05/25/2023 03:37:33 PM, Whittington, Scott) MSH DR2 DETECTOR Z NOT MOVING INITIALLY NO DET Z , BUT X, Y OK DISCONNECTED POWER TO Z 15 AMP CONTROLLER, LOST X, Y & WATCHDOG 36V PULLED DOWN TO 7V FUSE IN Z CONTROLLER BLOWN LOCATE 15A CONTROLLER IN GW, DIFFICULT TO EXTRACT PT USE CONTINUED IN DR2 USING MOBILE TO CLEAR PT BACKLOG CONTINUE WITH 36 V INVESTIGATION & ASSUME THAT 15A CONTROLLER REQ REPLACEMENT (TC 05/26/2023 01:04:49 PM, Whittington, Scott) MSH DR2 DETECTOR Z NOT MOVING RESOLVED 36V ISSUE LOCATED 15A CONTROLLER MOUNT ORIGINAL CONTROLLER HEAD ON FUNFTIONAL POWER PORTION (TC 05/26/2023 03:45:33 PM, Dolinar, John) MSH DR2 DETECTOR Z NOT MOVING (TC 05/25/2023 09:55:00 PM, Whittington, Scott) MSH DR2 DETECTOR Z NOT MOVING | MSH DR2 DETECTOR Z NOT MOVING
502044 | 194964 | CM | CMPLT | CMPLT | 2023-05-25 10:39 | 2023-09-13 19:42 | Brake Adjustment | (TC 05/25/2023 10:38:47 AM, Paz, Ramon (Tino)) Arm brake loose once more. Adjusted for now. Got an adjustment tip from Michael. We will try it today.
1059710 | 194965 | CM | CMPLT | NFF | 2023-05-25 11:29 | 2023-05-25 11:31 | ECG not working | (TC 05/25/2023 11:29:50 AM, Reid, Emmerson ) Connected to a simulator and it was working. I suspect that they were using bad electrodes. NFF
1061211 | 194966 | CM | ACTIV | ISSUE | 2023-05-25 11:38 |  | Monitor #2 briefly shut down in the middle of a case | (TC 05/25/2023 11:37:49 AM, Reid, Emmerson ) Reported monitor #2 shut down at critical point in the case, but only once. Will monitor and report to Olympus. (TC 05/25/2023 12:31:23 PM, Reid, Emmerson ) Was called again. This time all the monitors went down briefly. I checked to see if there were any loose cables and it didn't seem so.
1061204 | 194967 | CM | CMPLT | CMPLT | 2023-05-25 11:39 | 2023-05-25 11:40 | PiP on monitor showing | (TC 05/25/2023 11:39:18 AM, Reid, Emmerson ) Jessica asked for help becuase the corner of the screen was not showing. I turned the PiP off on the monitor and it was working.
1048523 | 194968 | PM | CMPLT | CMPLT | 2023-05-25 13:41 | 2023-05-25 13:56 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/25/2023 01:41:35 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1063746 | 194969 | PM | CMPLT | CMPLT | 2023-05-25 13:57 | 2023-05-25 14:39 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/25/2023 02:38:34 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
68840 | 194970 | PM | CMPLT | CMPLT | 2023-05-25 15:03 | 2023-05-25 15:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/25/2023 03:04:21 PM, Reid, Emmerson ) Replaced the battery. Ran ASM PM, all tests passed. PM Complete.
65954 | 194971 | PM | CMPLT | CMPLT | 2023-05-25 15:09 | 2023-05-25 15:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/25/2023 03:09:01 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1063708 | 194972 | CM | CMPLT | CMPLT | 2023-05-25 15:14 | 2023-07-31 10:32 | Error when starting up | (TC 05/25/2023 03:14:06 PM, Fan, Tak) Checked error but can't do anything. Called in for support. Called back by support. Windows access via "Ctrl+Alt+Shift+R" password "sierra". Can retrieve logs from D drive. Sent logs, for support. Suggested disabling the "Intel Audio Driver", not sure why. Made the change but the system was working. Restarted a few times, no problem. (TC 05/29/2023 08:14:00 AM, Fan, Tak) Error again on Fri. Retrieved logs again and uploaded. (TC 06/01/2023 11:00:31 AM, Fan, Tak) SW engineer from Fujifilm replied. Error is due to the system sending images and scanning at the same time (???). Work around is to disable Wifi when scanning then turn it back on when sending... Forwarded their reply to Faith. (TC 06/01/2023 12:12:00 PM, Fan, Tak) System was never on wifi. Confirmed. Kept system off wifi, Faith will inform the users not to connect to the network when scanning for now. No reply from Fujifilm.
68537 | 194973 | CM | CMPLT | CMPLT | 2023-05-26 07:52 | 2023-05-26 07:53 | monitor configuration | (TC 05/24/2023 07:52:28 AM, Scero, Tomlyn) changed the AGM layout for 3 month trial
1058056 | 194974 | CM | CMPLT | CMPLT | 2023-05-26 08:05 | 2023-05-26 08:05 | monitor config | (TC 05/24/2023 08:05:04 AM, Scero, Tomlyn) changed the AGM layout for 3 month trial
68374 | 194975 | CM | CMPLT | CMPLT | 2023-05-26 08:06 | 2023-05-26 08:06 | monitor config | (TC 05/24/2023 08:05:54 AM, Scero, Tomlyn) changed the AGM layout for 3 month trial
68368 | 194976 | CM | CMPLT | CMPLT | 2023-05-26 08:07 | 2023-05-26 08:07 | monitor config | (TC 05/24/2023 08:06:40 AM, Scero, Tomlyn) changed the AGM layout for 3 month trial
68799 | 194977 | CM | CMPLT | CMPLT | 2023-05-26 08:18 | 2023-07-04 10:45 | Bent power cord | (TC 07/04/2023 10:42:34 AM, Rhuland, Courtney) Power cord replaced Electrical safety - OK
1061968 | 194978 | CM | CMPLT | CMPLT | 2023-05-26 10:04 | 2023-05-26 10:07 | motion controls not working | (TC 05/26/2023 10:05:00 AM, Cortez, John) Confirmed issue, Power cycling bed seems to resolve issue, but this is a repeat problem from a month ago. Will monitor, but inform nurse to put back into service
1061205 | 194979 | CM | CMPLT | CMPLT | 2023-05-26 10:21 | 2023-05-26 10:24 | VMC monitor touch screen not working | (TC 05/26/2023 10:21:38 AM, Reid, Emmerson ) The touchscreen curser was stuck on one spot and would not let me press anywhere. I tried power cycling the monitor and reseating the cables but this did not fix it. I switched the monitor with the one from OR12 and then replace the OR12 monitor with a spare from biomed. Working now.
67358 | 194980 | CM | CMPLT | CMPLT | 2023-05-29 07:45 | 2023-05-29 07:46 | Turning off on its own | (TC 05/29/2023 07:44:47 AM, Rhuland, Courtney) Came down stating device was turning off on its own Ran warmer for 2 hours and no fault found Cracked front resevoir case replaced All alarms checked Functional check - OK
66464 | 194981 | PM | CMPLT | CMPLT | 2023-05-29 07:45 | 2023-05-26 07:45 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. No GS500. Tested. Functional. SW.2.51 Leakage: 72 uA and 86 ohm.
1059999 | 194982 | CM | CMPLT | CMPLT | 2023-05-29 08:44 | 2023-05-29 08:46 | Patient Safety Net Screen was blank | (TC 05/29/2023 08:44:15 AM, Reid, Emmerson ) Patient Safety net screen was blank grey and the mouse was not repoonding. I rebooted the monitor and it is working now. Message also came up that 11S monitor was down.
1060000 | 194983 | CM | CMPLT | CMPLT | 2023-05-29 08:48 | 2023-05-29 08:49 | Patient Safety net down | (TC 05/29/2023 08:47:43 AM, Reid, Emmerson ) Saw on 11N's monitor that 11S was down. I went over and rebooted it. Working now.
1067117 | 194984 | CM | CMPLT | CMPLT | 2023-05-29 08:57 | 2023-05-29 08:58 | No detected | (TC 05/23/2023 08:57:38 AM, Rendon Lopez , Elizabeth) Entered in TMS. Received from repair, incoming inspection-OK
1067118 | 194985 | CM | CMPLT | CMPLT | 2023-05-29 08:59 | 2023-05-29 08:59 | No detected | (TC 05/23/2023 08:59:00 AM, Rendon Lopez , Elizabeth) Entered in TMS. Received from repair, incoming inspection-OK
1064556 | 194986 | CM | CMPLT | CMPLT | 2023-05-29 09:00 | 2023-06-29 09:23 | No FHR | (TC 05/23/2023 09:00:11 AM, Rendon Lopez , Elizabeth) Under philips warranty. Requested a replacement (TC 06/28/2023 09:22:07 AM, Rendon Lopez , Elizabeth) Received replacement, changed SN in TMS, tested, OK
1066459 | 194987 | CM | CMPLT | CMPLT | 2023-05-29 09:01 | 2023-07-17 10:05 | No FHR | (TC 05/23/2023 09:01:08 AM, Rendon Lopez , Elizabeth) Confirmed issue, kept to send out for repair. (TC 07/17/2023 10:04:22 AM, Rendon Lopez , Elizabeth) Broken beyond repair. Retired from service
1061387 | 194988 | CM | CMPLT | CMPLT | 2023-05-29 10:36 | 2023-05-29 10:37 | Not turning on | (TC 05/29/2023 10:36:23 AM, Canjimavadivel, Kamallesh ) Replaced 3xAAA batteries
1061208 | 194989 | CM | CMPLT | CMPLT | 2023-05-29 10:38 | 2023-05-29 10:39 | Surgical Table power cord frayed | (TC 05/29/2023 10:38:34 AM, Reid, Emmerson ) Replaced the power cord with another one. This was not the original power cord for the table, unsure why that was.
| 194990 | CM | CMPLT | CMPLT | 2023-05-29 10:54 | 2023-05-29 10:55 | Replaced the SSD in Philips Web server | (TC 05/29/2023 10:54:16 AM, Samanian, Kamran) replaced the defective SSDD in Philips web server. RAID was built and server back to normal condition.
| 194991 | PA | CMPLT | CMPLT | 2023-05-29 10:55 | 2023-06-01 08:38 | Genius Probe Covers For Biomed Stock
| 194992 | CM | CMPLT | CMPLT | 2023-05-29 11:34 | 2023-05-29 11:35 | Installing clear sheet on 55" wall monitor by Karl Storz | (TC 05/29/2023 11:34:15 AM, Samanian, Kamran) installed the clear plexi-glass on the wall monitor tested. staff might provide some feedbacks.
67092 | 194993 | CM | CMPLT | CMPLT | 2023-05-29 15:34 | 2023-05-29 15:39 | "spo2 not working" | (TC 05/29/2023 03:37:19 PM, Cortez, John) Temp probe exposed wire, spo2 trunk faulty, ground pin on power cord bent Replaced power cord, spo2 trunk and temp probe. Secured power cord with tyrap
1061901 | 194994 | CM | CMPLT | CMPLT | 2023-05-29 15:44 | 2023-05-29 15:46 | "scale needs to be calibrated" | (TC 05/26/2023 03:44:13 PM, Cortez, John) Power cycled bed, issue seems to be resolved, functional test ok
1062002 | 194995 | CM | CMPLT | CMPLT | 2023-05-29 15:47 | 2023-05-29 15:48 | "scale needs to be calibrated" | (TC 05/26/2023 03:47:33 PM, Cortez, John) Power cycled bed, Issue seems to be resolved, functional test ok
| 194996 | CM | CMPLT | CMPLT | 2023-05-29 16:45 | 2023-05-29 16:45 | Site Coverage
1057937 | 194997 | CM | CMPLT | CMPLT | 2023-05-30 09:14 | 2023-07-07 08:05 | Otoscope head not working | (TC 05/30/2023 09:55:15 AM, Canjimavadivel, Kamallesh ) Bulb burnt out, need to replace. Requested quote from vendor. (TC 06/02/2023 03:13:50 PM, Canjimavadivel, Kamallesh ) Received quote, placed order (TC 06/20/2023 02:17:17 PM, Canjimavadivel, Kamallesh ) Went up to install otoscope head with new bulb. Discovered intermittent light. Brought down unit to biomed. Issue found with one connector on the board. (TC 06/21/2023 11:33:37 AM, Canjimavadivel, Kamallesh ) Emailed vendor to request repair next steps (TC 06/29/2023 11:18:56 AM, Canjimavadivel, Kamallesh ) Vendor to provide warranty replacement, requested PO to cover shipping cost (TC 07/07/2023 08:03:27 AM, Canjimavadivel, Kamallesh ) Received replacement. Updated serial number on asset and brought up device to unit. Waiting for instruction from vendor for the defective device
1057936 | 194998 | CM | CMPLT | CMPLT | 2023-05-30 09:59 | 2023-05-30 10:02 | Rubber collar on dispenser missing | (TC 05/30/2023 10:01:11 AM, Canjimavadivel, Kamallesh ) Installed rubber collar from stock parts
| 194999 | CM | CMPLT | CMPLT | 2023-05-30 11:03 | 2023-06-16 09:46 | Alaris PCU plastics order for biomed stock (recall items)
| 195000 | CM | CMPLT | CMPLT | 2023-05-30 12:24 | 2023-05-30 12:25 | huddle | (TC 05/30/2023 12:24:47 PM, Samanian, Kamran) done!
| 195001 | CM | CMPLT | CMPLT | 2023-05-30 12:26 | 2023-05-30 12:27 | destryed 2 x AF2b and POds from NICU | (TC 05/30/2023 12:26:23 PM, Samanian, Kamran) all three were distryed by Courtney, email was sent out to Laura as well. AF2B SN# 19173 , Pod SN# 19147 AF2B SN# 19149 , Pod SN# 19143 AF2B SN# 19136 , Pod SN# 19126
| 195002 | CM | CMPLT | CMPLT | 2023-05-30 12:42 | 2023-05-30 12:42 | Monday Huddle/PM Update
| 195003 | CM | CMPLT | CMPLT | 2023-05-30 12:54 | 2023-07-31 12:48 | XDS option for NICU research used in NICU dept | (TC 05/30/2023 12:53:59 PM, Samanian, Kamran) XDS option to purchased for NICU reseach (TC 07/31/2023 12:47:52 PM, Samanian, Kamran) received the license and installed in the NICU research laptop
64868 | 195004 | PM | CMPLT | CMPLT | 2023-05-30 14:49 | 2023-05-30 15:14 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 05/30/2023 03:13:58 PM, Reid, Emmerson ) PM Complete, see proceedure for details.
| 195005 | CM | CMPLT | CMPLT | 2023-05-30 16:00 | 2023-05-30 16:00 | Site coverage
1062099 | 195006 | CM | CMPLT | CMPLT | 2023-05-31 07:59 | 2023-06-06 08:27 | Screen is dim. No back light? | Warranty repair. Contacted Draeger. Waititng for parts. Screen is dark red.Draeger replaced C500 display MK31421 under warranty. Tested. Functional. " Replaced bad LCD with a new one.Did necessary checks and tests as per IPM-L and recorded the results in the test card. All checks and tests passed OK.Handed over the device in good working order."
1066460 | 195007 | FC | CMPLT | CMPLT | 2023-05-31 08:01 | 2023-05-31 08:01 | Tested. Functional. | QC test. Tested. Functional.
1059778 | 195008 | FC | CMPLT | CMPLT | 2023-05-31 08:02 | 2023-05-31 08:02 | Tested. Functional. | QC test. Tested. Functional.
1049878 | 195009 | CM | CMPLT | CMPLT | 2023-05-31 08:26 | 2023-06-08 09:51 | Equipment malfunction | (TC 05/29/2023 08:26:07 AM, Rendon Lopez , Elizabeth) Removed equipment from L&D, reached out to Dornier for tech support (TC 05/30/2023 08:26:07 AM, Rendon Lopez , Elizabeth) Received call from tech support, I need key to turn on device and troubleshoot. Reached out to clinical staff to provide key (TC 05/31/2023 03:40:52 PM, Rendon Lopez , Elizabeth) Troubleshooting with tech on the phone, unable to find issue. Tech support wirll come on site next week (TC 06/05/2023 09:48:48 AM, Rendon Lopez , Elizabeth) Service rep on site - unable to find issue. Contacted clinical team and discussed together with rep. No fault found at the time, all functions OK
| 195010 | CM | CMPLT | CMPLT | 2023-05-31 08:27 | 2023-05-31 08:37 | Proceedure room 1 overhead light not working | (TC 05/31/2023 08:35:38 AM, Reid, Emmerson ) Attached invoice to WO to charge back to Women's unit for work done on overhead light in proceedure room 1
1049416 | 195011 | CM | CMPLT | CMPLT | 2023-05-31 08:27 | 2023-05-31 08:29 | Low battery | (TC 05/29/2023 08:27:42 AM, Rendon Lopez , Elizabeth) Device sent down from 15M-Antenatal. Replaced batteries, tested with weights-OK
52470 | 195012 | CM | CMPLT | CMPLT | 2023-05-31 08:31 | 2023-05-31 08:33 | System leaking | (TC 05/30/2023 08:31:16 AM, Rendon Lopez , Elizabeth) Device missing O-rings, replaced O-rings, during testing device alarms High temp and shuts down. Unable to turn back on, open device, troubleshooted, unable to repair. Retire from service, requested capital replacement
1059656 | 195013 | CM | CMPLT | CMPLT | 2023-05-31 08:35 | 2023-05-31 08:35 | Not turning On | (TC 05/30/2023 08:34:39 AM, Rendon Lopez , Elizabeth) power cord unplugged, reseated. Tested, all functions-OK
1067164 | 195014 | ACC | CMPLT | CMPLT | 2023-05-31 08:36 | 2023-05-31 08:37 | Incoming inspection | (TC 05/30/2023 08:35:52 AM, Rendon Lopez , Elizabeth) Incoming inspection, dead weight tester for Bunnell HFV. Entered device in TMS, stored in 18th Fl Biomed
63303 | 195015 | CM | CMPLT | CMPLT | 2023-05-31 08:37 | 2023-06-05 08:45 | O2 knob loose | (TC 05/30/2023 08:37:16 AM, Rendon Lopez , Elizabeth) Removed device from L&D, confirmed issue. Blender knob has to be adjusted. (TC 06/02/2023 08:42:05 AM, Rendon Lopez , Elizabeth) Adjusted blender, O2 reading okay. Tested all functions-OK
| 195016 | CM | CMPLT | CMPLT | 2023-05-31 08:38 | 2023-12-12 09:01 | Meetings | (TC 05/30/2023 08:38:11 AM, Rendon Lopez , Elizabeth) Huddle (TC 05/31/2023 03:39:13 PM, Rendon Lopez , Elizabeth) PM commitee (TC 05/31/2023 03:39:13 PM, Rendon Lopez , Elizabeth) Draeger AGM (TC 06/05/2023 09:33:36 AM, Rendon Lopez , Elizabeth) Huddle (TC 06/15/2023 09:10:29 AM, Rendon Lopez , Elizabeth) 15M Masimo deployment (TC 06/15/2023 09:10:29 AM, Rendon Lopez , Elizabeth) PM meeting (TC 06/16/2023 09:10:29 AM, Rendon Lopez , Elizabeth) IT meeting (TC 06/19/2023 08:23:37 AM, Rendon Lopez , Elizabeth) Huddle (TC 06/21/2023 08:23:37 AM, Rendon Lopez , Elizabeth) P3A meeting (TC 06/27/2023 09:07:23 AM, Rendon Lopez , Elizabeth) PM meeting (TC 07/17/2023 10:53:55 AM, Rendon Lopez , Elizabeth) P3A (TC 07/17/2023 02:13:45 PM, Rendon Lopez , Elizabeth) huddle (TC 07/20/2023 02:54:12 PM, Rendon Lopez , Elizabeth) Cleaning (TC 07/21/2023 03:15:03 PM, Rendon Lopez , Elizabeth) 1:1 (TC 07/24/2023 08:44:02 AM, Rendon Lopez , Elizabeth) Huddle (TC 07/25/2023 09:19:20 AM, Rendon Lopez , Elizabeth) PM update (TC 07/27/2023 01:29:03 PM, Rendon Lopez , Elizabeth) Rapid infusor meeting (TC 08/01/2023 06:08:56 PM, Rendon Lopez , Elizabeth) AGM meeting (TC 08/03/2023 09:25:19 AM, Rendon Lopez , Elizabeth) PM committee (TC 08/14/2023 02:29:05 PM, Rendon Lopez , Elizabeth) huddle (TC 08/14/2023 02:31:37 PM, Rendon Lopez , Elizabeth) WOs input (TC 08/21/2023 02:29:51 PM, Rendon Lopez , Elizabeth) huddle (TC 08/28/2023 02:22:18 PM, Rendon Lopez , Elizabeth) huddle (TC 11/29/2023 02:29:42 PM, Rendon Lopez , Elizabeth) Fetal monitor information collect (TC 11/29/2023 02:29:42 PM, Rendon Lopez , Elizabeth) Fetal monitor information collect (TC 12/05/2023 04:53:28 PM, Rendon Lopez , Elizabeth) WOs (TC 12/05/2023 04:53:28 PM, Rendon Lopez , Elizabeth) WOs
1055789 | 195017 | CM | CMPLT | CMPLT | 2023-05-31 08:40 | 2023-05-31 08:43 | No pads connected warning | (TC 05/31/2023 08:41:06 AM, Reid, Emmerson ) Went to check the defib because the dashboard gave a warning but the unit had already managed the issue.
1055792 | 195018 | CM | CMPLT | CMPLT | 2023-05-31 08:45 | 2023-05-31 08:49 | Pads needed message on dashboard | (TC 05/31/2023 08:48:31 AM, Reid, Emmerson ) Went to check the defib and it was missing a pad. I went to the storage room, connected the pad and did a 30J test on it.
1061203 | 195019 | CM | ACTIV | ISSUE | 2023-05-31 09:40 |  | OR2 Video Issues with DVI | (TC 05/31/2023 09:40:12 AM, Reid, Emmerson ) The VMC needed to be restarted. The MX800 video signal is having a syncing problem. The HD scope was also having this sync issue with video capture and image capture. Spoke with derrick and he thinks it may be an issue with the DVI extendor within the boom or with the connection to the Corio matrix. Heis going to speak with Jaron to have a time to come in ASAP.
67103 | 195020 | CM | CMPLT | CMPLT | 2023-05-31 10:05 | 2023-05-31 10:22 | BP not working | (TC 05/31/2023 10:20:19 AM, Reid, Emmerson ) The connector next to the cuff was damaged I replaced this. I was getting a low battery alert when doing BP and it was not successful. I saw the battery was not replaced ever, so I replaced the battery. Working now.
| 195021 | PA | CMPLT | CMPLT | 2023-05-31 10:18 | 2023-07-14 09:04 | VSM 300 Batteries | (TC 05/31/2023 10:19:49 AM, Reid, Emmerson ) Placed the order using an old shopping cart. (TC 07/14/2023 08:59:58 AM, Reid, Emmerson ) Recieved
| 195022 | CM | ACTIV | ISSUE | 2023-05-31 13:01 |  | PM Committee | (TC 05/31/2023 01:01:01 PM, Chukri, Samir) PM Comittee (TC 08/03/2023 12:09:33 PM, Chukri, Samir) meeting (TC 11/30/2023 12:12:58 PM, Chukri, Samir) meeting (TC 01/22/2024 10:04:31 AM, Chukri, Samir) PM committee add a functional check for the non-PM equipment (TC 02/20/2024 02:38:42 PM, Chukri, Samir) attending PM committee meeting (TC 02/27/2024 03:11:12 PM, Chukri, Samir) PM meeting (TC 03/13/2024 09:22:00 AM, Chukri, Samir) PM committee (TC 04/29/2024 02:34:58 PM, Chukri, Samir) meeting (TC 07/24/2024 01:11:06 PM, Chukri, Samir) PM meeting .
| 195023 | CM | CMPLT | CMPLT | 2023-05-31 13:05 | 2023-05-31 13:07 | Draeger Meeting | (TC 05/31/2023 01:05:07 PM, Chukri, Samir) completed
| 195024 | CM | CMPLT | CMPLT | 2023-05-31 13:38 | 2023-06-16 16:20 | MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 | (TC 05/31/2023 01:37:52 PM, Whittington, Scott) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 (TC 06/05/2023 04:40:27 PM, Whittington, Scott) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 (TC 06/07/2023 03:14:59 PM, Whittington, Scott) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 EPIQ SN# US622B1025 CONNECT TO PACS REQ WHITELIST OF MAC (TC 06/08/2023 02:27:51 PM, Whittington, Scott) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 EPIQ SN# US622B1025 AFFINITI AS AT TGH IRUS MEET WITH AMY PROULX AND ANDREA HOLLINGSHEAD RE PLANS FOR THE SECOND WEEK OF PHILIPS AND FIRST WEEK OF GE DEMO (TC 06/12/2023 03:50:04 PM, Whittington, Scott) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 EPIQ SN# US622B1025 AFFINITI AS AT TGH IRUS AFFINITI SW CRASH NETWORK & VIDEO SETUPS (TC 06/14/2023 03:20:38 PM, Whittington, Scott) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 EPIQ SN# US622B1025 AFFINITI AS AT TGH IRUS SHIFTING SYSTEMS AROUND FOR L&L, AND BACK AGAIN (TC 06/14/2023 03:21:21 PM, Chau, Ken) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 EPIQ SN# US622B1025 AFFINITI AS AT TGH IRUS SHIFTING SYSTEMS AROUND FOR L&L, AND BACK AGAIN (TC 06/16/2023 04:13:15 PM, Whittington, Scott) MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 ALL 3 SYSTEMS , EPIQ, AFFINITI, COMPACT 5500 MOVED TO MSH OPG 3 AND PICKED UP BY MACKIE | MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO MSH OPG SPP PHILIPS DEMO AFFINITI ROOM C2 JUNE 5- JUNE 9 EPIQ SN# US622B1025 AFFINITI AS AT TGH IRUS
67153 | 195025 | CM | CMPLT | CMPLT | 2023-05-31 13:38 | 2023-06-12 10:37 | BP not working | (TC 05/31/2023 01:38:37 PM, Reid, Emmerson ) E41 came up, refered to the manual and it was replated to the BP module. I tried reseating, and replacing the module but neither was fixing the issue. I closed up the unit with all original parts and left for Kamran to repair when he returns. (TC 06/12/2023 10:36:57 AM, Samanian, Kamran) due to the cost monitor was retired. it was replaced with another spare one taken from 11N
| 195026 | CM | CMPLT | CMPLT | 2023-05-31 13:42 | 2023-07-11 07:42 | MSH OPG SPP GE DEMO ROOM C2 | (TC 05/31/2023 01:42:22 PM, Whittington, Scott) MSH OPG SPP GE DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO Christina is available Monday June 12’th and June 13’th . We can have both systems delivered Friday June 9’th . 1. Monday June 12’th Demo the Voluson S8 . 2. Tuesday June 13’th Demo the Voluson S10 Expert (TC 06/09/2023 02:03:06 PM, Whittington, Scott) MSH OPG SPP GE DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO Christina is available Monday June 12’th and June 13’th . We can have both systems delivered Friday June 9’th . 1. Monday June 12’th Demo the Voluson S8 . 2. Tuesday June 13’th Demo the Voluson S10 Expert (TC 06/12/2023 03:40:16 PM, Heron, Chris) MSH OPG SPP GE DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO Christina is available Monday June 12’th and June 13’th . We can have both systems delivered Friday June 9’th . 1. Monday June 12’th Demo the Voluson S8 . 2. Tuesday June 13’th Demo the Voluson S10 Expert SWAP S10 & S8 OPG<>MSH 15 (TC 06/12/2023 03:49:38 PM, Whittington, Scott) SWAP S10 & S8 OPG<>MSH 15 (TC 06/14/2023 03:24:16 PM, Whittington, Scott) MSH OPG SPP GE DEMO AFFINITI ROOM C2 SHIFITING SYSTEMS, DELETE PT & NW INFO | MSH OPG SPP GE DEMO AFFINITI ROOM C2 LATE NOTICE OF DEMO Christina is available Monday June 12’th and June 13’th . We can have both systems delivered Friday June 9’th . 1. Monday June 12’th Demo the Voluson S8 . 2. Tuesday June 13’th Demo the Voluson S10 Expert
1067169 | 195027 | CM | CMPLT | CMPLT | 2023-05-31 13:48 | 2023-05-31 13:47 | Incoming Instllation | (TC 05/31/2023 01:47:47 PM, Chukri, Samir) Installed and removed the old system SN 7000637 returned to department to ship it our back to company.
1066475 | 195028 | CM | CMPLT | CMPLT | 2023-05-31 13:48 | 2023-10-30 13:41 | Neptune Cleaning Failed | (TC 05/31/2023 01:48:38 PM, Reid, Emmerson ) This morning Asset# 1066475 was used in an ortho case where a triple wash was done. After the case the SA went to run the cleaning cycle but it failed and produced an error. I tried cleaning on both docks and got the same issue, the cleaning cycle would stop at step# 5 after about 1 or 2 minutes. What it looked like was that it was trying to drain but was not able to fully drain the unit. I looked in the canister and found a solid left over that was not able to pass through the drain. I reached out to Chris who said that it would be impossible for a solid to be sucked into the canister and that it must be a result of a reaction with the chemicals in the triple wash. He is reaching out to his team about how to get the unit up and running again. (TC 06/27/2023 03:28:52 PM, Reid, Emmerson ) Brought the two stryker techs in to work on the 2 neptune systems and set them up with everything they need
1046196 | 195029 | CM | CMPLT | CMPLT | 2023-05-31 14:32 | 2023-05-31 14:32 | CT table horizontal movement error | CT table horizontal movement error -Lifted table top and made sure nothing was caught underneath -Placed table top back on and confirmed that it was moving freely with no errors. -Returned system for use
| 195030 | CM | CMPLT | CMPLT | 2023-05-31 15:37 | 2023-06-05 09:04 | Spare fetal monitor parts | (TC 05/31/2023 03:37:46 PM, Rendon Lopez , Elizabeth) Spare fetal monitor NIBP assembly. Will be charged when used
1060518 | 195031 | CM | CMPLT | CMPLT | 2023-05-31 16:09 | 2023-05-31 16:08 | No power to the system. Room is DOWN. | Reset the UPS by going into the service menu, restarted the invertor and cleared the error messages.
25987 | 195032 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26152 | 195033 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2024-10-31 10:38 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 10/31/2024 10:38:15 AM, Rhuland, Courtney) PM completed by Spectramedx Returned and functional check completed 21% O2 - 20.9% O2 40% O2 - 39.3% O2 60% O2 - 61.9% O2 80% O2 - 81.9% O2 100% O2 - 99.6% O2 Returned to RT
50518 | 195034 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-17 14:55 | PM-MAJOR - Annual PM For: Vacuum Extractors, Obstetric | (TC 08/16/2023 02:54:15 PM, Rendon Lopez , Elizabeth) Performance verification-OK
51601 | 195035 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Scales, Infant
62919 | 195036 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
62966 | 195037 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-07-07 13:44 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/07/2023 01:40:58 PM, Rhuland, Courtney) Wrong side door came down, replaced wall latch with spare Broken door hinge snap replaced Side wall hinge replaced 2 porthole and 2 ventilator tube grommets replaced Canopy membrane switch replaced Scale calibrated and verified LV and main voltage calibrated PM complete All tests passed Ground leakage - 82.6 uA Resistance - 27 mOhms
62972 | 195038 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62973 | 195039 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2024-06-26 13:00 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/26/2024 12:59:39 PM, Rhuland, Courtney) PM completed All tests passed Irradiance check - OK Electrical safety - OK
62984 | 195040 | PM | CANCL | CANCL | 2023-06-01 03:20 | 2024-07-23 09:50 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62988 | 195041 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-06-20 07:38 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/20/2023 07:37:50 AM, Rhuland, Courtney) PM complete Filters cleaned Light intensity check - OK Electrical safety - OK
63060 | 195042 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63302 | 195043 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-25 17:42 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 08/09/2023 05:42:19 PM, Huang, Jessica) took the device from MBU, PM completed
63410 | 195044 | PM | CANCL | DUP | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | PM done every 3 yrs, postponed
63412 | 195045 | PM | CANCL | DUP | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | PM done every 3 yrs, postponed
63835 | 195046 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-11-16 15:43 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 09/20/2023 03:43:44 PM, Huang, Jessica) worked with Shermaine to take the warmer for PM completed it, sent it back to MBU
63870 | 195047 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
63871 | 195048 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-25 17:13 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 08/16/2023 05:12:29 PM, Huang, Jessica) took the warmer from MBU. PM completed.
64821 | 195049 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-11-03 08:55 | PM-MAJOR - Annual PM For: Insufflators | (TC 11/03/2023 08:54:22 AM, Reid, Emmerson ) Turned on and tested for flow and occlusion. PM Complete.
65121 | 195050 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65129 | 195051 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-06-21 12:29 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/21/2023 12:28:40 PM, Reid, Emmerson ) PM Complete, see proceedure for details.
65132 | 195052 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-06-19 11:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/19/2023 11:32:10 AM, Reid, Emmerson ) PM Cmplete, see proceedure for details.
65133 | 195053 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65375 | 195054 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-25 17:57 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 08/04/2023 05:56:57 PM, Huang, Jessica) took the device from MBU and PM completed
65879 | 195055 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-04 08:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/04/2023 08:56:27 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65916 | 195056 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2024-03-28 12:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 12:56:20 PM, Rhuland, Courtney) PM completed by BD
65928 | 195057 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2024-04-05 15:18 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/05/2024 03:17:28 PM, Huang, Jessica) replaced the rear case and the handle completed PM
65931 | 195058 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-17 15:12 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/17/2023 03:11:51 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
65935 | 195059 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65937 | 195060 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65938 | 195061 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65940 | 195062 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-07-12 09:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/12/2023 09:05:54 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65944 | 195063 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-07-12 14:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/12/2023 02:33:07 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65982 | 195064 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-07-13 09:12 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/13/2023 09:11:37 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
66001 | 195065 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-04 08:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/04/2023 08:36:23 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 54 mohms Leakage current: 77.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66008 | 195066 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66013 | 195067 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-24 09:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/21/2023 05:40:04 PM, Rendon Lopez , Elizabeth) Running battery conditioning (TC 08/22/2023 09:15:16 AM, Rendon Lopez , Elizabeth) Front and rear case cracked. Replaced. PM completed, all tests passed
66015 | 195068 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66016 | 195069 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-07-18 06:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/17/2023 08:52:56 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 64 mohms Leakage current: 82.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 07/18/2023 06:54:45 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4029 mamphr PM completed.
66023 | 195070 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-07-20 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/20/2023 11:33:12 AM, Rhuland, Courtney) Rear and front case replaced PM complete All tests passed Refer to ASM for details
66044 | 195071 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66066 | 195072 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-09 09:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/08/2023 11:58:45 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 130 mohms Leakage current: 81.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/09/2023 09:08:17 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3558 mamphr New battery capacity: 4063 mamphr PM completed.
66402 | 195073 | PM | CANCL | CANCL | 2023-06-01 03:20 | 2024-09-12 14:59 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
66468 | 195074 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-06-14 10:23 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. No GS500. Functional. SW.2.51 Leakage: 81 uA and 88 ohm.
66471 | 195075 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-06-19 11:33 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. NO PS500. Replaced GS500 filters. Functional. SW.2.42 Leakage: 82 uA and 87 ohm.
66494 | 195076 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-17 14:43 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 08/16/2023 02:42:18 PM, Rendon Lopez , Elizabeth) Performance tested-OK
66805 | 195077 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
66815 | 195078 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
66820 | 195079 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
66821 | 195080 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
66824 | 195081 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
66849 | 195082 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
66861 | 195083 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
66864 | 195084 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-14 09:40 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/11/2023 09:39:57 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed
66869 | 195085 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-17 14:38 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/16/2023 02:37:28 PM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
66875 | 195086 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2024-07-24 14:28 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/24/2024 02:28:15 PM, DeGuzman, Leila) PM passed and completed. See procedure for details.
67204 | 195087 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | BP Check
67382 | 195088 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67384 | 195089 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67386 | 195090 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67893 | 195091 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-06-26 08:34 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/21/2023 08:33:50 AM, Rendon Lopez , Elizabeth) Upgraded software, downloaded library. PM completed, all tests passed
67932 | 195092 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-09-01 08:26 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/01/2023 08:25:12 AM, Baghirzade, Nara) PM completed. VRef, Mains V, and LV were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C. Functional tests passed. Electrical safety results - Leakage current = 84.2 uA Power cord resistance = 164 mohms
68380 | 195093 | PM | ACTIV | ISSUE | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
68405 | 195094 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM MAJOR - Delivery Module DESO | Device anot in current use
68615 | 195095 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68630 | 195096 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68633 | 195097 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-10 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/10/2023 02:15:39 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
68638 | 195098 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68647 | 195099 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-06-08 14:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 02:07:38 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68649 | 195100 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68665 | 195101 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68681 | 195102 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-09-05 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 09:24:41 AM, Reid, Emmerson ) Ran optimal bnattery conditioning test, new battery capacity 3983 mamphrs, good. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68685 | 195103 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2024-03-14 10:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 10:02:55 AM, Huang, Jessica) replaced the power cord wrap and the female iui connector completed PM
68693 | 195104 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68699 | 195105 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68703 | 195106 | PM | CNP | ENL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68738 | 195107 | PM | CANCL | CANCL | 2023-06-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68752 | 195108 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-07-12 09:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 09:52:34 AM, Cortez, John) Cleaned unit, Replaced battery, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4162 mAhrs
68778 | 195109 | PM | CMPLT | CMPLT | 2023-06-01 03:20 | 2023-08-29 15:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/29/2023 03:45:47 PM, Canjimavadivel, Kamallesh ) Replaced battery Cleaned PC PM tests pass
68798 | 195110 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-22 12:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/22/2023 12:50:38 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 70 mohms Leakage current: 56.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/23/2023 08:44:41 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4036 mamphr PM completed.
68804 | 195111 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68809 | 195112 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-10-23 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/23/2023 10:35:58 AM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioning battery overnight (TC 10/25/2023 12:56:44 PM, Cortez, John) 4077 mAhs
68826 | 195113 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-24 07:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 08:49:27 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 65 mohms Leakage current: 70.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/24/2023 07:35:16 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4047 mamphr PM completed.
68827 | 195114 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-04-03 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:41:32 PM, Rhuland, Courtney) PM completed by BD
68828 | 195115 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68864 | 195116 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68872 | 195117 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-27 12:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/27/2023 12:58:40 PM, Rhuland, Courtney) Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
68878 | 195118 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-08 09:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 09:34:32 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed all tests passed
68883 | 195119 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68926 | 195120 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-17 09:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/17/2023 09:27:32 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 124 mohms Leakage current: 68.9 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68947 | 195121 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-24 09:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/21/2023 05:38:46 PM, Rendon Lopez , Elizabeth) battery replaced, charging (TC 08/23/2023 09:08:32 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
68961 | 195122 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-13 08:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 11:26:49 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 30 mohms Leakage current: 71.8 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 07/13/2023 08:01:55 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4049 mamphr PM completed.
68973 | 195123 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-10 11:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/10/2023 11:02:28 AM, Cortez, John) Cleaned unit, Replaced power cord, damaged Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: (TC 08/10/2023 11:02:28 AM, Cortez, John) Cleaned unit, Replaced power cord, damaged Downloaded error/event logs, PM completed successfully, see ASM for results Conditioned battery overnight, Battery capacity: 4081mAhs
68982 | 195124 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-12 14:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/12/2023 02:12:10 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 142 mohms Leakage current: 70.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68988 | 195125 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-11-15 08:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/15/2023 08:29:29 AM, Campbell, Connor) battery replaced with a spare that was already charged power cord replaced due to fraying of the old one PM complete/passed, refer to ASM refer to procedure for ESA
69016 | 195126 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-09 14:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:45:11 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 242 mohms Leakage current: 71.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
69029 | 195127 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69044 | 195128 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69048 | 195129 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69053 | 195130 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-04-09 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/09/2024 02:54:58 PM, Reid, Emmerson ) Calibrated and PM with ASM, all tests passed. PM Complete.
69071 | 195131 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69072 | 195132 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-19 21:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69092 | 195133 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-24 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 09:57:58 AM, Canjimavadivel, Kamallesh ) Performed rate calibration PM tests pass
69112 | 195134 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-27 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/27/2023 11:33:35 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69114 | 195135 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-08 09:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 09:05:35 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69116 | 195136 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-21 18:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69127 | 195137 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69147 | 195138 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-04-02 16:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69156 | 195139 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-28 12:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/25/2023 08:25:13 AM, Campbell, Connor) PM passed/complete, refer to ASM
69161 | 195140 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69164 | 195141 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-01 13:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 01:26:05 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69177 | 195142 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69179 | 195143 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69180 | 195144 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-14 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69196 | 195145 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-04-26 09:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69221 | 195146 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-11-01 16:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/01/2023 03:57:38 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69223 | 195147 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-28 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 09:21:28 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69226 | 195148 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69232 | 195149 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69250 | 195150 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-20 10:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/20/2023 10:38:30 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69288 | 195151 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-08 13:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:48:45 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69292 | 195152 | PM | CMPLT | PMFAIL | 2023-06-01 03:21 | 2023-10-31 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/31/2023 09:01:01 AM, Baghirzade, Nara) Calibrated pressure and rate. During calibration, both passed. During PM, the rate was reading low (11.505). Will recalibrate and will assess the pressure sensors.
69303 | 195153 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69339 | 195154 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-22 12:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/22/2023 12:25:06 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69344 | 195155 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69352 | 195156 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-10-31 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/31/2023 09:32:21 AM, Reid, Emmerson ) Replaced the display board. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69355 | 195157 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69356 | 195158 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69359 | 195159 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-18 15:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/18/2024 03:13:49 PM, Reid, Emmerson ) ASM PM Complete.
69360 | 195160 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-10-31 09:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/31/2023 09:16:02 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69379 | 195161 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69384 | 195162 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-10-05 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 10:22:59 AM, Campbell, Connor) PM passed/complete, refer to ASM
69385 | 195163 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-21 20:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69400 | 195164 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-07 07:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 07:48:18 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69403 | 195165 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-08 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:40:03 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69431 | 195166 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-19 20:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69444 | 195167 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-09 12:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 12:05:52 PM, Baghirzade, Nara) Latch was broken; replaced latch. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69449 | 195168 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-08 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/08/2023 10:16:28 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69468 | 195169 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-04 09:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/03/2023 09:22:15 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69478 | 195170 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-26 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/26/2023 02:57:59 PM, Canjimavadivel, Kamallesh ) Cleaned pump PM tests pass
69493 | 195171 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69495 | 195172 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69496 | 195173 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-18 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/15/2023 10:43:08 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69502 | 195174 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-14 09:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 09:04:39 AM, Rendon Lopez , Elizabeth) Dimmed display during PM. Replaced display under recall, PM completed, all tests passed
69503 | 195175 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-24 11:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 11:52:21 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69518 | 195176 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-25 13:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69526 | 195177 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-09 15:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 03:02:44 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69549 | 195178 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69561 | 195179 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-04-30 13:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69582 | 195180 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-01 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/01/2023 10:40:47 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69586 | 195181 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-04-24 09:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69591 | 195182 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-09 14:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:51:02 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69592 | 195183 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-10-03 09:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/02/2023 09:36:34 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69594 | 195184 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-01 10:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/01/2023 10:16:05 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69600 | 195185 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-21 17:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69602 | 195186 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-12 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 02:55:18 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69609 | 195187 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-09 13:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 01:22:11 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69613 | 195188 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-10-20 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/20/2023 02:35:18 PM, Campbell, Connor) PM complete/passed, refer to ASM
69631 | 195189 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-11 07:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/11/2023 07:50:35 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69636 | 195190 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69668 | 195191 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-10-20 12:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/20/2023 11:59:37 AM, Campbell, Connor) PM complete/passed, refer to ASM
69676 | 195192 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
530502 | 195193 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-20 11:16 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 07/20/2023 10:26:18 AM, Ferreira, Qwynn) Obtained device from NICU Removed old fluid cleaned with 50/50 mix of alcohol and water Tested alarms - OK (TC 07/20/2023 11:15:40 AM, Ferreira, Qwynn) Flushed cleaning solution Cleaned the water chamber Added hydrogen peroxide mixing solution with distilled water Testing temperature with the analyzer
32540 | 195194 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-10 15:40 | PM - Monthly | PM completed
32560 | 195195 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-04 15:32 | PM - Monthly | PM completed Mechanical PM done SC-1222738 TOS phantom a(-989.4,9.8) b(334.2,14.2) c(133.8,13.5) d(95.9,13.5) e(-106.7,11.5) f(0.9,13.7) -passed all phantom tests
32580 | 195196 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-05 14:52 | PM - Monthly | PM Major done Scan PC system disk duplicated Slice count - 69940 TOS phantom a(-990.6, 10.1) b(333.9, 15.0) c(134.4, 14.4) d(98.3, 14.4) e(-106.5, 11.8) f(0.8, 14.7) -System passed all tests
501974 | 195197 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-11-28 10:57 | PM - Annual
502366 | 195198 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-08-29 14:53 | PM-Annual | (TC 07/31/2023 03:50:25 PM, Fan, Tak) System not available (esp since other system is waiting on parts). Panel up down might have issue. (TC 08/29/2023 02:51:14 PM, Fan, Tak) PM Completed. Gnd R = 141 mohm Max A = 0.3 / 156.9 / 304 uA ECG = 0.5 / 2.0 uA ISO = 6.3 uA M5S = 0.2 / 113.5 uA Probe has crack around lens. Needs replacement. XY brake assembly not working. All the brakes are open for some reason. (TC 08/31/2023 03:01:52 PM, Fan, Tak) Replaced cracked probe. 189191YP0 replaced by 458680YP6.
1046203 | 195199 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-26 10:55 | PM-Annual | (TC 07/26/2023 10:53:38 AM, Fan, Tak) PM Completed. Gnd R = 183 mohm Max A = 0.3 / 160.9 / 312 uA ECG = 0.4 / 2.8 uA ISO = 7.6 uA M5S = 0.7 / 1.9 uA 4Vc = 0.5 / 2.6 uA
1047034 | 195200 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2024-03-21 14:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/21/2024 01:56:36 PM, Reid, Emmerson ) Brand new R-IUI was not put in properly, the gasket was blocking the pins and did not allow secondary pump to connect. I was able to adjust the gasket and it is working now. Ran Calibration and ASM PM. PM Complete.
1047471 | 195201 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-10 12:30 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 6th year PM after 5th year PM last year. No GS500. Replaced PS500 batteries, V-battery, O2 regulator, O2 filter, air filter. Calibrated and tested. Functional. SW.2.51 Leakage: 76uA and 84 ohm.
1047473 | 195202 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-06-05 08:33 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 6th year PM after 5th year PM last year. Replaced PS500 batteries, V-battery, air filter, O2 filter, O2 presure regulator. No GS500. calibrated. Tested. Functional. SW.2.51 Leakage: 75 uA and 86 ohm.
1047833 | 195203 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-07 11:41 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/07/2023 11:40:20 AM, Baghirzade, Nara) PM completed. See procedure document for PM results.
1047841 | 195204 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-18 09:26 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/18/2023 09:24:15 AM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1047845 | 195205 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-14 13:09 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/14/2023 01:08:45 PM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1047861 | 195206 | PM | CANCL | CANCL | 2023-06-01 03:21 | 2024-06-14 09:27 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe
1047865 | 195207 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-06 12:51 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 09/06/2023 12:45:28 PM, Campbell, Connor) replaced plunger head assembly due to broken standoff PM completed, see procedure worked under nara's supervision
1047867 | 195208 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-07-13 08:21 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/13/2023 08:21:04 AM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1048530 | 195209 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1050035 | 195210 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050264 | 195211 | PM | CMPLT | CMPLT | 2023-06-01 03:21 | 2023-09-01 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/31/2023 11:33:09 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1050265 | 195212 | PM | CNP | ENL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050274 | 195213 | PM | CANCL | CANCL | 2023-06-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050616 | 195214 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-07-11 07:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/06/2023 12:55:28 PM, Rhuland, Courtney) Rear case replaced (TC 07/11/2023 07:56:12 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
1051374 | 195215 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051383 | 195216 | PM | CANCL | CANCL | 2023-06-01 03:22 | 2024-07-23 09:05 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
1051614 | 195217 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-06-03 08:43 | PM-MAJOR - Annual PM For: Testers | (TC 06/03/2023 08:41:19 AM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached.
1051615 | 195218 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-06-03 08:47 | PM-MAJOR - Annual PM For: Testers | (TC 06/03/2023 08:45:24 AM, Chukri, Samir) Completed PM Performed by Getinge. Certifiate of Calibration attached .
1055105 | 195219 | PM | CANCL | CANCL | 2023-06-01 03:22 | 2024-03-12 12:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055111 | 195220 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055126 | 195221 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-06-13 08:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/13/2023 08:30:03 AM, Reid, Emmerson ) Ran optimal battery conditioning test, new battery capacity 4026 mamphr. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM complete.
1055160 | 195222 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2024-04-26 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055747 | 195223 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2024-03-25 09:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 09:36:45 AM, Rhuland, Courtney) Calibration and PM complete All tests passed Refer to ASM for details
1055748 | 195224 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-07-18 07:04 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/18/2023 07:03:47 AM, Baghirzade, Nara) Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1055799 | 195225 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-09-26 11:29 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 09/25/2023 11:25:50 AM, Scero, Tomlyn) PM tests started (TC 09/26/2023 11:26:10 AM, Scero, Tomlyn) All PM tests complete and passed.
1055801 | 195226 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-10-17 10:28 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 10/17/2023 10:25:20 AM, Scero, Tomlyn) All PM test complete and passed. See procedure for results.
1055838 | 195227 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055839 | 195228 | PM | CANCL | CANCL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055840 | 195229 | PM | CANCL | CANCL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055841 | 195230 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055842 | 195231 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1057390 | 195232 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1057743 | 195233 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-12-28 08:49 | PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 12/28/2023 08:48:36 AM, Whittington, Scott) PM, SEMI-ANNUAL, ULTRASOUND GENERAL AGREEMENT NOT REACHED | PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1058236 | 195234 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-09-22 15:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/22/2023 03:04:28 PM, Campbell, Connor) PM complete/passed, refer to ASM
1058238 | 195235 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058241 | 195236 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2024-03-14 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:43:49 PM, Huang, Jessica) replaced the female iui connector and completed PM
1058246 | 195237 | PM | CNP | ENL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058252 | 195238 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-09-08 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 09:50:04 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1058605 | 195239 | PM | CANCL | CANCL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: Waste-Disposal Units
1058606 | 195240 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-07-05 09:25 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 07/05/2023 09:25:09 AM, Reid, Emmerson ) See WO 195576
1059882 | 195241 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-08-16 09:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 09:14:40 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1060913 | 195242 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-06-09 14:52 | Calibration Defibrillator Analyzer | (TC 06/09/2023 02:50:32 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached.
1060914 | 195243 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-06-09 15:13 | Calibration Defibrillator Analyzer | (TC 06/09/2023 03:11:42 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration atttached .
1060974 | 195244 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1061554 | 195245 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-07-05 09:27 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 07/05/2023 09:26:49 AM, Reid, Emmerson ) See WO 195575
1061830 | 195246 | PM | CANCL | CANCL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: Waste-Disposal Units
1061831 | 195247 | PM | CANCL | CANCL | 2023-06-01 03:22 |  | PM-MAJOR - Annual PM For: Waste-Disposal Units
1062181 | 195248 | PM | CNP | ENL | 2023-06-01 03:22 | 2024-09-11 12:55 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
1064401 | 195249 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064402 | 195250 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064403 | 195251 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064404 | 195252 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064502 | 195253 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-07-07 11:13 | PM - Annual | (TC 07/07/2023 11:12:11 AM, Green, Paul) Service provided by OEM.
1064511 | 195254 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064512 | 195255 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064513 | 195256 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064514 | 195257 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064515 | 195258 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064516 | 195259 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | PM MAJOR - ANNUAL PM for: ERBE
1064974 | 195260 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | Annual PM | (TC 09/25/2023 11:01:32 AM, Ferreira, Qwynn) PM in progress Firmware is updated Unit is levelled Need to clean laser, and work through rest of PM
1064975 | 195261 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | Annual PM
1064976 | 195262 | PM | ACTIV | ISSUE | 2023-06-01 03:22 |  | Annual PM
1065513 | 195263 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-10-30 10:10 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 10:09:12 AM, Reid, Emmerson ) PM Complete, see attached PM Report
1065514 | 195264 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-10-30 10:12 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 10:11:22 AM, Reid, Emmerson ) PM Complete, see attached PM Report
1065516 | 195265 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-10-30 13:38 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 01:37:23 PM, Reid, Emmerson ) PM Complete, see attached PM Report
1066474 | 195266 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-10-30 13:40 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 01:40:00 PM, Reid, Emmerson ) PM Complete, see attached PM Report
1066475 | 195267 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-10-30 13:43 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 01:42:27 PM, Reid, Emmerson ) PM Complete, see attached PM Report
1066476 | 195268 | PM | CMPLT | CMPLT | 2023-06-01 03:22 | 2023-10-30 13:57 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 01:56:29 PM, Reid, Emmerson ) PM Complete, see attached PM Report
68989 | 195269 | PM | CMPLT | CMPLT | 2023-06-01 08:03 | 2023-06-01 08:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/01/2023 08:03:17 AM, Rhuland, Courtney) Showing battery issue error Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
| 195270 | CM | CMPLT | CMPLT | 2023-06-01 08:50 | 2023-06-01 08:51 | Ordered Cast Iron weight | (TC 06/01/2023 08:50:01 AM, Chukri, Samir) ordered items from Ulin
66516 | 195271 | CM | CMPLT | CMPLT | 2023-06-01 08:56 | 2023-06-01 08:58 | BP not inflating | (TC 05/31/2023 08:56:45 AM, Rendon Lopez , Elizabeth) Confirmed issue, removed monitor from triage. Replaced BP assembly and recorder assy. Tested all functions-OK
1055177 | 195272 | CM | CMPLT | CMPLT | 2023-06-01 08:59 | 2023-06-02 09:16 | Broken | (TC 05/31/2023 08:58:56 AM, Rendon Lopez , Elizabeth) Running on user mode, no alarms (TC 06/01/2023 09:15:07 AM, Rendon Lopez , Elizabeth) Downloaded error log, error 240.4150.5, bottle PS malfunction. Replaced PS, calibrated, Completed check in tests, all passed
69208 | 195273 | CM | CMPLT | CMPLT | 2023-06-01 09:00 | 2023-06-02 09:15 | Door latch broken | (TC 05/31/2023 09:00:43 AM, Rendon Lopez , Elizabeth) Unlatched handle from front door. Adjusted, replaced screw. Running on user mode, no issues (TC 06/01/2023 09:14:28 AM, Rendon Lopez , Elizabeth) Completed check in tests, all passed
| 195274 | PA | CMPLT | CMPLT | 2023-06-01 09:25 | 2023-06-08 14:00 | Portable doppler order | (TC 06/01/2023 09:27:38 AM, Scero, Tomlyn) received quote from stat medical. ordered doppler (TC 06/06/2023 01:57:45 PM, Scero, Tomlyn) doppler recieved. 1067216.
501224 | 195275 | CM | CMPLT | CMPLT | 2023-06-01 10:18 | 2023-06-01 10:24 | Syngo updated, requires IP change and net config | (TC 05/31/2023 02:18:00 PM, Fan, Tak) Updated but did not work. Rob sent network request to MSH. (TC 05/31/2023 02:18:00 PM, Fan, Tak) Needed additional system info. Office was closed, could not access room. Searched for old IP address. They will configure using old info. (TC 06/01/2023 08:18:00 AM, Fan, Tak) Checked system. Connected in the morning. System is DHCP. IP address is still 172.19.75.243. Sent test study, successful.
1067170 | 195276 | CM | CMPLT | CMPLT | 2023-06-01 11:41 | 2023-06-01 11:42 | Incoming inspection | (TC 06/01/2023 11:40:56 AM, Rhuland, Courtney) Unboxed and built device Functional check - OK
1067171 | 195277 | CM | CMPLT | CMPLT | 2023-06-01 11:42 | 2023-06-01 11:43 | Incoming inspection | (TC 06/01/2023 11:42:05 AM, Rhuland, Courtney) Unboxed and built device Functional check - OK
1067172 | 195278 | CM | CMPLT | CMPLT | 2023-06-01 11:43 | 2023-06-01 11:44 | Incoming inspection | (TC 06/01/2023 11:43:11 AM, Rhuland, Courtney) Unboxed and built device Functional check - OK
| 195279 | CM | ACTIV | ISSUE | 2023-06-01 13:48 |  | Parts recoveries Monthly report | (TC 06/01/2023 01:47:56 PM, Chukri, Samir) Completed parts recoveries report for May 23 (TC 07/06/2023 02:14:08 PM, Chukri, Samir) completed the manthly parts recovieries report and emailed to Raymundo. (TC 08/04/2023 11:01:25 AM, Chukri, Samir) completed parts recoveries report and emailed to Raymundo (TC 08/17/2023 03:04:19 PM, Chukri, Samir) Meeting with Nima and andrew for the O.R surgical table . (TC 08/17/2023 03:07:08 PM, Chukri, Samir) filter the Parts ordered from Hill-Rom . regarding the O.R surgical table (TC 09/06/2023 03:15:30 PM, Chukri, Samir) completed parts recoveries for Aug 2023 (TC 12/04/2023 02:05:06 PM, Chukri, Samir) Completed Nov 23, parts recoveries and emailed to Raymundo. (TC 01/05/2024 03:18:24 PM, Chukri, Samir) completed the parts recoveries report and email it to Raymundo. (TC 02/07/2024 11:49:43 AM, Chukri, Samir) completed Parts recoveries report for Jan 2024 and emailed to Raymundo (TC 03/05/2024 01:11:25 PM, Chukri, Samir) Completed the parts recoveries reports for Feb 24 (TC 04/04/2024 12:12:56 PM, Chukri, Samir) completed parts recoveries and emailed to Raymundo. (TC 05/01/2024 02:01:47 PM, Chukri, Samir) completed parts recoveries for Apr 2024, and emailed to Raymudo. (TC 07/10/2024 08:01:16 AM, Chukri, Samir) completed parts recoveries for Jun 2024
| 195280 | CM | ACTIV | ISSUE | 2023-06-01 13:55 |  | Weekly Meetings for Equipment Replacement Planning Project Team
1066114 | 195281 | FC | CMPLT | CMPLT | 2023-06-01 14:44 | 2023-06-01 14:46 | QC check | (TC 06/01/2023 02:44:35 PM, Canjimavadivel, Kamallesh ) QC check complete
1061333 | 195282 | FC | CMPLT | CMPLT | 2023-06-02 07:30 | 2023-06-02 07:31 | QC checklist | (TC 06/02/2023 07:30:28 AM, Rhuland, Courtney) QC checklist complete Returned to 11 North
502044 | 195283 | CM | CMPLT | CMPLT | 2023-06-02 07:39 | 2023-06-02 07:40 | brake for arm rotate
1048390 | 195284 | CM | CMPLT | CMPLT | 2023-06-02 07:50 | 2023-06-02 07:55 | PM | (TC 06/02/2023 07:50:58 AM, Chukri, Samir) Completed PM performed by Getinge certificate of Calibration attached
| 195285 | OC | CMPLT | CMPLT | 2023-06-02 08:06 | 2023-06-02 08:09 | Kamallesh On-Call; May 26 - June 1, 2023 | (TC 06/01/2023 05:15:00 PM, Canjimavadivel, Kamallesh ) Got call from 11S regarding leaking oxygen from room 15 and requested someone come tighten with wrench. Advised this call should be directed to building services. Did not need to come in and no further follow-up.
| 195286 | CM | CMPLT | CMPLT | 2023-06-02 08:48 | 2023-07-05 15:36 | Spare ECG patient data cables | (TC 06/02/2023 08:50:31 AM, Rendon Lopez , Elizabeth) Placing order for spare cables Received
66475 | 195287 | CM | CMPLT | CMPLT | 2023-06-02 09:12 | 2023-06-02 09:14 | Leads not picking up signal | (TC 06/01/2023 09:12:33 AM, Rendon Lopez , Elizabeth) Found damaged patient data cable, replaced. Testing (TC 06/02/2023 09:12:33 AM, Rendon Lopez , Elizabeth) Tested-OK. Replaced some leads labels
1060596 | 195288 | CM | CMPLT | CMPLT | 2023-06-02 10:45 | 2023-06-02 10:47 | PM-MAJOR- Annual PM For: Multimeter | (TC 06/02/2023 10:45:17 AM, Chukri, Samir) PM completed performed by Getinge Certificate of Calibration attached
1062028 | 195289 | CM | CMPLT | CMPLT | 2023-06-02 10:52 | 2023-06-02 12:17 | Controls not working on R side or at bottom | (TC 06/02/2023 12:16:07 PM, Scero, Tomlyn) Took bed up to the 12th floor for assesment. Power cycled the bed and tested all the functions. Ned working properly. Took it back to the 10th floor and notified the CCU.
53078 | 195290 | CM | CMPLT | NFF | 2023-06-02 10:53 | 2023-06-02 10:55 | PM-MAJOR- Annual PM For: Pressure Meter, Digital | (TC 06/02/2023 10:53:16 AM, Chukri, Samir) Completed PM performed by Getinge Certificate of Calibration attached
1054464 | 195291 | CM | CMPLT | CMPLT | 2023-06-02 10:53 | 2023-06-02 10:59 | side rail defective | (TC 04/04/2023 10:55:49 AM, Cortez, John) service was requested from hillrom (TC 05/23/2023 10:57:29 AM, Cortez, John) Bed repaired by Vince from Hillrom Checked functionality before turning to service
52232 | 195292 | PM | CMPLT | CMPLT | 2023-06-02 14:48 | 2023-06-06 09:49 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/02/2023 02:48:29 PM, Rhuland, Courtney) Sidewall and porthole latches replaced (recall) Missing corner grommet replaced LV and main voltage calibrated (TC 06/06/2023 09:43:41 AM, Rhuland, Courtney) Elevating platform plate tightened Exam light replaced PM complete All tests passed Leakage - 46.1uA Ground resistance - 170mOhms
1062312 | 195293 | CM | CMPLT | CMPLT | 2023-06-02 15:58 | 2023-06-02 15:59 | Patient helper bracket installation | (TC 06/02/2023 03:58:13 PM, Cortez, John) Installed bracket as requested
1062031 | 195294 | CM | CMPLT | CMPLT | 2023-06-02 15:59 | 2023-06-02 16:01 | "not turning on" | (TC 06/02/2023 03:59:38 PM, Cortez, John) Footboard not seated correctly, linens caught underneath Reseated and verified functionality
| 195295 | CM | CMPLT | CMPLT | 2023-06-03 07:51 | 2023-06-03 12:01 | 11N Decanting | (TC 06/05/2023 11:58:48 AM, Chukri, Samir) completed the 11N Decanting to 18th floor
| 195296 | CM | CMPLT | CMPLT | 2023-06-03 09:49 | 2023-06-03 12:12 | Replace the head tope pice | (TC 06/03/2023 12:11:52 PM, Chukri, Samir) Replaced the tope pice.
| 195297 | CM | CMPLT | CMPLT | 2023-06-04 08:43 | 2023-06-04 08:39 | Replace the head tope pice | (2023-06-03 8:39 AM Samir Chukri) Completed and replaced the top mattress.
66521 | 195298 | CM | CMPLT | CMPLT | 2023-06-05 08:37 | 2023-06-05 08:38 | Recorder not feeding paper | (TC 06/02/2023 08:36:53 AM, Rendon Lopez , Elizabeth) Removed and replaced fetal monitor with spare Replaced fetal recorder with new one. Calibrated, tested allfunctions-OK. Returned to ODU L
1061204 | 195299 | CM | CMPLT | CMPLT | 2023-06-05 08:46 | 2023-06-05 08:57 | OR Lights not connecting to WiFi | (TC 06/05/2023 08:47:34 AM, Reid, Emmerson ) The iPad had a message that the lights were down because of no wifi. I turned off and on the wifi on light 1 (it was set to be on). Light 1 started working. I repeated this for the other two lights and they are all working now.
| 195300 | CM | CMPLT | CMPLT | 2023-06-05 09:10 | 2023-06-05 09:13 | Masimo safety Net, not working in 18th | (TC 06/05/2023 09:10:33 AM, Chukri, Samir) working with Nara and Naz to solve the problem found root # 11-6 used in room 1115B . Naz will complete the work with Masimo Tech .
67944 | 195301 | CM | CANCL | CANCL | 2023-06-05 09:28 |  | Wrong serial number
67943 | 195302 | CM | CMPLT | CMPLT | 2023-06-05 09:28 | 2023-07-14 13:59 | Wrong serial number | (TC 07/14/2023 01:58:28 PM, Rhuland, Courtney) Found incubator and updated serial number
66683 | 195303 | CM | CMPLT | CMPLT | 2023-06-05 09:38 | 2023-08-21 07:17 | Wrong serial number | (TC 08/21/2023 07:16:16 AM, Rhuland, Courtney) Serial number updated
| 195304 | CM | CMPLT | CMPLT | 2023-06-05 09:45 | 2023-06-16 13:45 | misc biomed - June | (TC 06/02/2023 09:44:39 AM, Cortez, John) Spoke with Natalia (interim PCM) about equipment in 10N that can be thrown out She asked about a Hover Mat, and if we look at it. Jessica says we have this model in WIH Natalie got back to me and said they want to keep older bladder scanners. Also said that some older RNs like using Hover Mat for falls (TC 06/02/2023 09:49:11 AM, Cortez, John) Helped Naz with Masimo central move (TC 06/07/2023 04:32:58 PM, Cortez, John) update WOs on TMS (TC 06/07/2023 04:33:18 PM, Cortez, John) Nurse from CCU accidently discharged wrong patient. They want to verify if patient data during a code blue between 1:17-1:40pm is still accessible Checked clinical audit and confirmed patient data for MRN in question is complete (with exception between 3:30 to 3:40pm, when MRN was admitted into wrong room) (TC 06/07/2023 04:38:30 PM, Cortez, John) Emergency Picix upgrades with kam. 2 of 3 centrals failed and we had to setup backups Issue resolved, but updates will be halted in the meantime (TC 06/06/2023 04:38:30 PM, Cortez, John) Emergency Picix upgrades with kam. 2 of 3 centrals failed and we had to setup backups Issue resolved, but updates will be halted in the meantime (TC 06/05/2023 09:30:26 AM, Cortez, John) Interventional radiology called saying their portable monitors arent saving NIBP interval settings and alarm limits. After looking into it, found out they were set to manual mode and not auto. Changed setting and saved config (TC 06/05/2023 09:33:12 AM, Cortez, John) Spoke with Natalia (10N PCM) about hovermat, and she said she will hold onto it for now. She will reach out to the vendor to look into training for newer nurses (TC 06/09/2023 01:44:37 PM, Cortez, John) Katie Lai (CNS) requested training AED, stopped by office to drop it off (TC 06/05/2023 03:01:50 PM, Cortez, John) huddle (TC 06/09/2023 03:03:13 PM, Cortez, John) Kam and I talked with Vince from baxter about WO processes (TC 06/09/2023 03:03:13 PM, Cortez, John) Kam and I talked with Vince from baxter about WO processes (TC 06/13/2023 10:59:07 AM, Cortez, John) Organized 10S equip room Setup/supplied spare masimo monitors with missing accessories. Most accessories (sensors, cuffs, power cords) were already scattered around equip room Missing 3 adult cuffs (TC 06/12/2023 11:02:12 AM, Cortez, John) Grabbed pumps due for PM from ICU materials room (TC 06/12/2023 11:03:02 AM, Cortez, John) ICU rounds to serach for 2 defibs. Found in pod B and PMH (TC 06/14/2023 01:05:11 PM, Cortez, John) 11S wasn't able to admit patients on Masimo Roots. One root didn't even have option to admit. Wifi was switched off on the one root. Showed CNS how to admit correctly (TC 06/15/2023 03:38:54 PM, Cortez, John) Retrieved AED + training pads from 10S, searched for 8 microstreams due for PM in 6th floor storage. 6 were found, 1 is down in biomed (TC 06/15/2023 04:02:12 PM, Cortez, John) TMS cleanup (TC 06/16/2023 01:06:30 PM, Cortez, John) ICU called saying there was a test patient admitted into monitor of C35, and they couldn't remove it. Went to C35 to discharge test profile (TC 06/16/2023 01:44:38 PM, Cortez, John) PACU called saying ECG signal noisy in bay 8 Went up and replaced leads, electrodes ok
1056425 | 195305 | CM | CMPLT | CMPLT | 2023-06-05 09:54 | 2023-06-05 09:55 | Radical module error | (TC 06/02/2023 09:54:10 AM, Cortez, John) Replaced Radical module All parameters tested ok
1061215 | 195306 | CM | ACTIV | ISSUE | 2023-06-05 11:10 |  | Vitals not working on wall monior | (TC 06/05/2023 11:10:07 AM, Reid, Emmerson ) Issue with the DVI on the vitals showing on the wall monitor.
1054813 | 195307 | CM | CMPLT | CMPLT | 2023-06-05 11:17 | 2023-06-05 11:20 | Membrane capacity 15% | (TC 06/05/2023 11:17:40 AM, Canjimavadivel, Kamallesh ) Replaced expiratory cassette membrane and reset counter
| 195308 | CM | CMPLT | CMPLT | 2023-06-05 12:09 | 2023-06-05 12:09 | TMS- Repair Cost. | (TC 06/05/2023 12:09:26 PM, Chukri, Samir) Completed.
1049905 | 195309 | CM | CMPLT | CMPLT | 2023-06-05 13:25 | 2023-06-05 13:26 | Battery charger slot damaged | (TC 06/05/2023 01:25:08 PM, Fan, Tak) Tech reported charger not working. Confirmed not charging. Replaced top charger module (could not disassemble module due to the scerws).
| 195310 | CON | CMPLT | CMPLT | 2023-06-05 14:26 | 2023-06-02 18:27 | MSH SONOSITE CLINICAL VISITS AND PANEL DISCUSSION | (TC 06/05/2023 02:26:33 PM, Whittington, Scott) MSH SONOSITE CLINICAL VISITS AND PANEL DISCUSSION (TC 06/02/2023 06:26:00 PM, Whittington, Scott) MSH SONOSITE CLINICAL VISITS AND PANEL DISCUSSION | MSH SONOSITE CLINICAL VISITS AND PANEL DISCUSSION
1062254 | 195311 | CM | CMPLT | CMPLT | 2023-06-05 14:28 | 2023-06-06 09:38 | Motion controls not working | (TC 06/06/2023 09:37:06 AM, Baghirzade, Nara) Motion controls on patient side, side panels, and footboard were not working. Power cycled and issue was resolved.
| 195312 | CM | ACTIV | ISSUE | 2023-06-05 14:49 |  | Monday Huddle/PM meeting | (TC 06/05/2023 02:49:14 PM, Chukri, Samir) Completed. (TC 06/19/2023 02:22:08 PM, Chukri, Samir) meeting (TC 06/26/2023 02:37:47 PM, Chukri, Samir) Biomed week meeting and lunch. (TC 07/10/2023 02:15:09 PM, Chukri, Samir) Monday Huddle meeting (TC 07/17/2023 02:27:51 PM, Chukri, Samir) completed. (TC 07/31/2023 01:50:30 PM, Chukri, Samir) meeting (TC 08/14/2023 02:59:12 PM, Chukri, Samir) Monday huddle/PM meeting . (TC 08/17/2023 03:02:49 PM, Chukri, Samir) PM department meeting (TC 08/28/2023 08:46:10 AM, Chukri, Samir) Department Meeting (TC 11/20/2023 02:44:15 PM, Chukri, Samir) Monday meeting (TC 11/29/2023 02:31:49 PM, Chukri, Samir) Meeting (TC 12/13/2023 02:13:52 PM, Chukri, Samir) Monday Huddle meeting (TC 12/13/2023 02:14:31 PM, Chukri, Samir) Peroformance and future co-op meeting (TC 01/22/2024 02:22:08 PM, Chukri, Samir) Monday Meeting (TC 01/29/2024 02:03:48 PM, Chukri, Samir) attending metting. (TC 02/05/2024 07:56:36 AM, Chukri, Samir) Attending Meeting (TC 02/12/2024 02:28:39 PM, Chukri, Samir) Huddle/PM meeting (TC 02/26/2024 02:19:03 PM, Chukri, Samir) meeting (TC 03/04/2024 03:01:31 PM, Chukri, Samir) Monday Huddle/pm meeting (TC 03/18/2024 02:20:32 PM, Chukri, Samir) attending meeting (TC 04/16/2024 11:15:06 AM, Chukri, Samir) Monday Huddle. (TC 04/29/2024 02:34:26 PM, Chukri, Samir) meeting (TC 06/03/2024 03:15:02 PM, Chukri, Samir) Meeting (TC 06/21/2024 11:25:08 AM, Chukri, Samir) IS/CI/Biomed meeting (TC 08/19/2024 08:05:14 AM, Chukri, Samir) Meeting (TC 10/07/2024 03:07:51 PM, Chukri, Samir) Monday Meeting huddle
| 195313 | MEP | CMPLT | CMPLT | 2023-06-05 15:02 | 2023-06-14 14:33 | Echopac Standalone SW survey | (TC 06/05/2023 03:46:00 PM, Fan, Tak) Meeting with Dianne. Fellows might have a dongle for meetings. (TC 06/06/2023 09:27:14 AM, Fan, Tak) Found 6 of the dongles. Likely 416596EPSW1 through 416596EPSW6. Asked Grant for Site ID confirmation. 0x47997085 0x5B8FC1F3 0x7E79E361 0x69700A76 0x3D0476E7 0x11C73B46 (TC 06/06/2023 01:07:09 PM, Fan, Tak) Sorted through some old boxes. 416596EPSW5 is unaccounted for. May be 0x3FAB2D1, confirming with Grant. Confirmed that NICU isn't on a contract for ECHOPAC. Forwarded info to mangers/purchasing team.
| 195314 | CM | ACTIV | ISSUE | 2023-06-06 08:25 |  | Shopping Cart approval | (TC 06/06/2023 08:25:25 AM, Chukri, Samir) Approved shopping cart for Emmerson (TC 06/08/2023 10:27:54 AM, Chukri, Samir) approved shopping cart for Jessica which missing the supplier name. jessica will correct it with plexxuse (TC 06/13/2023 01:43:08 PM, Chukri, Samir) Ordered tools for Olcay from Canucktools (TC 06/13/2023 02:05:06 PM, Chukri, Samir) pproving shopping cart for Kamran after received approval of purchasing from Reeba O.R (TC 06/21/2023 11:01:43 AM, Chukri, Samir) Approved shopping carts for Jessica and emmerson. (TC 06/21/2023 02:37:11 PM, Chukri, Samir) Ordering Chairs and goggel for biomed department (TC 06/26/2023 02:48:19 PM, Chukri, Samir) Approvided shoping carts for Qwynn, Elizabeth and kam (TC 07/10/2023 09:55:43 AM, Chukri, Samir) Review and approved shopping cart (TC 07/14/2023 02:58:19 PM, Chukri, Samir) approvied shopping cart and ordered cable from Amazon to Kam for Alaris pump calibration. (TC 08/28/2023 11:25:26 AM, Chukri, Samir) Ordered Flow tester for Kamran from Omega. (TC 11/28/2023 08:48:52 AM, Chukri, Samir) Review and approved shopping cart from Courteny (TC 11/28/2023 02:27:45 PM, Chukri, Samir) review and approved shopping cart form Emmerson (TC 11/29/2023 02:33:52 PM, Chukri, Samir) Review and approved shoping cart for Elizabeth. (TC 12/05/2023 12:51:23 PM, Chukri, Samir) Review and approved shopping cart for Emmerson (TC 12/06/2023 08:27:53 AM, Chukri, Samir) Review shopping cart from Courtney and rejected asking her to correct the Qyt and approvded . also Review and checked the WO and approvded the shopping cart from Elizabith (TC 12/08/2023 02:09:38 PM, Chukri, Samir) Review and approved 4 shopping carts for Olympus Light source and email to Claire for rush PO. (TC 12/11/2023 09:43:37 AM, Chukri, Samir) review and approved some shopping carts. (TC 01/09/2024 09:24:20 AM, Chukri, Samir) Review and apparoved shoping carts. (TC 01/24/2024 09:13:49 AM, Chukri, Samir) Review and approved shoping cart for Nara. (TC 01/24/2024 09:36:37 AM, Chukri, Samir) review and approved shopping cart and talked to Kam to correct the WO charge back with 8 kits. (TC 01/24/2024 11:25:38 AM, Chukri, Samir) Review and Signed of for two ICU invoices (TC 01/29/2024 09:16:42 AM, Chukri, Samir) Review and approved two shopping carts for John and Elizabeth. (TC 02/01/2024 09:31:49 AM, Chukri, Samir) Approved shopping carts from yesterday. (TC 03/12/2024 11:38:34 AM, Chukri, Samir) Review and aproved shopping cart for Rutvik and ordering Amazon iteams (TC 03/19/2024 09:23:30 AM, Chukri, Samir) Review and approved shopping cart for Rutvik (TC 03/20/2024 02:27:23 PM, Chukri, Samir) search and ordered power strip for Rutvik. (TC 03/25/2024 08:31:37 AM, Chukri, Samir) Review shopping cart and talked to Rutvik to add the ordered parts with shopping cart # to the WO. (TC 03/26/2024 02:46:02 PM, Chukri, Samir) Review shopping carts and approved it (TC 04/04/2024 12:17:16 PM, Chukri, Samir) review and approved shopping carts (TC 04/08/2024 11:00:15 AM, Chukri, Samir) working with Rutvik correct his shopping cart and approved
| 195315 | PA | CMPLT | CMPLT | 2023-06-06 08:53 | 2023-07-07 13:01 | LVP display boards | (TC 06/06/2023 08:53:19 AM, Reid, Emmerson ) Placing an order for 10x LVP display boards under the following SNs: 14972065 14972079 14972152 14972082 14972066 14965393 14965724 14964839 14972003 14693629 (TC 06/27/2023 01:31:45 PM, Reid, Emmerson ) Followed up and was redirected to a different email that requested I fill out an intake complaint form. I filled this in and requested a no charge PO
1061214 | 195316 | CM | CMPLT | CMPLT | 2023-06-06 08:57 | 2023-06-06 09:03 | Light #3 needed adjustments | (TC 06/06/2023 08:56:58 AM, Reid, Emmerson ) Amanda from OR13 called about light #3 "floating". I saw the light arm was set too tight so i losened the spring in the arm to balance the light.
| 195317 | PA | CMPLT | CMPLT | 2023-06-06 09:19 | 2023-06-07 13:19 | Water for Pump PMs | (TC 06/06/2023 09:21:24 AM, Reid, Emmerson ) Ordered water for pump PMs (TC 06/07/2023 01:18:50 PM, Reid, Emmerson ) water recieved and put away
| 195318 | CM | CMPLT | CMPLT | 2023-06-06 09:35 | 2023-06-06 09:36 | Retiring devices and disposal | (TC 06/06/2023 09:35:36 AM, Baghirzade, Nara) Retired 8 filac thermometers; decomissioned from 11N move to 18th fl. Disposed of thermometers to e-waste.
| 195319 | CM | CMPLT | CMPLT | 2023-06-06 09:39 | 2023-06-05 14:00 | Biomed huddle | (TC 06/05/2023 02:00:00 PM, Baghirzade, Nara) Biomed huddle
68530 | 195320 | CM | CMPLT | CMPLT | 2023-06-06 09:41 | 2023-06-06 09:42 | Battery replacement | (TC 06/06/2023 09:41:38 AM, Baghirzade, Nara) Replaced battery for digital display. Battery PO already charged by Samir.
1053391 | 195321 | CM | CMPLT | CMPLT | 2023-06-06 09:43 | 2023-06-27 11:07 | Does not charge, AC light does not come on when plugged in | (TC 06/06/2023 02:36:08 PM, Baghirzade, Nara) When plugging the device to AC, the orange LED indicator does not light up and the device is not charging. Opened the pump to check the connector on the board; connector is good. Checked the adapter and found that when wiggling the cable around, it intermittently works. Called Ifrah at 3M and cancelled an earlier service request regarding this issue (RMA 12688000). Will order a new adapter and battery as battery is due for a change. (TC 06/27/2023 10:20:02 AM, Baghirzade, Nara) Replaced battery. Replaced power supply. Power supply connector on the board had a damaged pin and was not making good contact with the PCB. Samir desoldered the connector, replaced the damaged pin with a wire, and soldered the connector back on the PCB. This is a temporary solution until a new connector comes in. Communicated this with Sri. Physical inspection good. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 124 mmHg Set 200 mmHg - Measured 199 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time and date. Device is functional. Returned to 11S.
67761 | 195322 | PM | CMPLT | CMPLT | 2023-06-06 10:15 | 2023-06-06 10:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/06/2023 10:57:03 AM, Reid, Emmerson ) Had to run the volume test 3 times. PM Complete, see proceedure for details.
1061203 | 195323 | CM | CMPLT | NFF | 2023-06-06 11:16 | 2023-06-06 11:17 | SCD issue resolved before i got there | (TC 06/06/2023 11:16:39 AM, Reid, Emmerson ) Went up but before arriving at the OR i was told they fixed it. NFF
52252 | 195324 | CM | CMPLT | CMPLT | 2023-06-06 11:23 | 2023-06-06 11:27 | Broken flip door | (TC 06/06/2023 11:23:37 AM, Rhuland, Courtney) Broken flip door replaced Recall parts replaced sidewall bottom latch replaced Functional check - OK
1061333 | 195325 | FC | CMPLT | CMPLT | 2023-06-06 12:57 | 2023-06-06 13:02 | QC checklist | (TC 06/06/2023 01:01:24 PM, Baghirzade, Nara) QC checklist completed. Returned to 18 South surgical unit
1061202 | 195326 | CM | CMPLT | CMPLT | 2023-06-06 13:04 | 2023-06-06 13:05 | Bronch scope not connecting to overhead monitors | (TC 06/06/2023 01:03:57 PM, Reid, Emmerson ) Scope was connected to the HD port on 4K1. It was showing fine on the wall monitor but not on the other monitors for some reason. I moved it to AUX 1 and it is working on all monitors.
1061204 | 195327 | CM | CMPLT | CMPLT | 2023-06-06 13:06 | 2023-06-06 13:06 | Microscope not showing video | (TC 06/06/2023 01:05:43 PM, Reid, Emmerson ) The camera module was turned off on the microscope. I turned it on, working now.
1055515 | 195328 | CM | CMPLT | NFF | 2023-06-06 14:02 | 2023-06-06 14:03 | Reported that it's measuring low | (TC 06/06/2023 02:01:51 PM, Baghirzade, Nara) Reported that the thermometer was measuring on the lower side; compared it to other thermometers and no differences in temperatures. Device is in good physical condition. Cleaned the probe and ensured that it is seated properly. No fault found.
68373 | 195329 | CM | CMPLT | CMPLT | 2023-06-06 14:09 | 2023-06-29 13:59 | Gas measurement failure during case | (TC 06/06/2023 02:10:35 PM, Scero, Tomlyn) Email from Peter said that device stopped working during the case. Attached a picture. I sent the picture and the DSR to Syed from Draeger for advice. Initial response was that it it could be a blocked sample line or water trap. They said they replaced the water trap when it happened and I replace the sample line and tubing when I ran the self test and Volume control mode with no errors. Waiting for a response from draeger. (TC 06/08/2023 01:01:56 PM, Scero, Tomlyn) Checked the logs again, redownloaded twice. PGA errors still present. Syed advised that I replace the Pneumatic complete in the PGA. Emailed Tanya for a price quote and ordered through ereq (TC 06/19/2023 02:05:10 PM, Scero, Tomlyn) Spare part was received. Brought the unit down to RT/biomed for repair. (TC 06/27/2023 02:55:47 PM, Scero, Tomlyn) PGA pnuematics complete replaced. self test complete and passed. PGA/SGA IPM-L tests complete. (TC 06/29/2023 01:57:53 PM, Scero, Tomlyn) Ran device on different vent modes to ensure proper gas measurements. Downloaded logs to check for any errors. None found.
1048500 | 195330 | CM | CMPLT | CMPLT | 2023-06-06 14:18 | 2023-06-16 09:59 | MSH 5 US EPIQ PROBE CABLE BUNGEE CORDS | (TC 06/06/2023 02:55:18 PM, Whittington, Scott) MSH 5 US EPIQ PROBE CABLE BUNGEE CORDS site 75025255 REF# case #0121796770 PN# (TC 06/16/2023 09:56:56 AM, Whittington, Scott) MSH 5 US EPIQ PROBE CABLE BUNGEE CORDS site 75025255 REF# case #0121796770 PN# 453561759491 assy, EASY-CLIP CABLE MANAGEMENT MOBILE RECEIVED 3 SETS, INSTALLED 2 ON EPIQ IN MSH 5 US AND 1 ON MSH ANGIO EPIQ | MSH 5 US EPIQ PROBE CABLE BUNGEE CORDS site 75025255 REF# case #0121796770 PN# 453561759491 assy, EASY-CLIP CABLE MANAGEMENT MOBILE
| 195331 | CM | CMPLT | CMPLT | 2023-06-06 14:20 | 2023-06-06 14:20 | Communications Committee meeting | (TC 06/06/2023 02:20:09 PM, Baghirzade, Nara) Communications Committee meeting
| 195332 | CM | ACTIV | ISSUE | 2023-06-06 14:21 |  | Medical device acquistion and acceptance policy | (TC 06/06/2023 02:21:04 PM, Baghirzade, Nara) Working on medical device acquisition and acceptance policy; researching other hospital policies for reference.
| 195333 | CM | ACTIV | ISSUE | 2023-06-06 14:31 |  | Communications Committee | (TC 06/06/2023 02:30:42 PM, Chukri, Samir) Meeting (TC 02/21/2024 02:18:45 PM, Chukri, Samir) Meeting (TC 04/03/2024 12:35:18 PM, Chukri, Samir) Meeting with Adeel and Tmolyn to change and renew the RFP documents
| 195334 | CM | ACTIV | ISSUE | 2023-06-06 14:32 |  | MRI Incubator Discussion Points | (TC 06/06/2023 02:32:09 PM, Chukri, Samir) MRI Meeting (TC 06/07/2023 07:55:16 AM, Chukri, Samir) Meeting with Michael and AME Manufacture manager . (TC 06/13/2023 01:41:29 PM, Chukri, Samir) Contact Nav Biomed tech and Jason R.T from Sick kits asking for MRI transort incubator, there is no Blender use in sick kits . (TC 06/19/2023 09:32:15 AM, Chukri, Samir) Review Maisha letter and the Manual disinfection recomended in the manual . (TC 07/13/2023 02:42:19 PM, Chukri, Samir) Meeting with SA Tabish regarding the cleaning procedure (TC 12/11/2023 09:41:07 AM, Chukri, Samir) Removed Ventilators from old MRI incubator for disposal next week (TC 02/23/2024 01:34:58 PM, Chukri, Samir) checked with Adeel and qwynn, regarding the Fan motor defective parts price and replied to Michael. (TC 02/28/2024 02:45:13 PM, Chukri, Samir) meeting with Nima and John refe to MRI incubator issues.
1053391 | 195335 | CM | CMPLT | CMPLT | 2023-06-06 14:40 | 2023-06-26 08:21 | Ordering parts | (TC 06/06/2023 02:50:59 PM, Baghirzade, Nara) SC 1002973992 (TC 06/26/2023 08:19:31 AM, Baghirzade, Nara) PO 7202226823
| 195336 | CM | ACTIV | ISSUE | 2023-06-06 14:40 |  | MRI Blender checked | (TC 06/06/2023 02:40:28 PM, Chukri, Samir) checked and explain to Qwynn, if need to adjust the Air and O2 pressure .
1057758 | 195337 | CM | CMPLT | CMPLT | 2023-06-06 15:03 | 2023-06-07 09:42 | Created by Failed PM WO # 194599 : PM MAJOR-Monitors, Physiologic, Central Station | (TC 06/06/2023 03:04:44 PM, Canjimavadivel, Kamallesh ) Swapped out PIC with biomed spare. Reconfigured and verified working as expected (TC 06/07/2023 09:40:42 AM, Canjimavadivel, Kamallesh ) Reinstalled PICiX software from USB key. Configured back to original settings in the unit. Reconfingured spare PICiX back to original settings | (TC 06/06/2023 03:02:14 PM, Canjimavadivel, Kamallesh ) Installed patches failed. PC auto shutdown prior to complete
1056190 | 195338 | CM | CMPLT | CMPLT | 2023-06-06 15:08 | 2023-06-06 15:11 | Created by Failed PM WO # 194569 : PM MAJOR-Monitors, Physiologic, Central Station | (TC 06/06/2023 03:08:42 PM, Canjimavadivel, Kamallesh ) Swapped out with spare PIC from biomed shop. Reinstalled PICiX software from usb key. Reconfigured to original settings in ER. Working as expected. | (TC 06/06/2023 03:06:31 PM, Canjimavadivel, Kamallesh ) Patches installion failed. PC auto shutdown prior to completion.
67961 | 195339 | CM | CMPLT | CMPLT | 2023-06-07 08:10 | 2023-06-07 08:53 | Hood cracked | (TC 06/07/2023 08:10:28 AM, Rhuland, Courtney) Hood replaced Porthole grommet replaced Air probe connector wires reseated Back air probe cover replaced Missing screws replaced Melted light replaced with new light Functional check - OK
1059828 | 195340 | PM | CMPLT | CMPLT | 2023-06-07 08:43 | 2023-06-07 09:40 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 06/07/2023 09:39:25 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1061209 | 195341 | CM | CMPLT | CMPLT | 2023-06-07 09:02 | 2023-06-07 09:03 | ESU was not connected properly | (TC 06/07/2023 09:01:48 AM, Reid, Emmerson ) The ESU was on a mobile cart and was not connected properly to the smoke evac. I fixed this and synced the devices. working now.
1061210 | 195342 | CM | CMPLT | CMPLT | 2023-06-07 09:03 | 2023-06-07 09:05 | Vitals not working on monitors | (TC 06/07/2023 09:03:17 AM, Reid, Emmerson ) Vitals was not showing up on any monitor. I checked the DVI cable and it was a little lose so i tightened it and it came up on the screens. Note that the image on the wall monitor was having the sync issues also, but working for now.
1059161 | 195343 | CM | CMPLT | NFF | 2023-06-07 09:06 | 2023-06-07 09:54 | X3 Not communicating and having oncoing issues | (TC 06/07/2023 09:05:49 AM, Reid, Emmerson ) Was told the X3 was not communicating to the main monitor. I tried and it was working normally. Then I recieved an email from Wendy saying they have been having ongoing issues with this bay/monitor invoving not reading arrhythmias properly. Going to look at again with a simulator. (TC 06/07/2023 09:52:40 AM, Reid, Emmerson ) Checked the monitor with the simulator. Everything was working fine. Note that the ABP has not been zeroed.
65868 | 195344 | PM | CMPLT | CMPLT | 2023-06-07 09:08 | 2023-06-07 09:10 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 06/07/2023 09:08:34 AM, Baghirzade, Nara) PM completed. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
63276 | 195345 | PM | CMPLT | CMPLT | 2023-06-07 09:11 | 2023-06-07 09:13 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 06/07/2023 09:13:18 AM, Baghirzade, Nara) PM completed. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
65864 | 195346 | PM | CMPLT | CMPLT | 2023-06-07 09:15 | 2023-06-07 09:15 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 06/07/2023 09:14:52 AM, Baghirzade, Nara) PM completed. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
1067216 | 195347 | ACC | CMPLT | CMPLT | 2023-06-07 09:43 | 2023-06-08 14:00 | Incoming Inspection | (TC 06/08/2023 02:00:06 PM, Scero, Tomlyn) incoming inspection complete
1057722 | 195348 | CM | CMPLT | CMPLT | 2023-06-07 09:49 | 2023-06-07 09:59 | battery malfunction alert on defib | (TC 06/02/2023 09:50:58 AM, Cortez, John) battery fault led lit up reconditioned battery, fault led cleared
1059532 | 195349 | CM | CMPLT | CMPLT | 2023-06-07 09:51 | 2023-06-07 11:20 | Blue screen | (TC 06/07/2023 11:18:48 AM, Canjimavadivel, Kamallesh ) Reinstalled PICiX software for USB. Reconfigured back to original PACU settings. Put back in storage
1056454 | 195350 | CM | CMPLT | NFF | 2023-06-07 10:13 | 2023-06-07 10:14 | BP cuff not registering | (TC 06/07/2023 10:13:14 AM, Rhuland, Courtney) Ran functional check - could not replicate error Functional check - OK No fault found
| 195351 | CM | CMPLT | CMPLT | 2023-06-07 11:17 | 2024-01-05 14:46 | Training, Learning, and Other Works | (TC 06/09/2023 11:27:52 AM, Ferreira, Qwynn) Removed prints Cleaned them Measured some features Cleaned gaps Test fitted (TC 06/09/2023 02:40:11 PM, Ferreira, Qwynn) Measured the model files for reference Packaged models for hand off Explained models and guides to surgeons for planned VSP Worked on test coupon and testing sheet for future testing prints. (TC 06/12/2023 11:11:36 AM, Ferreira, Qwynn) Recevied information for VR software and contact for in person meeting or questions. Worked on obtaining some clinical examples to try and print. Rib and tumor model will pass to emmerson to try and clean up/print if he has the time. Worked on other models including anerysms, heart models, and other artery models Checked on 12th floor for rounds - saw bed was still there - followed up with stryker and the request was not put in so they will look into it this week. (TC 06/12/2023 02:33:41 PM, Ferreira, Qwynn) Worked on more models to print in both clear and elastic. Created a renal artery model, splenic artery model, and based splenic artery model Need to repair and edit the models to get them ready for printing. Currently creating 3 different models to print in clear. (TC 06/13/2023 10:40:12 AM, Ferreira, Qwynn) Meeting to view viable cases and discuss conditions (TC 06/13/2023 10:40:50 AM, Ferreira, Qwynn) Worked on medical models - spenic and renal arteries. Cleaned up TMS for work orders that need to be cancelled. (TC 06/14/2023 01:52:27 PM, Ferreira, Qwynn) MRI Incubator discussion (TC 06/14/2023 01:56:54 PM, Ferreira, Qwynn) Set up meeting for VR Read papers relating to IP, patents, trademarks, copyright for 3d printing medical components and parts Learned some differences between reciprocating saws vs oscillating saws (TC 06/15/2023 02:54:44 PM, Ferreira, Qwynn) PM Meeting (TC 06/16/2023 01:08:29 PM, Ferreira, Qwynn) Team Meeting (TC 06/19/2023 12:21:32 PM, Ferreira, Qwynn) Pharmacy meeting to go over solutions for the speaker issue (TC 06/19/2023 02:01:58 PM, Ferreira, Qwynn) General PM Meeting (TC 06/20/2023 11:21:00 AM, Ferreira, Qwynn) OMF/ENT Meeting (TC 06/21/2023 09:58:13 AM, Ferreira, Qwynn) Discussed with Tanya about research and whe nwe can order products and the REB review (TC 06/21/2023 09:59:02 AM, Ferreira, Qwynn) Removed supports and cleaned off rough edges of the prostate model. Filled hollow cavity with resin for colour distinction Clear coated one of the models to see if it makes a difference Reviewed the study submission for details while clear coat was settling (TC 06/21/2023 12:36:07 PM, Ferreira, Qwynn) Worked with Janet to replace all of the transport numbers from 1-7 as well as update the transport incubator list. All renumbering is complete. (TC 06/21/2023 02:39:51 PM, Ferreira, Qwynn) Fleet meeting/P3A Discuss new masimo alerting (TC 06/22/2023 11:05:49 AM, Ferreira, Qwynn) Finished draft of the quality control test coupon excel sheet (TC 06/22/2023 02:54:32 PM, Ferreira, Qwynn) VR Meeting - Elucis (TC 06/23/2023 09:27:02 AM, Ferreira, Qwynn) Assembled new chairs that Samir purchased (TC 06/26/2023 09:33:00 AM, Ferreira, Qwynn) Took inventory and ordered additional parts for incubator (TC 06/27/2023 08:55:51 AM, Ferreira, Qwynn) ENT/OMF meeting (TC 06/27/2023 01:22:25 PM, Ferreira, Qwynn) Put an order together for the research project to fund. Communciated with supervisor on how to reimburse the amount spent. Worked with Tanya on repaying strategy. Placed order and will wait for it to arrive for the invoice (TC 06/27/2023 01:23:57 PM, Ferreira, Qwynn) Worked on a Kidney model to print for Dr. Swallow to see potential interest. (TC 06/28/2023 10:04:01 AM, Ferreira, Qwynn) Cleaned and filled print with resin and left before polishing (TC 06/28/2023 10:04:32 AM, Ferreira, Qwynn) TEsted the abdominal depressor to 50lbs - 25lbs per side Noticed the top was wedging apart and added zipties to the top to better secure them for testing. No other visible defects. (TC 06/28/2023 02:01:35 PM, Ferreira, Qwynn) Meeting with Dr. Haider Abdominal Depressor went over design, changes, future work, patents, REB, and setting up time to get the coil back and run tests. Made notes for the work being done - and will discuss further during 3D printing meeting. Communicated with Dr. Warren that we will be needing to borrow the coil once again. (TC 06/28/2023 02:42:54 PM, Ferreira, Qwynn) Received a call from rex stating a cmac blade was blurry Used a magnifying lense to check and noticed the camera lense was scratched Reacched out to Storz to see what the options are, until I noticed that the device had been retired. Removing from service and notified Rex that the camera is retired. (TC 06/29/2023 02:00:00 PM, Ferreira, Qwynn) Maisha held a meeting for onerecall (TC 06/29/2023 12:26:00 PM, Ferreira, Qwynn) worked with emmerson to go over design options and thought process for the depressor Started making ultrasound clip for the doppler as well as found a way so if the clip breaks on the physical doppler, we can mount it to the post of the stand. (TC 06/30/2023 09:58:39 AM, Ferreira, Qwynn) Test fit the ultrasound clip and it fit/worked. MAde 3 different version as OR, L&D, and mother baby use different configurations. Made one for the roll stand, one to replace a broken clip, and one for the back mount screw hole. Waiting for the new batch to finish printing to test fit new designs for mounting. (TC 06/30/2023 01:56:03 PM, Ferreira, Qwynn) test fit the new printed devices, two worked, one did not print well, however, it still functioned as it should. Some edits will be made to the pole clamp design, and will reprint that for the carted dopplers. (TC 07/04/2023 10:51:28 AM, Ferreira, Qwynn) ENT/Head/Neck Meeting for cases (TC 07/05/2023 08:31:50 AM, Ferreira, Qwynn) Helped Tomi assemble her desk Moved everything into the shop/cleared shop Set up my desk with 3 tear shelves Organized my area to get rid of old papers, equipment, etc (TC 07/05/2023 08:33:35 AM, Ferreira, Qwynn) Meeting to discuss the Boom PMs (TC 07/06/2023 10:56:51 AM, Ferreira, Qwynn) removed welch allyn stands Cleared printer and printed parts (TC 07/07/2023 09:42:24 AM, Ferreira, Qwynn) Discused the process for setting up the VR software and how many days we will get 60 days total - 30 on one pc and 30 for both devices (TC 07/07/2023 09:43:07 AM, Ferreira, Qwynn) Worked with Welch Allyn on trouble shooting the intermittent issue on the VSM. So far no result as it is still booting on and off (TC 07/07/2023 09:43:55 AM, Ferreira, Qwynn) Researched filament and straps for the depressor Sent to adeel for ordering through amazon (TC 07/07/2023 01:11:50 PM, Ferreira, Qwynn) 1 on 1 meeting (TC 07/07/2023 01:12:12 PM, Ferreira, Qwynn) Brought air tank to shipping and swapped the tanks. Also installed quick connects for easier using for everyone. (TC 07/07/2023 01:33:09 PM, Ferreira, Qwynn) Rounds for unit as well as for transport incubators (TC 07/10/2023 02:19:24 PM, Ferreira, Qwynn) cybersecurity meeting (TC 07/10/2023 02:19:39 PM, Ferreira, Qwynn) Monday Huddle (TC 07/10/2023 02:19:58 PM, Ferreira, Qwynn) Obtaining PM equipment for myself and courtney to work on. Grabbed two incubators, some lights, and a broken endoscope (TC 07/11/2023 09:20:18 AM, Ferreira, Qwynn) ENT/Head?Neck meeting to discuss possible cases (TC 07/11/2023 03:43:11 PM, Ferreira, Qwynn) NICU Meeting (TC 07/12/2023 12:44:16 PM, Ferreira, Qwynn) training for vr software (TC 07/12/2023 12:44:59 PM, Ferreira, Qwynn) searching uline products for the repair shop. Worked with samir to select cabinets and the layout for the shop. (TC 07/12/2023 02:38:44 PM, Ferreira, Qwynn) NICU Upgrades to the Overview stations 3O, 4O, and 5O. (TC 07/13/2023 12:41:10 PM, Ferreira, Qwynn) Edit the email to send to all clinical staff for 3d printing Showed Courtney how to order through plexxus to be familiar with the steps Discussed the poster design with emmerson (TC 07/13/2023 01:35:36 PM, Ferreira, Qwynn) MRI Incubator cleaning meeting (TC 07/13/2023 03:48:53 PM, Ferreira, Qwynn) NICU Meeting with John, Also discussed Boom project and spoke with Naz to figure out the panda warmer and omni bed situation (TC 07/14/2023 01:02:38 PM, Ferreira, Qwynn) Worked with emmerson on the poster to finalize it Worked with him to generate the email for reeba and my email as well. Looked into form for the VSP outlines Had some ideas for labelling the prints (TC 07/14/2023 02:01:29 PM, Ferreira, Qwynn) Performed the upgrade for the overview stations 6-8 Perfrormed a count of incubator lights for John Grtabbed a incubator from M&B that had a damaged light, porthole door, and due for PM (TC 07/14/2023 02:02:54 PM, Ferreira, Qwynn) Performed the Overivew upgrades 6-8 Performed incubator light count for John Grabbed damaged incbuator that was due for PM (TC 07/14/2023 02:02:54 PM, Rhuland, Courtney) Performed the Overivew upgrades 6-8 Performed incubator light count for John Grabbed damaged incbuator that was due for PM (TC 07/14/2023 02:01:29 PM, Ferreira, Qwynn) Performed the upgrade for the overview stations 6-8 Perfrormed a count of incubator lights for John Grtabbed a incubator from M&B that had a damaged light, porthole door, and due for PM (TC 07/18/2023 11:01:53 AM, Ferreira, Qwynn) Recevied case from joel obtained cd from robin started segmentation - bad scans - 2mm will continue segmentations (TC 07/18/2023 03:25:27 PM, Ferreira, Qwynn) Worked with Kam to clean shop and organize what is to keep and what is not. Threw out old cables and parts no longer in use. Placed shelves in respective places (TC 07/18/2023 03:32:29 PM, Ferreira, Qwynn) Discussed the demonstration for 3D printing/VR tables setup, poster setup, and what to have ready for next week. (TC 07/20/2023 10:20:44 AM, Ferreira, Qwynn) Worked on segmations for the case - davies will be meeting with me at 1130 to discuss (TC 07/20/2023 10:21:15 AM, Ferreira, Qwynn) Created labels for the models that will be showcased for the demo (TC 07/21/2023 09:23:26 AM, Ferreira, Qwynn) Finalized guides and sent them for printing (TC 07/24/2023 11:01:07 AM, Ferreira, Qwynn) Post processing guides Cleaned them again to remove resin Sanded Washed Cured for 15min just to ensure no remaining resin (TC 07/26/2023 10:08:25 AM, Ferreira, Qwynn) spoke with raayed to discuss 3d scanners, and which one would be a suitable option. looking to setup a virtual demo to showcase the scanner (TC 07/24/2023 09:25:48 AM, Ferreira, Qwynn) 3D stand in lobby (TC 07/25/2023 09:26:10 AM, Ferreira, Qwynn) 3D stand in lobby (TC 07/26/2023 09:26:29 AM, Ferreira, Qwynn) 3d stand in lobby (TC 07/27/2023 09:26:49 AM, Ferreira, Qwynn) 3d stand in lobby (TC 07/28/2023 09:53:57 AM, Ferreira, Qwynn) Performed the upgrades for NICU overview 10 and Resus Surveillance (TC 07/31/2023 09:19:22 AM, Ferreira, Qwynn) 3d printing stand (TC 08/01/2023 09:19:41 AM, Ferreira, Qwynn) 3d printing stand (TC 08/03/2023 08:44:13 AM, Ferreira, Qwynn) 3d printing setup (TC 08/02/2023 08:44:13 AM, Ferreira, Qwynn) 3d printing setup (TC 08/03/2023 11:38:21 AM, Ferreira, Qwynn) 3d scanner meeting (TC 08/03/2023 11:38:38 AM, Ferreira, Qwynn) removed 3d printed model, added clear coat and reached out to person who ordered. (TC 08/03/2023 11:40:24 AM, Ferreira, Qwynn) set up printer for new print of mandible, change printing trays as old one expired. print should take 10 hours (TC 08/08/2023 09:08:31 AM, Ferreira, Qwynn) ENT meeting for the cases (TC 08/08/2023 12:49:31 PM, Ferreira, Qwynn) recevied imaging from E-imaging Uploaded files to software Started various segmentations for the cases Showed Maisha the intro to 3D slicer and how to navigate it. Worked on mandible and maxilla case (TC 08/09/2023 08:04:54 AM, Ferreira, Qwynn) removing print, washing, and removing supports from mandible print (TC 08/09/2023 02:12:15 PM, Ferreira, Qwynn) vr/ar meeting with ryerson u (TC 08/09/2023 02:12:39 PM, Ferreira, Qwynn) obtaining patient cds discussing the cases model cleaning/removal (TC 08/09/2023 02:54:00 PM, Ferreira, Qwynn) modelling the maxilla and finalizing prints (TC 08/14/2023 10:43:23 AM, Ferreira, Qwynn) Prepared the model for the meeting at 1215 Removed prints and instpected Removed supports (TC 08/14/2023 02:29:07 PM, Ferreira, Qwynn) worked on maxilla case (TC 08/15/2023 02:29:35 PM, Ferreira, Qwynn) Finalized the model for the case and sent to printer Will clean up wednesday and have them ready for MDRD (TC 08/16/2023 02:30:21 PM, Ferreira, Qwynn) 3d printing meeting (TC 08/16/2023 02:30:36 PM, Ferreira, Qwynn) surgical case meeting to go over prints and setup (TC 08/16/2023 02:31:11 PM, Ferreira, Qwynn) cleaned prints for the case Post-process the prints Handed the print over to ENT Worked on second case Retrived the CD for gynecology for cases provided emmerson with a case to work on (TC 08/21/2023 03:09:24 PM, Ferreira, Qwynn) PM Meeting (TC 08/21/2023 03:09:42 PM, Ferreira, Qwynn) Discussed the cut planes and procedure for thursdays surgery Decided to use the fibula instead of the scapula (TC 08/21/2023 03:10:33 PM, Ferreira, Qwynn) Worked on models for the recon surgery Created the fib models and adjusted the plane level and symmetry of the face (TC 08/22/2023 02:50:05 PM, Ferreira, Qwynn) ENT Meeting to discuss cases (TC 08/22/2023 02:50:38 PM, Ferreira, Qwynn) Obtained dicom from eimaging for a case next week, needed the MRI and Extremity CTA (TC 08/23/2023 12:41:59 PM, Ferreira, Qwynn) Removed build plate, removed prints from plate. Washed prints and removed them from supports very carefully Cured the parts Final post processing was done to remove and finish the guides. Meaasured accuracy of the prints to the model files Cleaned slots fo rcutting (TC 08/23/2023 12:43:50 PM, Ferreira, Qwynn) verfiied print geometry with the model file Test fit the guide to the anatomy as well as the maxilla guides Cleaned and rewashed with alcohol to remove dust Final dry/cure was done to remove any excess resin Met with Joel to discuss the case, what was printed, the outline for the guides, and what side decisisions were made. (TC 08/23/2023 02:44:04 PM, Ferreira, Qwynn) Worked on next case Got the manidlbe done and measured Segmetned both Right and Left Fibula for blood supply. Filled both models to prepare for surgical planning (TC 08/28/2023 09:52:22 AM, Ferreira, Qwynn) Worked on modelling and creating guides as well as other projects around the biomed area (TC 08/29/2023 09:53:00 AM, Ferreira, Qwynn) Worked on the surgical case as some edits were required from the physician (TC 08/30/2023 09:53:43 AM, Ferreira, Qwynn) Worked on VR, modelling, and presentation (TC 08/31/2023 09:54:09 AM, Ferreira, Qwynn) Worked on VR, and presentation, and other 3d projects. Also met with Sandy to discuss another 3d project (TC 09/01/2023 09:59:09 AM, Ferreira, Qwynn) Worked on obtaining licencne for the VR software, Worked with Andre to see if extension was possible Worked on presentation and tried to finalize it - edits were made after hours Joel said OR might be able to pitch in some Andre extended licence on saturday (TC 09/05/2023 10:00:52 AM, Ferreira, Qwynn) Worked on previous case to test dental preplanning Obtained two CDs for next cases. Infomred Dr.Haider that the print was done Reached out to Dr. Foltz about the coil (TC 09/05/2023 01:55:38 PM, Ferreira, Qwynn) Worked on the VR presentation for Joel to set up what we will be presenting. Worked on the dental implants case to see what we can do for printing these Viewed the pelvis case - started looking to see if feasible to print (TC 09/06/2023 01:58:04 PM, Ferreira, Qwynn) Worked on transcutaneous calbe for rex along with EC10 modules to check for the updated configuration files. Replaced 3 covers for them as they were mising Emailed Rex, Michelle, Steve, that the cable should not be repaired for a few reasons - shielding exposed, noise, infection etc (TC 09/06/2023 01:59:25 PM, Ferreira, Qwynn) Met with Joel to discuss new edits for a case Showed him the VR that we will use for the presentation Got him to demo the program to get himself comfortable in using it. (TC 09/07/2023 07:14:59 AM, Ferreira, Qwynn) Created script for the presentaiton Managed to get the maxilla case working for the recon struction Meeting with developers to fix extensions (TC 09/06/2023 07:24:12 AM, Ferreira, Qwynn) Worked with Joel for VR to test the demo and go over the presentation. Worked on maxilla case to get an understanding of what is expected and what we can do to create the guide he needs He provided and drawing to help, used it to get the rough model down Worked with Maisha to show the difference between purchased VR and open source/free (TC 09/07/2023 07:14:59 AM, Ferreira, Qwynn) Created script for the presentaiton Managed to get the maxilla case working for the recon struction Meeting with developers to fix extensions (TC 09/11/2023 02:48:59 PM, Ferreira, Qwynn) Went through PM with Samir and coop student to assit with his understanding of the procedure (TC 09/11/2023 02:49:33 PM, Ferreira, Qwynn) Worked on the models for the surgery thursday. This includes post processing, measurments, and fitment. Also sent new print of the reconstruction cuts to print Will meet with Joel tomorrow to further discuss Also worked on the n the print for the camera head. Found new solution for the NICU to use - Should be a great improvement (TC 09/11/2023 02:51:58 PM, Ferreira, Qwynn) General Monday Meeting 1pm (TC 09/18/2023 02:12:59 PM, Ferreira, Qwynn) monday huddle (TC 09/18/2023 02:13:19 PM, Ferreira, Qwynn) created model for next case - very messy scans Lots of clean up required Reached out to developers to see if they know whats up (TC 09/21/2023 02:25:17 PM, Ferreira, Qwynn) Meeting with Dr. Kim to go over Emmerson and Adeel's work Meeting with Joel/Karl regarding imaging (TC 09/21/2023 02:26:01 PM, Ferreira, Qwynn) Created surgical guides for the case on monday - printing now (TC 09/22/2023 03:20:15 PM, Ferreira, Qwynn) 1 on 1 meeting, meeting with Dr. Kim, meeting with Karl/Dr.Leila (TC 09/22/2023 03:20:43 PM, Ferreira, Qwynn) modelled jaw guides and took photos. Post-process the guides Cleaned the guides Will hand off shortly today (TC 09/25/2023 11:02:35 AM, Ferreira, Qwynn) Worked on dental model Printed the fibula segment with implants Cleaned print (TC 09/27/2023 10:54:29 AM, Ferreira, Qwynn) Worked with Dr. Laila, Karl, Eric, and Joel to change the imaging protocol for Head/Neck (TC 09/27/2023 10:56:25 AM, Ferreira, Qwynn) Emailed, and worked with Dr.Haider for the depressor and received data from Dr. Pia for the rhinoplasty model (TC 09/27/2023 10:57:03 AM, Ferreira, Qwynn) philips simulation (TC 09/27/2023 12:29:04 PM, Ferreira, Qwynn) staff meeting (TC 09/28/2023 12:01:42 PM, Ferreira, Qwynn) Started working on the rhinoplasty model Also received some photos for the placenta case as well as additonal info that will aid in the the design Discussed further with CT tech about the protocol for Head/Neck and which protocol would be best/easiest for both parties. Went to inspect MRI incubator but was not there (TC 10/02/2023 01:29:24 PM, Ferreira, Qwynn) Modelled the placenta case - start post processed the guides for tuesdays case Handed them to MDRD Picked up imaging for next case on Thursday replaced MRI doors (TC 10/03/2023 01:32:17 PM, Ferreira, Qwynn) modelled 2 ideas for the placenta stand Started designingthe guides for thursday case Went to OR to confirm fitment of the guides Meeting with NICU/John Worked with Joel to place new cut planes (TC 10/05/2023 11:50:03 AM, Ferreira, Qwynn) Worked on placenta stand model, fixing the model to fit within our printer limits. Nearly done first version to print (TC 10/12/2023 02:53:39 PM, Ferreira, Qwynn) Retrieved printed model from FDM for placenta project Started doing presentation for SoC Worked on rhinoplasty model Sent print for placenta project Went over one of emmersons prints - solid Received another case for monday morning Worked with imaging to change the imaging set for the imaging protocol we need Emergency shut down huddle for NICU and L&D (TC 10/16/2023 05:37:29 PM, Ferreira, Qwynn) Worked on placenta stand as the printer was having issues feeding the filament Took apart printer to troubleshoot and thought it was the feed tube. Did not solve the problem Will continue to work on Tuesday Recevied another case from Karl for wednesday, very rushed project, but doable. Worked on patient scans and setup inital model (TC 10/17/2023 05:39:28 PM, Ferreira, Qwynn) Worked on creating guides for the case on wednesday Worked on transport deck but did not create WO for it yet as very busy. Modelled the guides and recon the maxilla as well as copied the segments added labels to the pieces Sent to printer but some reason failed - did not cure to the plate Used new rsin and tank and build plate to try once again Printing now again for tomorrows case - should be done around 9pm Worked on Dremel printer as it was not working Took apart motor, and feeder Plastic particels were on the printer feed wheel Cleaned and assembled again to run test print (TC 10/23/2023 03:08:04 PM, Ferreira, Qwynn) Worked on Placenta Stand project to assemble, post porcess, and work on the printer (TC 10/23/2023 03:08:35 PM, Ferreira, Qwynn) Worked on the Mandible case for Nov 1, made a printable file for tomorrow to work on ad modelled th fibula (TC 10/23/2023 03:09:14 PM, Ferreira, Qwynn) Worked on a sentec device for Rex to see if we can change the memebrane interval time - was able to successfully change the interval date - but the repair of the cable was not done as it is not safe. Will inform Rex tomorrow when he is back and available (TC 10/24/2023 01:24:39 PM, Ferreira, Qwynn) Printed the Uterus model for Rachel Removed and waiting to post process Worked on mandible case and sent file to the printer Worked on the dremel to come up with new filament hanging idea as we wait for parts Worked with Joel and Karl to set up meeting date (TC 10/26/2023 07:31:02 AM, Ferreira, Qwynn) Worked on the mandible case for November 1st Worked on the 3d printed model Worked on the placenta stand sanding and fitting the segments together Worked with Maisha on the document for cost per OR minute (TC 10/27/2023 02:50:24 PM, Ferreira, Qwynn) Worked on the mandible case Created fibula guide templates and dental implant templates Started another mandible print as th epatient requested a model Obtained more imaging for another case occuring nov 6 and placed onto the drive Worked with Connor showing him the program (fusion 360) and got him familiar with the workflow and printing topics Meeting with Karl to discuss the cut angles and placement further and decided on a set location Assembled half of the placenta stand (TC 10/30/2023 12:45:05 PM, Ferreira, Qwynn) Worked on the surgical case for wednesday, finalized the guides Printed everything needed. One print failed Showed Connor the high level workflow of 3D printing Cleaned some of the models (TC 11/01/2023 12:01:42 PM, Ferreira, Qwynn) Worked with Connor all day Worked on the placenta stand and assembling the peices Worked on the mandible model for the patient and continuing the filling process Worked on the november 1st case Viewed the surgery with connor so he can see the gudies being used Had a meeting with Dr. Paul and emmerson to discuss ortho use cases (TC 10/31/2023 10:26:21 AM, Ferreira, Qwynn) Worked on the guides for the case november 1st Had connors help for the guides and models Worked on the placenta stand clamping, gluing and sanding Showed connor some modelling techniques for fusion and slicer Created another coloured model for us to use for demos (TC 11/02/2023 10:28:15 AM, Ferreira, Qwynn) Worked on next case for Nobember 6th Created the models and got the print setup Changed the clamps for the placenta stand and had connors help Showed connor some modelling tips for the case and what to look for/include Had a meeting with joel to discuss the case (TC 11/06/2023 09:20:58 AM, Ferreira, Qwynn) Worked with connor finalizing the placenta stand and adding the coatings Started researching ortho guides to see what we can do to design Showed Maisha the OR room and surgery for the fibula reconstruction Discussed with Nima the purchasing issue with VR Worked with Maisha to review the current document (TC 11/07/2023 09:24:42 AM, Ferreira, Qwynn) Head/Neck meeting (TC 11/07/2023 03:10:41 PM, Ferreira, Qwynn) Worked with Connor on the placenta stand adding some of the final touches including epoxy for the nozzle joints, and other areas where some water was able to escape from. Fixed the dremel printer and was able to show connor how to use the printer, take apart the printer head, and set up a print for the printer. Connor modelled and printed his first object with the dremel. Cleaned the mandible model from the previous case. Did research for the silicone materials and what would be needed to start creating silicone models for surgical practice/training. Created a list of required materials for the silicone molds Did research for the hip case wednesday for feasibility on creating a guide. (TC 11/07/2023 03:44:18 PM, Ferreira, Qwynn) worked on the drain spout model for the placenta stand to hopefully print tomorrow. Will show connor the work being done on fusion. (TC 11/10/2023 03:17:24 PM, Ferreira, Qwynn) Worked with connor showing him how to use 3d slicer Sprayed the mandible model Worked with connor again on 3d slicer Showed courtney how to change the membrane date for the sentec monitors. Picked up imaging for the pelvis case next week Had a meeting with emmerson and ortho team regarding the case (TC 11/20/2023 02:23:04 PM, Ferreira, Qwynn) monday huddle (TC 11/20/2023 02:23:24 PM, Ferreira, Qwynn) Modelled screws for pelvis case to import into VR. Set up silicone testing tray for silicone molding Meeting with placenta group to show the final product - happy with it - should be doing xray testing in the next few weeks Did some reading on iso 13485 (TC 11/21/2023 08:49:59 AM, Ferreira, Qwynn) Meeting with Placenta group to go over the stand with final comments - no changes, happy with overall design, waiting for approval to test (TC 11/21/2023 08:51:24 AM, Ferreira, Qwynn) Worked on the segmentation in VR for the hip case to see if the screws work being placed. Re-did the case using only VR and created a printable file by the end and told Arieh about the progress Went to 12 south to inform them about a bed call wasnt alarming, but the patient was not weighed so the alarm could not be set up Did some reading on teh ISO 13485 for QMS Post processed the print tray that was made for our test silicone mold tray (TC 11/22/2023 09:54:06 AM, Ferreira, Qwynn) Handed over the placenta stand to emily from the research group. Explained the proper way to setup and use (TC 11/22/2023 09:55:49 AM, Ferreira, Qwynn) Reorganized the 3d printing documentation Assited Emmerson with the table and Samir Meeting with Joel and Emanuela to discuss potential news article Printed Adapters for PFT lab Worked on collecting images from past cases for the research paper and potential media article (TC 11/23/2023 10:14:59 AM, Ferreira, Qwynn) Meeting with Joel to discuss capitol and options to consider for submission (TC 11/23/2023 03:50:15 PM, Ferreira, Qwynn) printed the heel pricker holder for the NICU and let the staff know. Had a discussion with imaging team and Eric about the imaging protocol and had Matt redo the reconstruction for imaging (TC 11/23/2023 03:51:51 PM, Ferreira, Qwynn) Worked on the paper for CMBES for ideas and photos, while creating the abstract (TC 11/23/2023 03:56:20 PM, Ferreira, Qwynn) Organized more of the T drive for 3D printing to ensure we have as much information as possible as some prints were not being captured. including redonwloading files, renaming files, verifiying prints and bilds correspond to the print file (TC 11/23/2023 03:57:53 PM, Ferreira, Qwynn) Recevied new case for Dec 7, only head/neck CT Downloaded and started making notes about the case (TC 11/24/2023 03:30:44 PM, Ferreira, Qwynn) Worked on segmentating the models for the case dec 7. Tumor was hard to see, MRI comes in nov 25 (TC 11/24/2023 03:31:27 PM, Ferreira, Qwynn) Worked with Connor creating the silicone model for testing the skin We were able to create the epidermis, dermis, fat layer and muscle layer Some improvements noted for next time including smaller dermis layer, larger fat layer, and shortening the curing times for the muscle and fat for better adhesion (TC 11/24/2023 03:54:15 PM, Ferreira, Qwynn) Read more into the 3BL printer and how to efficiently use the printer to aid in what we do. Read into the Fuse 1 and incldued some documents on the T drive for this printer. (TC 11/28/2023 10:12:05 AM, Ferreira, Qwynn) cybersecurity meeting (TC 11/28/2023 10:12:27 AM, Ferreira, Qwynn) read through the ortho fda guidance document to see what we should be capturing at all steps of the process to start forming a managment system for our setup. (TC 11/29/2023 10:13:23 AM, Ferreira, Qwynn) Worked with Connor creating another silicone pad reformatted th eimaging for the next case to create a more accurate tumor. Started segmenting the tumor (TC 11/29/2023 01:56:44 PM, Ferreira, Qwynn) Huddle Meeting Also meeting with materialise for 15min (TC 11/29/2023 01:57:23 PM, Ferreira, Qwynn) Worked on the case file for dec 7 Was updated more cases for January Worked with Connor on the silicone molds - great progress Had some troubleshooting issues with the dremel and tested with emmerson (TC 11/29/2023 03:54:24 PM, Ferreira, Qwynn) Poured another silicone mold with Connor - need a total of 5 pads Read more on ISO 13485 and the clauses as well as shared some resources with Maisha Made some notes for the high level meeting for our quality in 3d printing (TC 11/30/2023 03:41:34 PM, Ferreira, Qwynn) reviewed more research papers for the quality and what metrics we should be capturing. Made notes and spoke with Maisha to have a meeting tomorrow to put our thoughts together Read the technical considerations, interpretation of regulations, and ortho guidance document for guides (TC 11/30/2023 03:44:39 PM, Ferreira, Qwynn) modelled the left/right fibula for the case and set a time with joel to plan the cuts and orientation Read more on the documentation of guides and how we should be capturing - lots of information but what is relevant is challenging Reached out to some folks to follow up on printing ideas Received information on the certification for additive manufacturing through SME and the fee is only for the certificate and exam. we must read the textbook (TC 12/04/2023 02:44:18 PM, Ferreira, Qwynn) Meeting with adeel and maisha to go over the general plan for quality and how we can start to develop a plan to use tms and word for policies (TC 12/04/2023 02:45:10 PM, Ferreira, Qwynn) Worked on guides for the case on thursday Sent a print for the case dec 7 including the reconstructed mandible Created the fibula guides Working on the mandible guides Worked on quality document to see what is required and discussed with maisha what the next steps should be (TC 12/05/2023 01:45:15 PM, Ferreira, Qwynn) head and neck meeting (TC 12/05/2023 01:45:35 PM, Ferreira, Qwynn) finalized the guides for thursday case sent to the printer (TC 12/05/2023 01:47:43 PM, Ferreira, Qwynn) Worked on paper for CMBES Post process the mandibles for thursday case with connor Received new model for research to try and print, few different methods to try (TC 12/06/2023 02:27:50 PM, Ferreira, Qwynn) worked on hte case report for joel Sent over the report to joel and eric Handed over the guides to MDRD Post processed the printed guides (TC 12/06/2023 02:33:02 PM, Ferreira, Qwynn) Went to UHN to return the coil and spoke with Dr. Folts in imaging regarding the use of 3d printing in the lab (TC 12/06/2023 02:34:40 PM, Ferreira, Qwynn) received a call from ED regarding rectal probes. They only have 2 prong cables but this device required 1 prong. They were referring to the philips machine, not altrix. Tried to fin
| 195352 | CM | CMPLT | CMPLT | 2023-06-07 11:37 | 2023-06-26 10:53 | Ordering centroid sensors; no charge PO | (TC 06/07/2023 11:41:01 AM, Baghirzade, Nara) SC 1002974727 No Charge PO (TC 06/26/2023 10:52:32 AM, Baghirzade, Nara) PO 7202228008
1066462 | 195353 | FC | CMPLT | CMPLT | 2023-06-07 12:00 | 2023-06-07 12:02 | QC checklist | (TC 06/07/2023 12:01:44 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 North
32520 | 195354 | CM | CMPLT | CMPLT | 2023-06-07 12:01 | 2023-06-07 12:02 | MSH DR2 Xray RM - Grid Damaged | (TC 06/07/2023 12:00:44 PM, Chau, Ken) Grid guide pins broken and requires replacement -rethreaded with 3/32 -made new guide pins, tested okay
1046035 | 195355 | CM | CMPLT | CMPLT | 2023-06-07 12:51 | 2023-07-11 07:43 | MSH ER DRX EVO E5-02(ER) ON WALL STAND, WILL NOT COLLIMATE AUTOMATICALLY | (TC 06/07/2023 12:51:18 PM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) ON WALL STAND, WILL NOT COLLIMATE AUTOMATICALLY ADVISE FULL SHUTDOWN, INCLUDING UPS | MSH ER DRX EVO E5-02(ER) ON WALL STAND, WILL NOT COLLIMATE AUTOMATICALLY
| 195356 | CM | CMPLT | CMPLT | 2023-06-07 12:56 | 2023-07-11 07:42 | MSH DR ORDER QTY 4 SPAJ2900 DETECTOR BATTERIES FOR DISTRIBUTION AMONG SYSTEMS | (TC 06/07/2023 12:55:46 PM, Whittington, Scott) MSH DR ORDER QTY 4 SPAJ2900 DETECTOR BATTERIES FOR DISTRIBUTION AMONG SYSTEMS | MSH DR ORDER QTY 4 SPAJ2900 DETECTOR BATTERIES FOR DISTRIBUTION AMONG SYSTEMS
1054221 | 195357 | CM | CMPLT | CMPLT | 2023-06-07 13:46 | 2023-06-07 13:48 | Security mode activated | (TC 06/07/2023 01:46:44 PM, Rhuland, Courtney) Thermometer stuck in security mode and was not functional Placed thermometer into wall mount to clear mode Functional check - OK
68600 | 195358 | PM | CMPLT | CMPLT | 2023-06-07 14:23 | 2023-06-07 14:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/07/2023 02:23:52 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
| 195359 | PA | CMPLT | CMPLT | 2023-06-07 15:12 | 2023-06-07 15:15 | CVSM right side panel kit | (TC 06/07/2023 03:13:23 PM, Canjimavadivel, Kamallesh ) Part originally ordered from different WO. To be kept in stock until used. | Biomed stock, charge back when used
1066945 | 195360 | CM | CMPLT | CMPLT | 2023-06-07 16:29 | 2023-06-07 16:32 | Missing one of rubber feet | (TC 06/07/2023 04:29:37 PM, Cortez, John) Natalia (PCM) mentioned one of the feet was missing, but they later found it Told them to let me know if the feet get loose again, I may need to apply Loctite to better secure it
1061898 | 195361 | CM | CMPLT | NFF | 2023-06-07 16:54 | 2023-06-07 16:55 | "foot controls dont work" | (TC 06/06/2023 04:53:53 PM, Cortez, John) Couldn't confirm issue. Resetting bed may have resolved issue
1055802 | 195362 | PM | CMPLT | CMPLT | 2023-06-08 09:21 | 2023-06-08 12:34 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/08/2023 12:33:06 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure for results
| 195363 | CM | CMPLT | CMPLT | 2023-06-08 09:33 | 2023-08-28 09:04 | ordering OR table remote
| 195364 | CM | CMPLT | CMPLT | 2023-06-08 09:54 | 2023-08-15 14:30 | projects/repairs | (TC 06/05/2023 10:02:00 AM, Rendon Lopez , Elizabeth) Repairing 2X fetal monitor recorder for spare. Calibrated, tested-OK (TC 06/19/2023 08:23:11 AM, Rendon Lopez , Elizabeth) 15M masimo deployment (TC 06/20/2023 08:34:55 AM, Rendon Lopez , Elizabeth) LVL 7 AGM breathing block issues with three of them. (TC 06/21/2023 08:34:55 AM, Rendon Lopez , Elizabeth) Tested them, once block had large leak, needs PM, two other blocks okay, no issues found (TC 07/17/2023 08:37:01 AM, Rendon Lopez , Elizabeth) L&D OR3 surgical lights kept moving side to side, adjusted screws, Ok (TC 07/24/2023 10:20:19 AM, Rendon Lopez , Elizabeth) Rounds (TC 07/24/2023 09:02:32 AM, Rendon Lopez , Elizabeth) L&D OR 3 light not going high enough. Adjusted lights, noticed lights were flashing, breaked needs to be reset, placed request to access breaked panel (TC 07/25/2023 09:02:32 AM, Rendon Lopez , Elizabeth) Electrician opened breaker panel and reset breaked for lights, working properly now. Adjusted height movement again. 3-LED light, one LED is not adjusting properly, will follow up with olympus for repair (TC 07/26/2023 09:25:54 AM, Rendon Lopez , Elizabeth) Admin/WOs (TC 08/01/2023 06:26:41 PM, Rendon Lopez , Elizabeth) Testing/troubleshooting, wireless transducers. Testing new Cardinal spiral electrode compatibility (TC 08/02/2023 08:48:12 AM, Rendon Lopez , Elizabeth) Installed new CCE in Triage F for testing with patients (TC 08/08/2023 08:56:43 AM, Rendon Lopez , Elizabeth) ROunds (TC 08/08/2023 08:56:43 AM, Rendon Lopez , Elizabeth) Olympus L&D OR boom check with technician for screws recall
1057393 | 195365 | CM | CMPLT | CMPLT | 2023-06-08 10:03 | 2023-06-08 10:06 | Not working | (TC 06/06/2023 10:03:29 AM, Rendon Lopez , Elizabeth) Downloaded error logs, none were found. Running on user mode not issues (TC 06/07/2023 10:03:29 AM, Rendon Lopez , Elizabeth) Device keeps failing patient side occlusion, replaced patient PS, performed check in tests, all passed
1064398 | 195366 | CM | CMPLT | CMPLT | 2023-06-08 10:04 | 2023-06-08 12:18 | Low pressur alarm delay | (TC 06/08/2023 10:03:46 AM, Chukri, Samir) Checked with Qwynn and R.T Sandy. refere to Operating Manual low pressure should alarming in 10s. no problem found.
1064398 | 195367 | CM | CMPLT | CMPLT | 2023-06-08 10:08 | 2023-06-08 10:11 | Not alarming during disconnect/ takes long time | (TC 06/08/2023 10:08:44 AM, Ferreira, Qwynn) Sandy from RT reached out regarding this issue. Went upstaris to test the circuit and there was no alarm for a few seconds Timed it to be roughly 9-10 seconds. Tried to troubleshoot and decided to see if there was a note on disconnect alarms which there was in the Vent user manual. disconnect alarms will occur after 10 seconds when below 3cmH2O of pressure. The vent is working, just a long alarm wait time. Sandy from RT was notified of the disconnect alarm times
67934 | 195368 | CM | CMPLT | CMPLT | 2023-06-08 11:33 | 2023-06-08 13:31 | Latch recall | (TC 06/08/2023 11:36:58 AM, Rhuland, Courtney) Hood replaced Recall parts changed 4 porthole grommets and ventilator slot grommet changed Functional check - OK
| 195369 | CM | ACTIV | ISSUE | 2023-06-08 12:20 |  | Weekly meeting for equipment replacment Planning Project Team | (TC 06/08/2023 12:19:46 PM, Chukri, Samir) Meeting (TC 06/13/2023 01:29:40 PM, Chukri, Samir) Zoom: Kick off-CMMS Data meeting (TC 06/15/2023 12:50:13 PM, Chukri, Samir) Meeting (TC 06/21/2023 02:35:38 PM, Chukri, Samir) Bi-weekly department meeting : fleet managment (TC 06/27/2023 12:58:13 PM, Chukri, Samir) Weekly-CMMS data quality project (TC 06/29/2023 11:25:27 AM, Chukri, Samir) weekly meeting for equipment
66007 | 195370 | CM | CMPLT | CMPLT | 2023-06-08 13:05 | 2023-06-08 13:06 | Returned from SickKids | (TC 06/08/2023 01:05:34 PM, Rhuland, Courtney) Device returned from SickKids Check-in completed Returned to NICU
1050590 | 195371 | CM | CMPLT | NFF | 2023-06-08 13:36 | 2023-06-08 13:38 | Not working | (TC 06/08/2023 01:35:51 PM, Reid, Emmerson ) Checked the unit and it was turning on and weighing normally. Spoke with the unit and they said the weighing was inconsistent, but it seemed to be working fine. NFF
| 195372 | CM | CMPLT | CMPLT | 2023-06-08 13:42 | 2023-08-01 09:53 | Tank regulators for new NICU vents | Received
69267 | 195373 | PM | CMPLT | CMPLT | 2023-06-08 13:45 | 2023-06-08 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:45:20 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69067 | 195374 | PM | CMPLT | CMPLT | 2023-06-08 13:54 | 2023-06-08 13:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:54:07 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69413 | 195375 | PM | CMPLT | CMPLT | 2023-06-08 13:56 | 2023-06-08 13:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/08/2023 01:55:49 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
500857 | 195376 | CM | CMPLT | CMPLT | 2023-06-08 14:20 | 2023-07-11 07:44 | MSH-5-US RM 8 HAS BLUE SCREEN HARDWARE MALFUNCTION PARITY CHECK/MEMORY PARITY ERROR | (TC 06/08/2023 02:21:23 PM, Whittington, Scott) MSH-5-US RM 8 HAS BLUE SCREEN HARDWARE MALFUNCTION PARITY CHECK/MEMORY PARITY ERROR REBOOT OK | MSH-5-US RM 8 HAS BLUE SCREEN HARDWARE MALFUNCTION PARITY CHECK/MEMORY PARITY ERROR
1054381 | 195377 | CM | CMPLT | CMPLT | 2023-06-08 14:31 | 2023-06-08 14:32 | BP cuff leaking | (TC 06/08/2023 02:30:56 PM, Rhuland, Courtney) BP cuff leaking Replaced with new cuff Functional check - OK
1054383 | 195378 | CM | CMPLT | CMPLT | 2023-06-08 14:33 | 2023-06-08 14:34 | BP cuff leaking | (TC 06/08/2023 02:33:17 PM, Rhuland, Courtney) BP cuff leaking and not taking pressure Replaced with new cuff Functional check - OK
| 195379 | CM | CMPLT | CMPLT | 2023-06-08 15:05 | 2023-12-12 09:45 | JIQ medification | (TC 06/08/2023 03:05:26 PM, Chukri, Samir) Made some modification on the form. (TC 06/09/2023 08:56:30 AM, Chukri, Samir) completed the form and emailed to Adeel.
1062099 | 195380 | PM | CMPLT | CMPLT | 2023-06-09 07:58 | 2023-06-08 07:57 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 2nd year PM after 1st year PM last year. Replaced V-battery, air filter. No GS500. Did battery check. 60 Ah, 180 minutes. Tested. Functional. SW.2.51 Leakage: 79 uA and 91 ohm.
1055809 | 195381 | PM | CMPLT | CMPLT | 2023-06-09 10:17 | 2023-06-09 13:13 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/09/2023 01:04:39 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure for results
66438 | 195382 | CM | CMPLT | CMPLT | 2023-06-09 10:52 | 2023-06-09 10:53 | Battery adaptor | (TC 06/09/2023 10:51:53 AM, Scero, Tomlyn) I was using the telepak for testing the signals on the 16th floor for deant. Noticed that the battery adaptor was in terrible shape. Replaced it and proceeded with testing.
1056427 | 195383 | CM | CMPLT | CMPLT | 2023-06-09 13:42 | 2023-06-09 13:44 | "thermometer doesn't work" | (TC 06/09/2023 01:43:20 PM, Cortez, John) Replaced temp probe Functional check ok
1058530 | 195384 | CM | CMPLT | NFF | 2023-06-09 13:48 | 2023-06-09 13:51 | "unable to do preop check, no prompts" | (TC 06/09/2023 01:47:40 PM, Cortez, John) Unable to replicate issue. Powered on device and completed preop check as normal (TC 06/09/2023 01:48:53 PM, Canjimavadivel, Kamallesh ) Unable to replicate issue. Powered on device and completed preop check as normal
1046952 | 195385 | CM | CMPLT | CMPLT | 2023-06-09 14:13 | 2023-07-11 07:44 | MSH ICU POD A SPARQ CRASHED | (TC 06/20/2023 05:55:00 PM, Whittington, Scott) MSH ICU POD A SPARQ CRASHED ABLE TO RESTART | MSH ICU POD A SPARQ CRASHED
1062778 | 195386 | CM | CMPLT | CMPLT | 2023-06-09 14:28 | 2023-06-09 14:28 | NIBP values not populating in Mortara software | (TC 06/09/2023 02:28:11 PM, Baghirzade, Nara) Jen reported that the NIBP values are not populating on the Mortara software; Switched off the computer, NIBP module, bike, and the Medisoft box. Upon booting the computer, I switched on the bike and NIBP module. Jen took another NIBP measurement on the patient and at first the NIBP values was not populating in Mortara. After the patient began the spirometry test and began breathing through the analyzer, the NIBP values were automatically populating in Mortara. Stayed for the first stress portion of the test and NIBP values were populating through successfully. Unsure of what caused the issue, could have been a communication error between the 4 different modules.
67129 | 195387 | CM | CMPLT | NFF | 2023-06-09 14:30 | 2023-06-09 14:31 | BP cuff not inflating | (TC 06/09/2023 02:30:32 PM, Baghirzade, Nara) Reported cuff not inflating; took repeated BP measurements, no issue with inflation of cuff. NIBP cable, connectors, and cuff are good/no physical damage. SpO2 and temperature module are functional. Adjusted the date and time. No fault found.
52178 | 195388 | CM | CMPLT | CMPLT | 2023-06-09 14:33 | 2023-06-09 14:40 | Latch recall | (TC 06/09/2023 02:32:45 PM, Rhuland, Courtney) Recall latches replaced Broken light replaced Bent power cord replaced Missing plastic plate replaced Functional check - OK Leakage current - 60 mA Ground resistance - 6mOhms
1060333 | 195389 | CM | CMPLT | CMPLT | 2023-06-09 14:37 | 2023-06-09 14:49 | PM-MAJOR- Annual PM For: Patient Simulator | (TC 06/09/2023 02:36:48 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached.
68607 | 195390 | PM | CMPLT | CMPLT | 2023-06-09 14:38 | 2023-06-09 14:54 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/09/2023 02:39:21 PM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68807 | 195391 | PM | CMPLT | CMPLT | 2023-06-09 14:44 | 2023-06-12 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:43:50 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 243 mohms Leakage current: 70.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 06/12/2023 02:09:59 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4054 mamphr Replaced power cord strap. PM completed.
69012 | 195392 | PM | CMPLT | CMPLT | 2023-06-09 14:47 | 2023-06-09 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:47:30 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 216 mohms Leakage current: 70.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
69227 | 195393 | PM | CMPLT | CMPLT | 2023-06-09 14:54 | 2023-06-09 14:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:54:34 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68632 | 195394 | PM | CMPLT | CMPLT | 2023-06-09 14:56 | 2023-06-09 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:55:48 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69568 | 195395 | PM | CMPLT | CMPLT | 2023-06-09 15:00 | 2023-06-09 15:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/09/2023 02:59:54 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1064572 | 195396 | CM | CMPLT | CMPLT | 2023-06-09 15:23 | 2023-06-09 15:25 | PM-MAJOR- Annual PM for: Multimeter | (TC 06/09/2023 03:23:27 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration attached.
1060834 | 195397 | CM | CMPLT | CMPLT | 2023-06-09 15:34 | 2023-06-09 15:36 | PM-MAJOR- Annual PM For: Scale | (TC 06/09/2023 03:34:00 PM, Chukri, Samir) Completed PM performed by Getinge. Certificate of Calibration Attached.
32520 | 195398 | CM | CMPLT | CMPLT | 2023-06-11 10:26 | 2023-07-11 07:46 | MSH DR2 DETECTOR Z RAISED TO HIGHEST EVEL, UNALBE TO MOVE | (TC 06/11/2023 04:44:00 PM, Whittington, Scott) MSH DR2 DETECTOR Z RAISED TO HIGHEST LEVEL, UNALBE TO MOVE JEANEETE DESCRIBED AS DETECTOR Z MOVED BY ITSELF POSSIBLE THAT COMPUTER SIDE OF MODULE IS FAULTY AND CAUSED PREVIOUS FAULTS ONLY POWER SIDE REPLACED IN MAY 25 INCIIDENT (TC 06/12/2023 09:45:00 AM, Whittington, Scott) MSH DR2 DETECTOR Z RAISED TO HIGHEST LEVEL, UNALBE TO MOVE JEANEETE DESCRIBED AS DETECTOR Z MOVED BY ITSELF POSSIBLE THAT COMPUTER SIDE OF MODULE IS FAULTY AND CAUSED PREVIOUS FAULTS ONLY POWER SIDE REPLACED IN MAY 25 INCIIDENT (TC 06/12/2023 09:45:17 AM, Heron, Chris) MSH DR2 DETECTOR Z RAISED TO HIGHEST LEVEL, UNALBE TO MOVE JEANEETE DESCRIBED AS DETECTOR Z MOVED BY ITSELF POSSIBLE THAT COMPUTER SIDE OF MODULE IS FAULTY AND CAUSED PREVIOUS FAULTS ONLY POWER SIDE REPLACED IN MAY 25 INCIIDENT REVIEW ARISTOS REPAIR METHODS (TC 06/12/2023 03:38:41 PM, Heron, Chris) REPAIR STUCK GRID (TC 06/11/2023 04:44:00 PM, Whittington, Scott) MSH DR2 DETECTOR Z RAISED TO HIGHEST LEVEL, UNALBE TO MOVE JEANEETE DESCRIBED AS DETECTOR Z MOVED BY ITSELF POSSIBLE THAT COMPUTER SIDE OF MODULE IS FAULTY AND CAUSED PREVIOUS FAULTS ONLY POWER SIDE REPLACED IN MAY 25 INCIIDENT (TC 06/11/2023 04:44:00 PM, Whittington, Scott) MSH DR2 DETECTOR Z RAISED TO HIGHEST LEVEL, UNALBE TO MOVE JEANEETE DESCRIBED AS DETECTOR Z MOVED BY ITSELF POSSIBLE THAT COMPUTER SIDE OF MODULE IS FAULTY AND CAUSED PREVIOUS FAULTS ONLY POWER SIDE REPLACED IN MAY 25 INCIIDENT | MSH DR2 DETECTOR Z RAISED TO HIGHEST LEVEL, UNALBE TO MOVE JEANEETE DESCRIBED AS DETECTOR Z MOVED BY ITSELF POSSIBLE THAT COMPUTER SIDE OF MODULE IS FAULTY AND CAUSED PREVIOUS FAULTS ONLY POWER SIDE REPLACED IN MAY 25 INCIIDENT
| 195399 | CM | ACTIV | ISSUE | 2023-06-11 11:29 |  | Cannot do pre-use check
1058530 | 195400 | CM | CMPLT | CMPLT | 2023-06-11 11:30 | 2023-06-09 08:54 | Cannot do pre-use check | (TC 06/12/2023 08:53:18 AM, Chukri, Samir) Chedcked and performed Pre-use chack by Kam and John. pass no problem found.
1055947 | 195401 | CM | CMPLT | CMPLT | 2023-06-12 07:51 | 2023-06-10 07:51 | Worked on 14N Decant | Worked on 14N Decant
66158 | 195402 | CM | CMPLT | CMPLT | 2023-06-12 08:13 | 2023-06-12 08:15 | No temperature reading | (TC 06/12/2023 08:13:59 AM, Ferreira, Qwynn) There was a tear in the cable Replaced temp probe Tested 37 = 37 deg
68027 | 195403 | CM | CMPLT | CMPLT | 2023-06-12 09:09 | 2023-06-12 09:10 | Displays error | (TC 06/12/2023 09:09:03 AM, Canjimavadivel, Kamallesh ) Performed calibration, passed
65898 | 195404 | CM | CMPLT | CMPLT | 2023-06-12 09:59 | 2023-06-12 10:04 | Channel error | (TC 06/12/2023 09:59:30 AM, Rhuland, Courtney) Plunger system error in error logs Performed check in and plunger force accuracy test failed Opened device and replaced worn wiper assembly Reassembled device and performed calibration and check in All tests passed Refer to ASM for details
65964 | 195405 | CM | CMPLT | NFF | 2023-06-12 10:08 | 2023-06-12 10:10 | Channel error | (TC 06/12/2023 10:08:26 AM, Rhuland, Courtney) Error logs checked and no error found Performed check-in and no fault found Refer to ASM for details
1058237 | 195406 | CM | CMPLT | CMPLT | 2023-06-12 10:14 | 2023-06-12 10:34 | Door does not close properly | (TC 06/12/2023 10:33:07 AM, Rhuland, Courtney) Latch pivot screw tightened Functional check - OK
1061209 | 195407 | CM | CMPLT | CMPLT | 2023-06-12 10:25 | 2023-06-12 10:26 | Table issue | (TC 06/12/2023 10:24:43 AM, Reid, Emmerson ) Was called to check the table, when I got there I was told to wait 5 minutes before going in for them to finsh. When I went in I saw the remote was flickering on/off. I switch the remote with another table and it was working. I brought the flickering remote to the other table and it was working. Both remotes working now on different tables.
67123 | 195408 | CM | CMPLT | CMPLT | 2023-06-12 10:41 | 2023-06-12 10:41 | functional test and relocation | (TC 06/12/2023 10:41:03 AM, Samanian, Kamran) tested all functions and passed. adjusted time. alocated this device to 9S due to a defective VSM300 in the unit.
| 195409 | PA | CMPLT | CMPLT | 2023-06-12 11:03 | 2023-06-21 08:27 | Masimo forehead sensors | (TC 06/12/2023 11:07:36 AM, Canjimavadivel, Kamallesh ) ED requested forehead sensors for stock. Placed order for 1 box and provided info for further orders.
| 195410 | CM | CMPLT | CMPLT | 2023-06-12 11:49 | 2023-06-12 11:49 | KVM issue in biomed lab | (TC 06/12/2023 11:48:56 AM, Samanian, Kamran) resolved the issue, tested and passed
1046195 | 195411 | CM | CMPLT | CMPLT | 2023-06-12 12:41 | 2023-07-10 09:41 | device not turning on | (TC 06/14/2023 02:21:56 PM, Scero, Tomlyn) ordered batteries for replacement (TC 07/10/2023 09:39:25 AM, Scero, Tomlyn) Replaced battery. tested unit and it's turning on and taking a BP correctly. Tested plugged in and on battery.
| 195412 | CM | CMPLT | CMPLT | 2023-06-12 13:02 | 2023-07-10 09:38 | batteries for WA BP | (TC 07/10/2023 09:33:27 AM, Scero, Tomlyn) Batteries received
1066454 | 195413 | FC | CMPLT | CMPLT | 2023-06-12 14:09 | 2023-06-12 14:09 | QC checklist | (TC 06/12/2023 02:09:04 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
| 195414 | CM | CMPLT | CMPLT | 2023-06-12 14:23 | 2023-06-27 11:07 | EMG Installation | (TC 06/12/2023 02:22:36 PM, Baghirzade, Nara) Natus installation of EMG devices; replacement of old EMG devices.
1061208 | 195415 | CM | CMPLT | CMPLT | 2023-06-12 14:56 | 2023-06-12 15:01 | PACS PC not connecting to Olympus | (TC 06/12/2023 02:55:48 PM, Reid, Emmerson ) Was called and told this was not a PACS PC issue, but it was. I said I would help since I was there. The monitor was turned off, so i turned it on and they signed in. It still was not connecting to Olympus because the cable was damaged. I went to OR 12 but the cable was already taken. Then I went to OR 17 and took the cable from there. This fixed the issue. I let Maureen know to call help desk to get them to replace the other cables. After I left Maureen called back to get me to help explain to IT the issue and the cable, so I came back and explained it to them. They said they would be there to fix within the hour.
1055982 | 195416 | CM | CMPLT | CMPLT | 2023-06-13 08:34 | 2023-06-12 15:00 | In use on 14N PSN | (TC 06/12/2023 03:00:00 PM, Baghirzade, Nara) Radical is being used on a 14N PSN root asset #1056470 (16N-03). (TC 06/12/2023 03:00:00 PM, Baghirzade, Nara) Battery was replaced and upgraded.
1056470 | 195417 | CM | CMPLT | CMPLT | 2023-06-13 08:36 | 2023-06-12 15:00 | Radical associated with this Root does not turn on; constant alarming on PSN | (TC 06/12/2023 03:00:00 PM, Baghirzade, Nara) Radical (asset #1056528, SN: 1000206465) would not turn on, constant alarming on PSN. Replaced the radical with a spare radical belonging to 11South (asset #1055982). Probe well was also missing, replaced the probe well kit. After replacing the radical, the alarming stopped and the PSN root is operating as it should. Touched base with Amanda to let her know that device is now functional. Returned the defective Radical to Olcay for repair.
1061202 | 195418 | CM | CMPLT | CMPLT | 2023-06-13 09:53 | 2023-06-13 10:00 | OR 1 lights dimming | (TC 06/13/2023 09:53:38 AM, Reid, Emmerson ) Called reporting that the lights were at 100% but were getting dim intermittently. I asked them to check the anti-shadow settings and they said it was off. I went up to check and it seemed to be working normally, and the anti-shadow was off for all the lights. Stayed for a few minutes to watch and nothing happened. I left to return some pumps and check the cables replaced by IT. When I came back still nothing had happened with the lights. I told them to call again if anything happens.
| 195419 | CM | CMPLT | CMPLT | 2023-06-13 10:48 | 2023-06-13 10:52 | Accessories for ED defibs | (TC 06/13/2023 10:51:52 AM, Canjimavadivel, Kamallesh ) Provided to adapters to ED defibs from stock
1056425 | 195420 | CM | CMPLT | CMPLT | 2023-06-13 11:05 | 2023-06-13 11:06 | "needs to be calibrated" | (TC 06/12/2023 11:04:46 AM, Cortez, John) NIBP needs calibration notification when powered on Accessed NIBP calibration menu and proceeded through tests. Leak test and accuracy passed but, calibrated NIBP anyway All parameters ok
1055947 | 195421 | CM | CMPLT | CMPLT | 2023-06-13 11:13 | 2023-06-12 11:13 | BP cuff not reading | replaced adult cuff. tested BP, temp. and SpO2. functional.
1056466 | 195422 | CM | CMPLT | CMPLT | 2023-06-13 11:15 | 2023-06-12 11:14 | Cuff leaking | replaced adult cuff. tested BP, temp. and SpO2. functional.
| 195423 | CM | ACTIV | ISSUE | 2023-06-13 11:45 |  | CMMS Clean up committee | (TC 06/13/2023 11:45:39 AM, Scero, Tomlyn) Kick off zoom meeting (TC 06/27/2023 11:02:15 AM, Scero, Tomlyn) meeting (TC 07/04/2023 11:50:34 AM, Scero, Tomlyn) Weekly meeting discussing categories (TC 07/11/2023 11:34:20 AM, Scero, Tomlyn) Weekly meeting to discuss new asset categories (TC 07/18/2023 11:35:31 AM, Scero, Tomlyn) Weekly meeting (TC 07/25/2023 11:08:14 AM, Scero, Tomlyn) Weekly meeting (TC 08/01/2023 10:47:50 AM, Scero, Tomlyn) weekly meeting (TC 08/15/2023 11:47:21 AM, Scero, Tomlyn) Weekly CCMS clean up meeting (TC 09/12/2023 11:00:11 AM, Scero, Tomlyn) Zoom meeting (TC 09/19/2023 11:11:43 AM, Scero, Tomlyn) weekly meeting
1062417 | 195424 | CM | CMPLT | CMPLT | 2023-06-13 13:25 | 2023-06-13 13:26 | No power on accessories outlets | (TC 06/13/2023 01:25:23 PM, Samanian, Kamran) verified the problem. checked the power cord and it was fine. fuse tripped. put it back and all came up. communicated with the unit in case it happens then we will look into that.
1056454 | 195425 | CM | CMPLT | NFF | 2023-06-13 13:38 | 2023-06-13 13:39 | BP not working | (TC 06/13/2023 01:38:01 PM, Rhuland, Courtney) Came down stating the cuff continues to inflate past 300 mmHg Ran functional check and multiple BP readings No fault found
65246 | 195426 | CM | CMPLT | CMPLT | 2023-06-13 14:19 | 2023-06-13 14:24 | Battery replacement | (TC 06/13/2023 02:19:13 PM, Rhuland, Courtney) Battery replaced temperature check failed, temperature reading off Device out of support, retired
32560 | 195427 | CM | CMPLT | CMPLT | 2023-06-13 15:10 | 2023-06-12 20:00 | Protocols not up to date | Missing recently updated protocols This happened after scan PC was replaced and system drive was not compatible. New drive was installed and protocols were not saved on DVD ram because of a faulty drive. For the time being old protocols were restored but doses were not correct Original drive was installed on 522 CT and protocols retrieved and restored on 512 CT.
1061316 | 195428 | CM | CMPLT | CMPLT | 2023-06-13 15:24 | 2023-06-13 15:25 | TG Syngo IP address needs to be updated on Viewpoint | (TC 06/13/2023 09:24:00 AM, Fan, Tak) Logged in but DBA user was logged in already on another client, read-only. Could not make changes, referred to Jose. Jose away, Nixon couldn't log in either, GE remotely logged in to make the IP change.
63437 | 195429 | CM | CMPLT | NFF | 2023-06-13 15:49 | 2023-06-13 15:51 | "error message" | (TC 06/12/2023 03:48:44 PM, Cortez, John) Pump cleaned Ran functional tests (accuracy, alarms) successfully Unable to replicate issue
63222 | 195430 | CM | CMPLT | CMPLT | 2023-06-13 15:51 | 2023-06-13 15:52 | "system not running, always 'clogged downstream'" | (TC 06/12/2023 03:51:32 PM, Cortez, John) Pump cleaned Ran functional tests (accuracy, alarms) successfully Unable to replicate issue
1046035 | 195431 | CM | CMPLT | CMPLT | 2023-06-13 23:27 | 2023-06-16 15:25 | MSH ER DRX EVO E5-02(ER) COLLIMATOR NO ILLUMINATION | (TC 06/13/2023 03:21:36 PM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) COLLIMATOR NO ILLUMINATION (TC 06/13/2023 03:21:36 PM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) COLLIMATOR NO ILLUMINATION | MSH ER DRX EVO E5-02(ER) COLLIMATOR NO ILLUMINATION K5800-8895 SPAC4868 COLLIMATOR R221
| 195432 | PA | CMPLT | CMPLT | 2023-06-14 09:51 | 2023-07-04 10:03 | Aspirator batteries | (TC 06/14/2023 10:03:30 AM, Canjimavadivel, Kamallesh ) Placed order for batteries | Biomed stock, to be charged back as used
1067241 | 195433 | CM | CMPLT | CMPLT | 2023-06-14 11:21 | 2023-06-14 14:25 | 2 year PM | (TC 06/14/2023 02:24:38 PM, Scero, Tomlyn) 2 year parts replaced. Device passed self test.
1062281 | 195434 | CM | CMPLT | CMPLT | 2023-06-14 14:05 | 2023-07-19 11:49 | Multiple Error Codes - 3009-6-5-1503, 3009-1-23-1308, 3009-1-15-1308 | (TC 06/14/2023 02:06:35 PM, Ferreira, Qwynn) Got a call indicating the bed will not turn on. Initial troubleshooting and I was able to get the bed working. Cleaned the pins at first as they were dirty and the bed had a communication error. Errors were still present, but no motion of the bed was present. Mark was on the floor and I showed him this issue. Initially thought it was the fowler motor but seems to be more issues that suspected. Mark did not have parts for this, so he will come back to work on the issue. The bed has been removed from service and the unit is aware. (TC 07/11/2023 10:53:35 AM, Ferreira, Qwynn) Followed up with mark from stryker, however, no resolution and no parts or product available for the repair. Will follow up next week if still no solution (TC 07/19/2023 11:47:50 AM, Ferreira, Qwynn) Worked with Mark to fix the bed and complete the functional check. Mark supplied documentation and attached to the WO
| 195435 | CM | CMPLT | CMPLT | 2023-06-14 14:20 | 2023-06-14 14:22 | Artifact on Telepack 5 | (TC 06/14/2023 02:20:08 PM, Canjimavadivel, Kamallesh ) Brought up second set of leads and asked nurse to swap out. Issue was still observed. Asked nurse to switch out electrodes. Issue was cleared.
| 195436 | CON | CMPLT | CMPLT | 2023-06-14 14:36 | 2023-06-14 14:36 | Moving systems for demo.
1055774 | 195437 | CM | CMPLT | CMPLT | 2023-06-14 15:28 | 2023-08-02 13:25 | Created by Failed PM WO # 194551 : PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/16/2023 09:05:06 AM, Cortez, John) Contacted Zoll tech support, unit under warranty Will ship device for repair (TC 06/16/2023 09:05:06 AM, Cortez, John) Contacted Zoll tech support, unit under warranty Will ship device for repair RMA 1008449 (TC 06/16/2023 03:38:34 PM, Cortez, John) shipped to zoll (TC 08/02/2023 01:13:42 PM, Cortez, John) Came back from service Completed PM | (TC 06/14/2023 03:27:03 PM, Cortez, John) Discovered faulty speaker during PM Will call tech support
1053348 | 195438 | CM | CMPLT | CMPLT | 2023-06-14 15:39 | 2023-06-14 15:40 | MSH BI 1272 H3 12th Floo FACE SHIELD ASSY NOT HANGING CORRECTLY | (TC 06/14/2023 03:38:58 PM, Whittington, Scott) MSH BI 1272 H3 12th Floo FACE SHIELD ASSY NOT HANGING CORRECTLY FOUND PLATEON BACK OF ASSSEMBLY WAS LOOSE AND 2 SCREWS MISS. TIGHTENED OK | MSH BI 1272 H3 12th Floo FACE SHIELD ASSY NOT HANGING CORRECTLY 1/16 SAE HEX KEY
| 195439 | CM | CMPLT | CMPLT | 2023-06-14 17:08 | 2023-06-22 09:02 | working on EMG system installation | (TC 06/14/2023 05:08:25 PM, Samanian, Kamran) checked the errors. talked to IT. mapped drives and partially resolved the issue. system is not configured yet. (TC 06/15/2023 08:08:33 AM, Samanian, Kamran) worked with Christina remotely 12-1PM to get setting back up from old system and duplicate to the new system. it seems to be working now but Vinod needs to confirm it. (TC 06/23/2023 08:58:33 AM, Samanian, Kamran) Worked with Christinato finish the second system. all settings and configuration synced and copied. the only thing is that the in this version report is in DB. if Vinod needs it in MS then he needs to save it somewhere ( not doable in server any more ) and then take it away. (TC 06/22/2023 08:58:33 AM, Samanian, Kamran) Worked with Christinato finish the second system. all settings and configuration synced and copied. the only thing is that the in this version report is in DB. if Vinod needs it in MS then he needs to save it somewhere ( not doable in server any more ) and then take it away.
66006 | 195440 | CM | CMPLT | CMPLT | 2023-06-15 07:52 | 2023-06-15 07:54 | Broken case | (TC 06/15/2023 07:52:44 AM, Rhuland, Courtney) Front and rear case replaced (recall) PM conducted All tests passed Refer to ASM for details
| 195441 | CM | CMPLT | CMPLT | 2023-06-15 10:15 | 2023-06-15 10:20 | 12S called regarding suction was not working | (TC 06/15/2023 10:15:21 AM, Ferreira, Qwynn) Called regarding oxygen and suction not working Tested the oxygen on all four beds, was working fine. Tested suction and the three pods were working, 4th pod was difficult to see if it was a wall issue or regualtor issue. Told them to contact engineering to see if they can check the wall outlet for suction.
| 195442 | CM | CMPLT | CMPLT | 2023-06-15 10:23 | 2023-06-15 11:00 | Transport Incubator Light | (TC 06/15/2023 10:23:11 AM, Ferreira, Qwynn) Recieved call from NICU regarding the light on the transport being broken. Went to inspect and noticed the cable was cut from the transport Removed old threaded cable and removed light from hood. Worked with Samir to remove the cable and assembled new light with cable Installed the light and tested - Working
1057608 | 195443 | CM | CMPLT | CMPLT | 2023-06-15 10:30 | 2023-06-15 10:31 | Location change | (TC 06/15/2023 10:30:23 AM, Rhuland, Courtney) Brought over from BH Location changed to 16M and asset tag changed Functional check - OK
1056402 | 195444 | CM | CMPLT | CMPLT | 2023-06-15 10:32 | 2023-06-15 10:33 | Location change | (TC 06/15/2023 10:31:46 AM, Rhuland, Courtney) Brought over from BH Location changed to 15M and asset sticker changed Functional check - OK
1057616 | 195445 | CM | CMPLT | CMPLT | 2023-06-15 10:34 | 2023-06-15 10:35 | Location change | (TC 06/15/2023 10:34:20 AM, Rhuland, Courtney) Brought over from BH Location changed to 15M and asset tag changed Functional check - OK
1057119 | 195446 | CM | CMPLT | CMPLT | 2023-06-15 10:37 | 2023-06-21 09:05 | Location change | (TC 06/15/2023 10:36:48 AM, Rhuland, Courtney) Brought over from BH Location changed to 16M andasset tag changed Functional check - OK
| 195447 | CM | CMPLT | CMPLT | 2023-06-15 12:47 | 2023-06-15 12:49 | Repaired Exam Light on Transport Incubator. | (TC 06/15/2023 12:47:32 PM, Chukri, Samir) Replaced exam light cable
| 195448 | CM | CMPLT | CMPLT | 2023-06-15 12:52 | 2023-06-19 09:40 | ordered three exam light cable from RS Allied | (TC 06/15/2023 12:52:12 PM, Chukri, Samir) Ordered three cable from RS-Allied. (TC 06/19/2023 09:39:44 AM, Chukri, Samir) Received three cables used two and keep one in stock.
| 195449 | CM | CMPLT | CMPLT | 2023-06-15 13:23 | 2023-06-21 12:39 | Transport Incubator Light Damaged #2 | (TC 06/15/2023 01:23:34 PM, Ferreira, Qwynn) Received a call from NICU about a light not working Found the transport and the light was broken in a very similar way compared to another call previously today. Emailed Janet to ask her team if anyone noticed anything as well as Tabish. Cables will be ordered to replace. (TC 06/21/2023 12:37:32 PM, Ferreira, Qwynn) Replaced cable Added shrink tube for support and electrical tape for strain relief Tested - OK
64953 | 195450 | CM | CMPLT | CMPLT | 2023-06-15 14:01 | 2023-06-15 14:09 | error E7 | (TC 06/15/2023 02:04:33 PM, Scero, Tomlyn) Tried to take a blood pressure and error code E7 was produced. power cycled the unit and rested the unit. able to take 2 blood pressures without errors.
1046035 | 195451 | CM | CMPLT | CMPLT | 2023-06-15 16:05 | 2023-07-11 07:48 | MSH ER DRX EVO ARTIFACT | (TC 06/15/2023 04:14:22 PM, Whittington, Scott) MSH ER DRX EVO ARTIFACT (TC 06/16/2023 10:00:18 AM, Whittington, Scott) MSH ER DRX EVO ARTIFACT INSTALL COLLIMATOR 1715262 MFG NOV 2017 FROM LOCAL STOCK ORIGINAL SIZE CONING, SINGLE SPACER REQ OPERATING CORRECTLY (TC 06/16/2023 10:03:00 AM, Whittington, Scott) MSH ER DRX EVO ARTIFACT INSTALL COLLIMATOR 1715262 MFG NOV 2017 FROM LOCAL STOCK ORIGINAL SIZE CONING, SINGLE SPACER REQ (TC 06/16/2023 06:03:00 PM, Whittington, Scott) MSH ER DRX EVO ARTIFACT INSTALL COLLIMATOR 1715262 MFG NOV 2017 FROM LOCAL STOCK ORIGINAL SIZE CONING, SINGLE SPACER REQ | MSH ER DRX EVO ARTIFACT
1053347 | 195452 | CM | CMPLT | CMPLT | 2023-06-15 16:08 | 2023-07-11 07:48 | MSH BI 1270 H1 QUERY/RETRIEVE | (TC 06/15/2023 04:13:52 PM, Whittington, Scott) MSH BI 1270 H1 QUERY/RETRIEVE (TC 06/19/2023 01:19:39 PM, Whittington, Scott) MSH BI 1270 H1 QUERY/RETRIEVE CHARPELLE FROM HOLOGIC ON SITE LAST FRIDAY. ON THE INSTALL OF NEW CPU, THERE WAS A PT NAME PROGRAMMED INTO THE QUERY, SO WOULD NOT RETURN ANY RESULTS | MSH BI 1270 H1 QUERY/RETRIEVE
| 195453 | CM | CMPLT | CMPLT | 2023-06-16 08:10 | 2023-06-15 08:10 | Biomed Masimo Root reconfigured for biomed test | (TC 06/15/2023 08:10:08 AM, Samanian, Kamran) done
| 195454 | RDS | CMPLT | CMPLT | 2023-06-16 08:12 | 2023-06-15 08:12 | Meeting for PM june | (TC 06/15/2023 08:11:44 AM, Samanian, Kamran) done!
63435 | 195455 | CM | CMPLT | CMPLT | 2023-06-16 09:11 | 2024-02-13 15:30 | Buttons not responding | (TC 06/12/2023 09:10:44 AM, Cortez, John) Confirmed issue, will contact Cardinal Health for exchange (TC 06/12/2023 09:10:44 AM, Cortez, John) Confirmed issue, will contact Cardinal Health for exchange SC 1002982331 (TC 06/20/2023 03:45:03 PM, Cortez, John) shipped to cardinal health (TC 02/13/2024 03:13:07 PM, Cortez, John) Replacement pump came in, Reinstalled power adapter and clamp, Tested ok
65936 | 195456 | CM | CANCL | CANCL | 2023-06-16 09:49 |  | Channel error
| 195457 | CM | ACTIV | ISSUE | 2023-06-16 12:43 |  | Full team Meeting | (TC 06/16/2023 12:43:09 PM, Chukri, Samir) Meeting attended (TC 12/15/2023 11:19:34 AM, Chukri, Samir) full meeting (TC 01/19/2024 01:06:01 PM, Chukri, Samir) completed. (TC 02/16/2024 11:15:08 AM, Chukri, Samir) Meeting
501794 | 195458 | CM | CMPLT | CMPLT | 2023-06-16 12:52 | 2023-06-16 12:53 | MSH OPG CEOU - TV Probe Knob Issues | (TC 06/16/2023 12:52:41 PM, Chau, Ken) Screw loose, retightened, tested okay, reprocessed probe
501066 | 195459 | CM | CMPLT | CMPLT | 2023-06-16 12:54 | 2023-06-19 15:05 | MSH OPG RM#4 - IU22 - Not Working | (TC 06/16/2023 12:54:13 PM, Chau, Ken) System shuts off immediately and cannot grab any logs -case #0121851687 for power supply replacement (TC 06/19/2023 03:03:27 PM, Fan, Tak) Replaced power supply. Tested system, OK. Cody took PS back. | case #0121851687
502089 | 195460 | RQ | CMPLT | CMPLT | 2023-06-16 13:18 | 2023-09-13 19:41 | Symbia 2 message to contact service rep. Uncommanded Motion on H1 | (TC 06/16/2023 06:30:40 AM, Paz, Ramon (Tino)) Persistent error caused by uncommanded motion on Head1. System needs to be looked at ASAP. | error cleared by Tino remotely ~2 hrs ago
| 195461 | CM | CMPLT | CMPLT | 2023-06-16 14:49 | 2023-06-16 14:49 | IT Monthly meeting | (TC 06/16/2023 02:49:29 PM, Samanian, Kamran) done!
| 195462 | CM | CMPLT | CMPLT | 2023-06-16 15:34 | 2023-07-05 15:34 | misc biomed - june 2023 con't | (TC 06/15/2023 03:34:22 PM, Cortez, John) pm meeting (TC 06/16/2023 03:34:45 PM, Cortez, John) IS meeting (TC 06/20/2023 03:59:05 PM, Cortez, John) Ivy from interventional radiology requested that settings be changed on new vitals monitors in RM 514 Waiting for them to respond with which settings they want set (TC 06/19/2023 04:03:24 PM, Cortez, John) Monday huddle (TC 06/22/2023 01:26:31 PM, Cortez, John) Dropped off X3 to skills lab, swapped infusion pumps due for PM in pod C (TC 06/22/2023 01:28:24 PM, Cortez, John) Changed settings on patient monitors in Room 514 interventional radiology -SPO2 limits to 100 high, 90 low -NIBP interval 10 minutes (TC 06/22/2023 01:31:22 PM, Cortez, John) Edited PIICiX PM procedure (TC 06/21/2023 01:38:16 PM, Cortez, John) CE meeting (TC 06/21/2023 01:38:40 PM, Cortez, John) Systems committee meeting (TC 06/21/2023 01:38:40 PM, Cortez, John) Systems committee meeting (TC 06/22/2023 03:41:41 PM, Cortez, John) There were errors on Masimo Safetynet regarding several of the clients being down, and one of the Masimo Roots is alarming (SN 2000082114) Restarted most of the centrals. Unable to locate SN 2000082114. Will need to ask Naz and Becky if they know (TC 06/22/2023 03:47:54 PM, Cortez, John) Linda Fuller from Roxon called about needing an update a Bioscan bladder scanner Looked into it and Jessica is the tech who is working on it (TC 06/22/2023 04:15:43 PM, Cortez, John) Revised Masimo daily/weekly check instructions (TC 06/22/2023 05:33:55 PM, Cortez, John) Checked 10S equip room to verify if all vitals monitors were fully equipped, as requested by Barb (PCM) Added from spares -2 probe wells -2 spo2 sensors -1 welch allyn reg cuff -1 masimo external fan filter -1 dual lumen hose (TC 06/21/2023 05:35:44 PM, Cortez, John) Missing headboard in room 35A 10S. Borrowed from bed in hallway. (TC 06/27/2023 10:53:37 AM, Cortez, John) Contacted Murtala from Plexxus about status of various POs (TC 06/27/2023 10:59:43 AM, Cortez, John) TMS cleanup with Samir (TC 06/27/2023 11:00:35 AM, Cortez, John) PFT meeting (TC 06/30/2023 02:48:49 PM, Cortez, John) Delivered SPO2 LNCS Neo sensors to Vincy(CNS) in 10S (TC 06/30/2023 02:49:29 PM, Cortez, John) Responded to call in ICU podA rm10, Art waveform not showing on monitor. Reseated cable and resetted profile, issue resolved (TC 07/04/2023 02:37:13 PM, Cortez, John) Kamran and I responded to a call from Core lab about a stainer that was leaking. Attempted to troubleshoot it but concluded that it would be best if the lab contacted the vendor (TC 07/04/2023 03:00:51 PM, Samanian, Kamran) Kamran and I responded to a call from Core lab about a stainer that was leaking. Attempted to troubleshoot it but concluded that it would be best if the lab contacted the vendor
1046273 | 195463 | CM | CMPLT | CMPLT | 2023-06-16 16:53 | 2023-06-16 16:53 | 3D probes not being recognized. |  | Only 2 units are licenced for the 3D probes. Clinical were trying to use an unlicenced Aplio. Hardware Control Error TUS_IMG_118
67918 | 195464 | CM | CMPLT | CMPLT | 2023-06-19 07:18 | 2023-06-19 07:20 | System failure 25 | (TC 06/19/2023 07:18:33 AM, Rhuland, Courtney) System failure 25 alarm presented when powered on Fan was out of place, reseated fan and error was cleared Replaced broken scale with spare Functional check and scale check - OK
| 195465 | CM | CMPLT | CMPLT | 2023-06-19 07:46 | 2023-06-29 07:34 | Sterile water
1062121 | 195466 | PM | CMPLT | CMPLT | 2023-06-19 08:27 | 2023-06-19 08:30 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/14/2023 08:29:56 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67871 | 195467 | PM | CMPLT | CMPLT | 2023-06-19 08:31 | 2023-06-19 08:34 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/14/2023 08:32:43 AM, Rendon Lopez , Elizabeth) Updated software and library. PM completed, all tests passed
1062119 | 195468 | PM | CMPLT | CMPLT | 2023-06-19 08:34 | 2023-06-19 08:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/13/2023 08:35:58 AM, Rendon Lopez , Elizabeth) Updated software and library. PM completed, all tests passed
66555 | 195469 | PM | CMPLT | CMPLT | 2023-06-19 08:47 | 2023-06-19 09:04 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 06/13/2023 09:03:42 AM, Rendon Lopez , Elizabeth) handle cracked, repalced with spare. battery replaced, PM completed, all tests passed
1051701 | 195470 | CM | CMPLT | CMPLT | 2023-06-19 09:09 | 2023-06-19 09:10 | Innacurate temperature | (TC 06/12/2023 09:08:46 AM, Rendon Lopez , Elizabeth) Replaced temp probe, tested-OK
1061214 | 195471 | CM | CMPLT | CMPLT | 2023-06-19 09:12 | 2023-06-19 09:19 | Light iPad has no battery | (TC 06/19/2023 09:15:54 AM, Reid, Emmerson ) iPad had no battery charge left. I reseated the connection and it turned on and started charging. I tested to make sure the iPad worked to control the lights and turned off auto shadow. Working now.
1057279 | 195472 | CM | CMPLT | CMPLT | 2023-06-19 10:06 | 2023-06-19 10:07 | BP cuff missing | (TC 06/19/2023 10:05:53 AM, Rhuland, Courtney) BP cuff and hose replaced Temperature probe and well replaced functional check - OK
| 195473 | CM | CMPLT | CMPLT | 2023-06-19 10:43 | 2023-06-22 08:50 | PICiX Part - SSD-512GB and RAM 8 G for new PICiX - Flex Pro | (TC 06/19/2023 10:48:46 AM, Samanian, Kamran) placed the order S/C#1002983548 (TC 06/22/2023 08:49:48 AM, Samanian, Kamran) Received all parts. PO# 7202234022
65952 | 195474 | PM | CMPLT | CMPLT | 2023-06-19 11:37 | 2023-06-20 07:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 06/20/2023 07:54:22 AM, Rhuland, Courtney) Came down stating channel error Calibrated and performed check in to clear error PM complete All tests passed Refer to ASM for details
| 195475 | CM | ACTIV | ISSUE | 2023-06-19 12:08 |  | Pro-card expense report | (TC 06/19/2023 12:07:48 PM, Chukri, Samir) completed and managed Pro-Card expenses for Jun. (TC 08/21/2023 08:37:36 AM, Chukri, Samir) completed epense report for July. (TC 09/21/2023 02:45:05 PM, Chukri, Samir) submit Aug - Sep report (TC 09/22/2023 02:58:54 PM, Chukri, Samir) Replaced order for Jesica and kamran from Amazon used my Pro-Card. (TC 12/21/2023 01:29:19 PM, Chukri, Samir) completed the Pro-card reports for Nov 23 (TC 01/09/2024 09:25:28 AM, Chukri, Samir) Ordered video Cables for Kamarn (TC 02/05/2024 09:01:13 AM, Chukri, Samir) Ordered tools from Amazon to Elizabeth (TC 02/20/2024 09:53:21 AM, Chukri, Samir) working and download invoices and submit the report to be approved . (TC 03/19/2024 09:22:37 AM, Chukri, Samir) completed the Pro-card monthely report (TC 04/02/2024 09:21:55 AM, Chukri, Samir) completed the last invoice from Mar 2024
| 195476 | CM | CMPLT | CMPLT | 2023-06-19 14:38 | 2023-06-19 14:38 | huddle meeting | (TC 06/19/2023 02:38:32 PM, Samanian, Kamran) done!
1066851 | 195477 | CM | CMPLT | CMPLT | 2023-06-19 14:49 | 2023-06-19 14:50 | returned to company | (TC 06/15/2023 02:48:55 PM, Samanian, Kamran) Rob brought our back after repair and installed in EMRG cart. so he arrranged this loaner to be picked up by Olympus.
1046035 | 195478 | CM | CMPLT | CMPLT | 2023-06-19 17:06 | 2023-06-19 17:06 | Upright and table bucky off centre. Detector not charging in table connector. | (TC 06/26/2023 11:05:55 AM, Fan, Tak) Replaced tether cable. Detector needed a recalibration, calibration OK. | Re-centred table and upright bucky. Ordered new tether cable.
1064352 | 195479 | CM | CMPLT | CMPLT | 2023-06-19 17:11 | 2023-06-28 08:45 | Unit switches modes and screens sponaneously. |  | Reset unit. Called Philips Ref# 0121859432
| 195480 | CM | CMPLT | CMPLT | 2023-06-19 17:13 | 2023-06-19 17:13 | Site coverage.
62979 | 195481 | PM | CMPLT | CMPLT | 2023-06-20 07:30 | 2023-06-20 07:31 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/20/2023 07:30:09 AM, Rhuland, Courtney) PM complete Light intensity check - OK Electrical safety - OK
62978 | 195482 | PM | CMPLT | CMPLT | 2023-06-20 07:35 | 2023-06-20 07:35 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/20/2023 07:34:40 AM, Rhuland, Courtney) PM complete Filters cleaned Light intensity check - OK Electrical safety - OK
1061214 | 195483 | CM | CMPLT | CMPLT | 2023-06-20 08:12 | 2023-06-20 08:14 | ForceTriad not turning on | (TC 06/20/2023 08:12:39 AM, Reid, Emmerson ) ESU was not plugged in, plugged it in and it is working now.
1061210 | 195484 | CM | CMPLT | CMPLT | 2023-06-20 08:15 | 2023-06-20 08:34 | TS7000 remote issue | (TC 06/20/2023 08:14:53 AM, Reid, Emmerson ) Had to switch the remote from OR12, both and wokring now.
1055947 | 195485 | CM | CMPLT | CMPLT | 2023-06-20 08:19 | 2023-06-20 08:19 | BP cuff not working | Tested BP, SpO2 and temp. Functional.
66468 | 195486 | CM | CMPLT | CMPLT | 2023-06-20 08:34 | 2023-07-07 12:18 | Shipping C500 for repair | Shipping C500 for repair. The fixed-rate cost to repair a C500 is $2606.00. received C500. Draeger replaced service pcb and main(basis) PCB. Unit came damaged. rotary knob damaged and left USB out of place. installed and tested. functional
1061209 | 195487 | CM | CMPLT | CMPLT | 2023-06-20 09:21 | 2023-06-20 09:22 | TS7000 remote issue | (TC 06/20/2023 09:20:47 AM, Reid, Emmerson ) Reported that the remote was not working, when I get there they showed me but it was working. I told them to call me if there are more issue.
| 195488 | CM | CMPLT | CMPLT | 2023-06-20 11:22 | 2023-07-07 09:46 | Stryker Crib Handle | (TC 06/20/2023 11:22:12 AM, Ferreira, Qwynn) Received a call from John regarding a crib that almost injured a nurse Mentioned how we do not look after cribs - they are mechanical This crib had a scale and maybe thought it was us. Took a look at the crib and noticed the bearing sliders were loose. Tried to see if there is anyway we could repair this but we cannot. Reached out to stryker to see what can be done with the crib - waiting for response. (TC 07/07/2023 09:44:45 AM, Ferreira, Qwynn) No repsonse from stryker, worked with building services to ensure it gets fixed and is sent back to NICU for use. Tthe bearing lock is worn out - building services will be looking after this from now on
| 195489 | CM | CMPLT | CMPLT | 2023-06-20 12:08 | 2023-06-20 12:09 | Repaired Transport Incubator Light | (TC 06/20/2023 12:08:22 PM, Ferreira, Qwynn) Removed torn wiring from the light Rplaced cable and soldered Added heat shrink to protect cable and electrical tape for strain relief in the slot Tested - Working
1067294 | 195490 | ACC | CMPLT | CMPLT | 2023-06-20 12:48 | 2023-06-20 13:44 | Incoming inspection | (TC 06/20/2023 01:42:39 PM, Canjimavadivel, Kamallesh ) Assembled and added to database. All functions pass
| 195491 | CM | CMPLT | CMPLT | 2023-06-20 13:27 | 2023-06-20 13:28 | Image Capture in ENDO#4 | (TC 06/20/2023 01:27:16 PM, Samanian, Kamran) verified the problem ( not able to see the images in PowerChart ) checked the Hyland and device was OK. Checked the server and all was good. Advised Anita to open the tikcet to help desk for more investigation. Katrina indicated it is on the citrix for the client in room 4 as other rooms are OK and working fine
| 195492 | CM | CMPLT | CMPLT | 2023-06-20 13:30 | 2023-06-23 08:57 | order for Kingston SSD 256GB for old PICiX
68136 | 195493 | CM | CMPLT | CMPLT | 2023-06-20 13:45 | 2023-06-28 13:33 | Not turning on | (TC 06/20/2023 01:45:13 PM, Canjimavadivel, Kamallesh ) Attemped to replace batteries. Discovered corrosion on contact points. Cleaned and inserted batteries. Still not turning on. Issue might be with main board or keypad. Unable to get replacement parts, need to replace unit. (TC 06/23/2023 10:11:11 AM, Canjimavadivel, Kamallesh ) Confirmed issue is defective keypad. Quote has been requested for replacement. This asset to be retired
1049216 | 195494 | CM | CMPLT | CMPLT | 2023-06-20 14:43 | 2023-07-04 11:33 | Replacing HD and assign it to biomed | (TC 06/20/2023 02:42:50 PM, Samanian, Kamran) replaced the HD due to a failure, set up the PICiX fro biomed spare named as MSHSPAREPICIX2 with 172.31.16.53 IP address. (TC 07/04/2023 11:32:52 AM, Samanian, Kamran) connected to KVM. tested and passed
1046196 | 195495 | CM | CMPLT | CMPLT | 2023-06-20 15:46 | 2023-06-20 15:49 | Table movement error | Intermittent table movement error -Table top encoder replaced and adjustments made -Vertical movement adjustment failed - Micro limit switch adjusted -System checkup done and passed (TC 06/16/2023 07:00:00 PM, Mctaggart, Doug) Horizontal table movement error - "check if strap is caught under table top" -System has to be restarted to clear error -Worked with Siemens online support and it wa suggested to replace the pulley mechanism. -Prem ordered parts
66429 | 195496 | CM | CMPLT | CMPLT | 2023-06-20 15:51 | 2023-06-20 15:52 | Exposed leads | (TC 06/20/2023 03:51:32 PM, Cortez, John) Swapped leads with one from biomed stock
1062260 | 195497 | CM | CMPLT | CMPLT | 2023-06-20 15:55 | 2023-06-20 15:57 | Head section not moving up or down | (TC 06/20/2023 03:54:51 PM, Cortez, John) Head section controls were set to locked position Talked with ward clerk about lock setting
1056212 | 195498 | CM | CMPLT | CMPLT | 2023-06-20 15:58 | 2023-06-20 15:58 | Monitor loose on rollstand | (TC 06/20/2023 03:57:49 PM, Cortez, John) Tightened monitor on rollstand Cleaned external fan filter
1046275 | 195499 | CM | CMPLT | CMPLT | 2023-06-20 16:52 | 2023-06-20 16:53 | Broken TGC switch |  | Relaced TGC Slider top
| 195500 | CM | CMPLT | CMPLT | 2023-06-20 16:55 | 2023-06-20 16:55 | Site Coverage
65897 | 195501 | CM | CMPLT | CMPLT | 2023-06-21 07:33 | 2023-06-21 14:55 | Channel error | (TC 06/21/2023 02:53:30 PM, Rhuland, Courtney) Pressure sensor test failed during check in Replaced pressure sensor Calibrated and performed check in All tests passed Refer to ASM for details
1054612 | 195502 | CM | CMPLT | CMPLT | 2023-06-21 08:38 | 2023-06-21 08:39 | Low Temp Reading | (TC 06/21/2023 08:37:43 AM, Reid, Emmerson ) Temp was reported to be reading low. Ran calibration and it passed first attempt. Working now.
| 195503 | CM | ACTIV | ISSUE | 2023-06-21 08:55 |  | Working in TMS and Alert Recall system | (TC 06/20/2023 08:54:53 AM, Chukri, Samir) working on Msimo TC-I Tip Clip sensor recall emailed it to Peter and Steve. (TC 06/20/2023 08:56:48 AM, Chukri, Samir) Completed and emailed to Anita the Hemospray Endoscopic Hemostat recall from Cook Medical . (TC 06/21/2023 03:54:48 PM, Chukri, Samir) called Masimo service tech wairing RMA . (TC 06/22/2023 08:23:13 AM, Chukri, Samir) Recall for TC-I TipClip sensor RMA 1035821 return 4 sensor SC# 1002986597 (TC 06/23/2023 12:41:08 PM, Chukri, Samir) Alert and recall distrbuited (TC 06/23/2023 12:41:08 PM, Chukri, Samir) Alert and recall distrbuited (TC 06/26/2023 02:39:21 PM, Chukri, Samir) completed Incubators alert and PCiX rev 4 . (TC 06/28/2023 08:17:17 AM, Chukri, Samir) Received 4 TC-I tipclip sensors from Masimo,recall .PO # 7202236714 no charge. (TC 06/28/2023 01:29:14 PM, Chukri, Samir) correcting the new incoming equipment data for May 2023 (TC 06/29/2023 03:06:24 PM, Chukri, Samir) OneRecall introduction (TC 07/04/2023 09:25:53 AM, Chukri, Samir) Weekly-CMMS quality Project (TC 07/06/2023 02:13:25 PM, Chukri, Samir) OneRecall meeting (TC 07/11/2023 10:31:30 AM, Chukri, Samir) Correcting and approvied new equipment . (TC 07/12/2023 09:50:39 AM, Chukri, Samir) working with Emmerson and Jessica to correct the Neptune loaner asset to Hospital owned and returned the three defective to the Stryker (TC 07/12/2023 01:12:15 PM, Chukri, Samir) clear alert and recal from OneRecall system. (TC 07/18/2023 11:04:38 AM, Chukri, Samir) Pharmacy Team Inventory Meeting with OneRecall (TC 08/03/2023 10:29:57 AM, Chukri, Samir) Meeting OneRecall with O.R and Radiology department. (TC 08/02/2023 10:29:57 AM, Chukri, Samir) Meeting OneRecall with O.R and Radiology department. (TC 08/10/2023 02:30:49 PM, Chukri, Samir) distrbuted three recall from OneRecall (TC 08/16/2023 08:04:34 AM, Chukri, Samir) working in OneCall and reassigned some Recall (TC 08/16/2023 08:05:56 AM, Chukri, Samir) Meeting with OneCall with O.R team. (TC 08/17/2023 03:05:30 PM, Chukri, Samir) working on the OneRecall alert. (TC 08/25/2023 02:55:27 PM, Chukri, Samir) working on SAFER for Incident ID # 62467. (TC 09/01/2023 10:08:25 AM, Chukri, Samir) Helping Nara, to complete the safer regarding the Omni bed exam light issue. (TC 09/01/2023 10:09:39 AM, Chukri, Samir) Create PM schedules for the rest of Med-Surg bed. (TC 09/22/2023 02:55:03 PM, Chukri, Samir) completed the Radical-7 Batteries recall, and working on OneRecall system (TC 11/16/2023 11:38:19 AM, Chukri, Samir) Meeting with Becky working on the TMS Account correction. (TC 12/06/2023 08:19:13 AM, Chukri, Samir) working on One Recall system and distributed recall and alerts (TC 12/05/2023 02:15:00 PM, Chukri, Samir) working on One Recall system and distributed recall and alerts (TC 12/12/2023 08:34:42 AM, Chukri, Samir) Working on TMS and approved new asset and mass edit (TC 12/13/2023 02:18:34 PM, Chukri, Samir) working on TMS and add a new PM schedules for Hemodialysis unit in ICU (TC 12/15/2023 02:56:46 PM, Chukri, Samir) working on One-Recall system and completed some alert and recall (TC 01/05/2024 03:22:44 PM, Chukri, Samir) working with Maisha to update the Matrial biomed stock on the TMS (TC 01/11/2024 02:03:56 PM, Chukri, Samir) working On OneRecall completed Alert and recall (TC 01/22/2024 08:17:41 AM, Chukri, Samir) working in total weekly hours for SH biomed query (TC 01/23/2024 11:21:44 AM, Chukri, Samir) completed the query for the SHS weekly work time (TC 01/23/2024 02:24:59 PM, Chukri, Samir) Working on TMS account update with Becky (TC 01/24/2024 02:20:58 PM, Chukri, Samir) Activated asset tags (TC 01/25/2024 02:58:22 PM, Chukri, Samir) working on alert onerecall (TC 01/29/2024 11:59:37 AM, Chukri, Samir) Update the Biomed work time Query and donloaded on excel sheet with Pavit table emaileded to Adeel, wating to review it. (TC 01/30/2024 11:00:59 AM, Chukri, Samir) Completed and review the asset waiting for the approve and approved equipment from emmerson and kam. (TC 02/05/2024 07:55:14 AM, Chukri, Samir) Review and distrbute alert on Onerecall (TC 02/12/2024 09:58:23 AM, Chukri, Samir) change the time charges-weekly data sources (TC 02/13/2024 02:35:54 PM, Chukri, Samir) distributed urgent alerts, and working in TMS PM schedule (TC 02/14/2024 11:11:42 AM, Chukri, Samir) Working on the OneRecall system and distrbuted alert and recall . (TC 02/20/2024 02:40:33 PM, Chukri, Samir) Working on the medfusion pump recall and emailed the respond to Sedgwick (TC 02/26/2024 11:21:24 AM, Chukri, Samir) working on OneRecall distrbuted and closed the complete recall. (TC 03/19/2024 02:08:27 PM, Chukri, Samir) Completed batch asset template for 20 Assets for Symphoney Breast pumps (TC 03/21/2024 02:50:52 PM, Chukri, Samir) review and distributed recal and alert in OneRecall system. (TC 03/25/2024 02:27:45 PM, Chukri, Samir) working on One-Recall system and distrbuted Alert and Recalls (TC 03/26/2024 02:49:53 PM, Chukri, Samir) checked and activated the Asset from Rutvik (TC 04/02/2024 11:00:39 AM, Chukri, Samir) working in OneRecall system and distributed recalla and alerts (TC 04/03/2024 09:57:50 AM, Chukri, Samir) Mass edit for new 22 Breast feeding Pump and printe 22 asset numbers (TC 04/18/2024 01:41:16 PM, Chukri, Samir) working on the OneRecall alert (TC 06/11/2024 02:34:13 PM, Chukri, Samir) working in OneRecall (TC 08/20/2024 02:33:13 PM, Chukri, Samir) distrbuted Alert and working on Onerecall .
1057607 | 195504 | CM | CMPLT | CMPLT | 2023-06-21 09:02 | 2023-06-21 09:03 | Missing BP hose | (TC 06/21/2023 09:02:17 AM, Rhuland, Courtney) Replaced blood pressure hose Functional check with simCube - OK
| 195505 | CM | CMPLT | CMPLT | 2023-06-21 10:15 | 2023-08-28 14:02 | OR headlight warranty repair
| 195506 | CM | ACTIV | ISSUE | 2023-06-21 10:17 |  | CO2 laser repair service
| 195507 | CM | CMPLT | CMPLT | 2023-06-21 10:18 | 2023-06-27 09:26 | ordering parts for MBU | (TC 06/27/2023 09:24:11 AM, Chukri, Samir) Parts received by Naz
| 195508 | PA | CMPLT | CMPLT | 2023-06-21 10:51 | 2023-07-05 13:33 | Ordering Alcohol wipes for Biomed | (TC 06/21/2023 10:54:12 AM, Reid, Emmerson ) Placed an order for 20 boxes of alcohol wipes. (TC 07/04/2023 01:45:55 PM, Reid, Emmerson ) had to pick up a box from materials, order quantity was not correct. Placed another order.
1061204 | 195509 | CM | CMPLT | CMPLT | 2023-06-21 14:20 | 2023-06-21 14:21 | Monitor showing half of the screen | (TC 06/21/2023 02:20:09 PM, Reid, Emmerson ) Monitor was only displaying on the left half of the screen. I adjusted the display settings and it was normal now. Someone must have played with the monitor.
1055585 | 195510 | CM | CMPLT | CMPLT | 2023-06-21 14:37 | 2023-06-22 10:57 | MX board not communicating - no pulse oximetry | (TC 06/22/2023 10:55:32 AM, Ferreira, Qwynn) Tested and had the error dispaly on the radical. Removed battery and reset radical and then reinstalled Tested - working normally now, no errors displayed on te screen. Tested SPo2 and was working correctly.
1057758 | 195511 | CM | CMPLT | CMPLT | 2023-06-21 14:38 | 2023-06-26 14:23 | Delay in boot up | (TC 06/21/2023 02:38:06 PM, Samanian, Kamran) the host was patched and crashed. rebuilt by Kam. taken very lomng time to boot up. troublshooted, SSD and RAM both are fine. found out that Philips USB flash on the motherboard is corrupted. Rev C.o3.06 does not need the USB. it was removed from motherboard and host came up in normal working condition, running test in biomed shop. (TC 06/26/2023 02:21:51 PM, Samanian, Kamran) returned the host to the location, turned on, all looks good.
1055042 | 195512 | CM | CMPLT | CMPLT | 2023-06-21 14:39 | 2023-06-22 10:59 | MX board not communicating - no pulse oximetry | (TC 06/22/2023 10:57:00 AM, Ferreira, Qwynn) Reset the radical 7 as the error message was displayed at the top of the device. Reinstalled the radical 7 and tested Tested spO2 as they noted it was not working, but once the radical was reset, it was working correctly.
| 195513 | CM | CMPLT | CMPLT | 2023-06-21 15:03 | 2023-06-21 15:03 | meeting-Phase 3A update | (TC 06/21/2023 03:03:24 PM, Samanian, Kamran) done!
1062257 | 195514 | CM | CMPLT | CMPLT | 2023-06-21 15:18 | 2023-06-21 15:20 | Bent Aux Power Cord | (TC 06/21/2023 03:18:34 PM, Reid, Emmerson ) Aux power was bent and needed to be replaced. Working now.
1055788 | 195515 | CM | CMPLT | CMPLT | 2023-06-22 08:57 | 2023-06-22 08:59 | R Series not checked in for 48 hours | (TC 06/22/2023 08:57:18 AM, Reid, Emmerson ) Defib was left on pacer mode. I ran a 30J test and turned off. WIll see if it checks in tomorrow.
66766 | 195516 | PM | CMPLT | CMPLT | 2023-06-22 09:06 | 2023-06-22 09:20 | PM MAJOR: Annual PM - For CADD Pump | (TC 06/22/2023 09:18:24 AM, Reid, Emmerson ) PM complete, see proceedure for details.
1057027 | 195517 | CM | CMPLT | CMPLT | 2023-06-22 09:17 | 2023-06-22 09:20 | BP not taking reading | (TC 06/22/2023 09:18:05 AM, Canjimavadivel, Kamallesh ) BP measurement taking too long. Found leak in cuff. Replaced cuff, tested ok
1057021 | 195518 | CM | CMPLT | CMPLT | 2023-06-22 09:20 | 2023-06-22 09:26 | Not reading accurately | (TC 06/22/2023 09:20:11 AM, Canjimavadivel, Kamallesh ) Reading high, calibrated with tester. Tested ok
1057018 | 195519 | CM | CMPLT | CMPLT | 2023-06-22 09:27 | 2023-07-14 07:49 | Not taking reading | (TC 06/22/2023 10:57:17 AM, Canjimavadivel, Kamallesh ) When attempting to initiate calibration, device did not behave as expected. Removed rechargeable battery and tested with AA batteries, was able to initiate calibration. (TC 06/22/2023 01:02:33 PM, Canjimavadivel, Kamallesh ) Calibrated device with tester. Rechargeable batteries will need to be replaced but device can go back with AA for now. (TC 07/14/2023 07:48:56 AM, Canjimavadivel, Kamallesh ) Installed rechargeable battery
1057026 | 195520 | CM | CMPLT | CMPLT | 2023-06-22 09:41 | 2023-07-14 07:45 | Not taking readings | (TC 06/22/2023 10:52:41 AM, Canjimavadivel, Kamallesh ) When attempting to initiate calibration, device did not behave as expected. Removed rechargeable battery and tested with AA batteries, was able to initiate calibration. (TC 06/22/2023 01:01:23 PM, Canjimavadivel, Kamallesh ) Calibrated device with tester. Rechargeable batteries will need to be replaced but device can go back with AA for now. (TC 07/14/2023 07:43:44 AM, Canjimavadivel, Kamallesh ) Replaced rechargeable battery pack
| 195521 | PA | CMPLT | CMPLT | 2023-06-22 10:29 | 2023-07-14 07:43 | Rechargeable battery for braun PRO 6000 | (TC 06/22/2023 10:30:40 AM, Canjimavadivel, Kamallesh ) Received quote from vendor, placed order for parts. | Biomed stock, charge back when used
1064418 | 195522 | CM | CMPLT | CMPLT | 2023-06-22 10:59 | 2023-06-22 11:05 | Disassembly and adding coils | (TC 06/22/2023 10:59:38 AM, Ferreira, Qwynn) Sandy called regarding the coils were stuck. Removed the old table and installed spine coil and head coil. Cleaned MRI incubator as well as it was dirty.
1055940 | 195523 | CM | CMPLT | CMPLT | 2023-06-22 13:14 | 2023-06-22 13:14 | temp probe not working | tested BP, temp and SpO2. Functional.
1061454 | 195524 | CM | CMPLT | CMPLT | 2023-06-22 13:16 | 2023-06-22 13:16 | bed exit alarm not working | tested bed. alarm not functioning. talked to Mark from Stryker. printed user manual. reset scale. started alarm. functional.
1055940 | 195525 | CM | CMPLT | CMPLT | 2023-06-22 13:35 | 2023-06-22 13:36 | temp probe well blocked | (TC 06/22/2023 01:34:52 PM, Cortez, John) Amanda Kerr (CNS) called about several monitors with blocked temp probe wells Replaced wells on these Roots
1055945 | 195526 | CM | CMPLT | CMPLT | 2023-06-22 13:37 | 2023-06-22 13:38 | Temp probe well blocked | (TC 06/21/2023 01:37:20 PM, Cortez, John) Amanda Kerr (CNS) called about several monitors with blocked temp probe wells Replaced wells on these Roots
1062099 | 195527 | CM | CMPLT | CMPLT | 2023-06-22 13:41 | 2023-06-22 13:41 | Replaced PS500 batteries | replaced PS500 batteries. tested. Functional.
1063709 | 195528 | CM | CMPLT | CMPLT | 2023-06-22 13:50 | 2023-07-11 07:49 | MSH OR SPARQ LOCATE FOR IN SERVICE | (TC 06/23/2023 08:38:00 AM, Whittington, Scott) MSH OR SPARQ LOCATE FOR IN SERVICE SETUP & PREP | MSH OR SPARQ LOCATE FOR IN SERVICE
1057060 | 195529 | CM | CMPLT | CMPLT | 2023-06-22 14:23 | 2023-07-11 07:50 | MSH BI HOLOGIIC SVDX MGR SERVER MAKE BACKUP | (TC 06/22/2023 07:23:00 AM, Whittington, Scott) MSH BI HOLOGIIC SVDX MGR SERVER MAKE BACKUP | MSH BI HOLOGIIC SVDX MGR SERVER MAKE BACKUP
1061205 | 195530 | CM | CMPLT | CMPLT | 2023-06-23 09:38 | 2023-06-23 09:41 | Pxyis Medstation not turning on | (TC 06/23/2023 09:39:29 AM, Reid, Emmerson ) Went up to check OR 4 and see the Pxyis. When I got there it was turned on. They said it was not before. I checked the power cable and it was barely in the wall because one of the waste bins was pulling it out. I securely connected the AC power. Working now.
66709 | 195531 | CM | CMPLT | CMPLT | 2023-06-23 12:59 | 2023-06-23 13:05 | Broken door latch | (TC 06/23/2023 01:04:06 PM, Rhuland, Courtney) Repaired broken door latch Replaced two missing porthole seals Functional check - OK
1049362 | 195532 | CM | CMPLT | CMPLT | 2023-06-23 16:40 | 2023-06-23 16:40 | A side piston issued due to contrast pill | A side piston issued due to contrast pill Removed covers and took apart piston A cleaned all contrast spill and tested functionality after. In the mean time spare injector was moved down from fifth floor and setup for use.
1067317 | 195533 | CM | CMPLT | CMPLT | 2023-06-26 07:34 | 2023-06-26 11:16 | Incoming inspection | (TC 06/26/2023 11:15:28 AM, Rhuland, Courtney) Unboxed and performed functional check Ran test feed for 30 minutes
| 195534 | OC | CMPLT | CMPLT | 2023-06-26 07:55 | 2023-06-25 12:30 | Red light on black box in doctor's room zone c | (TC 06/25/2023 12:30:00 PM, Baghirzade, Nara) Christine ext 103298 Dr.'s room in zone c and by christine, there is a black box connected to computer blinking red light. not a medical device. call got disconnected and when i was transferred to christine again, the call was disconnected. tried to call back 3 times but nobody would answer. remoted in and checked all NICU centrals, all functional. Will follow up on Monday morning regarding this device.
68755 | 195535 | CM | CMPLT | NFF | 2023-06-26 07:58 | 2023-06-26 07:59 | Not charging | (TC 06/26/2023 07:58:13 AM, Rhuland, Courtney) Came down stating battery would not charge Charged device over night and battery run time showed full charge No fault found
| 195536 | PA | CMPLT | CMPLT | 2023-06-26 08:02 | 2023-07-06 08:48 | Incubator Parts
68820 | 195537 | PM | CMPLT | CMPLT | 2023-06-26 08:15 | 2023-06-26 08:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/26/2023 08:15:40 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 200 mohms Leakage current: 73.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. Battery conditioning completed. Old battery capacity: 2139 mamphr New battery capacity: 4002 mamphr PM completed.
67874 | 195538 | PM | CMPLT | CMPLT | 2023-06-26 08:17 | 2023-06-26 08:20 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/19/2023 08:18:44 AM, Rendon Lopez , Elizabeth) Upgraded software, downloaded library. PM completed, all tests passed
1050098 | 195539 | CM | CMPLT | CMPLT | 2023-06-26 08:18 | 2023-08-03 11:17 | Rear mount broke | (TC 06/26/2023 11:12:51 AM, Canjimavadivel, Kamallesh ) Contacted vendor for part and pricing info for mount. Placed order. (TC 08/03/2023 11:16:04 AM, Canjimavadivel, Kamallesh ) Received mounting hardware. Replaced mount and returned device to ED
67887 | 195540 | PM | CMPLT | CMPLT | 2023-06-26 08:20 | 2023-06-26 08:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/19/2023 08:21:28 AM, Rendon Lopez , Elizabeth) Upgraded software, downloaded library. PM completed, all tests passed
1059744 | 195541 | CM | CMPLT | CMPLT | 2023-06-26 08:22 | 2023-06-26 08:23 | Turns off by itself | (TC 06/19/2023 08:22:19 AM, Rendon Lopez , Elizabeth) Low battery. Replaced battery, tested-OK
67877 | 195542 | PM | CMPLT | CMPLT | 2023-06-26 08:25 | 2023-06-26 08:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/20/2023 08:26:33 AM, Rendon Lopez , Elizabeth) Upgraded software, downloaded library. PM completed, all tests passed
| 195543 | CM | CMPLT | CMPLT | 2023-06-26 08:37 | 2023-10-30 09:03 | L&D OR 5 - Touchscreen monitor blank | (TC 06/21/2023 08:37:08 AM, Rendon Lopez , Elizabeth) Troubleshooting monitor, tried new video calbe, RS-232 cable, not resolved. (TC 06/22/2023 08:37:08 AM, Rendon Lopez , Elizabeth) Swapped monitor, same issue. Found issue may be with cabled behind wall, contacted Olympus, they will have to come in and repair. (TC 06/26/2023 09:29:12 AM, Rendon Lopez , Elizabeth) Requested service, created PO for service (TC 07/17/2023 08:36:15 AM, Rendon Lopez , Elizabeth) Talking to olympus tech support. They will come in July 25th to diagnose (TC 07/25/2023 09:15:59 AM, Rendon Lopez , Elizabeth) working with olympus tech to teoubleshoot issue. Found video/ touchscreen processor to be the problem. System is no longer supported, integration team will send us software to reload, or we will have to replace system (TC 08/03/2023 09:24:25 AM, Rendon Lopez , Elizabeth) Received software to reload on AV interface, downloaded software, in process (TC 08/16/2023 01:17:00 PM, Rendon Lopez , Elizabeth) Try to reload software and config, unable to connect to main processor. System is out of support, received letter from Olympus, notified unit. (TC 10/30/2023 09:02:58 AM, Rendon Lopez , Elizabeth) Waiting on decision to repalced system.
65307 | 195544 | CM | CMPLT | CMPLT | 2023-06-26 08:40 | 2023-06-26 08:41 | Missing power cord | (TC 06/23/2023 08:40:38 AM, Rendon Lopez , Elizabeth) Replaced power cord with spare, charging, tested-OK
1056540 | 195545 | CM | CMPLT | CMPLT | 2023-06-26 08:42 | 2023-06-26 08:42 | Alarming replace battery | (TC 06/23/2023 08:41:44 AM, Rendon Lopez , Elizabeth) Replaced battery, self test performed, all passed
1067302 | 195546 | CM | CMPLT | CMPLT | 2023-06-26 09:55 | 2023-06-26 09:58 | Malfunction | (TC 06/23/2023 09:55:32 AM, Rendon Lopez , Elizabeth) Breathing block with AGM 68554 has residual ventilation with no patient. Found high leak on breathing block. 6 year PM parts replaced, same issue, replaced APL valve, resolved. Tested, leak-Ok, self test-OK. Entered new asset in TMS.
| 195547 | CM | CMPLT | CMPLT | 2023-06-26 09:59 | 2023-08-11 15:29 | Zeus AGM spare parts | (TC 06/26/2023 10:02:46 AM, Rendon Lopez , Elizabeth) Zeus IE AGM spare parts ordering Received APL valve 07/25 received Oxytrace 08/11 Rec'd safety valve
1061210 | 195548 | CM | CMPLT | CMPLT | 2023-06-26 10:43 | 2023-06-26 10:47 | VMC error | (TC 06/26/2023 10:46:41 AM, Reid, Emmerson ) VMC had an error message and was not letting the nurses change the monitors in OR9. I rebooted the VMC from the closet. Working now.
1055914 | 195549 | CM | CMPLT | CMPLT | 2023-06-26 11:05 | 2023-06-26 11:06 | Temperature probe not working | (TC 06/26/2023 11:05:20 AM, Baghirzade, Nara) Confirmed that probe is defective; repeated error when taking the probe out of the well, does not take any measurements. Replaced temperature probe with a spare probe. All functional tests passed.
| 195550 | PA | CMPLT | CMPLT | 2023-06-26 11:27 | 2023-07-11 09:19 | Formlabs - Research Materials | (TC 07/11/2023 09:19:15 AM, Ferreira, Qwynn) Received materials and organized the room with the supplies. sent the invoice to Tanya for funding
1057761 | 195551 | CM | CMPLT | CMPLT | 2023-06-26 14:28 | 2023-06-29 09:32 | Created by Failed PM WO # 194601 : PM MAJOR-Monitors, Physiologic, Central Station | (TC 06/26/2023 02:28:58 PM, Canjimavadivel, Kamallesh ) Swapped host with spare PICiX and attempted to configure as previous setup. PC comes up with unexpected error and does not work. Swapped with second spare PICiX. Brought both PCs with errors to Biomed for further investigation. (TC 06/26/2023 08:10:39 AM, Samanian, Kamran) tried to do a few things to resolve the issue and fix the windows but no success. started rebuilding the PICIX host from scratch (TC 06/27/2023 08:12:00 AM, Samanian, Kamran) finished the host and ready to put it back to the unit (TC 06/29/2023 09:31:09 AM, Canjimavadivel, Kamallesh ) Installed host back to original location and completed setup | (TC 06/26/2023 02:25:04 PM, Canjimavadivel, Kamallesh ) Cleaned PC and restarted PC to complete OS patch install. Host keeps shutting down on its own
| 195552 | CM | CMPLT | CMPLT | 2023-06-27 07:36 | 2023-07-26 08:50 | Transport Incubator Light | (TC 06/26/2023 10:38:00 AM, Ferreira, Qwynn) Light did not work after testing - opened light to see if there was a short. No short - but 4 of the 6 lights are not working Might need new board. (TC 07/07/2023 09:52:58 AM, Ferreira, Qwynn) Ordered new light bar for the device Could not troubleshoot the broken light - Samir agreed we dispose and purchase new. (TC 07/26/2023 08:48:47 AM, Ferreira, Qwynn) Received light bar and closed WO Will keep the light as spare as well as store the replacement cable.
1061210 | 195553 | CM | CMPLT | CMPLT | 2023-06-27 08:52 | 2023-06-27 08:53 | Vitals not showing on Monitors | (TC 06/27/2023 08:52:15 AM, Reid, Emmerson ) Vitals was not able to display on the monitors, it showed a bnlue screen. I checked the cable from the monitor to the boom and tightened the end connecting to the boom. Picture came up, working now.
1056463 | 195554 | CM | CMPLT | CMPLT | 2023-06-27 09:04 | 2023-06-27 09:04 | no "stand by" option on main screen | no "stand by" option on main screen. enabled "stand by" shortcut. Tested BP. SpO2 and temp. Functional.
1056466 | 195555 | CM | CMPLT | CMPLT | 2023-06-27 09:05 | 2023-06-27 09:05 | not working | Tested BP. SpO2 and temp. Functional.
1055947 | 195556 | CM | CMPLT | CMPLT | 2023-06-27 09:06 | 2023-06-27 09:06 | temp prob stuck | repmaoved probe and cover. Tested BP. SpO2 and temp. Functional.
1049608 | 195557 | CM | CMPLT | CMPLT | 2023-06-27 09:19 | 2023-06-27 09:21 | MSH SPP - No Network Connectivity | (TC 06/27/2023 09:20:08 AM, Chau, Ken) Network got stuck and cannot connect to network despite removing from jack and restarting. Only resolution was to contact IT to restart the network jacks. -Tested okay after | Case#NSD034801144
1057570 | 195558 | PM | CMPLT | CMPLT | 2023-06-27 09:45 | 2023-06-27 10:20 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 06/27/2023 10:09:02 AM, Reid, Emmerson ) Was called about the cart. Two of the wheels were bent. I replaced one caster with a spare and the other I was able to bend into place. Completed ESU PM as well and PM checklist is attached to the WO. Also went over FT10 proceedure with Kamran.
1056125 | 195559 | PM | CMPLT | CMPLT | 2023-06-27 09:57 | 2023-06-27 09:58 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 06/27/2023 09:57:24 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 195560 | CM | CMPLT | CMPLT | 2023-06-27 10:32 | 2023-06-27 10:32 | gathering information on ICU hosts and IT closet | (TC 06/27/2023 10:32:03 AM, Samanian, Kamran) gathered informatioon and updated excel sheet
1061033 | 195561 | CM | CMPLT | CMPLT | 2023-06-27 11:08 | 2023-07-27 10:24 | Switched off and alarming on another view station | (TC 06/27/2023 11:08:45 AM, Baghirzade, Nara) Found the view station to be switched off and unplugged; switched the display back on and the booting cycle was successful; no errors. After a few minutes, the second view station for 16N was alarming that the first view station is down. Unsure of why it is alarming at this point since the display is plugged in and is live. Naz will be connecting with the remote tech support from Masimo to help figure out why this keeps alarming. Informed Amanda to ensure that the view stations remain on.
1056470 | 195562 | CM | CMPLT | CMPLT | 2023-06-27 11:12 | 2023-06-27 11:13 | Defective power cord | (TC 06/27/2023 11:12:09 AM, Baghirzade, Nara) Root was not charging; replaced the power cord. Battery is charging and the charging symbol on the battery is present.
1056459 | 195563 | CM | CMPLT | CMPLT | 2023-06-27 11:13 | 2023-06-27 11:13 | Defective power cord | (TC 06/27/2023 11:13:08 AM, Baghirzade, Nara) Root was not charging; replaced the power cord. Battery is charging and the charging symbol on the battery is present.
1056463 | 195564 | CM | CMPLT | CMPLT | 2023-06-27 11:14 | 2023-06-27 11:16 | Beeping when on standby | (TC 06/27/2023 11:13:54 AM, Baghirzade, Nara) Clinical staff complained that when putting the Root on standby, it alarms/beeps every few seconds. Checked the settings on the Root and found that the standby tone interval was set to 30 seconds. Took another Patient Safety Net Masimo to compare settings with. Switched off the standby tone interval and data collections was switched off also as per the settings with another PSN Masimo. Functionally ok.
501583 | 195565 | CM | CMPLT | CMPLT | 2023-06-27 13:17 | 2023-07-11 07:51 | MSH SPP 15-217 ROSE L&D Hallway (IU22) 501583 OVERHEATING ERROR, SHUTDOWN | (TC 06/27/2023 01:26:58 PM, Whittington, Scott) MSH SPP 15-217 ROSE L&D Hallway (IU22) 501583 OVERHEATING ERROR, SHUTDOWN SITE 52668740 REF# 0121900949 | MSH SPP 15-217 ROSE L&D Hallway (IU22) 501583 OVERHEATING ERROR, SHUTDOWN SITE 52668740 REF# 0121900949
1061210 | 195566 | CM | CMPLT | CMPLT | 2023-06-27 13:34 | 2023-06-27 13:35 | Monitor not working | (TC 06/27/2023 01:33:40 PM, Reid, Emmerson ) One of the monitors was not showing the image. I checked it out and "port B" was selected instead of "port A". I swithced it and locked the controls. Working now.
67393 | 195567 | CM | CMPLT | CMPLT | 2023-06-27 14:12 | 2023-07-21 15:15 | Scale not working | (TC 06/27/2023 09:06:06 AM, Rendon Lopez , Elizabeth) Replaced scale with one from spare warmer. Troubleshooting scale, unable to calibrate and get load cell readings, ordered load cell replacements (TC 07/20/2023 02:54:47 PM, Rendon Lopez , Elizabeth) Replaced top load cell, chek, issue not resolved. (TC 07/21/2023 03:13:41 PM, Rendon Lopez , Elizabeth) Replaced bottom load cell, working properly. Returned old top cell, tested-OK. Calibrated scale, tested different times, working. All warmer functions OK, returned to L&D
| 195568 | CM | CMPLT | CMPLT | 2023-06-27 14:30 | 2023-06-27 15:12 | Expair meeting | (TC 06/27/2023 03:12:06 PM, Baghirzade, Nara) Expair meeting
1066941 | 195569 | CM | CMPLT | CMPLT | 2023-06-27 14:37 | 2023-07-11 10:35 | Incoming inspection | (TC 07/11/2023 10:34:21 AM, Chukri, Samir) Incoming Inspection and switched with old unit.
1053348 | 195570 | CM | CMPLT | CMPLT | 2023-06-27 17:02 | 2023-07-11 07:51 | MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING | (TC 06/27/2023 05:02:24 PM, Whittington, Scott) MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING (TC 06/30/2023 01:35:48 PM, Whittington, Scott) MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING FOOT PEDAL RECEIVED (TC 06/27/2023 05:02:24 PM, Whittington, Scott) MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING (TC 06/27/2023 05:02:24 PM, Whittington, Scott) MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING (TC 06/30/2023 05:35:00 PM, Whittington, Scott) MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING FOOT PEDAL RECEIVED FOOT PEDAL REPLACED (TC 06/30/2023 01:35:00 PM, Whittington, Scott) MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING FOOT PEDAL RECEIVED (TC 06/30/2023 05:24:00 PM, Whittington, Scott) MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING FOOT PEDAL REPLACED | MSH BI 1272 H3 ON ONE FOOTPEDAL ASSLY THGEGANTRY DOWN PEDAL NOT WORKING
| 195571 | CM | CMPLT | CMPLT | 2023-06-28 07:30 | 2023-06-27 15:35 | NICU Huddle | (TC 06/27/2023 03:35:00 PM, Baghirzade, Nara) NICU huddle
65934 | 195572 | FC | CMPLT | CMPLT | 2023-06-28 07:49 | 2023-06-28 07:52 | Functional check | (TC 06/28/2023 07:49:14 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
63439 | 195573 | FC | CMPLT | CMPLT | 2023-06-28 07:52 | 2023-06-28 07:56 | Functional check | (TC 06/28/2023 07:52:42 AM, Baghirzade, Nara) Ran feeding pump on various rates and volumes; feed completed successfully, no errors or issues. Physical inspection passed.
1062322 | 195574 | CM | CMPLT | CMPLT | 2023-06-28 08:28 | 2023-06-27 15:00 | Bracket installation for patient helper | (TC 06/27/2023 03:00:00 PM, Baghirzade, Nara) Installed bracket on bed for patient helper. (TC 06/27/2023 03:00:00 PM, Rhuland, Courtney) Installed bracket on bed for patient helper.
1061554 | 195575 | CM | CMPLT | CMPLT | 2023-06-28 08:40 | 2023-07-05 09:29 | Neptune Total Overhaul by Stryker | (TC 06/28/2023 08:40:08 AM, Reid, Emmerson ) Styker came in yesterday to replace the contaminated components. Service report will be added when they send. (TC 07/05/2023 09:28:09 AM, Reid, Emmerson ) Stryker came in and replaced many components to the neptune and it is working now. See attached service report.
1058606 | 195576 | CM | CMPLT | CMPLT | 2023-06-28 08:43 | 2023-07-05 09:25 | Neptune Total Overhaul by Stryker | (TC 06/28/2023 08:42:58 AM, Reid, Emmerson ) Styker came in yesterday to replace the contaminated components. Service report will be added when they send. (TC 07/05/2023 09:23:51 AM, Reid, Emmerson ) Stryker came in and replaced many components to the neptune and it is working now. See attached service report.
1067320 | 195577 | ACC | CMPLT | CMPLT | 2023-06-28 09:08 | 2023-06-28 14:44 | Incoming inspection | (TC 06/28/2023 01:31:10 PM, Canjimavadivel, Kamallesh ) Added to database and assembled scale. Verified functions and set up defaults. (TC 06/28/2023 02:43:23 PM, Canjimavadivel, Kamallesh ) Brought scale up to the floor and installed in corridor
66432 | 195578 | CM | CMPLT | CMPLT | 2023-06-28 09:17 | 2023-06-28 11:24 | Not turning on | (TC 06/28/2023 11:22:48 AM, Scero, Tomlyn) Replaced battery adaptor and tested device using a patient simulator. Device is functional. Returned to CCU
66431 | 195579 | CM | CMPLT | CMPLT | 2023-06-28 09:22 | 2023-08-28 11:06 | Not turning on | (TC 06/28/2023 09:30:19 AM, Scero, Tomlyn) Replaced the battery adaptor and device is turning on. While testing I noticed that after a minute or so of being conected to the network, the telemetry pack would power cycle on its own and then reconnect. And continue to do this. (TC 06/28/2023 11:21:22 AM, Scero, Tomlyn) removed tele 1 and replaced with tele 3 (TC 08/28/2023 11:04:31 AM, Scero, Tomlyn) Telepack tested again and it's no longer restarting. Once connected to the network I will test again.
1060030 | 195580 | CM | CMPLT | CMPLT | 2023-06-28 09:29 | 2023-06-28 09:32 | Heater does not heat properly | (TC 06/28/2023 09:28:50 AM, Baghirzade, Nara) Heater was cold after the system has been on for an hour. Removed the heater and measured resistance, within range. Good continuity between with the heater wire input and output. There was a short in the heater wire connector on the bodybox side; this was resolved. Heater started to heat slowly, will measure the temperature after a couple of patient tests to confirm that it is staying warm. (TC 06/28/2023 09:31:32 AM, Cortez, John) Heater was cold after the system has been on for an hour. Removed the heater and measured resistance, within range. Good continuity between with the heater wire input and output. There was a short in the heater wire connector on the bodybox side; this was resolved. Heater started to heat slowly, will measure the temperature after a couple of patient tests to confirm that it is staying warm.
1051527 | 195581 | CM | CMPLT | CMPLT | 2023-06-28 10:59 | 2023-06-28 14:06 | Calibration verification | (TC 06/28/2023 01:54:12 PM, Rhuland, Courtney) Verification complete using SimCube failed, but accuracy test with Welch Allyn software passed Changed tester to Rigel and verification passed 120/80 = 115/81 140/100 = 134/95 180/140 = 170/145 Verification within testing limits
1051526 | 195582 | CM | CMPLT | CMPLT | 2023-06-28 11:05 | 2023-06-28 11:38 | Calibration verification | (TC 06/28/2023 11:36:28 AM, Rhuland, Courtney) Temperature verified with tester 36.0 C = 36.0 C 38.5 C = 38.5 C 41.0 C = 41.0 C Functional check passed
1062174 | 195583 | CM | CMPLT | CMPLT | 2023-06-28 11:18 | 2023-06-28 11:21 | No functionality - no buttons and raise/lift does not work | (TC 06/28/2023 11:18:03 AM, Ferreira, Qwynn) Tested the bed, it was not working or lighting up Reset the bed and was able to get it to power on but no functionality Presented error code regarding communication checked the pins on the foot board and noticed 1 or 2 were crossed/corroded Cleaned pins, straightened the bent pins and cleaned the female pins as well Reseated the footboard and reset the bed. Buttons were working Informed nurse to enusre matress and sheets are not in the way
63347 | 195584 | CM | CMPLT | CMPLT | 2023-06-28 13:05 | 2023-06-28 13:07 | V1-V3 wires damaged | (TC 06/28/2023 01:05:36 PM, Baghirzade, Nara) Replaced V1-V3 ECG lead wires. Checked ECG with simulator; functional tests passed. ECG in good physical condition. (TC 06/28/2023 01:05:36 PM, Baghirzade, Nara) Replaced V1-V3 ECG lead wires. Checked ECG with simulator; functional tests passed. ECG in good physical condition.
| 195585 | CM | CMPLT | CMPLT | 2023-06-28 13:06 | 2023-06-28 13:06 | MSHTEST1S and IBE set up for test | (TC 06/28/2023 01:06:36 PM, Samanian, Kamran) done, checked and passed
1054310 | 195586 | CM | CMPLT | CMPLT | 2023-06-28 13:18 | 2023-06-28 13:29 | Low battery error | (TC 06/28/2023 01:28:25 PM, Canjimavadivel, Kamallesh ) Replaced batteries, tested ok
1067324 | 195587 | CM | CMPLT | CMPLT | 2023-06-28 13:22 | 2023-06-28 13:22 | Damaged probe | (TC 06/28/2023 01:22:23 PM, Baghirzade, Nara) Replaced probe. Functional tests passed. Thermometer was not in TMS. Created asset on TMS and new asset sticker for device.
1051528 | 195588 | CM | CMPLT | CMPLT | 2023-06-28 13:24 | 2023-06-28 13:24 | Damaged probe, functional check | (TC 06/28/2023 01:24:02 PM, Baghirzade, Nara) Replaced probe. Functional tests passed.
1051639 | 195589 | FC | CMPLT | CMPLT | 2023-06-28 13:25 | 2023-06-28 13:26 | Functional check | (TC 06/28/2023 01:24:53 PM, Baghirzade, Nara) Doppler passed functional tests. Physical condition is ok.
1067037 | 195590 | CM | CMPLT | CMPLT | 2023-06-28 13:32 | 2023-06-28 13:48 | Incoming inspection | (TC 05/08/2023 01:34:19 PM, Cortez, John) Incoming inspection for cart passed
1067038 | 195591 | ACC | CMPLT | CMPLT | 2023-06-28 13:45 | 2023-06-28 13:47 | Incoming inspection | (TC 05/08/2023 01:45:55 PM, Cortez, John) Incoming inspection of cart passed
1061210 | 195592 | CM | CMPLT | CMPLT | 2023-06-28 14:04 | 2023-06-28 14:07 | Cover on Boom Came Off | (TC 06/28/2023 02:05:01 PM, Reid, Emmerson ) The bottom cover on the boom fell off somehow. It looked like one of the snaps that connect the two pieces of the cover had broken but it still held together. I placed the cover back into place, it seemed fairly secure. If it comes off again it will need to be replaced, but it is working now.
| 195593 | CM | CMPLT | CMPLT | 2023-06-28 14:30 | 2023-06-27 14:30 | meeting for EXPAIR discussion | (TC 06/27/2023 02:30:29 PM, Samanian, Kamran) done!
1051031 | 195594 | CM | CMPLT | CMPLT | 2023-06-28 15:34 | 2023-07-31 10:30 | NICU purchased new EchoPAC plug in software for their laptops | (TC 06/28/2023 03:34:26 PM, Fan, Tak) Contacted IT for admin rights. Tony from GE came in for the install and added two sites for the new software licenses to the existing dongles. Installed on one PC but did not attempt other PC since it was a low spec windows 11 PC. Dr Jain wants to participate in active echo research with GE, try out the new temporary demo options (contact Jean Nash), possibly purchase a Viewpoint license to Q/R between the one PC and viewpoint, and potentially upgrade the Vivid e95 to v204 or higher.
1056430 | 195595 | CM | CMPLT | CMPLT | 2023-06-28 15:56 | 2023-06-28 15:57 | "O2 sat monitor not working" | (TC 06/27/2023 03:56:10 PM, Cortez, John) Tested all parameters, no issue found
1055902 | 195596 | CM | CMPLT | NFF | 2023-06-28 16:06 | 2023-06-28 16:07 | "not turning on when plugged in" | (TC 06/27/2023 04:06:08 PM, Cortez, John) Powering up normally, and all parameters tested ok
| 195597 | CM | CMPLT | CMPLT | 2023-06-28 17:32 | 2023-06-28 17:32 | Site coverage
1067326 | 195598 | FC | CMPLT | CMPLT | 2023-06-29 07:29 | 2023-06-28 13:00 | Functional check and asset tagging | (TC 06/28/2023 01:00:00 PM, Baghirzade, Nara) Functional tests passed. Good physical condition. Doppler was not in TMS; added to TMS and asset tagged.
| 195599 | CM | CMPLT | CMPLT | 2023-06-29 07:31 | 2023-06-28 12:30 | Leaking, not properly inflating cuff | (TC 06/28/2023 07:35:53 AM, Baghirzade, Nara) Manual sphyg. was not properly inflating cuff, also leaking from one of the tubings. Connector between hose and cuff was also broken. Replaced tubing and connectors. Took multiple measurements; no leaking and is fully functional.
68279 | 195600 | CM | CMPLT | CMPLT | 2023-06-29 07:38 | 2023-06-28 15:00 | Remote troubleshooting with Natus | (TC 06/28/2023 03:00:00 PM, Baghirzade, Nara) 1. During sensory acquisition, the table that displays acquired data during measurements was blown out of proportion and table was not properly visible for the tech. Natus remoted in and resolved the issue by changing the size of the table and font. This change was applied to all tests/test reports. 2. When inputting the patient ID/MRN to view past visits/reports, the reports are displayed on the software but the data takes too long to populate. This issue was not resolved and Natus tech (Tina) will confirm with their server/IT team about this. 3. During actual acquistion and recording, the patient full name, age, and gender does not show up at the top of the screen. This is important to have to make sure the right patient is being recorded. There are too many icons at the top of the screen which limits the amount of characters allowed within those fields. Tina will confirm with clinical apps team to make these changes. Vinod confirmed that the icons at the top of the screen are not needed and can be removed as he is able to access these icons through his control panel. Thursday, July 6th will follow up with Tina. (TC 06/29/2023 01:18:13 PM, Baghirzade, Nara) Amplifier was not being detected by the EMG system; after multiple reboots of the computer, the amplifier was finally detected. Checked cables, all good. Vinod had reseated all the cables and there were no changes to the machine since the day before so the root cause of the amplifier detection is unknown. Connected with Christina from Natus for remote troubleshooting; we had noticed that the EMG server had only 5GB space left. Sent a request to Welsey for drive space increase; increase was granted by 20GB. This did not resolve the issue of slow data population in old reports; issue is still persistent. Also noted that there was a red "x" on the shared drive however the "x" disappeared after clicking on the drive. During patient test, another error occurred "not able to sync local data with server data, continuing will cause patient data loss"; relayed this information to Natus however root cause is unknown. EMG had troubles communicating with the server, but all patient data and reports are on the server and accessible from both EMG stations. Christina passed my contact to a technical service rep from Natus who is more specialized in the hardware and he will contact us for more troubleshooting of the systems. (TC 08/01/2023 08:31:50 AM, Baghirzade, Nara) With the new server upgrade, the issues encountered during acquisition as well as format of the tests should be resolved. Upgrade to occur August 15th
1057607 | 195601 | CM | CMPLT | CMPLT | 2023-06-29 08:37 | 2023-06-29 08:38 | Alarming on 16N view stations | (TC 06/29/2023 08:37:04 AM, Baghirzade, Nara) Masimo was alarming on both viewstations on 16North even though this Root is not configured for PSN and is not on the domain. Rebooted the Root, all functional tests passed for all parameters. (TC 06/29/2023 08:37:04 AM, Baghirzade, Nara) Masimo was alarming on both viewstations on 16North even though this Root is not configured for PSN and is not on the domain. Rebooted the Root, all functional tests passed for all parameters.
1059999 | 195602 | CM | CMPLT | CMPLT | 2023-06-29 08:39 | 2023-06-29 08:42 | Alarming that a Root is down | (TC 06/29/2023 08:41:04 AM, Baghirzade, Nara) A serial number for a Root was shown to be alarming on the view stations. Located the Root in mother and baby and rebooted it which caused the error to be resolved. Rebooted both view stations on 16N and ensured that all errors/alarms were resolved.
| 195603 | CM | CMPLT | CMPLT | 2023-06-29 08:43 | 2023-06-28 11:00 | Meeting with Mara and Nelia | (TC 06/28/2023 11:00:00 AM, Baghirzade, Nara) Meeting with Mara and Nelia regarding: 1. state of the adult scales and searching for an old part; if not found, i let them know that they would need to buy a new scale. 2. state of the oto-ophthalmoscopes in all rooms. 3. retiring of old welch-allyns 4. replacing BPtrus with 2 Masimos 5. list of equipment they wanted me to check
69650 | 195604 | CM | CMPLT | CMPLT | 2023-06-29 08:56 | 2023-06-29 09:05 | Channel error, cannot program | (TC 06/26/2023 09:04:08 AM, Rendon Lopez , Elizabeth) Downloaded error logs, no errors, run in user mode, no alarms. Check in tests-all passed
67393 | 195605 | CM | CANCL | CANCL | 2023-06-29 09:05 |  | Scale not working | suplicate by mistake
1053434 | 195606 | CM | CMPLT | CMPLT | 2023-06-29 09:23 | 2023-06-29 09:26 | Doesn't pick up FHR | (TC 06/28/2023 09:23:21 AM, Rendon Lopez , Elizabeth) previously repaired, not needed, retired
1060134 | 195607 | CM | CMPLT | NFF | 2023-06-29 09:36 | 2023-06-29 09:37 | Doesn't pick up FHR | (TC 06/29/2023 09:36:24 AM, Rendon Lopez , Elizabeth) Tested, NFF
1059659 | 195608 | CM | CMPLT | CMPLT | 2023-06-29 09:43 | 2023-07-19 09:03 | Doesn't pick up FHR | (TC 06/28/2023 09:43:26 AM, Rendon Lopez , Elizabeth) equipment malfunction, sending out for repair (TC 07/17/2023 09:02:34 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1055561 | 195609 | CM | CMPLT | NFF | 2023-06-29 09:44 | 2023-06-29 09:45 | Doesn't work | (TC 06/28/2023 09:44:23 AM, Rendon Lopez , Elizabeth) Tested, NFF
1067322 | 195610 | CM | CMPLT | CMPLT | 2023-06-29 09:45 | 2023-07-19 08:59 | Not working | (TC 06/28/2023 09:45:24 AM, Rendon Lopez , Elizabeth) Added to TMS No reading FHR, sending out for repair (TC 07/17/2023 08:58:22 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1055669 | 195611 | CM | CMPLT | CMPLT | 2023-06-29 09:47 | 2023-07-19 09:02 | Signal cuts out | (TC 06/28/2023 09:47:10 AM, Rendon Lopez , Elizabeth) Confirmed issue, sending out under warranty (TC 07/17/2023 09:01:39 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1067323 | 195612 | CM | CMPLT | CMPLT | 2023-06-29 09:48 | 2023-07-19 09:01 | not working | (TC 06/28/2023 09:48:38 AM, Rendon Lopez , Elizabeth) Added to TMS Weak signal, physical damage. Send out for repair (TC 07/17/2023 09:00:30 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1057550 | 195613 | CM | CMPLT | CMPLT | 2023-06-29 09:49 | 2023-06-29 09:55 | SPO2 not working | (TC 06/29/2023 09:55:13 AM, Scero, Tomlyn) SPO2 tested with a brand new Probe on simulator. All measurements complete and passed with no issues. Also tested SPO2 on myself, all values showing as they should. No issues were found with the SPO2 port. Possible issue with whatever finger probe or truck cable was being used in the OR. Returned device to the OR.
1057545 | 195614 | CM | CMPLT | CMPLT | 2023-06-29 09:50 | 2023-06-29 09:55 | SPO2 not working | (TC 06/29/2023 09:52:28 AM, Scero, Tomlyn) SPO2 tested with a brand new Probe on simulator. All measurements complete and passed with no issues. Also tested SPO2 on myself, all values showing as they should. No issues were found with the SPO2 port. Possible issue with whatever finger probe or truck cable was being used in the OR. Returned device to the OR.
1047071 | 195615 | CM | CMPLT | CMPLT | 2023-06-29 09:50 | 2023-06-29 09:51 | No picking up | (TC 06/28/2023 09:50:26 AM, Rendon Lopez , Elizabeth) Confirmed issue, previously repaired, not needed, retire
1048003 | 195616 | CM | CMPLT | CMPLT | 2023-06-29 09:52 | 2023-07-19 09:06 | no picking up contractions | (TC 06/28/2023 09:52:41 AM, Rendon Lopez , Elizabeth) Readings are off, weak signal. Send out under warranty (TC 07/18/2023 09:05:35 AM, Rendon Lopez , Elizabeth) Returned from repair, no issues were found. No repairs incoming inspection-OK
1054196 | 195617 | CM | CMPLT | CMPLT | 2023-06-29 09:54 | 2023-07-19 09:08 | Physical damage | (TC 06/28/2023 09:54:19 AM, Rendon Lopez , Elizabeth) Send out for repair (TC 07/18/2023 09:06:52 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
| 195618 | CM | CMPLT | CMPLT | 2023-06-29 10:11 | 2023-06-28 10:11 | Masimo Meeting with NAZ | (TC 06/28/2023 10:11:17 AM, Samanian, Kamran) done!
66763 | 195619 | CM | CMPLT | CMPLT | 2023-06-29 10:25 | 2023-09-21 08:46 | error code 41171 | (TC 06/29/2023 10:24:40 AM, Cortez, John) error code 41171, smiths medical recommends to send pump for repair (TC 06/30/2023 02:51:29 PM, Cortez, John) Requested RMA from smiths medical (TC 07/19/2023 10:34:59 AM, Cortez, John) Received RMA (315417) shipped via KJV (TC 08/29/2023 04:02:02 PM, Cortez, John) ICUMED asking for PO again, emailed it (TC 09/20/2023 08:44:53 AM, Cortez, John) Pump returned from vendor PM was completed by vendor Added batteries and updated drug library
| 195620 | CM | CMPLT | CMPLT | 2023-06-29 10:35 | 2023-06-29 10:36 | Dropped off repaired equipment | (TC 06/29/2023 10:35:35 AM, Baghirzade, Nara) Dropped off repaired equipment with Mara. Installed an otoscope/ophthalmscope wall unit and did functional checks.
| 195621 | CM | CMPLT | CMPLT | 2023-06-29 12:38 | 2023-07-10 09:54 | Order desk for biomed | (TC 06/29/2023 12:40:30 PM, Chukri, Samir) order table bench (TC 07/10/2023 09:54:34 AM, Chukri, Samir) Received .
| 195622 | CM | CMPLT | CMPLT | 2023-06-29 13:09 | 2023-06-29 13:12 | parts charge back, bulk spare parts | (TC 06/29/2023 01:09:37 PM, Scero, Tomlyn) philips x3 bezel parts for OR. one time charge for spare parts.
1050074 | 195623 | CM | CMPLT | CMPLT | 2023-06-29 13:33 | 2023-06-29 13:52 | Not turning on | (TC 06/29/2023 01:34:52 PM, Canjimavadivel, Kamallesh ) Opened up case and discovered power supply wires were not connected correctly from previous repair. Reconnected all wires to correct connections. Power supply working, device charging
1062175 | 195624 | CM | CMPLT | CMPLT | 2023-06-29 13:42 | 2023-08-01 09:55 | Head portion very loud when raising/lowering | (TC 06/29/2023 01:41:47 PM, Baghirzade, Nara) Head portion of the bed very loud when raising/lowering. Created a service ticket with Stryker. (TC 07/06/2023 11:18:47 AM, Baghirzade, Nara) Bed was still occupied by patient when Mark from Stryker had come in last week for service. Bed has now been moved to 18S hallway. Placing service request with Stryker. (TC 08/01/2023 09:53:57 AM, Baghirzade, Nara) Bed was serviced by Mark. Attached service report.
1061424 | 195625 | CM | CMPLT | NFF | 2023-06-29 13:43 | 2023-06-29 13:44 | Loud noise when moving bed | (TC 06/29/2023 01:43:13 PM, Baghirzade, Nara) Nurses reported that the bed is very noisy when lowering/raising however they werent sure which motion would cause the noise. Checked all motions, none of them created noise. Checked all motion controls, all are functional. No fault found.
| 195626 | CM | CMPLT | CMPLT | 2023-06-29 14:45 | 2023-06-29 14:45 | OneRecall Meeting | (TC 06/29/2023 02:45:12 PM, Baghirzade, Nara) OneRecall Meeting
64871 | 195627 | CM | CMPLT | CMPLT | 2023-06-29 14:51 | 2023-07-20 14:42 | APC UPS "change battery" and voice analyzing program was not working | (TC 06/29/2023 02:50:43 PM, Baghirzade, Nara) Met with Elena, she mentioned that when they had a patient for voice analyzing, the APC UPS had an amber LED alerting "change battery"; will look into ordering a new UPS battery. Elena asked to confirm the price with them; possibly would not want to replace the battery if cost is too high. (TC 07/14/2023 02:13:20 PM, Baghirzade, Nara) SC1003003475 (TC 07/20/2023 02:41:02 PM, Baghirzade, Nara) PO 7202249312
1059795 | 195628 | CM | CMPLT | CMPLT | 2023-06-29 20:58 | 2023-06-29 20:57 | Touch screen not working. | Parts on order.
502044 | 195629 | CM | CMPLT | CMPLT | 2023-06-29 21:35 | 2023-06-29 21:39 | Reoved Head to tighten Brakes and fix collimator lock mechanism. | (TC 06/29/2023 09:34:45 PM, Paz, Ramon (Tino)) Brakes still lose, disengaging. Removed Head to access brake handle and tightened and adjusted brake handle assembly. Played around with the position and also the tension of the brake wire. Also adjusted brake pads. Tested a few times to make sure brake handle does not release prematurely. Also replaced collimator locking mechanism to fix lose collimator. Adjusted accordingly.
| 195630 | CM | CMPLT | CMPLT | 2023-06-30 08:02 | 2023-06-29 08:02 | Olympus Nessus scanning in OR12 | (TC 06/29/2023 08:02:05 AM, Samanian, Kamran) done!
| 195631 | CM | CMPLT | CMPLT | 2023-06-30 08:03 | 2023-06-29 08:08 | One Recall session meeting | (TC 06/29/2023 08:03:02 AM, Samanian, Kamran) done!
| 195632 | CM | CMPLT | CMPLT | 2023-06-30 08:11 | 2023-06-29 08:12 | Work on EMG to request space on shared drive in server - reading DB slow | (TC 06/30/2023 08:11:17 AM, Samanian, Kamran) verified the problem. requested to increase the space in the shared drive and applied but it did not resolve the issue. it seems either security or incompatibility is the root cause (TC 06/29/2023 08:11:17 AM, Samanian, Kamran) verified the problem. requested to increase the space in the shared drive and applied but it did not resolve the issue. it seems either security or incompatibility is the root cause
1050074 | 195633 | CM | CMPLT | CMPLT | 2023-06-30 10:04 | 2023-08-31 09:03 | Error with communication module | (TC 06/30/2023 11:07:20 AM, Canjimavadivel, Kamallesh ) Reseated all usb cables, error still present. Tried swapping communication module with spare parts. The board versions are not the same. Contacted vendor for quote. Items are non stock special order and will take a few weeks to deliver. (TC 07/26/2023 10:45:52 AM, Canjimavadivel, Kamallesh ) Swapped damaged cable that powers communication module with spare (TC 07/28/2023 02:36:47 PM, Canjimavadivel, Kamallesh ) Swapped out comm board and internal usb cables. Same alarm displaying. Spoke to tech support and they suggested connecting to WA support tool to reinitialize board. (TC 08/04/2023 09:29:46 AM, Canjimavadivel, Kamallesh ) Tried to use Gold license with support tool to recognize board. Was not able to figure out. Called tech support for further support. Another option is to replace main board, if not can send it in for further investigation. Awaiting recommendation from unit whether to proceed with repair or not. (TC 08/31/2023 09:02:40 AM, Canjimavadivel, Kamallesh ) Decision made to retire device and not pursue with repair.
1057888 | 195634 | CM | CMPLT | CMPLT | 2023-06-30 10:42 | 2023-06-30 10:45 | Mistral Air not turning on | (TC 06/30/2023 10:43:35 AM, Reid, Emmerson ) Went to PACU to pick up device. The back cover holding in the power cord was very loose. I opened it up and pushed the power cord in all the way then tightened the cover. Turning on now,
| 195635 | PA | CMPLT | CMPLT | 2023-06-30 10:46 | 2023-07-27 14:31 | Parts for stock | (TC 06/30/2023 11:06:35 AM, Canjimavadivel, Kamallesh ) Placed order for parts (TC 07/27/2023 02:25:26 PM, Canjimavadivel, Kamallesh ) Received all parts | Charge back as used
1061202 | 195636 | CM | CMPLT | CMPLT | 2023-06-30 12:51 | 2023-06-30 12:52 | Video tower not connecting | (TC 06/30/2023 12:50:55 PM, Reid, Emmerson ) Called me up to OR 1 to check the video tower. When I arrived I was told it was working...
1067334 | 195637 | ACC | CMPLT | CMPLT | 2023-06-30 13:32 | 2023-06-30 14:06 | Incoming inspection | (TC 06/30/2023 02:03:07 PM, Canjimavadivel, Kamallesh ) Added to database. Brought scale up to unit and installed. All functions ok
1049904 | 195638 | CM | CMPLT | CMPLT | 2023-06-30 13:33 | 2023-06-30 13:36 | Battery chargers not working | (TC 06/30/2023 01:34:14 PM, Fan, Tak) Replaced charging docks. (TC 06/30/2023 10:16:06 AM, Whittington, Scott) MSH R64 CHARGING BIN BOTH DAMAGED CONNECTORS (TC 07/04/2023 10:16:44 AM, Whittington, Scott) MSH R64 CHARGING BIN BOTH DAMAGED CONNECTORS ORDER PN# SPAA3097 CHARGER ASSLY DRX-1 BATT | Damaged charging pins
26125 | 195639 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26127 | 195640 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26130 | 195641 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26136 | 195642 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26154 | 195643 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26794 | 195644 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
50057 | 195645 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Transport
50997 | 195646 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
51571 | 195647 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52058 | 195648 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52178 | 195649 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-16 08:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/16/2023 08:01:46 AM, Baghirzade, Nara) PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C across all temperature probes. Functional tests passed. Electrical safety results - Leakage current = 63.1 uA Power cord resistance = 70 mohms
52466 | 195650 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
53004 | 195651 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
60089 | 195652 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
62623 | 195653 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-10-05 12:50 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 10/05/2023 12:48:49 PM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. RH failed 2 calibrations. Replaced the RH sensor. Calibrated the new RH sensor. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36, measured 36. Functional tests passed. Electrical safety results - Leakage current = 80.0 uA Power cord resistance = 170 mohms
62911 | 195654 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2024-10-22 14:36 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 10/22/2024 02:36:09 PM, Rhuland, Courtney) PM completed by Spectramedx 21% - 20.9% O2 40% - 39.2% O2 60% - 59.5% O2 80% - 79% O2 100% - 99.7% O2
63024 | 195655 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-19 14:55 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/19/2023 02:49:35 PM, Rhuland, Courtney) Enclosure case replaced PM complete Irradiance (large pad) - 49.6 uW/cm^2/nm Hours - 6152.3 Leakage current - 11.7uA Resistance - 126mOhms
63727 | 195656 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Testers
63875 | 195657 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-25 10:43 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/17/2023 11:51:53 AM, Ferreira, Qwynn) Located device in NICU Noticed the door was broken and the hood needed repair. Waiting on hood to come (TC 07/18/2023 10:50:58 AM, Ferreira, Qwynn) Replaced hood replaced two grommets replaced 2 port hinges Installed recall latches PM Passed Calibrated line voltage Calibrated main voltage Calibrated system voltage Calibrated scale Functional Test - OK Electrical Safety - OK Waiting on battery (TC 07/25/2023 10:42:28 AM, Ferreira, Qwynn) Replaced battery Tested again just to enusre everything is working Returned to unit
64211 | 195658 | PM | CNP | ENL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Microstream CO2 extension
65136 | 195659 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65170 | 195660 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2024-05-21 09:50 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/21/2024 09:49:46 AM, Reid, Emmerson ) PM Complete, see procedure for details. (TC 05/21/2024 09:49:46 AM, Reid, Emmerson ) PM Complete, see procedure for details.
65416 | 195661 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
65878 | 195662 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-12-05 13:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 12/05/2023 01:36:06 PM, Campbell, Connor) pm complete/passed, refer to ASM
65888 | 195663 | PM | CMPLT | PMFAIL | 2023-07-01 03:16 | 2023-09-28 09:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/28/2023 09:13:39 AM, Campbell, Connor) Failed plunger position accuracy calibration of plunger position failed fatal error in syringe motor, needs new syringe assembly refer to WO: 198087
65893 | 195664 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2024-08-09 11:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65911 | 195665 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65915 | 195666 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2024-03-28 13:05 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 01:04:43 PM, Rhuland, Courtney) PM completed by BD
65925 | 195667 | PM | CNP | ENL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65926 | 195668 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2024-04-05 09:19 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/05/2024 09:17:19 AM, Rhuland, Courtney) Replaced inside shield During calibration, disc sensor flag was stuck Open device and reseated and cleaned area Calibrated and checked in device All tests passed Refer to ASM for details
65945 | 195669 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-11 07:59 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/06/2023 12:47:51 PM, Rhuland, Courtney) Rear case replaced Left IUI connector replaced (TC 07/11/2023 07:58:02 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
65968 | 195670 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-14 13:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/14/2023 01:40:37 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
66014 | 195671 | PM | CNP | ENL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66021 | 195672 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-25 14:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 10:27:15 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 180 mohms Leakage current: 84.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/25/2023 02:08:46 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4063 mamphr PM completed.
66028 | 195673 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-21 07:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/18/2023 01:47:01 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 55 mohms Leakage current: 68.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/21/2023 07:21:31 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4049 mamphr PM completed.
66047 | 195674 | PM | CANCL | CANCL | 2023-07-01 03:16 | 2024-03-14 14:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66051 | 195675 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-09-20 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 08:50:39 AM, Rendon Lopez , Elizabeth) Battery conditioning done-Passed. PM completed, all tests passed.
66466 | 195676 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-06-30 07:39 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. Replaced PS500 batteries. No GS500. Functional. SW.2.51 Leakage: 87 uA and 85 ohm.
66491 | 195677 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-17 14:42 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 08/16/2023 02:41:15 PM, Rendon Lopez , Elizabeth) Performance tested-OK
66492 | 195678 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-17 14:56 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 08/16/2023 02:55:43 PM, Rendon Lopez , Elizabeth) Performance verification-OK
66510 | 195679 | PM | CMPLT | UNL | 2023-07-01 03:16 | 2023-07-19 11:05 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | Unable to locate for more than 3 years. Retired from TMS
66534 | 195680 | PM | CMPLT | UNL | 2023-07-01 03:16 | 2023-07-19 11:07 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | Unable to locate for more than 3 years. Retire from service
66552 | 195681 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-14 08:51 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/08/2023 08:50:09 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed
66562 | 195682 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-14 08:53 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/08/2023 08:51:47 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed
66708 | 195683 | PM | CANCL | DUP | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM done every two years, reschedule 2024
66722 | 195684 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66730 | 195685 | PM | CANCL | DUP | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM done every two years, reschedule 2024
66762 | 195686 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-26 15:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/26/2023 03:05:25 PM, Cortez, John) Completed PM, see procedure for results
66878 | 195687 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2024-05-31 14:51 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 05/31/2024 02:50:43 PM, DeGuzman, Leila) PM Completed. Go to procedures to see details.
67759 | 195688 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67890 | 195689 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-27 13:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/26/2023 01:30:03 PM, Rendon Lopez , Elizabeth) Upgraded software and library. PM completed, all tests passed.
67933 | 195690 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-17 16:25 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/17/2023 12:18:45 PM, Rhuland, Courtney) Recall parts replaced Broken porthole door replaced Tightened all porthole hinges Greased bed tilt mechanism Calibrated LV and main voltage Calibrated scale and verified scale (TC 07/17/2023 04:24:18 PM, Rhuland, Courtney) Calibrated RH using calibration vile Calibration passed Funcitonal check - OK Electrical safety - OK
67945 | 195691 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-11-28 09:50 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/28/2023 09:32:01 AM, Rhuland, Courtney) Calibrated LV and main voltage Calibrated and verified scale Replaced 9V rechargeable battery Replaced cracked inside door latch Functional check - 35C & 75% humidity Read 35C & 74% Electrical safety - 83.4uA, 175mOhms
67957 | 195692 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-18 12:06 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/18/2023 12:05:06 PM, Baghirzade, Nara) During PM, RH sensor was not reading the humidity correctly and was not increasing past 75%. Replaced RH sensor. Calibrated RH sensor. PM completed. 9V battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 77%. Functional tests passed. Electrical safety results - Leakage current = 36.5uA Power cord resistance = 200 mohms
68369 | 195693 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68373 | 195694 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Anesthesia Units | 6 year PM just completed July 2023. Rescheduled next PM for July 2024
68377 | 195695 | PM | ACTIV | ISSUE | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68542 | 195696 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-10-25 10:44 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 10/16/2023 10:41:23 AM, Rendon Lopez , Elizabeth) Removed from OR for PM (TC 10/17/2023 10:41:23 AM, Rendon Lopez , Elizabeth) 1 Year PM parts replaced. PM tasks started (TC 10/20/2023 10:41:23 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Returned to L&D OR as spare
68543 | 195697 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-10-10 10:56 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 09/29/2023 04:41:44 PM, Rendon Lopez , Elizabeth) Removed AGM from OR1 for PM. brought spare AGM from LVL7 as L&D spare (TC 10/03/2023 09:43:25 AM, Rendon Lopez , Elizabeth) 1 year PM started, replaced PM parts, started tests (TC 10/02/2023 09:43:25 AM, Rendon Lopez , Elizabeth) 1 year PM started, replaced PM parts, started tests (TC 10/05/2023 10:54:47 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. See attached results (TC 10/10/2023 10:54:47 AM, Rendon Lopez , Elizabeth) Returned to L&D OR, swapped in OR 4
68612 | 195698 | PM | CNP | ENL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68626 | 195699 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68637 | 195700 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2024-04-26 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68655 | 195701 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-09-26 12:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/26/2023 12:57:57 PM, Campbell, Connor) PM complete/passed, refer to ASM
68663 | 195702 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-10 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/10/2023 11:03:34 AM, Reid, Emmerson ) Replaced the display board. Downloaded and cleared log history. Ran ASM Calibration, all tests passed. Ran ASM PM, all tests passed. PM Complete.
68664 | 195703 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-09-14 14:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/14/2023 02:27:32 PM, Cortez, John) Cleaned, Downloaded error/event logs, LVP display dim, replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
68675 | 195704 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-08-18 13:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/18/2023 01:27:10 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 29 mohms Leakage current: 68.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68691 | 195705 | PM | CMPLT | CMPLT | 2023-07-01 03:16 | 2023-07-07 09:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/07/2023 09:28:41 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 118 mohms Leakage current: 76.8 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68708 | 195706 | PM | CANCL | CANCL | 2023-07-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68720 | 195707 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68728 | 195708 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-10-27 13:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/27/2023 01:18:42 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 38 mohms Leakage current: 75.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68739 | 195709 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-21 09:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/18/2024 11:13:44 AM, Reid, Emmerson ) Giving an error with the network com board 600.6835 Tried swapping comm boards, but did not work. Trying to troubleshoot. (TC 03/18/2024 02:38:21 PM, Reid, Emmerson ) Found that is was the COM board - RJ45 connector that was the problem. Reaching out to BD to order the part. (TC 03/18/2024 11:13:44 AM, Reid, Emmerson ) Giving an error with the network com board 600.6835 Tried swapping comm boards, but did not work. Trying to troubleshoot. (TC 03/19/2024 03:11:32 PM, Reid, Emmerson ) Switched the com board. Unit not turning on now. Will look at further (TC 03/21/2024 09:33:06 AM, Reid, Emmerson ) Replaced the keypad, comm board and power cord. Pushed network config and did ASM PM. PM Complete.
68785 | 195710 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-02 11:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68805 | 195711 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68815 | 195712 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-05 09:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/05/2023 09:36:04 AM, Rhuland, Courtney) Power cord replaced Battery charged and conditioned PM complete All tests passed Refer to ASM for details
68822 | 195713 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68837 | 195714 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-05 14:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:45:43 PM, Rhuland, Courtney) PM completed by BD
68842 | 195715 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68847 | 195716 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-18 13:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68854 | 195717 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-14 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:17:57 PM, Huang, Jessica) replaced the power cord and completed PM
68855 | 195718 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-11-15 08:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/15/2023 08:38:26 AM, Campbell, Connor) battery replaced with a spare that was alreay charged PM complete/passed, refer to ASM refer to procedure for ESA
68884 | 195719 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-28 11:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 09:41:51 AM, Cortez, John) Cleaned unit, Downloaded error/event logs, PM completed successfully, see ASM for results Conditioning battery overnight (TC 08/28/2023 11:37:53 AM, Cortez, John) 4030 mAhs
68902 | 195720 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-24 14:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68910 | 195721 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-09-07 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 09:05:14 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed
68933 | 195722 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68941 | 195723 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69003 | 195724 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69025 | 195725 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-16 11:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 11:54:51 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 30 mohms Leakage current: 63.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
69030 | 195726 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-16 11:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 11:53:07 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 46 mohms Leakage current: 70.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
69037 | 195727 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-17 15:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 03:09:38 PM, Rendon Lopez , Elizabeth) Calibrated rate accuracy X3. PM completed, all tests passed
69041 | 195728 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-10-25 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/25/2023 10:40:36 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69049 | 195729 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-09-25 13:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/25/2023 01:44:44 PM, Campbell, Connor) PM complete/passed, refer to ASM
69061 | 195730 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-10-25 11:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/25/2023 11:10:35 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69096 | 195731 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-19 21:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69101 | 195732 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-14 15:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69123 | 195733 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-12 11:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 11:11:23 AM, Cortez, John) Cleaned, Downloaded error/event logs, Bottle-side pressure sensor errors were logged, Pressure sensor replaced, Replaced LED cover Calibrated infusion rate, PM completed successfully, results logged in ASM
69144 | 195734 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-23 10:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69145 | 195735 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69149 | 195736 | PM | ACTIV | ISSUE | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/06/2023 03:52:57 PM, Cortez, John) Channel error Binary sensor test failed (door ajar) Door latch sear damaged, was replaced
69152 | 195737 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-23 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69163 | 195738 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-12 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 02:24:38 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69169 | 195739 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-25 16:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 04:21:57 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69175 | 195740 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-10-05 10:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 10:28:14 AM, Campbell, Connor) PM complete/passed, refer to ASM
69199 | 195741 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-27 16:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/27/2023 04:26:48 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69209 | 195742 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-12-06 10:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 12/06/2023 10:56:23 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69220 | 195743 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69256 | 195744 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-26 09:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69268 | 195745 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69274 | 195746 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-14 16:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69277 | 195747 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69285 | 195748 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-11-06 16:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/06/2023 04:35:19 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69286 | 195749 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-27 15:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/27/2023 03:12:41 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69323 | 195750 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-23 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 10:40:44 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69326 | 195751 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-03 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:40:36 PM, Rhuland, Courtney) PM completed by BD
69346 | 195752 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-21 10:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/21/2023 10:45:00 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69374 | 195753 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-09-01 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 11:20:25 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
69383 | 195754 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69386 | 195755 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-25 11:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69388 | 195756 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69417 | 195757 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-21 20:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69421 | 195758 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69439 | 195759 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-10 16:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/10/2023 04:40:15 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69474 | 195760 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-09-07 14:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 02:36:58 PM, Campbell, Connor) calibrated pressure and rate PM completed, refer to ASM supervised by nara
69481 | 195761 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-28 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 01:59:05 PM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
69490 | 195762 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-01 18:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 06:14:06 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69492 | 195763 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-21 18:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69497 | 195764 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-26 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69498 | 195765 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-16 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 10:32:53 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69504 | 195766 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-31 10:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/31/2023 10:50:29 AM, Cortez, John) Cleaned, Downloaded error/event logs, Bottle-side pressure sensor errors were logged, Pressure sensor replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
69505 | 195767 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69506 | 195768 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69535 | 195769 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69540 | 195770 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-09-20 08:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/20/2023 08:14:07 AM, Campbell, Connor) PM complete/passed, refer to ASM
69546 | 195771 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69570 | 195772 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69612 | 195773 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-10 17:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/10/2023 04:59:14 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69621 | 195774 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-31 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/31/2023 10:22:29 AM, Cortez, John) Cleaned, Downloaded error/event logs, patient-side pressure sensor errors were logged, Pressure sensor replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
69623 | 195775 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-21 21:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69629 | 195776 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69640 | 195777 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-09-25 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/25/2023 01:40:15 PM, Campbell, Connor) PM complete/passed, refer to ASM
69653 | 195778 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-22 12:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/22/2023 12:00:17 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69654 | 195779 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-14 09:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 09:06:27 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69663 | 195780 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-04 12:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 12:49:50 PM, Rhuland, Courtney) PM completed by BD
69673 | 195781 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
531003 | 195782 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: BLADDER SCANNER
65610 | 195783 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: Smoke Evacuation Systems, Surgical
65616 | 195784 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: Stimulators, Electrical, Brain, Convulsive Therapy
32540 | 195785 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-11 15:42 | PM - Monthly | PM completed Mechanical PM done SC-616089 a(-992.1,10.2) b(332.4,13.7) c(132.2,13.0) d(95.9,13.8) e(-105.2,12.6) f(0.1,13.9) -passed all phantom tests
32560 | 195786 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-01 13:48 | PM - Monthly | PM completed Mechanical PM done TOS phantom a(-992.1,14.1) b(334.8,19.2) c(134.1,17.7) d(100.1,17.1) e(-104.9,14.7) f(0.1,18.9) -passed all phantom tests
32580 | 195787 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-02 21:38 | PM - Monthly | PM completed Mechanical PM done sc-85221 TOS phantom a(-992.9,10.7) b(333.7,16.1) c(133.9,14.2) d(98.6,13.7) e(-106.8,13.0) f(0.7,13.1) -passed all phantom tests
500320 | 195788 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM - Semi-Annual
501440 | 195789 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-10 23:08 | PM - Semi-Annual | PM
501440 | 195790 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-11 23:07 | HARP - Semi-Annual | HARP
502044 | 195791 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-11-28 11:04 | PM - Semi-Annual | (TC 11/03/2023 11:01:08 AM, Green, Paul) PM was done, replaced dessicant, topped up water, greased lead screw, adjusted brake and performed a full PM.
1046440 | 195792 | PM | ACTIV | READY | 2023-07-01 03:17 |  | Annual PM
1047015 | 195793 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047040 | 195794 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-23 13:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 01:46:30 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1047044 | 195795 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-25 11:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/25/2023 11:51:07 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1047098 | 195796 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-14 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/14/2023 10:06:43 AM, Reid, Emmerson ) Replaced the battery. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
1047831 | 195797 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-19 10:50 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
1047832 | 195798 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-09-12 14:59 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
1047849 | 195799 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-17 14:05 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/17/2023 02:04:18 PM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
1048486 | 195800 | PM | ACTIV | ISSUE | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: Insufflators
1048526 | 195801 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-04-10 14:50 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:49:22 PM, Rhuland, Courtney) PM completed by BD
1048642 | 195802 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048645 | 195803 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048658 | 195804 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048668 | 195805 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048670 | 195806 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048681 | 195807 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048693 | 195808 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050160 | 195809 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1050247 | 195810 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-28 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/28/2023 11:16:38 AM, Rhuland, Courtney) Battery replaced and conditioned PM complete All tests passed Refer to ASM for details
1050257 | 195811 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051165 | 195812 | PM | ACTIV | ISSUE | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1051767 | 195813 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-11-03 10:35 | PM MAJOR - ANNUAL PM for: Electrosurgical Unit | (TC 11/03/2023 10:34:33 AM, Reid, Emmerson ) PM Complete, see attached report.
1054460 | 195814 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-12 10:04 | Lubrication cycle | (TC 07/12/2023 10:03:46 AM, Canjimavadivel, Kamallesh ) Lubricated O-rings
1054462 | 195815 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-06 08:50 | Lubrication cycle | (TC 07/06/2023 08:49:41 AM, Canjimavadivel, Kamallesh ) Lubricated O-rings
1054468 | 195816 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-12 07:34 | Lubrication cycle | (TC 07/12/2023 07:33:52 AM, Cortez, John) Lubricated o-rings (TC 07/11/2023 07:33:52 AM, Cortez, John) Lubricated o-rings
1054813 | 195817 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-11 11:32 | PM-MAJOR - Annual PM For: Ventilators | (TC 07/11/2023 11:30:45 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 22848 hrs Total vent time: 12974 hrs PM complete
1054814 | 195818 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-07-11 11:28 | PM-MAJOR - Annual PM For: Ventilators | (TC 07/11/2023 11:12:39 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 18102 hrs Total vent time: 10314 hrs PM complete
1054815 | 195819 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-08-03 11:01 | PM-MAJOR - Annual PM For: Ventilators | (TC 08/03/2023 10:58:58 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 18097 hrs Total vent time: 11314 hrs
1055112 | 195820 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055123 | 195821 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055162 | 195822 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055163 | 195823 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055172 | 195824 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2024-03-15 08:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1056125 | 195825 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | Schedule shifted, PM to be performed in OCtober
1057512 | 195826 | PM | ACTIV | ISSUE | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057764 | 195827 | PM | CANCL | CANCL | 2023-07-01 03:17 |  | PM MAJOR-Monitors, Physiologic, Central Station | Storage
1057836 | 195828 | PM | ACTIV | ISSUE | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057842 | 195829 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-12-07 14:56 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 12/07/2023 02:55:27 PM, Reid, Emmerson ) Replaced the display board on previous WO. PM complete, see attached PM Report for details.
1058270 | 195830 | PM | CMPLT | CMPLT | 2023-07-01 03:17 | 2023-09-07 08:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 08:56:45 AM, Rendon Lopez , Elizabeth) battery conditioning. (TC 09/06/2023 08:56:45 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1058274 | 195831 | PM | CNP | ENL | 2023-07-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1061105 | 195832 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-27 10:10 | PM-MAJOR- for 9S Bed Elecric
1061106 | 195833 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-27 11:44 | PM-MAJOR- for 9S Bed Elecric
1061107 | 195834 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-28 09:19 | PM-MAJOR- for 9S Bed Elecric
1061108 | 195835 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-28 10:03 | PM-MAJOR- for 9S Bed Elecric
1061109 | 195836 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-09-13 16:18 | PM-MAJOR- for 9S Bed Elecric
1061110 | 195837 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-27 10:37 | PM-MAJOR- for 9S Bed Elecric
1061111 | 195838 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-09-13 16:20 | PM-MAJOR- for 9S Bed Elecric
1061112 | 195839 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-28 09:46 | PM-MAJOR- for 9S Bed Elecric
1061113 | 195840 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-09-13 16:21 | PM-MAJOR- for 9S Bed Elecric
1061114 | 195841 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-27 09:31 | PM-MAJOR- for 9S Bed Elecric
1061115 | 195842 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-08-27 11:15 | PM-MAJOR- for 9S Bed Elecric
1061116 | 195843 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-09-13 16:23 | PM-MAJOR- for 9S Bed Elecric
1061117 | 195844 | PM | CANCL | CANCL | 2023-07-01 03:17 | 2024-09-13 16:24 | PM-MAJOR- for 9S Bed Elecric
1061118 | 195845 | PM | CANCL | CANCL | 2023-07-01 03:18 | 2024-08-27 10:27 | PM-MAJOR- for 9S Bed Elecric
1061119 | 195846 | PM | CANCL | CANCL | 2023-07-01 03:18 | 2024-08-27 10:47 | PM-MAJOR- for 9S Bed Elecric
1061120 | 195847 | PM | CANCL | CANCL | 2023-07-01 03:18 | 2024-09-13 16:26 | PM-MAJOR- for 9S Bed Elecric
1061121 | 195848 | PM | CANCL | CANCL | 2023-07-01 03:18 | 2024-08-27 10:17 | PM-MAJOR- for 9S Bed Elecric
1061154 | 195849 | PM | CMPLT | CMPLT | 2023-07-01 03:18 | 2023-12-28 08:49 | PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 12/28/2023 08:49:29 AM, Whittington, Scott) PM, SEMI-ANNUAL, ULTRASOUND GENERAL AGREEMENT NOT REACHED | PM, SEMI-ANNUAL, ULTRASOUND GENERAL AGREEMENT NOT REACHED
1061192 | 195850 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | Device to be retired and replaced by Stryker
1061364 | 195851 | PM | CMPLT | CMPLT | 2023-07-01 03:18 | 2023-08-25 16:18 | HARP - Semi Annual | HARP
1061364 | 195852 | PM | CMPLT | CMPLT | 2023-07-01 03:18 | 2023-08-24 16:19 | PM - Semi Annual | PM
1064641 | 195853 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064642 | 195854 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064643 | 195855 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064644 | 195856 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064645 | 195857 | PM | CNP | ENL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1064646 | 195858 | PM | CNP | ENL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1064647 | 195859 | PM | CMPLT | CMPLT | 2023-07-01 03:18 | 2024-04-15 08:53 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 08:51:44 AM, Huang, Jessica) replaced the rear case and completed the PM
1064648 | 195860 | PM | CNP | ENL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1064649 | 195861 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064650 | 195862 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064651 | 195863 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064652 | 195864 | PM | CANCL | CANCL | 2023-07-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066756 | 195865 | PM | CMPLT | CMPLT | 2023-07-01 03:18 | 2023-07-12 07:43 | Lubrication cycle | (TC 07/11/2023 07:41:44 AM, Cortez, John) Lubricated o-rings
1066757 | 195866 | PM | CMPLT | CMPLT | 2023-07-01 03:18 | 2023-07-06 15:47 | Lubrication cycle | (TC 07/06/2023 03:46:42 PM, Cortez, John) Passed by ICU and lubricated o-rings
| 195867 | PA | CMPLT | CMPLT | 2023-07-04 08:11 | 2023-07-21 14:14 | PM Kit 5000hr | (TC 07/04/2023 08:14:45 AM, Canjimavadivel, Kamallesh ) Placed order for PM kits (TC 07/21/2023 02:13:29 PM, Canjimavadivel, Kamallesh ) Received PM kits
1067336 | 195868 | CM | CMPLT | CMPLT | 2023-07-04 08:15 | 2023-06-30 02:00 | SafetyNet appliance link down | (TC 06/30/2023 02:00:00 PM, Baghirzade, Nara) Worked with Kamran to resolve this issue. The PSN central was not connecting to the server; was unable to ping the server or any other viewstations. After multiple reboots, central still down. Replaced the ethernet cable but this did not resolve the issue. Checked other viewstations and was able to ping the server and other viewstations, the issue is only affecting the 18S central. The IT closet was too messy to be able to see the switch light for that specific port. After many attempts at trying to connect to the server, I found out that in the morning there was a power outage that affected all computers and phones on 18S. Connected with Mirza from IT and it turns out that when the power came back on, the switch did not retain its configuration settings for the port and as a result, was not able to connect to that port, resulting in the central losing connection to the VLAN. Mirza added the configuration settings back and the central was successfully connecting to the server. Pinged the server and several other viewstations; central is fully functional.
1067336 | 195869 | CM | CMPLT | CMPLT | 2023-07-04 08:50 | 2023-07-04 09:31 | SafetyNet appliance link down | (TC 07/04/2023 08:50:14 AM, Baghirzade, Nara) Unknown when this issue started, no complaints over the weekend. Restarted the PSN, still no connection the server. Tried to ping the server and other viewstations, no connection. Pinged the central itself and that was working. Shut down the PSN, issue still persists. Re-seated the ethernet cable on both ends, still no connection to the server. Contacted Mirza to check the switch ports for the central. (TC 07/04/2023 09:28:08 AM, Baghirzade, Nara) The switch had a failure the night before. IT replaced the switch but the port was not properly configured, causing the central to lose connection with the server. After configuring the switch, the central connected to the server. I was able to ping the server and all segments on the display were showing up. (TC 07/04/2023 09:30:05 AM, Samanian, Kamran) The switch had a failure the night before. IT replaced the switch but the port was not properly configured, causing the central to lose connection with the server. After configuring the switch, the central connected to the server. I was able to ping the server and all segments on the display were showing up.
| 195870 | CM | CMPLT | CMPLT | 2023-07-04 09:18 | 2023-07-10 09:04 | Tomi On Call - June 30-July7 | (TC 07/02/2023 09:18:18 AM, Scero, Tomlyn) Call from ICU. X3 wouldn't working properly with SPO2. I advised the nurse to change the cable but it didn't help. I then asked her to admit the patient to a spare X3 and either leave the "broken" X3 to dry or send it to biomed with a tag. Followed up with Kam on Tuesday. (TC 07/02/2023 01:10:00 PM, Scero, Tomlyn) Call from ED. X3 showing "battery malfunction" alarm on screen. I walked the nurse through changing the X3 with a spare unit. She said it didn't clear the yellow alarm message. I explained to her that as long as the monitor was pluuged into a red outlet it will operate fine, but the X3 can't be used for transport. I followed up with Kam on tuesday and he will check to see if he can recondition the battery. unfortunately X3 batteries are on back order so we have none to replace it at this time.
1050672 | 195871 | CM | CMPLT | CMPLT | 2023-07-04 09:19 | 2023-07-04 09:33 | RL lead broken | (TC 07/04/2023 09:19:00 AM, Canjimavadivel, Kamallesh ) Replaced broken RL lead Replaced ripped keyboard cover
68979 | 195872 | CM | CMPLT | CMPLT | 2023-07-04 09:26 | 2023-07-04 09:38 | AC light turning off | (TC 07/04/2023 09:28:31 AM, Rhuland, Courtney) faulty power cord caused charging to be inconsistent Replaced power cord Electrical safety - OK
1056125 | 195873 | PM | CMPLT | CMPLT | 2023-07-04 09:41 | 2023-07-04 11:16 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/04/2023 11:15:23 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water When putting hose away, noticed cracked connector Replaced hose with spare
1066114 | 195874 | FC | CMPLT | CMPLT | 2023-07-04 11:05 | 2023-07-04 11:09 | QC check | (TC 07/04/2023 11:08:21 AM, Canjimavadivel, Kamallesh ) QC pass, VAC returned to ICU
1059828 | 195875 | PM | CMPLT | CMPLT | 2023-07-04 11:14 | 2023-07-04 11:14 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/04/2023 11:13:57 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1060030 | 195876 | CM | CMPLT | CMPLT | 2023-07-04 12:24 | 2023-07-26 11:16 | Pneumotach not properly heating up | (TC 07/04/2023 12:25:04 PM, Baghirzade, Nara) Last week, the pneumotach was reported not to heat up properly. When I checked, it was cold. Using infrared thermometer, measured the temp to be 23C (room temp); measured other pneumotachs on other systems and it measures around 35-38C. Changed the heater wire and waited about 20 minutes, it started to heat to 29C. Will give it another hour and check the temperature then. Heater wire is suspected to be defective but not yet confirmed. (TC 07/04/2023 02:26:00 PM, Baghirzade, Nara) Measured the temp on pneumotach heater; it was heating up to 33C, but needs to be slightly higher. Adjustment will need to be made on the PCB potentiometer. Gaby had said that its heating up much better and quicker. Heater wire will need to get properly replaced by routing it through the umbilical. Will check the temperature again to make sure its still heating properly. (TC 07/05/2023 09:00:00 AM, Baghirzade, Nara) Measured temperature of pneumotach; maximum measurement i was able to obtain was 31C, not as hot as it should be which is around 35-38C. Will need to adjust the potentiometer on the board to send more voltage to the heater. Will organize this job for a time when the clinic is least busy or after it closes. (TC 07/06/2023 02:45:49 PM, Baghirzade, Nara) Measured temperature around 30-31C, can definitely confirm that it will not go higher than that. Noted that it takes much less time to heat up. (TC 07/06/2023 02:45:49 PM, Baghirzade, Nara) Measured temperature around 30-31C, can definitely confirm that it will not go higher than that. Noted that it takes much less time to heat up. PFT techs did not report any significant changes to their flow. (TC 07/26/2023 11:10:22 AM, Baghirzade, Nara) John and I worked on this together. When measuring the heater wire voltage, was getting a reading of 0V; checked the fuse on the heater PCB and the fuse was blown. Fuse specs 800 mA/250V. The only fuse we had on hand was 600mA/250V, will order the exact fuse. After fuse replacement, John adjusted the potentiometer, waited 15 minutes and the temperature was approximately 30C-32C; not high enough. He adjusted the potentiometer again and the temperature reached approximately 37C-38C. The only remaining issue is that the heater was not heating uniformly; it was a lot hotter on the top than the bottom of the heater. We also checked the old heater wire to ensure that the fuse was not the cause of the issue the first time but the heater wire is dead. Had Gaby take a look and she said the temperature seems adequate. Will be monitoring the temperature throughout the week to ensure that it is holding up. (TC 07/27/2023 10:28:33 AM, Baghirzade, Nara) On the outside of the heater, the temperature measured around 38C. The temperature if checking at the screen site is 28C; this matches with the other stations. Gaby also mentioned that it is much better. Put a ziptie to secure the heater wire to the arm.
1049902 | 195877 | CM | CMPLT | CMPLT | 2023-07-04 12:40 | 2023-07-05 12:36 | MSH R62 DETECTOR BATTERY CHARGER CONTACTS BROKEN, UNABLE TO CHARGE EFFICIENTLY K5804-4513 | (TC 07/04/2023 01:42:38 PM, Whittington, Scott) MSH R62 DETECTOR BATTERY CHARGER CONTACTS BROKEN, UNABLE TO CHARGE EFFICIENTLY K5804-4513 (TC 07/04/2023 03:16:58 PM, Dolinar, John) MSH R62 DETECTOR BATTERY CHARGER CONTACTS BROKEN, UNABLE TO CHARGE EFFICIENTLY K5804-4513 REPLACE CAHARGER PINS IN MODLES REPLACE MODULES IN SYSTEM | MSH R62 DETECTOR BATTERY CHARGER CONTACTS BROKEN, UNABLE TO CHARGE EFFICIENTLY K5804-4513
32520 | 195878 | CM | CMPLT | CMPLT | 2023-07-04 12:43 | 2023-07-10 07:22 | MSH DR2 DIFFICULT TO CALIBRATE DETECTOR | (TC 07/04/2023 12:43:01 PM, Whittington, Scott) MSH DR2 DIFFICULT TO CALIBRATE DETECTOR (TC 07/05/2023 07:59:00 AM, Whittington, Scott) MSH DR2 DIFFICULT TO CALIBRATE DETECTOR | MSH DR2 DIFFICULT TO CALIBRATE DETECTOR
1064418 | 195879 | CM | CMPLT | CMPLT | 2023-07-04 13:06 | 2023-08-21 09:52 | Loose Door and broken hinge | (TC 07/04/2023 01:06:53 PM, Ferreira, Qwynn) Received call that the door was not correctly working Went to inspect the MRI and noticed the door was hanging off the chassis Grabbed a new door for the Left side and replaced the door as well as cleaned the inside from dust/dirt. Replaced the coil back and assembled the unit Informed Samir/Adeel that the door was coming off again Will order two more doors (Left and Right) (TC 07/25/2023 10:43:39 AM, Ferreira, Qwynn) Replaced the Right side door Replaced the seal, and installed head coil
1064826 | 195880 | CM | CMPLT | CMPLT | 2023-07-04 13:52 | 2023-07-27 10:09 | SpO2 not working | (TC 07/04/2023 01:51:46 PM, Canjimavadivel, Kamallesh ) Tested with test cable and sensor, X3 not detecting cable is connected. Verified test cable is good on a spare X3. Sending spare X3 back to unit. Need to further test. (TC 07/05/2023 09:34:19 AM, Canjimavadivel, Kamallesh ) Verified issue is still present, created a case on Philips customer portal (TC 07/05/2023 02:22:33 PM, Canjimavadivel, Kamallesh ) Spoke to tech over phone, bench repair initiated. Requested no charge PO. (TC 07/06/2023 01:19:04 PM, Canjimavadivel, Kamallesh ) Received PO. Shipped for repair via KJV (TC 07/27/2023 10:08:47 AM, Canjimavadivel, Kamallesh ) Went to received and got package. Verified SpO2 working. See attached report
| 195881 | CM | CMPLT | CMPLT | 2023-07-04 13:53 | 2023-08-09 13:31 | UPS Batteries | (TC 07/04/2023 03:12:40 PM, Reid, Emmerson ) ordered (TC 08/09/2023 01:29:31 PM, Reid, Emmerson ) All recieved and on the shielf
1059532 | 195882 | CM | CMPLT | CMPLT | 2023-07-04 14:07 | 2023-07-20 07:50 | not able to register the PICiX in FP | (TC 07/04/2023 02:07:33 PM, Samanian, Kamran) PACU1S was in biomed shop. rebuilt and reconfigured it in FP
68880 | 195883 | CM | CMPLT | NFF | 2023-07-04 14:28 | 2023-07-04 15:23 | "Not working properly" | (TC 07/04/2023 03:22:11 PM, Rhuland, Courtney) Checked error logs and no error found Functional check - OK No fault found
1063746 | 195884 | PM | CMPLT | CMPLT | 2023-07-04 14:47 | 2023-07-04 14:48 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/04/2023 02:47:04 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1061204 | 195885 | CM | CMPLT | CMPLT | 2023-07-04 15:16 | 2023-08-11 08:31 | OR 3 Decant | (TC 07/04/2023 03:15:50 PM, Reid, Emmerson ) Met with Maureen to discuss details of the decant for tomorrow afternoon. (TC 07/05/2023 04:29:50 PM, Reid, Emmerson ) Cleared off a conmed cart from Biomed. Moved the equipment onto the cart from one boom and detached the shelf from the other and added labels stating "OR 3" The equipment was left in OR 15 (TC 08/11/2023 08:30:41 AM, Reid, Emmerson ) Complete
1061208 | 195886 | CM | CMPLT | CMPLT | 2023-07-04 15:17 | 2023-08-09 15:02 | OR 7 Decant | (TC 07/04/2023 03:16:49 PM, Reid, Emmerson ) Met with Maureen to discuss details of the decant for tomorrow afternoon. (TC 07/05/2023 04:39:38 PM, Reid, Emmerson ) Created labels and an epty cart to place equipment on. Moved equipment from the boom to the cart and left in the front hallway.
1061333 | 195887 | CM | CMPLT | CMPLT | 2023-07-05 07:59 | 2023-07-05 08:04 | QC checklist button was greyed out | (TC 07/05/2023 08:03:31 AM, Baghirzade, Nara) When the VAC was sent down for QC, the QC checklist button was greyed out. Went in to VAC therapy, started and stopped the therapy, switched off and on the VAC, but the QC button was still greyed out as well as utilities and other options. Tried this method multiple times and could not find anything in the manual for troubleshooting this. Held the question mark button down for a few seconds and that unlocked the utilities and other options. After switching it off and then on, the QC checklist button was available. Ran a QC successfully, switched off and then on, and ran a VAC therapy. Switched it off, and upon booting the QC checklist button was available.
1061333 | 195888 | FC | CMPLT | CMPLT | 2023-07-05 08:04 | 2023-07-05 08:09 | QC checklist | (TC 07/05/2023 08:04:28 AM, Baghirzade, Nara) QC checklist completed. Returned to 18S.
| 195889 | CM | CMPLT | CMPLT | 2023-07-05 08:26 | 2023-07-04 11:00 | 1-1 with Adeel | (TC 07/04/2023 11:00:00 AM, Baghirzade, Nara) 1-1 meeting with Adeel
1067344 | 195890 | CM | CMPLT | CMPLT | 2023-07-05 09:02 | 2023-07-05 09:02 | Probe cover stuck on probe, unable to place it back in the holder | (TC 07/05/2023 09:01:56 AM, Baghirzade, Nara) The probe was damaged and unable to take the probe cover off. Replaced the probe. Functional test passed. Thermometer was not in TMS, created a new asset and asset sticker.
65892 | 195891 | PM | CMPLT | CMPLT | 2023-07-05 09:09 | 2023-07-05 09:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/05/2023 09:10:51 AM, Baghirzade, Nara) Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. Replaced left handle.
66242 | 195892 | CM | CMPLT | CMPLT | 2023-07-05 09:38 | 2023-07-05 09:38 | not working on the BAT | (TC 07/05/2023 09:37:51 AM, Samanian, Kamran) NFF, BAT WAS CHARGED AND TESTED AND PASSED. USER ISSUE.
| 195893 | CM | ACTIV | ISSUE | 2023-07-05 10:15 |  | Communications Committee | (TC 07/05/2023 10:20:03 AM, Scero, Tomlyn) mothly meeting to go over Medical device aquisition policy (TC 11/01/2023 01:17:14 PM, Scero, Tomlyn) Met with Adeel. Spoke about creating a Biomedical Department Handbook with policies and procedures. (TC 11/08/2023 02:08:44 PM, Scero, Tomlyn) Worked on the Employee Handbook from home (TC 11/27/2023 02:24:37 PM, Scero, Tomlyn) researching current policies for the biomed handbook (TC 11/28/2023 02:25:25 PM, Scero, Tomlyn) biomed handbook research (TC 01/05/2024 11:38:15 AM, Scero, Tomlyn) Working on compiling information and creating the Employee resource manual (TC 02/21/2024 01:58:38 PM, Scero, Tomlyn) discussed yellow repair tags (TC 04/03/2024 11:03:50 AM, Scero, Tomlyn) Meeting to discuss RFP laguage around Biomed service repair (TC 05/01/2024 11:14:06 AM, Scero, Tomlyn) meeting to discuss RFP language when it comes to Biomed.
1057714 | 195894 | CM | CMPLT | CMPLT | 2023-07-05 10:29 | 2023-07-06 13:25 | Recalibration icon when tested by unit | (TC 07/05/2023 10:29:35 AM, Canjimavadivel, Kamallesh ) Swapped battery with spare and starting reconditioning (TC 07/06/2023 01:23:58 PM, Canjimavadivel, Kamallesh ) Recondition complete, capacity 3/5 LEDs
| 195895 | CM | CMPLT | CMPLT | 2023-07-05 12:17 | 2023-07-21 07:53 | Configuring Focal Point server for all Philips Supplied Network | (TC 07/05/2023 04:46:28 PM, Samanian, Kamran) in progress (TC 07/06/2023 02:08:04 PM, Samanian, Kamran) working on FP, switches config, OS pathces for ER PIciX (TC 07/11/2023 08:10:52 AM, Samanian, Kamran) switches in NICU and 18 and 5SE configured for SSH, all swictes added to FP and validated. (TC 07/12/2023 10:01:48 AM, Samanian, Kamran) rest of switches such as L11, L12, L4, L15, L16 and biomed
1056986 | 195896 | CM | CMPLT | CMPLT | 2023-07-05 13:22 | 2023-07-05 13:23 | Config update | (TC 07/05/2023 01:22:08 PM, Canjimavadivel, Kamallesh ) Update config to correct ED profile, swapped battery for reconditioning
1057758 | 195897 | CM | CMPLT | CMPLT | 2023-07-05 13:28 | 2023-07-05 13:28 | Install patch | (TC 07/05/2023 01:27:59 PM, Canjimavadivel, Kamallesh ) Finished installing latest patch, sent update to FP
68903 | 195898 | CM | CMPLT | CMPLT | 2023-07-05 14:33 | 2023-07-05 14:44 | Missing RJ45 connector plug | (TC 07/05/2023 02:43:58 PM, Reid, Emmerson ) Repalced the RJ45 connector plug
68945 | 195899 | CM | CMPLT | CMPLT | 2023-07-05 15:37 | 2023-07-05 15:43 | ground pin bent on power cord | (TC 07/05/2023 03:38:34 PM, Cortez, John) Replaced power cord electrical safety- 140 mohms, 79 uA
63928 | 195900 | CM | CMPLT | NFF | 2023-07-05 15:44 | 2023-07-05 15:50 | "not working" | (TC 07/05/2023 03:44:35 PM, Cortez, John) Functional test ok, unable to find issue with cart Secured power cord with zip tie Added labels to some of the leads
| 195901 | CM | ACTIV | ISSUE | 2023-07-06 07:07 |  | Draeger boom for NICU meeting | (TC 07/05/2023 07:06:42 AM, Chukri, Samir) Meeting with Draeger person for booms 12 years maintenance (TC 08/01/2023 02:11:58 PM, Chukri, Samir) meeting
1067349 | 195902 | CM | CMPLT | CMPLT | 2023-07-06 07:19 | 2023-07-06 09:15 | Incoming inspection | (TC 07/06/2023 09:13:49 AM, Rhuland, Courtney) Set up and calibrated scale Zero and span calibrated due to weight being off
1054542 | 195903 | CM | CMPLT | CMPLT | 2023-07-06 08:36 | 2023-07-06 08:36 | "service" on screen | Unit was an ICU lift. Chaged account to 14 North. Inspected. Functional.
1054462 | 195904 | CM | CMPLT | CMPLT | 2023-07-06 08:37 | 2023-07-06 08:49 | Orange light and constant alarm | (TC 07/06/2023 08:46:30 AM, Canjimavadivel, Kamallesh ) Connected test set and let device run. No alarms were triggered, all functions working normal. Potential user issue with improper installation. Was not able to contact requester for follow-up. Sent device back to unit. | Note indicating check disposable alarm was alarming was sent with0000000 device
1046035 | 195905 | CM | CMPLT | CMPLT | 2023-07-06 08:56 | 2023-07-06 08:57 | Wallstand off the rail | (TC 07/06/2023 08:56:53 AM, Fan, Tak) Repositioned back on tracks. Checked movement and alignments, OK. | Both sets of wheels off. Not sure how the issue was caused.
| 195906 | CM | CMPLT | CMPLT | 2023-07-06 09:32 | 2023-07-05 09:32 | Communications committee meeting | (TC 07/05/2023 10:00:00 AM, Baghirzade, Nara) Communications committee meeting
| 195907 | CM | CMPLT | CMPLT | 2023-07-06 09:34 | 2023-07-05 11:30 | Fleet management meeting | (TC 07/05/2023 09:34:26 AM, Baghirzade, Nara) Fleet management meeting
| 195908 | CM | CMPLT | CMPLT | 2023-07-06 09:36 | 2023-07-05 13:00 | Rhinomanometer not measuring any pressure and loose connector | (TC 07/05/2023 01:00:00 PM, Baghirzade, Nara) Opened the rhinomanometer, found twisted tubing as well as one of the tubing was completely disconnected. Reconnected the tubing and fixed the twisted tube. Tightened the connector from inside the device. Passed off the rhinomanometer to Carter for testing. Carter confirmed that the device is fully functional. Also showed him how to disconnect the tubing so that it would not cause the connector to loosen on the inside.
1056215 | 195909 | CM | CMPLT | CMPLT | 2023-07-06 10:31 | 2023-07-06 10:34 | Repurposed Masimo for clinic and training | (TC 07/06/2023 10:31:04 AM, Baghirzade, Nara) Repurposed Masimo for family practice clinic (with Nima's approval) to use for BP averaging. Affixed a new asset tag and updated all information for the asset on TMS. All functional tests passed. Trained Mara on how to operate the device for BP averaging, SpO2, temperature modules, and some settings adjustment. Went over a few scenarios to ensure she understands how to use it and is comfortable with it.
1056458 | 195910 | CM | CMPLT | CMPLT | 2023-07-06 10:34 | 2023-07-06 10:34 | Repurposed Masimo for clinic and training | (TC 07/06/2023 10:34:17 AM, Baghirzade, Nara) Repurposed Masimo for family practice clinic (with Nima's approval) to use for BP averaging. Affixed a new asset tag and updated all information for the asset on TMS. All functional tests passed. Trained Mara on how to operate the device for BP averaging, SpO2, temperature modules, and some settings adjustment. Went over a few scenarios to ensure she understands how to use it and is comfortable with it.
1057761 | 195911 | CM | CMPLT | CMPLT | 2023-07-06 10:40 | 2023-07-06 10:42 | Patch install | (TC 07/06/2023 10:40:34 AM, Canjimavadivel, Kamallesh ) Completed latest patch install successfully
1060149 | 195912 | CM | CMPLT | CMPLT | 2023-07-06 11:02 | 2023-07-06 13:24 | lead issues | (TC 07/06/2023 01:20:57 PM, Scero, Tomlyn) Replaced Long chest leas V1-V3 and both RA and LA limb leads ECG simulation complete and passed. returned unit to 16 ECG
65943 | 195913 | CM | CMPLT | CMPLT | 2023-07-06 11:26 | 2023-07-06 11:31 | Error 354.6770 | (TC 07/06/2023 11:26:36 AM, Baghirzade, Nara) Replaced pressure sensor. After calibrating the sensor and during check-in, noted that the claws do not close properly and get stuck in the open position. Opened the plunger, cleaned the gears for the claws and noted that the spring was not in properly, causing the claws to stay open. Seated the spring in propely. Performed a full calibration of: -pressure sensor disc -barrel clamp -plunger force -plunger position Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1067336 | 195914 | CM | CMPLT | CMPLT | 2023-07-06 12:28 | 2023-07-06 12:29 | TightVNC installation | (TC 07/06/2023 12:27:48 PM, Baghirzade, Nara) Installed TightVNC and applied the appropriate password for access. Rebooted PSN. Kamran checked for remote access and it was successful.
1053385 | 195915 | CM | CMPLT | NFF | 2023-07-06 12:29 | 2023-07-06 12:30 | Does not measure temperature and constantly prompts for a probe cover, even when its on | (TC 07/06/2023 12:29:18 PM, Baghirzade, Nara) Was unable to replicate the error that Iliana was getting. Cleaned the probe sensor. Cleaned the probe holders. Measured temperature multiple times; thermometer is functional. Put the cover on and off several times, each time the prompt was correct. No fault found.
1058276 | 195916 | CM | CMPLT | CMPLT | 2023-07-06 12:34 | 2023-07-06 12:38 | Damaged power cord | (TC 07/06/2023 12:36:34 PM, Baghirzade, Nara) Replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 78 mohms Leakage current: 65.5 uA
| 195917 | CM | CMPLT | CMPLT | 2023-07-06 14:07 | 2023-07-06 14:07 | meeting with Rhys for SES | (TC 07/06/2023 02:07:07 PM, Samanian, Kamran) meeting done and sent email with files
1057760 | 195918 | CM | CMPLT | CMPLT | 2023-07-06 14:38 | 2023-07-10 12:40 | Host auto shutdown issue | (TC 07/06/2023 02:38:02 PM, Canjimavadivel, Kamallesh ) Shutdown host and removed internal USB. Restarted PC and discovered OS stuck in auto shutdown loop. Replaced host with Spare PICiX (TC 07/07/2023 08:10:11 AM, Samanian, Kamran) rebuilt the PICiX from Philips USB, changed the name and Ip address. once it will be installed in the unit the host will be patched (TC 07/07/2023 09:25:22 AM, Canjimavadivel, Kamallesh ) Setup host with original configuration in unit. Installed VNC and added printer. Setup Spare PICiX back to original config in biomed. (TC 07/10/2023 12:39:09 PM, Canjimavadivel, Kamallesh ) Completed installing latest patch
1062141 | 195919 | CM | CMPLT | CMPLT | 2023-07-06 14:44 | 2023-07-06 14:45 | Patient helper bracket installation | (TC 07/06/2023 02:43:47 PM, Baghirzade, Nara) Installed patient bracket on bed. Placed patient helper in bracket. (TC 07/06/2023 02:44:21 PM, Rhuland, Courtney) Installed patient bracket on bed with Courtney's help. Placed patient helper in bracket.
| 195920 | CM | CMPLT | CMPLT | 2023-07-06 15:48 | 2023-08-02 13:34 | biomed misc - July 2023 | (TC 07/06/2023 03:49:49 PM, Cortez, John) Swapped out LVPs with PMs due in ICU storage room (TC 07/06/2023 03:50:34 PM, Cortez, John) Philips SCS meeting (TC 07/06/2023 03:55:24 PM, Cortez, John) bi-weekly meeting (TC 07/05/2023 03:55:24 PM, Cortez, John) bi-weekly meeting (TC 07/07/2023 09:04:44 AM, Cortez, John) TMS cleanup, followed up on open POs (TC 07/10/2023 03:09:54 PM, Cortez, John) monday huddle (TC 07/10/2023 03:10:57 PM, Cortez, John) Provided flowmeter ordering info to Katie (CNS) (TC 07/13/2023 03:30:47 PM, Cortez, John) Looked at issue with PFT PC in room 1615 with Nara. Couldn't replicate issue with error on app shutdown. Nara will reach out to PFT to schedule time with clinical staff (TC 07/19/2023 03:43:55 PM, Cortez, John) Philips SCS meetings (TC 07/19/2023 03:44:31 PM, Cortez, John) Was asked by Sherrye to verify SPO2 sensor price. Called Philips and confirmed price with her (TC 07/17/2023 03:48:13 PM, Cortez, John) Showed Samir how to reset Masimo viewstations remotely to clear errors (TC 07/17/2023 03:49:04 PM, Cortez, John) Helped Jessica install cover on OR 9 monitor (TC 07/17/2023 03:51:53 PM, Cortez, John) huddle (TC 07/19/2023 03:52:26 PM, Cortez, John) biweekly biomed meeting (TC 07/17/2023 03:52:26 PM, Cortez, John) biweekly biomed meeting (TC 07/17/2023 03:51:53 PM, Cortez, John) huddle (TC 07/20/2023 03:44:24 PM, Cortez, John) Philips biweekly meeting (TC 07/20/2023 03:45:05 PM, Cortez, John) systems committee (TC 07/20/2023 03:45:34 PM, Cortez, John) biomed cleanup (TC 07/21/2023 03:58:17 PM, Cortez, John) human factors survey at TGH (TC 07/25/2023 03:43:55 PM, Cortez, John) Put together list of equipment from Mississauga clinic for Kamran (TC 07/26/2023 03:33:24 PM, Cortez, John) Helped Kam with exam light (TC 07/27/2023 11:12:47 AM, Cortez, John) Looked into Masimo "16N1 client down" and "keylock authentication" errors. Apparently Naz contacted Masimo about it already (TC 07/27/2023 11:18:22 AM, Cortez, John) Looked into windows PC error in PFT with Nara (TC 07/27/2023 02:50:26 PM, Cortez, John) Rebooted Pod B Piicix due to patch (TC 07/27/2023 03:55:47 PM, Cortez, John) Recorder not working in Pod B nurse station Resetted Picix, all good (TC 07/24/2023 04:36:19 PM, Cortez, John) huddle (TC 07/24/2023 04:36:19 PM, Cortez, John) huddle (TC 07/28/2023 01:49:56 PM, Cortez, John) Looked into Masimo "keylock authentication error" with Kamran, Naz and Nara Unable to ping myUTN device/server Cables/port seem to be good. myUTN device may be faulty. Issue ongoing. (TC 07/28/2023 03:40:00 PM, Cortez, John) Restarted safetynet application on centrals, after Masimo made edit to script (TC 07/28/2023 03:40:57 PM, Cortez, John) Addressed call from breast imaging about "hologic" system. Confirmed that biomed doesn't look after it, and forwarded them to Scott at UHN (TC 07/31/2023 01:33:37 PM, Cortez, John) huddle
| 195921 | PA | CMPLT | CMPLT | 2023-07-07 07:31 | 2023-07-12 10:03 | Ordering clear mattress plates for Giraffe incubators | (TC 07/07/2023 07:36:09 AM, Baghirzade, Nara) SC 1002997369 (TC 07/12/2023 10:02:11 AM, Baghirzade, Nara) PO 7202245428
65166 | 195922 | CM | CMPLT | CMPLT | 2023-07-07 07:54 | 2023-07-07 10:41 | CADD Error: Remove and reattach casette | (TC 07/07/2023 10:40:37 AM, Reid, Emmerson ) emailed smiths about getting an RMA
1055749 | 195923 | PM | CMPLT | CMPLT | 2023-07-07 09:23 | 2023-07-07 09:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/07/2023 09:22:52 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1047833 | 195924 | CM | CMPLT | CMPLT | 2023-07-07 09:40 | 2023-07-07 11:39 | Config, time/date update | (TC 07/07/2023 11:38:25 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
1061205 | 195925 | CM | CMPLT | CMPLT | 2023-07-07 10:59 | 2023-07-07 15:00 | OR 4 DVI connector on foot boom is damaged | (TC 07/07/2023 02:58:52 PM, Reid, Emmerson ) Replaced with Jessica. Working now. Was told that OR 6 also had the same issue.
1067357 | 195926 | CM | CMPLT | CMPLT | 2023-07-07 11:19 | 2023-07-07 11:31 | Incoming Inspection | (TC 07/07/2023 11:27:27 AM, Reid, Emmerson ) Deviced used as a demo for one case on July 21st. Technical Safety form sent with device, I have physical copy. Electrical safety complete in proceedure. Asset created. Incoming complete.
1067360 | 195927 | ACC | CMPLT | CMPLT | 2023-07-07 11:41 | 2023-07-07 11:48 | Incoming Inspection | (TC 07/07/2023 11:46:24 AM, Reid, Emmerson ) Deviced used as a demo for one case on July 21st. Technical Safety form sent with device, I have physical copy. Electrical safety complete in proceedure. Asset created. Incoming complete.
| 195928 | CM | CMPLT | CMPLT | 2023-07-07 12:09 | 2023-07-07 12:10 | Laser meeting | (TC 07/07/2023 12:09:38 PM, Baghirzade, Nara) Laser committee meeting
1056585 | 195929 | CM | CMPLT | CMPLT | 2023-07-07 12:11 | 2023-07-07 12:11 | AED cabinet alarming | (TC 07/07/2023 12:10:49 PM, Baghirzade, Nara) AED cabinet was alarming; replaced the 9V battery. Activated the cabinet.
1066462 | 195930 | FC | CMPLT | CMPLT | 2023-07-07 12:29 | 2023-07-07 12:29 | device check | passed QC test. functional. returned.
| 195931 | CM | CMPLT | CMPLT | 2023-07-07 12:36 | 2023-07-07 12:36 | Updating network and system information for EMGs | (TC 07/07/2023 12:35:40 PM, Baghirzade, Nara) Updated all network and system documentation for 2 EMG units. Document has been uploaded to system documentation on T drive.
| 195932 | CM | CMPLT | CMPLT | 2023-07-07 12:37 | 2023-07-07 12:39 | Dropped off Olympus adapter for light source | (TC 07/07/2023 12:38:04 PM, Baghirzade, Nara) Dropped off an olympus light source adapter for Dr. Witterick; his complaint was that the light source moves around too much in the adapter. Incorrect placement of the light source; showed how to insert light source properly into adapter to Elena.
| 195933 | CM | CMPLT | CMPLT | 2023-07-07 12:40 | 2023-07-07 12:40 | working on switch config for 18 floor for SSH | (TC 07/07/2023 12:40:22 PM, Samanian, Kamran) In progress
1061898 | 195934 | CM | CMPLT | CMPLT | 2023-07-07 12:57 | 2023-07-07 13:01 | Report "footboard shocking patient upon contact" | (TC 07/07/2023 12:57:00 PM, Canjimavadivel, Kamallesh ) Went up to unit and identified bed. Waited for unit to switch patient to different bed. Performed electrical safety. All measurements within spec. Informed unit no issue found with bed Leakage: 130 uA Resistance: 100 mOhm
62700 | 195935 | CM | CMPLT | CMPLT | 2023-07-07 13:28 | 2023-07-10 14:08 | No note or tag came with device | (TC 07/07/2023 01:47:30 PM, Canjimavadivel, Kamallesh ) Further troubleshooting indicated issue is with the pins V1-V3. All leads okay/ No visible issue with any pins. (TC 07/07/2023 02:12:41 PM, Canjimavadivel, Kamallesh ) Batteries also not charging when plugged in. Same behaviour when swapping power supply and batteries. (TC 07/10/2023 02:05:50 PM, Scero, Tomlyn) Unable to fix V2 lead in PIM. I echanged the PIM with an old ICU cart no longer in service (62703) I plugged in the cart and the batteries charged. We'll keep an eye on this issue. Returned device to 16 ECG depart. | Connected with simulator and found intermittent issue with V2 lead
1061207 | 195936 | CM | CMPLT | CMPLT | 2023-07-07 15:00 | 2023-07-07 15:01 | OR 6 DVI connector broken | (TC 07/07/2023 03:00:35 PM, Reid, Emmerson ) Replaced with Jessica, had to replace one of the screws as well because it had a part broken inside of it. Working now.
1046274 | 195937 | CM | CMPLT | CMPLT | 2023-07-07 15:04 | 2023-07-11 07:52 | MSH IVF CFRH 7flr APLIO 300 TRACKBALL FOAM REPLACEMENT PN# BSM34-1630-02 FROM IVF STOCK | (TC 07/10/2023 04:37:05 PM, Whittington, Scott) MSH IVF CFRH 7flr APLIO 300 TRACKBALL FOAM REPLACEMENT PN# BSM34-1630-02 FROM IVF STOCK | MSH IVF CFRH 7flr APLIO 300 TRACKBALL FOAM REPLACEMENT PN# BSM34-1630-02 FROM IVF STOCK
1046275 | 195938 | CM | CMPLT | CMPLT | 2023-07-07 15:05 | 2023-07-11 07:52 | MSH IVF CFRH 7flr APLIO 300 TRACKBALL FOAM REPLACEMENT PN# BSM34-1630-02 FROM IVF STOCK | (TC 07/10/2023 04:36:43 PM, Whittington, Scott) MSH IVF CFRH 7flr APLIO 300 TRACKBALL FOAM REPLACEMENT PN# BSM34-1630-02 FROM IVF STOCK | MSH IVF CFRH 7flr APLIO 300 TRACKBALL FOAM REPLACEMENT PN# BSM34-1630-02 FROM IVF STOCK
1059180 | 195939 | CM | CMPLT | NFF | 2023-07-07 15:08 | 2023-07-07 15:10 | ECG not reading properly | (TC 07/07/2023 03:08:18 PM, Reid, Emmerson ) Said unable to show ECG rhythm. With the simulator it showed clearly. Likely were using expired electrodes. NFF
| 195940 | CON | CMPLT | CMPLT | 2023-07-07 15:09 | 2024-09-25 10:40 | MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM | (TC 07/11/2023 08:10:10 AM, Whittington, Scott) MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM ADVISED THAT UTILIZING A APLIO 500 PROGRAMMED FOR eIVF, AND STAFF FROM 250 DUNDAS BRINGING 781/661 TV PROBE AND NEEDLE GUIDE (TC 11/20/2023 07:18:13 AM, Whittington, Scott) MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM ADVISED THAT UTILIZING A APLIO 500 PROGRAMMED FOR eIVF, AND STAFF FROM 250 DUNDAS BRINGING 781/661 TV PROBE AND NEEDLE GUIDE (TC 11/20/2023 07:18:25 AM, Whittington, Scott) MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM ADVISED THAT UTILIZING A APLIO 500 PROGRAMMED FOR eIVF, AND STAFF FROM 250 DUNDAS BRINGING 781/661 TV PROBE AND NEEDLE GUIDE (TC 02/15/2024 03:55:34 PM, Whittington, Scott) MSH IVF ON MSH 15 SPP PROCEDURE PHWY9 (TC 02/29/2024 01:43:34 PM, Whittington, Scott) MSH IVF ON MSH 15 SPP PROCEDURE PHWY9 (TC 11/20/2023 07:18:25 AM, Whittington, Scott) MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM ADVISED THAT UTILIZING A APLIO 500 PROGRAMMED FOR eIVF, AND STAFF FROM 250 DUNDAS BRINGING 781/661 TV PROBE AND NEEDLE GUIDE (TC 06/04/2024 02:31:07 PM, Whittington, Scott) MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM SUUCESSFUL TEST OF PT CYCLE MWL TO STORAGE, SR MAY NEEED TO BE TWEAKED ON APLIO END MARK FROM eVF ON CALL OF DURATION (TC 06/04/2024 02:31:07 PM, Whittington, Scott) MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM SUUCESSFUL TEST OF PT CYCLE MWL TO STORAGE, SR MAY NEEED TO BE TWEAKED ON APLIO END MARK FROM eVF ON CALL OF DURATION (TC 06/05/2024 12:42:09 PM, Whittington, Scott) MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM ADVISED THAT UTILIZING A APLIO 500 PROGRAMMED FOR eIVF, AND STAFF FROM 250 DUNDAS BRINGING 781/661 TV PROBE AND NEEDLE GUIDE ASHLEY SISETO AND GOMA NARINE ON-SITE FOR CLINICAL SETUP OF SYSTEM ARRANGE OPERATOR PANEL FUNCTIONS TO MATCH APLIO 300 AT 250 DUNDAS (TC 06/06/2024 09:45:47 AM, Whittington, Scott) msh 15 ivf aplio 500 project 195940 2 2 jul 10 op APLIO 500 OP PANEL | MSH IVF IS PROPOSING TO PERFORM PROCEDURES ON FRAGILE PT IN THE MSH 15 SPP PROCEDURE ROOM ADVISED THAT UTILIZING A APLIO 500 PROGRAMMED FOR eIVF, AND STAFF FROM 250 DUNDAS BRINGING 781/661 TV PROBE AND NEEDLE GUIDE SYSTEM SN# W1E13X4656 MAC 00-60-0A-08-50-49 AE PHWY9 From: Wang, Wendy <Wendy.Wang@sinaihealth.ca> Sent: February 15, 2024 1:48 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: RE: [External] FW: Ticket # SINAI@175165 - Request to add and configure Ultrasound machine to EIVF system Last machine in Mississauga location was PHWY8 The one before that in Vaughan location was PHWY7 It is totally fine to go with the next one in line, like PHWY9 Wendy Wang Business Coordinator Women's & Infants Health Program. T 416-586-4800 ext. 7370 C 416-586-4686 wendy.wang@sinaihealth.ca 250 Dundas St W, Toronto, ON M5T 2Z5
32510 | 195941 | CM | CMPLT | CMPLT | 2023-07-07 15:10 | 2023-07-10 07:19 | MSH DR1 TUBE/DETECTOR ALIGNMENT ISSUES | (TC 07/08/2023 05:19:00 PM, Whittington, Scott) MSH DR1 TUBE/DETECTOR ALIGNMENT ISSUES COMPLETE CALIBRATION OF ROOM CORRECT MULTIPLE FAIOURES | MSH DR1 TUBE/DETECTOR ALIGNMENT ISSUES
67962 | 195942 | CM | CMPLT | CMPLT | 2023-07-10 07:25 | 2023-07-10 07:26 | Broken door latch | (TC 07/10/2023 07:25:23 AM, Rhuland, Courtney) Door latch replaced Functional check - OK
67956 | 195943 | CM | CMPLT | CMPLT | 2023-07-10 08:01 | 2023-07-10 08:02 | Broken - Tilt Bed | (TC 07/10/2023 08:01:30 AM, Ferreira, Qwynn) Reconnnected the tilt ball as the SA did not attach it correctly Functional Check - OK
1055792 | 195944 | CM | CMPLT | CMPLT | 2023-07-10 08:18 | 2023-07-10 08:22 | Defib not plugged in | (TC 07/10/2023 08:18:01 AM, Reid, Emmerson ) Went up and checked the defid. It was not plugged in, so i plugged it in.
| 195945 | PA | CMPLT | CMPLT | 2023-07-10 08:59 | 2023-07-18 09:52 | Battery for Transport Incubator
1054537 | 195946 | PM | CMPLT | CMPLT | 2023-07-10 09:23 | 2023-07-10 10:31 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/10/2023 10:30:18 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1057938 | 195947 | CM | CMPLT | CMPLT | 2023-07-10 09:38 | 2024-06-19 08:09 | TS7000 Segment Order/Repair | (TC 07/10/2023 10:06:54 AM, Reid, Emmerson ) Reached out to Reeba about purchasing. Entered in the parts and cost in the WO (TC 07/11/2023 01:13:20 PM, Reid, Emmerson ) Placed the order on erec (TC 07/20/2023 03:15:58 PM, Reid, Emmerson ) Head segment boar arrived. I also assembled from the old one. (TC 08/02/2023 09:56:03 AM, Reid, Emmerson ) Recieved 2 widening segment plates (TC 08/02/2023 10:42:38 AM, Reid, Emmerson ) Replaced the plates on the widening segments. Working now. (TC 08/10/2023 01:31:20 PM, Reid, Emmerson ) Recieved 2x widening segment screws (TC 08/10/2023 02:22:03 PM, Reid, Emmerson ) Both armboards recieved and brought to OR (TC 08/17/2023 10:35:59 AM, Reid, Emmerson ) Recieved the thigh plate, plate below knee, and handle (TC 08/17/2023 10:35:59 AM, Reid, Emmerson ) Recieved the thigh plate, plate below knee, and handle, but it did not include the velcrow. Need to email about havin that sent (TC 05/01/2024 01:58:09 PM, Reid, Emmerson ) Following up with this (TC 05/07/2024 02:35:52 PM, Reid, Emmerson ) Went through emails to follow up again and provide the parts information to MMC (TC 05/31/2024 02:49:27 PM, Reid, Emmerson ) Recieved the Holm, and Screw ISO10642 (TC 06/19/2024 08:09:13 AM, Reid, Emmerson ) All Parts recieved from this WO
68343 | 195948 | CM | CMPLT | CMPLT | 2023-07-10 09:45 | 2023-07-10 09:45 | not working | Scale not working. replaced batteries. checked calibration. functional
1059778 | 195949 | FC | CMPLT | CMPLT | 2023-07-10 09:46 | 2023-07-10 09:46 | QC check | passed QC check. functional.
| 195950 | CM | CMPLT | CMPLT | 2023-07-10 10:46 | 2023-07-10 10:47 | work with John on Alaris pump board. | (TC 07/10/2023 10:46:03 AM, Chukri, Samir) work with John to re-sold a cap. on thr Alarise pump board.
1048200 | 195951 | CM | CMPLT | CMPLT | 2023-07-10 14:18 | 2023-11-16 15:13 | bad connection | (TC 08/10/2023 08:57:15 AM, Huang, Jessica) received the calls from OR 10. the monitors went black a few times. checked the connection on the device. eveything seemed right. checked the connection between the scope and the vidoe system and it seemed right. satyed in OR 10 for awhile and it didnt happen again. Tory sent the email again regarding the same issue. spoke to Jaron and Maureen, swapped the video system on the ENT cart. sent the device for insoection.
1049046 | 195952 | CM | CMPLT | CMPLT | 2023-07-10 14:25 | 2023-08-28 08:56 | warranty replacement | (TC 07/10/2023 08:55:23 AM, Huang, Jessica) took the broken device from OR, sent the PO to the vendor and sent the replacement unit back to OR
1048018 | 195953 | CM | CMPLT | CMPLT | 2023-07-10 14:33 | 2023-07-10 14:35 | Not charging | (TC 07/10/2023 02:33:43 PM, Rhuland, Courtney) Device flagged by Rex that it was not charging Internal wires to pin were broken off Resoldered wires to pin and ground Functional check - OK
1062999 | 195954 | PM | CMPLT | CMPLT | 2023-07-10 14:37 | 2023-07-10 14:39 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/10/2023 02:36:55 PM, Rhuland, Courtney) Device would not turn on due to error 85 Booted up device in service mode and all functional operations were working Cleared error from log and device worked normally Device drained, disinfected, rinsed and refilled with sterile water
| 195955 | CM | CMPLT | CMPLT | 2023-07-10 14:58 | 2023-07-18 14:14 | OPG URODYNAMICS: Change bulb and order new ones | (TC 07/10/2023 02:57:57 PM, Reid, Emmerson ) Went over to OPG to change a light bulb. They requested me to order 6 new ones. (TC 07/12/2023 08:00:38 AM, Reid, Emmerson ) Placed and order on SAP
1054467 | 195956 | CM | CMPLT | CMPLT | 2023-07-10 15:06 | 2023-07-13 08:26 | "max-inflate not working, side rail bent" | (TC 07/10/2023 03:06:17 PM, Cortez, John) Assessed bed in room 12, and parked in ICU hallway Confirmed issue, no error codes Placed service request with Hillrom (TC 07/10/2023 03:08:19 PM, Canjimavadivel, Kamallesh ) Assessed bed in room 12, and parked in ICU hallway Confirmed issue, no error codes Placed service request with Hillrom (TC 07/11/2023 07:31:27 AM, Cortez, John) Took Vince (Hillrom) to ICU where bed was located (TC 07/12/2023 07:43:12 AM, Cortez, John) Vince repaired bed (TC 07/12/2023 07:43:12 AM, Cortez, John) Vince repaired bed. Service report attached
67762 | 195957 | CM | CMPLT | CMPLT | 2023-07-10 15:07 | 2023-07-10 15:07 | Library Pushed: APS EPIDURAL - 2023 JULY
67768 | 195958 | CM | CMPLT | CMPLT | 2023-07-10 15:07 | 2023-07-10 15:08 | Library Pushed: APS EPIDURAL - 2023 JULY
67769 | 195959 | CM | CMPLT | CMPLT | 2023-07-10 15:08 | 2023-07-10 15:08 | Library Pushed: APS EPIDURAL - 2023 JULY
1058153 | 195960 | CM | CMPLT | CMPLT | 2023-07-10 15:08 | 2023-07-10 15:09 | Library Pushed: APS EPIDURAL - 2023 JULY
67767 | 195961 | CM | CMPLT | CMPLT | 2023-07-10 15:09 | 2023-07-10 15:09 | Library Pushed: APS EPIDURAL - 2023 JULY
67758 | 195962 | CM | CMPLT | CMPLT | 2023-07-10 15:10 | 2023-07-10 15:10 | Library Pushed: APS EPIDURAL - 2023 JULY
1058170 | 195963 | CM | CMPLT | CMPLT | 2023-07-10 15:10 | 2023-07-10 15:10 | Library Pushed: APS EPIDURAL - 2023 JULY
67755 | 195964 | CM | CMPLT | CMPLT | 2023-07-10 15:11 | 2023-07-10 15:11 | Library Pushed: APS EPIDURAL - 2023 JULY
67765 | 195965 | CM | CMPLT | CMPLT | 2023-07-10 15:11 | 2023-07-10 15:12 | Library Pushed: APS EPIDURAL - 2023 JULY
1058178 | 195966 | CM | CMPLT | CMPLT | 2023-07-10 15:12 | 2023-07-10 15:12 | Library Pushed: APS EPIDURAL - 2023 JULY
1058175 | 195967 | CM | CMPLT | CMPLT | 2023-07-10 15:13 | 2023-07-10 15:13 | Library Pushed: APS EPIDURAL - 2023 JULY
67757 | 195968 | CM | CMPLT | CMPLT | 2023-07-10 15:13 | 2023-07-10 15:13 | Library Pushed: APS EPIDURAL - 2023 JULY
32580 | 195969 | CM | CMPLT | CMPLT | 2023-07-10 15:22 | 2023-07-10 15:22 | X-Ray condition prohibited | X-Ray condition prohibited -Oil flow error -Faulty Heat Exchanger -Ordered part and replaced it -Error cleared after and scanning was possible.
66003 | 195970 | CM | CMPLT | NFF | 2023-07-11 07:19 | 2023-07-11 07:22 | Will not connect to network | (TC 07/11/2023 07:20:03 AM, Rhuland, Courtney) Turned device on and network light was present Confirmed connectivity using the network connectivity test on ASM All tests passed No fault found
1048018 | 195971 | CM | CMPLT | CMPLT | 2023-07-11 10:17 | 2023-07-18 15:24 | Charging Cable | (TC 07/11/2023 10:17:55 AM, Ferreira, Qwynn) Worked with Kim to get the right power supply and input the order (TC 07/18/2023 03:24:03 PM, Ferreira, Qwynn) completed order and stored spare cables. Input po's as well.
| 195972 | PA | CMPLT | CMPLT | 2023-07-11 10:24 | 2023-07-18 15:22 | Endoscope CMAC Power Supply
| 195973 | CM | ACTIV | ISSUE | 2023-07-11 10:56 |  | test
1051384 | 195974 | CM | CMPLT | CMPLT | 2023-07-11 11:19 | 2023-07-12 08:54 | Fan not turing on, causing unit to power down | (TC 07/11/2023 11:18:43 AM, Ferreira, Qwynn) retireved the unit from the floor with tag on it. No fan Tested the device and noticed that the fan was not powering on. Fan is working, it was just out of place since it is held in place using a magnet. Will troubeshoot further (TC 07/11/2023 03:43:40 PM, Ferreira, Qwynn) made a new connection for the fan and added epoxy to hold the fan in place. Letting the fan cure to test further. (TC 07/12/2023 08:52:16 AM, Ferreira, Qwynn) tested the unit and is working well. Epoxy is holding the fan in place functional test - ok Electrical safety - ok
| 195975 | MEP | ACTIV | ISSUE | 2023-07-11 12:29 |  | Weekly- CMMS data quality project | (TC 07/11/2023 12:29:22 PM, Chukri, Samir) Weekly- CMMS data quality project (TC 07/11/2023 12:29:22 PM, Chukri, Samir) Weekly- CMMS data quality project (TC 07/18/2023 11:02:53 AM, Chukri, Samir) Meeting (TC 08/08/2023 01:27:00 PM, Chukri, Samir) meeting with Backy (TC 08/15/2023 11:36:51 AM, Chukri, Samir) Meeting (TC 08/17/2023 03:01:38 PM, Chukri, Samir) meeting with Becky clean the TMS account (TC 08/29/2023 03:05:25 PM, Chukri, Samir) working with Beky to update the department account (TC 08/31/2023 10:06:19 AM, Chukri, Samir) meeting with Becky correcting the TMS accounts (TC 09/12/2023 11:00:57 AM, Chukri, Samir) Zoom Meeting (TC 10/26/2023 11:05:40 AM, Chukri, Samir) working with Becky on account update (TC 02/02/2024 07:54:09 AM, Chukri, Samir) Meet with Beckay. (TC 02/16/2024 01:57:39 PM, Chukri, Samir) working with Becky on the CMMS acount (TC 02/22/2024 02:20:28 PM, Chukri, Samir) working on TMS accounts (TC 03/12/2024 02:30:00 PM, Chukri, Samir) CMMS Committee. (TC 06/03/2024 03:13:48 PM, Chukri, Samir) Meeting (TC 07/15/2024 07:44:40 AM, Chukri, Samir) meeting (TC 08/19/2024 01:10:00 PM, Chukri, Samir) CMMS meeting . (TC 10/07/2024 03:06:03 PM, Chukri, Samir) meeting
| 195976 | CM | CMPLT | CMPLT | 2023-07-11 12:34 | 2023-07-11 12:34 | working in two brest feeding pump | (TC 07/10/2023 12:33:47 PM, Chukri, Samir) working with Jessica repaired two Breast pump.
| 195977 | CM | CMPLT | CMPLT | 2023-07-11 12:35 | 2023-07-10 12:35 | Biomed meeting | Biomed meeting
530498 | 195978 | CM | CMPLT | CMPLT | 2023-07-11 13:04 | 2023-09-12 13:32 | Not working | (TC 07/11/2023 02:38:54 PM, Scero, Tomlyn) service repair PO has been requested. (TC 07/17/2023 11:37:47 AM, Scero, Tomlyn) Recieved quote from draeger. Had purchasing update the PO. Emailed Dreger to approve the quote and start the repair. (TC 09/12/2023 01:31:58 PM, Scero, Tomlyn) Returned from Draeger. Service report attached.
68389 | 195979 | CM | CMPLT | CMPLT | 2023-07-11 13:05 | 2023-09-12 13:32 | Not working | (TC 07/11/2023 02:38:18 PM, Scero, Tomlyn) service repair PO has been requested. (TC 07/17/2023 11:36:10 AM, Scero, Tomlyn) Recieved quote from draeger. Had purchasing update the PO. Emailed Dreger to approve the quote and start the repair. (TC 09/12/2023 01:31:10 PM, Scero, Tomlyn) Return from Draeger. Service report attached.
1057560 | 195980 | CM | CMPLT | CMPLT | 2023-07-11 13:05 | 2023-07-17 09:59 | BP module malfunction | (TC 07/11/2023 02:36:57 PM, Scero, Tomlyn) COnfirmed BP malfunction. Tested unit and BP couldn't reach high enough during test to read. In service mode a leak could be detected. A new BP module will have to be ordered. (TC 07/12/2023 10:24:07 AM, Scero, Tomlyn) Ordered BP pump from Philips (TC 07/17/2023 09:57:49 AM, Scero, Tomlyn) BP pump module received Friday. Replaced module monday. Tested BP and checked for leaks. No issues. BP reading 118/84
1067369 | 195981 | ACC | CMPLT | CMPLT | 2023-07-11 13:55 | 2023-07-11 13:57 | Incoming Inspection | (TC 07/11/2023 01:55:57 PM, Reid, Emmerson ) Incoming complete, electrical safety done in the proceedure. Created asset tag and labeled unit as a "Demo Unit"
32510 | 195982 | CM | CMPLT | CMPLT | 2023-07-11 16:58 | 2023-07-13 16:32 | MSH DR1 SYSTEM WILL NOT RESPOND; ERROR 30168 APID 22 ENCODER /POTENTIOMETER FAULTY | (TC 07/11/2023 04:57:54 PM, Whittington, Scott) MSH DR1 SYSTEM WILL NOT RESPOND; ERROR 30168 APID 22 ENCODER /POTENTIOMETER FAULTY (TC 07/11/2023 04:58:05 PM, Whittington, Scott) MSH DR1 SYSTEM WILL NOT RESPOND; ERROR 30168 APID 22 ENCODER /POTENTIOMETER FAULTY (TC 07/11/2023 09:58:00 PM, Whittington, Scott) MSH DR1 SYSTEM WILL NOT RESPOND; ERROR 30168 APID 22 ENCODER /POTENTIOMETER FAULTY | MSH DR1 SYSTEM WILL NOT RESPOND; ERROR 30168 APID 22 ENCODER /POTENTIOMETER FAULTY
69439 | 195983 | CM | CMPLT | CMPLT | 2023-07-12 07:30 | 2023-07-12 07:31 | LVP display | (TC 07/11/2023 07:30:09 AM, Cortez, John) Replaced LVP display
| 195984 | CM | CMPLT | CMPLT | 2023-07-12 08:09 | 2023-07-10 08:09 | cyber securoty meeting | (TC 07/10/2023 08:08:54 AM, Samanian, Kamran) done!
| 195985 | CM | CMPLT | CMPLT | 2023-07-12 08:10 | 2023-07-10 08:10 | huddle | (TC 07/10/2023 08:09:49 AM, Samanian, Kamran) done
1062277 | 195986 | CM | CMPLT | CMPLT | 2023-07-12 08:33 | 2023-07-12 08:39 | Error displayed and not weighing the patient | (TC 07/12/2023 08:37:25 AM, Ferreira, Qwynn) cleared the error checked the pins and cleaned them off with alcohol. Reset the bed and zeroed the scale. Error messaged gone and bed weighing again. Let the receptionist know that the bed is ready to go.
1051321 | 195987 | CM | CMPLT | CMPLT | 2023-07-12 08:58 | 2023-08-22 13:47 | Power plug connector damaged | (TC 07/12/2023 09:25:08 AM, Canjimavadivel, Kamallesh ) Plug connecter had a short with chassis and melted. Need to replace. Reached out to vendor for quote. Filed down chassis where damage occurred. (TC 07/13/2023 01:06:34 PM, Canjimavadivel, Kamallesh ) Received quote from vendor, emailed back to confirm as part might not be correct. (TC 07/17/2023 01:01:33 PM, Canjimavadivel, Kamallesh ) Placed order for part needed (TC 08/03/2023 08:18:56 AM, Canjimavadivel, Kamallesh ) Received part from Getinge, it is not the right part. Emailed back to the company get refund and correct part. (TC 08/10/2023 09:09:43 AM, Canjimavadivel, Kamallesh ) Reached out to Getinge again to follow up about refund. Ordered different part as provided by vendor. (TC 08/15/2023 01:12:36 PM, Canjimavadivel, Kamallesh ) Received shipping label from vendor, packaged and brought to receiving to be picked up by Purolator (TC 08/22/2023 01:45:48 PM, Canjimavadivel, Kamallesh ) Replaced power entry module. Secured power cord to chassis with ziptie. Plugged in to charge battery.
65949 | 195988 | CM | CMPLT | CMPLT | 2023-07-12 10:46 | 2023-07-12 10:48 | Barrel clamp stuck, does not go in "off" position | (TC 07/12/2023 10:46:17 AM, Baghirzade, Nara) When syringe pump was opened, noted that there were 3 screws missing on the bezel and PCB; the barrel clamp pin was loose and needed to be tightened, it was catching on a cable. Replaced the screws and tightened the barrel clamp pin. Barrel clamp moving to on and off positions properly.
65949 | 195989 | PM | CMPLT | CMPLT | 2023-07-12 10:48 | 2023-07-12 10:49 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/12/2023 10:48:24 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1067372 | 195990 | CM | CMPLT | CMPLT | 2023-07-12 11:34 | 2023-07-12 14:04 | PM | (TC 07/12/2023 01:58:28 PM, Scero, Tomlyn) All PM parts replaced
65862 | 195991 | CM | CMPLT | CMPLT | 2023-07-12 12:03 | 2024-04-03 10:00 | SAFER Report | (TC 05/17/2023 01:52:21 PM, Scero, Tomlyn) Received an email from Dianne regarding a safer report. Haven't recieved the actual safer report yet. From: Locke, Dianne Sent: May 17, 2023 2:21 PM To: Scero, Tomlyn <Tomlyn.Scero@sinaihealth.ca> Subject: RE: new ECG cart Safer file 56256 Hi Tomi, From the Safer file “Portable suction is kept on the crash cart, plugged when not in use and fully charged. During echo, portable suction malfunction in that it would not maintain constant suction for more then 10 seconds. Another portable suction unit was obtained from the 16 floor ICU, attached to the pt. Pt did not experience any desats/harm. Portable suction was cleaned, sent to biochem.” The asset tag is 65862 Biomed was copied on the SAFER. Please let me know who I should contact so that can close the file. Thanks for your help, Dianne (TC 04/03/2024 09:59:11 AM, Scero, Tomlyn) Closing this work order. The device was never found after it was removed from service.
| 195992 | CM | CMPLT | CMPLT | 2023-07-12 12:10 | 2023-07-12 12:11 | ENDO Room#4 , CO2 leaking from UCR | (TC 07/12/2023 12:10:24 PM, Samanian, Kamran) verified the problem, tube did not have the O-Ring, replaced with a new tube and resolved the issue
| 195993 | CM | CMPLT | CMPLT | 2023-07-12 12:12 | 2023-07-12 12:12 | ENDO Room#3 Removing ERBEs equipment from the BOOM to the cart plus Argon Tank | (TC 07/12/2023 12:12:28 PM, Samanian, Kamran) done!
1049903 | 195994 | CM | CMPLT | CMPLT | 2023-07-12 13:15 | 2023-07-12 13:15 | Damaged right outside cover | (TC 07/12/2023 01:14:55 PM, Fan, Tak) Replaced cover from stock.
1047842 | 195995 | CM | CMPLT | CMPLT | 2023-07-12 13:29 | 2023-07-12 13:30 | Date, time, and config update | (TC 07/12/2023 01:29:48 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1047859 | 195996 | CM | CMPLT | CMPLT | 2023-07-12 13:30 | 2023-07-12 13:31 | Date, time, and config update | (TC 07/12/2023 01:30:33 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1057033 | 195997 | CM | CMPLT | CMPLT | 2023-07-12 13:48 | 2023-09-13 10:27 | NIBP module defective | (TC 07/12/2023 01:47:56 PM, Baghirzade, Nara) NIBP starts inflation at 70mmHg (should be 0mmHg). Once it inflates, it does not deflate in step increments, instead it deflates and then re-inflates again. Calibrated zero and span; did not resolve the issue. Replaced cable and cuff; did not resolve the issue. Request RMA for service from Masimo. (TC 07/13/2023 09:59:18 AM, Baghirzade, Nara) SC1003001908 (TC 07/14/2023 09:33:02 AM, Baghirzade, Nara) Masimo RMA: 1036929 KJV waybill #380089 PO7202248513 (TC 08/18/2023 06:48:38 AM, Baghirzade, Nara) PO7202248513 Received (TC 08/18/2023 10:21:39 AM, Baghirzade, Nara) Tried to push the MSH VSM config using MICT; constantly receiving an "invalid profile setting" after rebooting the Root and the HR does not change to 45 bpm. Additionally, the Root software version is 2.1.7.0 US which is different than our fleet version, which is V 2.1.6.2 INT Gave the Root to John for further troubleshooting. (TC 08/21/2023 07:19:48 AM, Baghirzade, Nara) E-mailed Masimo tech service regarding not being able to push any configs and the different software version that was found on the replacement root. Requested a replacement or a tool to change the software version. (TC 08/22/2023 09:57:00 AM, Baghirzade, Nara) Received confirmation that we were sent a Root with the incorrect software version for our facility. RMA was created for a no-charge PO to send the Root back to Masimo. Masimo will send out a Root with the correct software version. (TC 09/13/2023 10:25:31 AM, Baghirzade, Nara) Replacement Root that was sent back still had the incorrect software version (US-NS). Used the USB with version: 2.1.4.9 and it was able to upgrade the Root to our current fleet version 2.1.4.9 INT. Acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1047864 | 195998 | PM | CMPLT | CMPLT | 2023-07-12 14:11 | 2023-07-12 14:13 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/12/2023 02:11:43 PM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
| 195999 | CM | CMPLT | CMPLT | 2023-07-12 14:16 | 2023-07-17 14:34 | Order iteams from Uline to Qwynn | (TC 07/12/2023 01:33:00 PM, Chukri, Samir) Replaced order for Qwynn from Uline. (TC 07/17/2023 02:33:33 PM, Chukri, Samir) received all.
1059150 | 196000 | CM | ACTIV | ISSUE | 2023-07-12 14:33 |  | X3 Battery Depleted | (TC 07/12/2023 02:33:39 PM, Reid, Emmerson ) Reported that the battery had failed and was alarming. I switched it with another battery and it is working. Brought the battery donw and trying to condition it.
1056665 | 196001 | CM | CMPLT | CMPLT | 2023-07-12 14:39 | 2024-03-28 08:46 | Leica Microscope ScaleOR power adaptor broken | (TC 07/12/2023 02:39:56 PM, Reid, Emmerson ) Replaced the power adaptor with another one that Jessica had. Working now. Looking to order a replacement.
1047867 | 196002 | CM | CMPLT | CMPLT | 2023-07-12 14:51 | 2023-07-12 14:52 | Time, date and config update | (TC 07/12/2023 02:51:20 PM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1046196 | 196003 | CM | CMPLT | CMPLT | 2023-07-12 16:05 | 2023-07-12 16:05 | X-ray tube Replacement - Tube B | X-ray tube B replaced Tune up done System passed all tests
1046035 | 196004 | CM | CMPLT | CMPLT | 2023-07-12 16:12 | 2023-08-24 08:54 | MSH ER EVO - Reported Coning Issues | (TC 07/12/2023 04:12:13 PM, Chau, Ken) Users reported coning issues -when set to 182cm WS, the collimation light does not go all the way out and is inside the square -require calibration and expansion of light, parts have been ordered (TC 07/13/2023 03:07:19 PM, Chau, Ken) Tried calibration to adjust but not working -cross section is off by .5"~ -will attempt again when room is free (TC 07/14/2023 12:10:51 PM, Chau, Ken) Tried calibration but did not succed -will require to replace collimator, ordering parts (TC 07/21/2023 10:09:16 AM, Chau, Ken) Repalced keypads and board
69106 | 196005 | FC | CMPLT | CMPLT | 2023-07-12 16:28 | 2023-07-12 16:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/12/2023 04:27:45 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1053391 | 196006 | FC | CMPLT | CMPLT | 2023-07-13 07:52 | 2023-07-13 07:53 | Functional check | (TC 07/13/2023 07:52:36 AM, Baghirzade, Nara) Physical inspection good. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
| 196007 | CM | CMPLT | CMPLT | 2023-07-13 07:59 | 2023-07-12 08:00 | ENDO Room#3 moving the shelves on the Boom a bit lower | (TC 07/12/2023 07:59:25 AM, Samanian, Kamran) per PCM rerquest , moved down the two shelves a bit lower so that staff would have better access to the Top Shelf
1066460 | 196008 | FC | CMPLT | CMPLT | 2023-07-13 08:01 | 2023-07-13 08:01 | QC checklist | (TC 07/13/2023 08:01:13 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
| 196009 | PA | CMPLT | CMPLT | 2023-07-13 08:59 | 2023-07-18 13:03 | Isopropyl 70% | (TC 07/13/2023 09:08:44 AM, Canjimavadivel, Kamallesh ) Placed order for isopropyl (TC 07/18/2023 01:02:09 PM, Canjimavadivel, Kamallesh ) Received stock, placed in cabinet | Department supplies
1048538 | 196010 | CM | CMPLT | CMPLT | 2023-07-13 09:52 | 2023-07-13 09:55 | Monitor Back Cover missing | (TC 07/13/2023 09:52:36 AM, Reid, Emmerson ) Amanda in OR5 told me that she found a monitor back cove and siad it did not belong to the room. I walked to where the mobile tower and found the monitor it belonged to and connected it.
1061206 | 196011 | CM | ACTIV | ISSUE | 2023-07-13 09:56 |  | Boom top cover falling off | (TC 07/13/2023 09:56:26 AM, Reid, Emmerson ) Top cover on the boom is losely hanging on. I am waiting for the case to end to fix it. (TC 07/14/2023 09:04:12 AM, Reid, Emmerson ) Went to Fix this and found that the adjacent boom, when moved, would brush against the cover. I put it back into position and it seems like it will stay despite the other boom, however I reached out to Jaron to find a better solution.
1056579 | 196012 | CM | CMPLT | CMPLT | 2023-07-13 10:51 | 2023-07-13 10:52 | Battery replacement | (TC 07/13/2023 10:51:20 AM, Baghirzade, Nara) Battery replaced. Unit is functional. Wi-fi check in completed.
67097 | 196013 | CM | CMPLT | CMPLT | 2023-07-13 11:01 | 2023-07-13 11:03 | C04 error | (TC 07/13/2023 11:01:31 AM, Cortez, John) Replaced female connector on BP hose issue resolved Supplied spare SPO2 sensor
66713 | 196014 | CM | CMPLT | CMPLT | 2023-07-13 11:25 | 2023-07-18 07:16 | Ordering parts | (TC 07/14/2023 10:20:09 AM, Rhuland, Courtney) Parts ordering (TC 07/18/2023 07:15:13 AM, Rhuland, Courtney) Parts delivered
69678 | 196015 | FC | CMPLT | CMPLT | 2023-07-13 11:31 | 2023-07-13 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/13/2023 11:32:22 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
| 196016 | CM | CMPLT | CMPLT | 2023-07-13 12:28 | 2023-07-17 14:32 | Order stryker Mattress from Stryker | (TC 07/13/2023 12:42:47 PM, Chukri, Samir) Ordered 5 mattress to L&D (TC 07/17/2023 02:32:27 PM, Chukri, Samir) Received One mattress and 4 BO 7-9-2023 (TC 11/23/2023 11:32:42 AM, Chukri, Samir) Received the BO 4 mattress emailed to Jonathan.
1064418 | 196017 | CM | CMPLT | CMPLT | 2023-07-13 12:43 | 2023-07-13 12:46 | Blender Repair for Alarm | (TC 07/13/2023 12:43:10 PM, Ferreira, Qwynn) Removed the blender from the MRI Incubator Removed alarm assembly Cleaned out dirt/debris/worn seals Regreased the o-rings in the alarm assembly to prevent the one sided alarms when adding gasses Tested with Samir and it was slghtly out of spec with the concentrations Adjusted the knob to provide accurate gas mix ratio. Tested at 21%, 30%, 40%, 50%, 60%, 70%, 80%, 90% and 100% Alarms are working now regardless which you add first, air or oxygen. (TC 07/13/2023 01:43:32 PM, Chukri, Samir) worked with qwynn to repair the blender
66713 | 196018 | CM | CMPLT | CMPLT | 2023-07-13 13:26 | 2023-07-13 13:27 | Latch recall | (TC 07/13/2023 01:26:38 PM, Baghirzade, Nara) Latch recall completed
| 196019 | CM | CMPLT | CMPLT | 2023-07-13 14:05 | 2023-08-28 08:53 | ordering the armboard | (TC 07/13/2023 08:53:26 AM, Huang, Jessica) received the part and sent it back to OR
66067 | 196020 | CM | CMPLT | CMPLT | 2023-07-13 14:33 | 2023-07-13 14:33 | Broken cases | (TC 07/13/2023 02:32:50 PM, Rhuland, Courtney) Replaced front and rear cases Replaced and conditioned battery Functional check - OK (TC 07/14/2023 10:53:39 AM, Rhuland, Courtney) PM completed PM schedule updated
531067 | 196021 | PA | CMPLT | CMPLT | 2023-07-13 14:36 | 2023-08-31 08:32 | Ordering parts for repair | (TC 07/13/2023 02:48:26 PM, Baghirzade, Nara) SC1003002502 (TC 08/31/2023 08:31:10 AM, Baghirzade, Nara) PO 7202248539
| 196022 | CM | ACTIV | ISSUE | 2023-07-13 14:45 |  | Co-op meeting | (TC 07/13/2023 02:45:27 PM, Chukri, Samir) Meeting with Mona from Centinial collage. (TC 08/02/2023 10:28:16 AM, Chukri, Samir) Arranged and schedule meeting with three Co-op students. (TC 08/08/2023 01:27:51 PM, Chukri, Samir) Co-op Student interview Abdul (TC 08/08/2023 01:28:35 PM, Chukri, Samir) Prepaired for zoom meeting interview with Kamran and Olcay . (TC 08/14/2023 02:56:27 PM, Chukri, Samir) completed coop student evaluation for courtney (TC 08/14/2023 02:57:18 PM, Chukri, Samir) emailed and signe a contract agremeent for coop student Conner Campbell. (TC 09/05/2023 01:11:11 PM, Chukri, Samir) working with new coop student Connor creat a network access and TMS sign in (TC 02/28/2024 12:31:10 PM, Chukri, Samir) zoom meeting and search coop student list. (TC 05/09/2024 09:19:31 AM, Chukri, Samir) completed a NERF for co-op student (TC 07/24/2024 01:09:55 PM, Chukri, Samir) Interview two student through zoom meeting
1066474 | 196023 | CM | CMPLT | CMPLT | 2023-07-13 14:48 | 2023-07-13 14:49 | From Loaner to Hospital Owned | (TC 07/13/2023 02:48:12 PM, Reid, Emmerson ) Unit is repalcing one of the neptune units that the vendor was not able to repair.
1062779 | 196024 | CM | CMPLT | CMPLT | 2023-07-13 14:50 | 2024-05-08 14:13 | PFT techs report a recurring error | (TC 07/13/2023 02:50:01 PM, Baghirzade, Nara) Gaby and other PFT techs have reported that there is a recurring error when closing Mortara and when closing the Expair application. Worked with John on troubleshooting this. We tried to re-create the issue but were unable to. Will touch base with PFT techs and sit in on one of the tests to see the sequence of actions that may cause the error to reoccur. We also checked Windows event viewer and noted one recurring error regarding the database but still unsure if it is linked to the ongoing intermittent error. (TC 07/27/2023 10:24:32 AM, Baghirzade, Nara) Recurring error after they try to close the patient test. John and I checked the event viewer for applications but were not able to correlate any errors to the error the techs are getting. Checked the logs on the Expair folder on C:/ and noted that it does log all the errors that they encounter. Ran the database integrity test and found 14 corrupted tests; could be another cause of the errors popping out. Will download the logs and send it to Medisoft for review. (TC 08/17/2023 02:23:06 PM, Baghirzade, Nara) Gaby and I deleted all 14 tests. Ran DB doctor again to ensure no tests are left. (TC 11/08/2023 08:48:18 AM, Baghirzade, Nara) Smain remoted to the CPET station: -looked through all user configurations -ran the paradox software to look for errors in the database (no errors or issues with the database) -looked through previous studies to see any discrepancies -created test patient/test study and found no errors during the test -upgraded the expair software to version 1.35.05 -waiting to see if the errors will reoccur during patient tests, upon closing the software, or during patient demographics input. (TC 12/21/2023 09:20:04 AM, Baghirzade, Nara) Gaby and other PFT techs have been reporting ongoing errors. The latest e-mail from Gaby regarding errors: "Hi Nara, I had a weird error come up when I was trying to save a 3 graph report for a CPET I did on today. I closed the system before it could send a bug report because it was taking a really long time( sorry I know that’s a no no). And when I went back to print my data was all off and my spirometry is now gone. Thank goodness I printed the test and had saved it on the usb. I thought it was weird that a whole section disappeared and the ht and weight was removed. I also had to go in and manually edit my VO2 and Max exercise as it was at the start of test.( usually the computer does this and we adjust it only if it looks off). Jen also had something strange happen on Wednesday. Her Ht and Wt was erased from her test and when she tried to save it as a pdf it would crash. I have cc’d her on this email so she can explain it more. I know we just updated the software, and we were hoping all the bugs would be fixed. I will take to you when I get back. Take care and have a good weekend. Gaby" Forwarded the e-mail to Antony from Novus and Smain from Medisoft; no response from either vendor or manufacturer. (TC 01/11/2024 10:25:55 AM, Baghirzade, Nara) Connected with Smain and Antony. Looking at past calibrations and current calibration, the O2 cell is starting to drift (which could cause a lower end tidal CO2/FiO2 value during patient tests). During next PM the O2 cell will be replaced. FiO2 is not fixed which is recommended and FiCO2 is in a good range. It is also possible that the patient was leaking. Asked Gaby to let me know the values of the next patient for FiO2 and FiCO2. For recurring errors: Smain updated the Expair software to the latest build 1.35.05 and turned the database from paradox to SQL, which should help mitigate the errors after patient tests. (TC 01/19/2024 07:25:23 AM, Baghirzade, Nara) After the change from the local to network and update to the databse from paradox to SQL; it had caused all the reports in the drop down menu to disappear. Connected with vendor and he walked me through reconfiguring the startup mode so that its not looking for a server. Removed the network ini file entirely and put it in a new folder (within the config folder). When booting up expair, the software is back on local mode. All reports are available in the drop down menu in final reports tab. When the study patient test was done on Wednesday, it was performed when expair was on "network" mode. When the techs tried to access reports, none of the reports were available. Gaby had copied the study patient file that was done on the "network" mode on an encrypted USB and restored the patient file when expair is in local mode. Study patient's report is now accessible. Will keep the network ini file on a different folder on the CPET computer until the system is ready to be put on the network. (TC 02/13/2024 11:27:14 AM, Baghirzade, Nara) A new issue occurred where the spirometry data disappeared after the patient test concluded. When Gaby went to print the report, expair froze and when the report finally loaded, the height and weight was gone off the report as well as all spirometry data. Another instance where Gaby went to print a report for a different patient and the height/weight was not on the report. Contacted Smain and Antony for support. (TC 02/14/2024 12:39:03 PM, Baghirzade, Nara) Downloaded and saved all logs from Jan 2024-Feb 2024. Found the patient where the spiro data was gone and confirmed with Gaby. Looked through the log to pinpoint the exact time of error occurrence however could not find the any errors on the log. Will send the logs to Smain and Antony for further assessment. (TC 02/22/2024 02:14:17 PM, Baghirzade, Nara) Smain remoted in to the CPET system. First we took a look at the patient where the spirometry data disappeared; he was able to retrieve the raw data but will take some time to manipulate the data in to the report. He will be remoting in on another day as it is a lengthy process. Second we looked at recurring error. I reported the errors I pinpointed during a normal protocol and then we looked at a study protocol; he suspects that the Mortara software is having issues communicating with the Expair software. He exported both the Expair logs (both normal and error) and Windows error logs to send to the Medisoft R&D team to mitigate the issue. Updated Gaby on what was done today. (TC 03/08/2024 02:11:06 PM, Baghirzade, Nara) After Gaby completed patient test and wanted to export the test to her USB, Expair froze and after Expair recovered, the patient pre+post spirometry data disappeared. I was able to locate the raw data however Medisoft will need to recover the data and place it back on the Expair software. Followed up with Adeel on this. Gaby to follow up with Steve. Sending email to Medisoft and Novus regarding this incident. (TC 03/08/2024 02:42:20 PM, Baghirzade, Nara) Sent email to Novus & Medisoft outlining the incident and included 2 pictures for reference. (TC 04/17/2024 10:54:02 AM, Baghirzade, Nara) Smain and his medisoft colleague remoted in to investigate what causes the constant expair errors and data loss. They found that there may be a communication error between the software and the encrypted USB that the techs use to save the reports. To rule out if this is the issue, they are asked to save the reports on the desktop or anywhere locally first, and then to copy that report on to the USB. If the issue still continues, then Medisoft will start looking elsewhere. Smain also copied the latest logs for further investigation. (TC 04/17/2024 10:55:30 AM, Baghirzade, Nara) Met with Gaby and Jen to update them on the information received from Medisoft and to change their workflow regarding report saving and printing. (TC 05/02/2024 11:03:18 AM, Baghirzade, Nara) Adjusted language settings for the windows lock screen as it kept reverting to french and causing issues for the techs at login. Checked several times and french is no longer an option on the lock screen. Updated Gaby and asked her to keep an eye on the keyboard language and if french appears as an option again. (TC 05/08/2024 02:13:25 PM, Baghirzade, Nara) Met with Gaby and Jen to discuss how CPET has been since the implementation of a new workflow. No issues or problems reported since the new workflow. CPET running smoothly since then.
1066475 | 196025 | CM | CMPLT | CMPLT | 2023-07-13 14:52 | 2023-07-13 14:55 | Change from Loaner to Hospital Owned | (TC 07/13/2023 02:54:47 PM, Reid, Emmerson ) Unit is repalcing one of the neptune units that the vendor was not able to repair.
1066476 | 196026 | CM | CMPLT | CMPLT | 2023-07-13 14:57 | 2023-07-13 14:57 | Changed from Loaner to Hospital Owned | (TC 07/13/2023 02:56:45 PM, Reid, Emmerson ) Unit is repalcing one of the neptune units that the vendor was not able to repair.
1066920 | 196027 | CM | CMPLT | CMPLT | 2023-07-13 15:04 | 2024-08-05 10:36 | Right side rail doesn't lock in upright position | (TC 07/13/2023 03:04:21 PM, Cortez, John) Confirmed issue with side rail, placed service request with Stryker (TC 08/03/2023 03:34:41 PM, Cortez, John) Received call from ICU regarding status of repair. Followed up with stryker (TC 08/05/2024 10:36:01 AM, Cortez, John) Repaired by Mark, report attached
| 196028 | CM | CMPLT | CMPLT | 2023-07-13 17:04 | 2023-07-27 14:49 | order for PCU batteries, biomed stock | (TC 07/13/2023 05:03:53 PM, Cortez, John) Ordered PCU batteries (TC 07/27/2023 02:48:28 PM, Cortez, John) order arrived
65639 | 196029 | PM | CMPLT | CMPLT | 2023-07-14 08:03 | 2023-07-12 08:04 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Calibrated and tested. Functional. SW.2.51 Leakage: 79uA and 87 ohm.
66461 | 196030 | PM | CMPLT | CMPLT | 2023-07-14 08:08 | 2023-07-13 08:08 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries, and left housing door. Tested. Functional. SW.2.51 Leakage: 71uA and 85 ohm.
| 196031 | CM | CMPLT | CMPLT | 2023-07-14 08:14 | 2023-07-14 08:14 | RT meeting with Rex | RT meeting with Rex
1061211 | 196032 | CM | CMPLT | CMPLT | 2023-07-14 09:10 | 2023-07-14 09:12 | Installed Monitor Protector | (TC 07/14/2023 09:10:16 AM, Reid, Emmerson ) Installed the monitor protector in OR10. Note the monitor had a number of scratches already.
1064495 | 196033 | CM | CMPLT | CMPLT | 2023-07-14 09:13 | 2023-07-14 09:14 | Defib alert | (TC 07/14/2023 09:13:14 AM, Reid, Emmerson ) Defib was not plugged in. I plugged it in.
1063000 | 196034 | PM | CMPLT | CMPLT | 2023-07-14 09:17 | 2023-07-14 11:13 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/14/2023 11:10:07 AM, Rhuland, Courtney) When completing cleaning cycle, water was leaking from the bottom Opened tray and found the bottom of the water reservoir was cracked and O-ring was loose Replaced water reservoir with spare and issue was resolved Drained, disinfected, rinsed and refilled with sterile water
1047844 | 196035 | CM | CMPLT | CMPLT | 2023-07-14 10:07 | 2023-07-14 10:07 | Time, date and config update | (TC 07/14/2023 10:06:59 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019.
1049218 | 196036 | CM | CMPLT | CMPLT | 2023-07-14 10:14 | 2023-07-14 10:15 | Patch update and restart | (TC 07/13/2023 10:14:20 AM, Scero, Tomlyn) Patch complete and update installed. Went to CCU and restarted the central station. complete
1058181 | 196037 | CM | CMPLT | CMPLT | 2023-07-14 10:17 | 2023-07-14 10:17 | Pushed APS EPIDURAL - 2023 JULY library | (TC 07/14/2023 10:16:45 AM, Reid, Emmerson ) Pushed APS EPIDURAL - 2023 JULY library
67760 | 196038 | CM | CMPLT | CMPLT | 2023-07-14 10:18 | 2023-07-14 10:18 | Pushed APS EPIDURAL - 2023 JULY library | (TC 07/14/2023 10:17:41 AM, Reid, Emmerson ) Pushed APS EPIDURAL - 2023 JULY library
1057547 | 196039 | CM | CMPLT | CMPLT | 2023-07-14 10:18 | 2023-07-20 07:51 | Temp not working | (TC 07/14/2023 10:43:24 AM, Scero, Tomlyn) Confirmed temperature malfunction with our temptest equipment. Contact Philips to order a replacement temp module. Case created with Philips 0121988048 (TC 07/18/2023 08:21:07 AM, Scero, Tomlyn) Received an email back from Philips. Confirmed which part I need and they gave me a quote. Now ordering the part from purchasing. (TC 07/19/2023 07:48:42 AM, Scero, Tomlyn) Temp board recieved and replaced. battery also replaced as it was malfunctioning too. Temperature test with test cable passed. Returned to service.
1058171 | 196040 | CM | CMPLT | CMPLT | 2023-07-14 10:18 | 2023-07-14 10:19 | Pushed APS EPIDURAL - 2023 JULY library | (TC 07/14/2023 10:18:41 AM, Reid, Emmerson ) Pushed APS EPIDURAL - 2023 JULY library
1058176 | 196041 | CM | CMPLT | CMPLT | 2023-07-14 10:20 | 2023-07-14 10:20 | Pushed APS EPIDURAL - 2023 JULY library | (TC 07/14/2023 10:20:00 AM, Reid, Emmerson ) Pushed APS EPIDURAL - 2023 JULY library
66713 | 196042 | CM | CMPLT | CMPLT | 2023-07-14 10:23 | 2023-07-14 10:24 | Incubator light | (TC 07/14/2023 10:23:22 AM, Baghirzade, Nara) Replaced old incubator light with new light. Light is functional.
| 196043 | CM | CMPLT | CMPLT | 2023-07-14 10:53 | 2023-07-14 10:54 | ENDO room#1 shelves on the BOOM lowered down | (TC 07/14/2023 10:53:33 AM, Samanian, Kamran) done, clables connected , tested and passed
| 196044 | CM | CMPLT | CMPLT | 2023-07-14 10:55 | 2023-07-14 10:56 | philips switch port mapping, ENDO room#1, | (TC 07/14/2023 10:55:35 AM, Samanian, Kamran) verified the data drop 6SE-241-D , patched and mapped to port Gi1/0/11 in ENDO switch , configured the PIciX, moved the monitor and plugged in to the jack, tested and all passed
65990 | 196045 | PM | CMPLT | CMPLT | 2023-07-14 11:43 | 2023-07-14 11:43 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/14/2023 11:42:47 AM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
501070 | 196046 | CM | CMPLT | CMPLT | 2023-07-14 12:13 | 2023-07-18 10:33 | MSH CEOU RM#10 - No Display | (TC 07/14/2023 12:12:45 PM, Chau, Ken) when booting up, main display and contorl panel display shows nothing. No power on control panel or keyboards. Power on/off button does not function. -requires an EBOX replacement -case#0121988229 (TC 07/17/2023 01:50:28 PM, Fan, Tak) Brought UMB to CEOU. UMB replaced and SW reloaded. Applied for MAC address whitelisting. (TC 07/18/2023 10:31:19 AM, Fan, Tak) Whitelisted. Verified worklist to enable it. Calculations were missing. Loaded from another room. OK. | case#0121988229
1062401 | 196047 | CM | CMPLT | CMPLT | 2023-07-14 12:14 | 2023-07-14 12:17 | Motion controls do not work | (TC 07/14/2023 12:14:28 PM, Baghirzade, Nara) Motion controls were locked out. Unlocked all motion controls. Checked that the power cord was plugged in. Showed Sara (nurse) and Sharmaine (PCM) how to lock and unlock the motion controls.
52230 | 196048 | PM | CMPLT | CMPLT | 2023-07-14 13:04 | 2023-07-14 13:05 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/14/2023 01:04:59 PM, Ferreira, Qwynn) PM Passed Latch recall completed Replaced port hole grommet Fixed pin in the back of the scale Calibrated Linevoltage Calibrated main voltage Calibrated system voltage Calibrated scale Electrical Safety - OK Functional Test - OK
1047866 | 196049 | PM | CMPLT | CMPLT | 2023-07-14 14:17 | 2023-07-14 14:18 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/14/2023 02:16:48 PM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
| 196050 | CM | CMPLT | CMPLT | 2023-07-14 14:19 | 2023-07-14 14:20 | Meeting with Henna regarding devices at clinic | (TC 07/14/2023 02:18:48 PM, Baghirzade, Nara) Meeting with Henna regarding: -crash cart -requirement for VSMs -repairs -old equipment found
1061219 | 196051 | CM | CMPLT | CMPLT | 2023-07-14 14:19 | 2023-07-14 14:19 | Installed Monitor Protector | (TC 07/14/2023 02:19:02 PM, Reid, Emmerson ) Installed Monitor Protector
1061218 | 196052 | CM | CMPLT | CMPLT | 2023-07-14 14:20 | 2023-07-14 14:20 | Installed Monitor Protector | (TC 07/14/2023 02:19:50 PM, Reid, Emmerson ) Installed Monitor Protector
1061217 | 196053 | CM | CMPLT | CMPLT | 2023-07-14 14:20 | 2023-07-14 14:21 | Installed Monitor Protector | (TC 07/14/2023 02:20:35 PM, Reid, Emmerson ) Installed Monitor Protector
1061209 | 196054 | CM | CMPLT | CMPLT | 2023-07-14 14:21 | 2023-07-14 14:22 | Installed Monitor Protector | (TC 07/14/2023 02:21:35 PM, Reid, Emmerson ) Installed Monitor Protector
1061206 | 196055 | CM | CMPLT | CMPLT | 2023-07-14 14:22 | 2023-07-14 14:23 | Installed Monitor Protector | (TC 07/14/2023 02:22:15 PM, Reid, Emmerson ) Installed Monitor Protector
1059596 | 196056 | CM | CMPLT | CMPLT | 2023-07-14 14:41 | 2023-07-14 14:41 | Not Working | (TC 07/14/2023 02:40:43 PM, Reid, Emmerson ) Ran Calibration, passed on first try. Working
| 196057 | CM | CMPLT | CMPLT | 2023-07-14 14:41 | 2023-07-14 14:41 | Masimo Safey Net | Checked Masimo Safey Net at 16 North. Restarted both terminals. Working fine.
| 196058 | CM | ACTIV | ISSUE | 2023-07-14 14:54 |  | ON-Call Daily Checked | (TC 07/14/2023 02:53:39 PM, Chukri, Samir) Daily checked for Patient monitor, masimo and Defibs. Call olcay to check safetyNetView 16N View1 and view2 clinent down Olcay turned on the checked working og but still showing 16N1 clinet down . we will check with Masimo on Mondy next week. (TC 07/17/2023 09:23:44 AM, Chukri, Samir) Completed and emailed the on-call daily checks (TC 07/19/2023 10:45:01 AM, Chukri, Samir) completed and emailed On-call Daily checks report. (TC 07/20/2023 08:32:09 AM, Chukri, Samir) On-Call daily checks report (TC 08/25/2023 02:53:23 PM, Chukri, Samir) completed the on-call daily report and emailed it. (TC 08/28/2023 11:24:29 AM, Chukri, Samir) completed the On-Call daily report. (TC 08/29/2023 08:50:10 AM, Chukri, Samir) completed the daily on-call for Aug 29 (TC 12/15/2023 02:55:59 PM, Chukri, Samir) Performed the On-Call daily report (TC 12/18/2023 10:30:13 AM, Chukri, Samir) completed the On-Call report for today. (TC 01/02/2024 11:46:50 AM, Chukri, Samir) performed the on-Call daily reports and contact Olcay to solve the Masimo turned off on16 N (TC 01/04/2024 07:55:01 AM, Chukri, Samir) On-Call Daily report . (TC 02/23/2024 07:59:58 AM, Chukri, Samir) On-Call daily report and email to Biomed (TC 02/26/2024 08:28:39 AM, Chukri, Samir) completed and emailed the ON-Call daily report (TC 02/28/2024 08:54:13 AM, Chukri, Samir) completed the ON-Call daily report and emailed it. (TC 03/06/2024 09:36:10 AM, Chukri, Samir) Performed On Call daily report and emailed it (TC 05/13/2024 09:18:44 AM, Chukri, Samir) completed the On-Cal daily check up.
1058586 | 196059 | CM | CMPLT | CMPLT | 2023-07-14 15:12 | 2023-07-18 10:20 | Thermometer Reading too high | (TC 07/14/2023 03:13:49 PM, Reid, Emmerson ) Hardware "error 8" is appearing every time I tried on the calibration checker. Will try again after the weekend. (TC 07/18/2023 10:19:54 AM, Reid, Emmerson ) Ran on the calibration checker and passed. Working now.
| 196060 | CM | CMPLT | CMPLT | 2023-07-14 16:37 | 2023-07-14 16:37 | Site Coverage.
| 196061 | RDS | CMPLT | CMPLT | 2023-07-17 07:32 | 2023-07-17 09:11 | Weekly Checks | (TC 07/17/2023 09:10:25 AM, Ferreira, Qwynn) weekly rounds
| 196062 | RDS | CMPLT | CMPLT | 2023-07-17 07:50 | 2023-07-17 08:27 | Weekly check for ED
| 196063 | CM | CMPLT | CMPLT | 2023-07-17 07:52 | 2023-07-18 15:25 | Sterile Water | for cleaning nicu altrix
| 196064 | OC | CMPLT | CMPLT | 2023-07-17 07:54 | 2023-07-16 11:24 | Samir Chukri On-Call; July 14 - July 20, 2023 | (TC 07/16/2023 11:30:00 AM, Chukri, Samir) Received a called from IT ( Problem in endo room#2 Video Capture) contact Olga from Endo, and contact Kamran and called Olga found that this Video Capture is connected with a green (ok) and can see patient ditals,. so it is Cerner issue. it is not argent and we promissed to look after this issue on Monday morning. | Received called from IT regarding Endo Video capture in room# 2 not connect .
69064 | 196065 | CM | CMPLT | NFF | 2023-07-17 08:20 | 2023-07-17 08:21 | Malfunctioning | (TC 07/17/2023 08:19:46 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. No fault found.
| 196066 | RDS | CMPLT | CMPLT | 2023-07-17 08:30 | 2023-07-25 15:44 | Weekly check for 10S - July 17 | (TC 07/17/2023 09:53:17 AM, Cortez, John) Checked hallways/equipment room. Patient rooms busy. Nothing to highlight
1047856 | 196067 | CM | CMPLT | CMPLT | 2023-07-17 08:41 | 2023-07-17 08:43 | Not working, states "biomed" message | (TC 07/17/2023 08:41:24 AM, Baghirzade, Nara) Primary Audible Alarm Post error upon booting. As per service manual, I opened the pump to check the connectors to the speakers; all looks fine. Went into diagnostics menu and performed speakers test; all tones and volume levels are good. Ran multiple feeds at various rates; no issues with alarms or speaker.
67895 | 196068 | CM | CMPLT | NFF | 2023-07-17 09:02 | 2023-07-17 09:03 | Keeps beeping/alarming | (TC 07/04/2023 09:02:29 AM, Rendon Lopez , Elizabeth) Testing continuously in all modes. NFF
69479 | 196069 | CM | CMPLT | CMPLT | 2023-07-17 09:04 | 2023-07-17 09:06 | Will not run oxytocin | (TC 07/04/2023 09:03:55 AM, Rendon Lopez , Elizabeth) Downloaded error log, Error 240.4150.0 - Faulty bottle pressure sensor. Replaced pressure sensor, calibrated, tested-OK
| 196070 | RDS | CMPLT | CMPLT | 2023-07-17 09:17 | 2023-07-17 09:20 | Weekly check for ICU
| 196071 | CM | CMPLT | CMPLT | 2023-07-17 10:06 | 2023-07-17 10:08 | Room 11 arm slipped down | (TC 07/17/2023 10:06:29 AM, Canjimavadivel, Kamallesh ) Found arm had slipped down column of boom during rounds. Removed damaged layer of anti slip pad and reinstalled arm at correct height.
| 196072 | RDS | CMPLT | CMPLT | 2023-07-17 10:55 | 2023-07-18 11:47 | Weekly check for CCU | (TC 07/18/2023 11:45:15 AM, Scero, Tomlyn) Weekly checks for CCU done. Bed locators not plugged in, outlets installed behind beds aren't working yet. Dianne is away, but I'll speak with her on her consumables ordering process when she returns.
| 196073 | CM | CMPLT | CMPLT | 2023-07-17 10:59 | 2023-07-17 10:59 | Biomed meeting | Biomed meeting
1056125 | 196074 | PM | CMPLT | CMPLT | 2023-07-17 11:12 | 2023-07-17 11:12 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/17/2023 11:11:53 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 196075 | CM | CMPLT | CMPLT | 2023-07-17 12:06 | 2023-07-17 12:07 | TGH, feedback on MR400 and MR5000 Portal | (TC 07/17/2023 12:07:02 PM, Samanian, Kamran) done!
1065418 | 196076 | CM | CMPLT | CMPLT | 2023-07-17 12:08 | 2023-07-17 12:12 | not working | (TC 07/15/2023 12:08:17 PM, Samanian, Kamran) on SAT JULY 15, it seems images were not exported to Cerner. Called Olga and she confirmed patinet list is shown in the device, that means network connection is good (TC 07/17/2023 12:10:51 PM, Samanian, Kamran) double checked all connection and networking and all is good. device was working on Friday patients but not SAT. Also checked the Hyland server and there is no error file. All seems to be working fine
| 196077 | CM | CMPLT | CMPLT | 2023-07-17 12:23 | 2023-07-17 12:23 | Capital meeting | (TC 07/17/2023 12:23:02 PM, Baghirzade, Nara) Capital meeting
67957 | 196078 | CM | CMPLT | CMPLT | 2023-07-17 12:36 | 2023-07-17 12:38 | Does not tilt | (TC 07/17/2023 12:36:17 PM, Baghirzade, Nara) Bed does not lock in to tilt ball properly and bed tilts without having to use the latch. Found the brake pad to be very worn out. Replaced the brake pad. Greased the tilt. Tilt is fully functional.
67957 | 196079 | CM | CMPLT | CMPLT | 2023-07-17 12:38 | 2023-07-17 12:39 | Latch recall | (TC 07/17/2023 12:38:32 PM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
| 196080 | RDS | ACTIV | ISSUE | 2023-07-17 12:39 |  | Weekly Check for ENDO | (TC 07/17/2023 12:39:15 PM, Samanian, Kamran) All looks good, no issue or concern. Image Capture in Room 2 checked that has its own WO (TC 07/24/2023 10:32:27 AM, Samanian, Kamran) nothing specific- All looks good (TC 07/31/2023 10:32:55 AM, Samanian, Kamran) found no issue or concern (TC 08/28/2023 04:25:59 PM, Samanian, Kamran) all is good (TC 09/06/2023 09:42:01 AM, Samanian, Kamran) done, no issue or concern (TC 09/11/2023 02:37:27 PM, Samanian, Kamran) All looks good. ENDO was returned back to the floor after Grey Code cleared (TC 09/19/2023 10:00:12 AM, Samanian, Kamran) nothing specific, all looks good
| 196081 | CM | CMPLT | CMPLT | 2023-07-17 14:03 | 2023-07-17 14:03 | Biomed meeting | Biomed meeting
| 196082 | CM | CMPLT | CMPLT | 2023-07-17 14:05 | 2023-07-17 14:06 | PM Huddle | (TC 07/17/2023 02:05:35 PM, Baghirzade, Nara) PM huddle
| 196083 | CM | CMPLT | CMPLT | 2023-07-17 14:18 | 2023-07-17 14:18 | meeting huddle | (TC 07/17/2023 02:18:18 PM, Samanian, Kamran) done!
| 196084 | CM | CMPLT | CMPLT | 2023-07-17 14:25 | 2023-07-17 14:26 | meeting | (TC 07/17/2023 11:00:00 AM, Rhuland, Courtney) Bi-weekly meeting
| 196085 | CM | CMPLT | CMPLT | 2023-07-17 14:27 | 2023-07-17 14:28 | Meeting | (TC 07/17/2023 02:00:00 PM, Rhuland, Courtney) Huddle
| 196086 | CM | ACTIV | ISSUE | 2023-07-17 14:31 |  | Bi-Weekly Department Meeting: Fleet Managment. | (TC 07/17/2023 02:31:07 PM, Chukri, Samir) completed. (TC 11/09/2023 09:37:54 AM, Chukri, Samir) Biomedical Engineering Planning (Bi-Weekly): CALM, ICAT, Compliance, Operations (TC 02/15/2024 09:41:39 AM, Chukri, Samir) Biomedical Engineering Planning (Bi-Weekly): CALM, ICAT, Compliance, Operations
| 196087 | CM | CMPLT | CMPLT | 2023-07-17 14:36 | 2023-07-17 14:36 | Flowmeter repair on Optiflows | (TC 07/17/2023 02:35:45 PM, Scero, Tomlyn) repair both flowmeters on the optiflows IV poles
| 196088 | CM | CMPLT | CMPLT | 2023-07-17 17:25 | 2023-07-17 17:25 | Site coverage.
65965 | 196089 | PM | CMPLT | CMPLT | 2023-07-18 07:17 | 2023-07-18 07:17 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/18/2023 07:16:42 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1046035 | 196090 | CM | CMPLT | CMPLT | 2023-07-18 08:06 | 2023-07-18 08:07 | MSH ER DRX EVO E5-02(ER alignment | (TC 07/18/2023 08:06:22 AM, Whittington, Scott) MSH ER DRX EVO E5-02(ER alignment (TC 07/18/2023 08:06:22 AM, Whittington, Scott) MSH ER DRX EVO E5-02(ER alignment | MSH ER DRX EVO E5-02(ER alignment
| 196091 | RDS | CMPLT | CMPLT | 2023-07-18 08:15 | 2023-07-17 08:23 | Weekly Check for 9S | (TC 07/17/2023 08:23:05 AM, Samanian, Kamran) done!
1063187 | 196092 | CM | CMPLT | CMPLT | 2023-07-18 08:28 | 2023-08-01 09:48 | Not picking up FHR | (TC 07/18/2023 08:28:28 AM, Rendon Lopez , Elizabeth) Confirmed issue, device under repair warranty. To be sent out (TC 07/28/2023 09:47:12 AM, Rendon Lopez , Elizabeth) Broken beyond repair. Retired from service
66533 | 196093 | CM | CMPLT | CMPLT | 2023-07-18 08:31 | 2023-07-20 09:49 | Recorder not printing FHR | (TC 07/17/2023 08:33:29 AM, Rendon Lopez , Elizabeth) Removed from room. Cleaned recorder printerhead, printing properly. Running on demo "check paper" alarm keeps showing, paper sensor is defective. Replaced recorder with spare, ordering exchange.
62792 | 196094 | PM | CMPLT | CMPLT | 2023-07-18 08:35 | 2023-07-18 08:36 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/18/2023 08:35:08 AM, Reid, Emmerson ) Ran and passed the functional test. PM Complete.
62633 | 196095 | CM | CMPLT | NFF | 2023-07-18 08:40 | 2023-07-18 08:41 | Warmer in OR5 not working | (TC 07/17/2023 08:39:55 AM, Rendon Lopez , Elizabeth) Checked all functions, no fault found, no errors in log. Run for a while, no issues.
67230 | 196096 | CM | CMPLT | CMPLT | 2023-07-18 08:41 | 2023-07-18 08:41 | Warmer in OR5 not working | (TC 07/17/2023 08:41:05 AM, Rendon Lopez , Elizabeth) Checked all functions, no fault found, no errors in log. Run for a while, no issues.
1051309 | 196097 | CM | CMPLT | CMPLT | 2023-07-18 08:42 | 2023-07-18 08:43 | not charging | (TC 07/17/2023 08:42:30 AM, Rendon Lopez , Elizabeth) Cleaned contacts, charged, tested-OK
| 196098 | CM | CMPLT | CMPLT | 2023-07-18 08:50 | 2023-07-18 08:51 | X3 battery charge back | (TC 07/18/2023 08:49:59 AM, Canjimavadivel, Kamallesh ) Replaced X3 battery from stock
62716 | 196099 | PM | CMPLT | CMPLT | 2023-07-18 08:52 | 2023-07-18 08:53 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/18/2023 08:52:21 AM, Reid, Emmerson ) Ran and passed the functional test. PM Complete.
1050618 | 196100 | PM | CMPLT | CMPLT | 2023-07-18 09:29 | 2023-07-18 09:30 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/18/2023 09:29:27 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1061205 | 196101 | CM | CMPLT | CMPLT | 2023-07-18 09:57 | 2023-07-18 10:00 | VMC touch screen frozen | (TC 07/18/2023 09:57:21 AM, Reid, Emmerson ) OR4 called that the VMC monitor was frozen. I went to check it out and it was frozen. I unplugged the power and restarted the monitor but this did not help, so i reseated the data cable and that worked.
1059746 | 196102 | CM | CMPLT | CMPLT | 2023-07-18 10:08 | 2023-07-20 14:49 | Unable to identify HR | (TC 07/18/2023 11:34:10 AM, Rendon Lopez , Elizabeth) Confirmed problem, weak heart rate signal, replaced probe with spare, working properly
1058584 | 196103 | CM | CMPLT | CMPLT | 2023-07-18 10:15 | 2023-07-18 10:16 | Inaccurate | (TC 07/18/2023 10:15:22 AM, Reid, Emmerson ) Sent down that it wasn't reading properly. Ran it on the calibration checker and passed first time.
| 196104 | CM | CMPLT | CMPLT | 2023-07-18 10:36 | 2023-07-18 10:36 | EMG server implementation pre-vendor meeting | (TC 07/18/2023 10:35:58 AM, Baghirzade, Nara) EMG server implementation pre-vendor meeting
| 196105 | PA | CMPLT | CMPLT | 2023-07-18 11:03 | 2023-07-26 08:48 | Incubator Parts
1054426 | 196106 | PM | CMPLT | CMPLT | 2023-07-18 11:08 | 2023-07-18 11:09 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/18/2023 11:08:17 AM, Canjimavadivel, Kamallesh ) PM tests pass
65865 | 196107 | PM | CMPLT | CMPLT | 2023-07-18 11:12 | 2023-07-18 11:13 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/18/2023 11:12:22 AM, Canjimavadivel, Kamallesh ) PM tests pass
1053208 | 196108 | PM | CMPLT | CMPLT | 2023-07-18 11:15 | 2023-07-18 11:16 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/18/2023 11:14:51 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
| 196109 | CM | CMPLT | CMPLT | 2023-07-18 11:22 | 2023-07-18 11:23 | Ground pin on crash cart power bar broken | (TC 07/18/2023 11:22:18 AM, Canjimavadivel, Kamallesh ) Found ground pin missing on power bar. Got replacement from biomed and installed onto crash cart.
1059272 | 196110 | CM | CMPLT | CMPLT | 2023-07-18 11:24 | 2023-08-18 10:24 | Pins on probe connector damaged | (TC 07/18/2023 11:24:09 AM, Canjimavadivel, Kamallesh ) Swapped out probe from other ED bladderscanner that is down for repair. Requested quote from vendor for probe. (TC 08/01/2023 02:18:49 PM, Canjimavadivel, Kamallesh ) Vendor provided case number and shipping instructions. Shipped both probe and console voa KJV. (TC 08/18/2023 10:21:21 AM, Canjimavadivel, Kamallesh ) Received probe and console back from repair. See attached service report for details. Assembled console back onto stand and sent to unit
| 196111 | CM | CMPLT | CMPLT | 2023-07-18 11:29 | 2023-07-18 11:29 | pre-meeting internally for EMG | (TC 07/18/2023 11:29:17 AM, Samanian, Kamran) done!
| 196112 | CM | CMPLT | CMPLT | 2023-07-18 11:37 | 2023-07-19 10:31 | Case# 0122002099 Philips Tech Support on FP | (TC 07/18/2023 11:37:04 AM, Samanian, Kamran) opened the case to philips to discuss few things about FP such as existing deleted units, updated OS however RED indicator in FP, ,,, (TC 07/19/2023 10:30:05 AM, Samanian, Kamran) Philips Tech support connected to the server and purged the old units. WSUS updated and server patched and rebooted. all looks good and server in normal condition
1062094 | 196113 | PM | CMPLT | CMPLT | 2023-07-18 12:17 | 2023-07-18 12:17 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 2nd year PM after 1st year PM last year. No GS500. Replaced air filter, V Battery and PS500 batteries. Tested. Functional. SW.2.51 Leakage: 77uA and 81 ohm.
1057537 | 196114 | CM | CMPLT | CMPLT | 2023-07-18 12:36 | 2023-07-18 12:39 | Battery not charging | (TC 07/18/2023 12:37:16 PM, Scero, Tomlyn) Battery malfunction confirmed, unable to recondition the battery. After batteries were ordered and received I replaced it in the unit and the device is fully functional.
66242 | 196115 | CM | CMPLT | CMPLT | 2023-07-18 12:52 | 2023-07-14 12:52 | BAT not working | (TC 07/18/2023 12:52:00 PM, Samanian, Kamran) replaced the battery, charged it and seems to be workinbg fine now, device was returned to the unit. asked to work it to see how it goes,
1047846 | 196116 | PM | CMPLT | CMPLT | 2023-07-18 13:50 | 2023-07-18 13:53 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/18/2023 01:50:32 PM, Baghirzade, Nara) PM completed. See procedure document for PM results. PM schedule updated.
| 196117 | RDS | CMPLT | CMPLT | 2023-07-18 13:54 | 2023-07-18 13:58 | Weekly round | (TC 07/18/2023 01:57:20 PM, Baghirzade, Nara) See procedure for weekly round results.
| 196118 | RDS | CMPLT | CMPLT | 2023-07-18 13:58 | 2023-07-18 14:00 | Weekly round | (TC 07/18/2023 01:58:32 PM, Baghirzade, Nara) See procedure for weekly round results.
| 196119 | RDS | CMPLT | CMPLT | 2023-07-18 14:03 | 2023-07-18 14:05 | Weekly round | (TC 07/18/2023 02:04:42 PM, Baghirzade, Nara) See procedure for weekly round results.
1057724 | 196120 | CM | CMPLT | CMPLT | 2023-07-18 14:19 | 2023-07-19 10:59 | Battery calibration recommended message | (TC 07/18/2023 03:16:20 PM, Canjimavadivel, Kamallesh ) Swapped out battery and started calibration (TC 07/19/2023 10:57:03 AM, Canjimavadivel, Kamallesh ) Conditioning complete, capacoity 3/5 LEDs
| 196121 | RDS | CMPLT | CMPLT | 2023-07-18 14:39 | 2023-07-18 14:51 | Weekly check for PACU | (TC 07/18/2023 02:50:11 PM, Reid, Emmerson ) Completed proceedure, all passed. Also turned on several MX800s and put on sandby
| 196122 | RDS | CMPLT | CMPLT | 2023-07-18 14:40 | 2023-07-18 14:55 | Weekly check for OR | (TC 07/18/2023 02:54:08 PM, Reid, Emmerson ) Proceedure complete. Not everything is plugged in that should be but there are not enough plugs for all the pumps
| 196123 | RDS | CANCL | CANCL | 2023-07-18 14:40 |  | Weekly check for PAU
1066409 | 196124 | CM | CMPLT | CMPLT | 2023-07-18 15:07 | 2024-02-13 10:10 | "not inflating" | (TC 07/18/2023 03:06:59 PM, Cortez, John) Functional test passed, unable to replicate issue
1054963 | 196125 | CM | CMPLT | CMPLT | 2023-07-18 15:26 | 2023-07-25 16:26 | "BP not accurate" | (TC 07/18/2023 03:25:57 PM, Cortez, John) air leak, calibration, and overpressure tests all passed in service mode Missing adult cuff, spo2 extension cable, and power cord (TC 07/25/2023 04:24:53 PM, Cortez, John) Supplied power cord, SPO2 sensor, adult cuff from spares All parameters tested ok
1055909 | 196126 | CM | CMPLT | NFF | 2023-07-18 15:55 | 2023-07-18 15:57 | "BP not working" | (TC 07/18/2023 03:55:24 PM, Cortez, John) BP cuff was worn out, will ask 10N to order more air leak, calibration, and overpressure tests all passed in service mode All parameters tested ok
1049901 | 196127 | CM | CMPLT | CMPLT | 2023-07-18 17:07 | 2023-07-21 08:42 | MSH R61 Column not going to park position. | (TC 07/19/2023 10:13:52 AM, Whittington, Scott) Column not going to park position. | MSH R61 Column not going to park position.
67129 | 196128 | CM | CMPLT | NFF | 2023-07-19 06:11 | 2023-07-19 06:15 | Oximetry not working | (TC 07/19/2023 06:13:44 AM, Baghirzade, Nara) All functional tests for SpO2, NIBP, and temp passed; no fault found. Measured SpO2 on a different device and was getting the same result; SpO2 finger probe and cable are both functional.
65924 | 196129 | PM | CMPLT | CMPLT | 2023-07-19 06:47 | 2023-07-19 06:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/19/2023 06:47:20 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 196130 | PA | CMPLT | CMPLT | 2023-07-19 08:22 | 2023-07-19 08:24 | Ordering LVP Display boards | (TC 07/19/2023 08:24:09 AM, Reid, Emmerson ) Created customer complaint intake form and submitted no charge PO request
1062120 | 196131 | CM | CMPLT | NFF | 2023-07-19 08:55 | 2023-07-19 08:57 | Cassette will not click into pump | (TC 07/18/2023 08:55:51 AM, Rendon Lopez , Elizabeth) Tested pump, latching mechanism working properly no issues. Ran on user mode for >1hr, NFF
1056425 | 196132 | CM | CMPLT | NFF | 2023-07-19 09:10 | 2023-07-19 09:12 | "BP not working" | (TC 07/19/2023 09:09:52 AM, Cortez, John) Air leak, calibration, and overpressure tests all passed in NIBP calibration menu No fault found All parameters tested ok
1054431 | 196133 | PM | CMPLT | CMPLT | 2023-07-19 09:40 | 2023-07-19 09:43 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 09:40:05 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
1054430 | 196134 | PM | CMPLT | CMPLT | 2023-07-19 09:46 | 2023-07-19 09:47 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 09:46:20 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
65856 | 196135 | PM | CMPLT | CMPLT | 2023-07-19 09:49 | 2023-07-19 09:50 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 09:48:47 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
1065022 | 196136 | PM | CMPLT | CMPLT | 2023-07-19 09:54 | 2023-07-19 09:55 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 09:54:10 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
1065012 | 196137 | PM | CMPLT | CMPLT | 2023-07-19 09:58 | 2023-07-19 09:59 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 09:58:13 AM, Canjimavadivel, Kamallesh ) PM tests pass
1065005 | 196138 | PM | CMPLT | CMPLT | 2023-07-19 10:00 | 2023-07-19 10:00 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 10:00:03 AM, Canjimavadivel, Kamallesh ) PM tests pass
1065019 | 196139 | PM | CMPLT | CMPLT | 2023-07-19 10:01 | 2023-07-19 10:02 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 10:01:39 AM, Canjimavadivel, Kamallesh ) PM tests pass
1065010 | 196140 | PM | CMPLT | CMPLT | 2023-07-19 10:04 | 2023-07-19 10:05 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 10:04:41 AM, Canjimavadivel, Kamallesh ) PM tests pass
1065418 | 196141 | CM | CMPLT | CMPLT | 2023-07-19 10:32 | 2023-07-19 10:34 | 4 patients on JULY 17 not sent to the Cerner | (TC 07/19/2023 10:32:20 AM, Samanian, Kamran) verified Hyland server and noticed there are 4 error files for these 4 patients. so Hyland was fine and has send images to Server. From Biomed side there is nothing we can do. Contacted Katrina and she found there was differences between patient info between Cerner and system. She is looking into that.
65852 | 196142 | PM | CMPLT | CMPLT | 2023-07-19 10:46 | 2023-07-19 10:47 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 10:46:31 AM, Canjimavadivel, Kamallesh ) Battery replaced PM tests pass
1063746 | 196143 | PM | CMPLT | CMPLT | 2023-07-19 11:02 | 2023-07-19 11:03 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/19/2023 11:02:15 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 196144 | PA | CMPLT | CMPLT | 2023-07-19 11:25 | 2023-07-21 14:12 | Aspirator batteries | (TC 07/19/2023 11:31:07 AM, Canjimavadivel, Kamallesh ) Consulted with Samir and decided more batteries are needed. Placed order for batteries (TC 07/21/2023 02:11:29 PM, Canjimavadivel, Kamallesh ) Received batteries, placed on shelf | Biomed stock, charge back as used
1047862 | 196145 | CM | CMPLT | CMPLT | 2023-07-19 11:28 | 2023-07-19 11:28 | Plunger coarse travel error during PM | (TC 07/19/2023 11:27:52 AM, Baghirzade, Nara) Clutch was worn out, replaced the clutch. Plunger travels smoothly.
1061333 | 196146 | CM | CMPLT | CMPLT | 2023-07-19 11:29 | 2023-07-28 12:30 | Does not switch on | (TC 07/19/2023 11:28:50 AM, Baghirzade, Nara) VAC does not switch on; left it overnight plugged to the wall outlet, but still does not switch on. Re-seated connector the battery but it did not resolve the issue. Initiated a service request with 3M for pickup of defective VAC and drop off of loaner. RMA for pickup: 12758579 Delivery confirmation: 12974018 (TC 07/19/2023 12:52:14 PM, Baghirzade, Nara) 3M picked up the defective VAC. Attached pickup confirmation. Rental received from 3M (SN: GAAC09831). Attached drop off confirmation. (TC 07/19/2023 12:52:14 PM, Baghirzade, Nara) 3M picked up the defective VAC. Attached pickup confirmation. Rental received from 3M (SN: GAAC09831). (TC 07/28/2023 12:28:28 PM, Baghirzade, Nara) Received 11N-owned VAC back. Attached QC from 3M. VAC powers on; issue was resolved by 3M. Loaner VAC still on patient, will get it back next week and organize a pickup with 3M. Returned to 18South.
1059828 | 196147 | PM | CMPLT | CMPLT | 2023-07-19 12:38 | 2023-07-19 13:16 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/19/2023 01:15:53 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
53010 | 196148 | CM | CMPLT | CMPLT | 2023-07-19 12:49 | 2023-12-08 08:11 | Unable to enter test mode | (TC 07/19/2023 12:49:38 PM, Baghirzade, Nara) Cannot enter test mode, unable to do PM. Swapped out canisters but did not resolve the issue. Swapped out the aspirator with a functional one from RT (asset # 65856) (TC 12/08/2023 08:08:58 AM, Baghirzade, Nara) Confirmed with Samir to retire the unit. Buttons are defective and unable to test.
65869 | 196149 | PM | CMPLT | CMPLT | 2023-07-19 12:52 | 2023-07-19 12:52 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 12:51:41 PM, Baghirzade, Nara) PM completed.
| 196150 | CM | CMPLT | CMPLT | 2023-07-19 12:54 | 2023-07-19 12:55 | Internal meeting prior to philips meeting | (TC 07/19/2023 12:54:18 PM, Samanian, Kamran) discussed what tasks need to be asked from Philips for SES
1067395 | 196151 | CM | CMPLT | CMPLT | 2023-07-19 12:58 | 2023-07-19 13:01 | Loaner received | (TC 07/19/2023 12:58:23 PM, Baghirzade, Nara) Received loaner from 3M while hospital owned VAC is getting serviced (asset#1061333) Attached drop off confirmation and inspection card. Dropped off at 18South.
65866 | 196152 | PM | CMPLT | CMPLT | 2023-07-19 13:44 | 2023-07-20 08:37 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 02:38:05 PM, Canjimavadivel, Kamallesh ) Found device under student desk area. Attempted to perform PM tests. Test button not working. Opened up device and reseated ribbon cable, did not resolve issue. (TC 07/20/2023 08:37:14 AM, Canjimavadivel, Kamallesh ) Tested again with full battery charge, still not working. Contacted tech support and opened case. This unit is out of warranty and can be sent in for repair for cost. Device retired and used for parts
69539 | 196153 | CM | CMPLT | CMPLT | 2023-07-19 15:25 | 2023-07-19 15:27 | Channel error | (TC 07/19/2023 03:26:12 PM, Rhuland, Courtney) Bottle side pressure sensor high failure Replaced pressure sensor Performed check-in All tests passed Refer to ASM for details
1065267 | 196154 | CM | CMPLT | CMPLT | 2023-07-19 15:42 | 2023-07-19 15:43 | Needs to be retired, not in use | (TC 07/19/2023 03:41:57 PM, Cortez, John) Matt (supervisor, imaging) asked about disposal of lift Asked him to send to biomed, unit to be retired
1061898 | 196155 | CM | CMPLT | CMPLT | 2023-07-19 15:54 | 2023-08-17 16:18 | AUX outlet not working | (TC 07/17/2023 03:54:11 PM, Cortez, John) Confirmed issue. Not cord or wall outlet related. Need to call Stryker to request service (TC 07/21/2023 09:22:45 AM, Cortez, John) Bed still in use by CCU. Informed nurse to swap out when possible, then I will request service from styker (TC 08/17/2023 04:11:22 PM, Cortez, John) Service manual mentions circuit breaker switch underneath fowler on patient left side. Reset switch and AUX power returned
69616 | 196156 | CM | CMPLT | CMPLT | 2023-07-20 08:32 | 2023-07-20 08:34 | LVP display to be replaced | (TC 07/20/2023 08:32:24 AM, Cortez, John) LVP display replaced
1061209 | 196157 | CM | CMPLT | CMPLT | 2023-07-20 08:36 | 2023-07-20 08:38 | Camera not showing up on monitors | (TC 07/20/2023 08:36:24 AM, Reid, Emmerson ) OR8 Called about the camera not showing up on the monitors. They were on the wrong input, was AUX1 needed 4K 1
1050672 | 196158 | CM | CMPLT | CMPLT | 2023-07-20 08:45 | 2023-07-20 09:34 | Half of screen not working | (TC 07/20/2023 09:28:11 AM, Canjimavadivel, Kamallesh ) Found intermittent issue where half screen not working. Closed screen down and opened up again. Turned on and working fine. Opened up display cover and reseated ribbon cable. To be monitored for future reoccurance. Tested ok, returned back to unit.
52181 | 196159 | CM | CMPLT | CMPLT | 2023-07-20 08:54 | 2023-07-20 08:56 | Blank screen | (TC 07/20/2023 08:54:16 AM, Rhuland, Courtney) Powered on and device worked as expected Functional check failed humidity test Performed calibration using calibration vile Tests passed
1061215 | 196160 | CM | CMPLT | CMPLT | 2023-07-20 10:02 | 2023-07-20 10:03 | Smoke evac not clearing smoke | (TC 07/20/2023 10:02:03 AM, Reid, Emmerson ) I saw the filter life was very low. I replaced the filter and waited for them to use the ESU again. They said it was working better after that.
69167 | 196161 | CM | CMPLT | CMPLT | 2023-07-20 10:09 | 2023-07-20 10:24 | Slight physical damage on latch/front case | (TC 07/20/2023 10:08:53 AM, Cortez, John) Attempted to file physical marks on case/latch
| 196162 | CM | CMPLT | CMPLT | 2023-07-20 12:47 | 2023-07-20 12:47 | patinet monitoring meeting | (TC 07/20/2023 12:46:51 PM, Samanian, Kamran) done!
66457 | 196163 | PM | CMPLT | CMPLT | 2023-07-20 13:49 | 2023-07-19 13:49 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteriesand left panel of C500. Tested. Functional. SW.2.51 Leakage: 79uA and 83 ohm.
66459 | 196164 | PM | CMPLT | CMPLT | 2023-07-20 13:52 | 2023-07-20 13:52 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteriesand left panel of C500. Tested. Functional. SW.2.51 Leakage: 73uA and 88 ohm.
| 196165 | CM | CMPLT | CMPLT | 2023-07-20 13:59 | 2023-07-20 13:59 | Biomed cleaning | Biomed cleaning
65416 | 196166 | CM | CMPLT | CMPLT | 2023-07-20 14:36 | 2023-07-20 14:40 | System failure 23 | (TC 07/20/2023 02:36:09 PM, Baghirzade, Nara) Switched on the incubator but did not get any system failure error. Checked the fan and it was properly seated, and was executing the correct sequence as per manual. Checked the error log and there were no errors listed. During functional check, the RH sensor was inaccurately displaying the humidity. After trying to calibrate the RH sensor twice and both times it failed to read the correct humidity, I calibrated using the salt calibration bottle and RH sensor was reading correctly. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% humidity. All motion controls and tilt function is working. Incubator is fully functional.
64871 | 196167 | CM | CMPLT | CMPLT | 2023-07-20 14:42 | 2023-07-20 14:46 | Battery repair and replacement | (TC 07/20/2023 02:42:11 PM, Baghirzade, Nara) After replacing the battery, I turned the UPS on but the "replace battery" LED was still flashing. Left it for about 1.5 hours to charge and when i came back to check, the UPS would not turn on. The battery cable had a defective wire terminal and was not making proper contact with the UPS. Took the connector off the old UPS battery and replaced it with the new one. Placed the battery back, let it charge for about an hour, and came back to check by unplugging it and turning on the UPS; UPS was turning on properly and the "on battery" LED was on. Replaced power cable for the voice tower cart. (TC 07/20/2023 02:42:11 PM, Baghirzade, Nara) After replacing the battery, I turned the UPS on but the "replace battery" LED was still flashing. Left it for about 1.5 hours to charge and when i came back to check, the UPS would not turn on. The battery cable had a defective wire terminal and was not making proper contact with the UPS. Took the connector off the old UPS battery and replaced it with the new one. Placed the battery back, let it charge for about an hour, and came back to check by unplugging it and turning on the UPS; UPS was turning on properly and the "on battery" LED was on. Replaced power cable for the voice tower cart.
1047862 | 196168 | CM | CMPLT | CMPLT | 2023-07-20 14:47 | 2023-07-20 14:48 | Position potentiometer replacement | (TC 07/20/2023 02:47:14 PM, Baghirzade, Nara) Could not calibrate the syringe, values were not within range. Opened the device and found that the position potentiometer was broken; replaced the position potentiometer.
| 196169 | CM | CMPLT | CMPLT | 2023-07-20 14:50 | 2023-07-20 14:51 | Biomed floor cleanup
1050480 | 196170 | CM | CMPLT | CMPLT | 2023-07-20 14:57 | 2023-07-27 09:25 | Expiratory cassette low capacity | (TC 07/20/2023 03:02:44 PM, Rendon Lopez , Elizabeth) Replaced expiratory cassette membrane, self checks passed. Two batteries are due for replacement, ordered (TC 07/20/2023 03:02:44 PM, Rendon Lopez , Elizabeth) Replaced expiratory cassette membrane SN: 102753, self checks passed. Two batteries are due for replacement, ordered (TC 07/27/2023 09:24:30 AM, Rendon Lopez , Elizabeth) Received batteries, deliver them to Rex
1054224 | 196171 | CM | CMPLT | CMPLT | 2023-07-20 15:06 | 2023-07-20 15:06 | Not turning on | (TC 07/20/2023 03:05:59 PM, Canjimavadivel, Kamallesh ) Replaced batteries, all functions ok
1050477 | 196172 | CM | CMPLT | CMPLT | 2023-07-20 15:06 | 2023-07-27 09:23 | batteries due for replacement | (TC 07/20/2023 03:06:36 PM, Rendon Lopez , Elizabeth) Notification both batteries due for replacement. ordered (TC 07/27/2023 09:22:41 AM, Rendon Lopez , Elizabeth) Received batteries, deliver to Rex
1064341 | 196173 | CM | CMPLT | CMPLT | 2023-07-20 15:09 | 2023-07-20 15:13 | OR 3 table not working | (TC 07/20/2023 03:10:42 PM, Rendon Lopez , Elizabeth) Damaged power cord. Replaced, tested-OK
| 196174 | CM | CMPLT | CMPLT | 2023-07-20 15:13 | 2023-07-20 15:13 | meeting for cleaning | (TC 07/20/2023 03:13:27 PM, Samanian, Kamran) done!
64995 | 196175 | CM | CMPLT | CMPLT | 2023-07-20 15:23 | 2023-07-20 15:24 | Not registering PCA button | (TC 07/19/2023 03:23:27 PM, Rendon Lopez , Elizabeth) Tested, no fault found. Possible PCA button malfunction, not sent down.
65863 | 196176 | PM | CMPLT | CMPLT | 2023-07-20 15:27 | 2023-07-20 15:28 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/19/2023 03:27:30 PM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
1055563 | 196177 | CM | CMPLT | CMPLT | 2023-07-20 15:29 | 2023-07-20 15:30 | Broken knob | (TC 07/19/2023 03:29:17 PM, Rendon Lopez , Elizabeth) Replaced button, tested-OK
| 196178 | CM | CMPLT | CMPLT | 2023-07-20 15:44 | 2023-07-20 15:44 | NICU huddle | (TC 07/20/2023 03:44:02 PM, Baghirzade, Nara) NICU Huddle
| 196179 | CM | CMPLT | CMPLT | 2023-07-20 16:01 | 2023-07-20 16:02 | department clean up
| 196180 | CM | CMPLT | CMPLT | 2023-07-20 16:14 | 2023-07-20 16:15 | EMG server implementation meeting w/ vendor and IT | (TC 07/20/2023 04:14:00 PM, Baghirzade, Nara) EMG server implementation meeting with vendor and IT
67988 | 196181 | CM | CMPLT | CMPLT | 2023-07-20 17:26 | 2023-07-20 17:26 | PM and battery | (TC 07/20/2023 05:26:11 PM, Samanian, Kamran) replaced battery, PM done and passed
1065008 | 196182 | CM | CMPLT | CMPLT | 2023-07-20 17:28 | 2023-07-20 17:28 | PM | (TC 07/20/2023 05:27:50 PM, Samanian, Kamran) Replaced battery, performed PM, tested and passed
| 196183 | CM | CMPLT | CMPLT | 2023-07-20 18:07 | 2023-07-20 11:12 | Order for HP server - PICiX server back up | (TC 07/21/2023 12:39:44 PM, Samanian, Kamran) S/C#1003008100 (TC 07/27/2023 11:10:26 AM, Samanian, Kamran) Cancelded the PO# 7202253319 as we will use the old WEB server which is exact the same model DL360 G10
| 196184 | CM | ACTIV | ISSUE | 2023-07-21 08:34 |  | Biomed Floor Organizing | (TC 08/03/2023 02:45:48 PM, Chukri, Samir) cleaning biomed work shop
1046273 | 196185 | CM | CMPLT | CMPLT | 2023-07-21 08:40 | 2023-07-21 08:41 | MSH CFRH 7flr APLIO 300 REPAIR TRACKBALL ASSEMBLY | (TC 07/21/2023 08:40:14 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 REPAIR TRACKBALL ASSEMBLY (TC 07/21/2023 08:40:14 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 REPAIR TRACKBALL ASSEMBLY | MSH CFRH 7flr APLIO 300 REPAIR TRACKBALL ASSEMBLY
| 196186 | CM | CMPLT | CMPLT | 2023-07-21 08:43 | 2023-07-21 08:44 | working with Nara NET UPS battery issue. | (TC 07/21/2023 08:43:32 AM, Chukri, Samir) working with Nara, solve the UPS battery issue.
68224 | 196187 | CM | CMPLT | CMPLT | 2023-07-21 08:49 | 2023-07-21 09:18 | Nurse saying ECG not working | (TC 07/21/2023 08:53:03 AM, Scero, Tomlyn) Nurse called and reported that the night nurse couldn't get an ECG reading for a patient on the monitor. I took a brand new set of leads and my simulator. The ECG and monitor was working perfectly. I showed the nurse and mentioned that possibly changing the patient electrodes would help.
1066107 | 196188 | CM | CMPLT | CMPLT | 2023-07-21 08:56 | 2023-07-21 09:19 | MSH 10 GIM GE POCUS CABLE LOCK, OBTAIN AND INSTALL | (TC 07/21/2023 08:56:14 AM, Whittington, Scott) MSH 10 GIM GE POCUS CABLE LOCK, OBTAIN AND INSTALL 22-55-28 | MSH 10 GIM GE POCUS CABLE LOCK, OBTAIN AND INSTALL
1065355 | 196189 | CM | CMPLT | CMPLT | 2023-07-21 09:22 | 2023-08-23 12:11 | MSH US PROBE PVT 661 XTAL AIRTAG 38 | (TC 07/21/2023 09:22:03 AM, Whittington, Scott) MSH US PROBE PVT 661 XTAL AIRTAG 38 | MSH US PROBE PVT 661 XTAL AIRTAG 38
| 196190 | CM | CMPLT | CMPLT | 2023-07-21 10:00 | 2023-07-21 10:03 | Weekly check for 10N - July 17 | (TC 07/17/2023 10:01:56 AM, Cortez, John) Patient rooms busy, checked storage and hallways. Nothing to note.
| 196191 | RDS | CMPLT | CMPLT | 2023-07-21 10:04 | 2023-07-21 10:10 | Weekly check for ICU - July 17 | (TC 07/21/2023 10:07:07 AM, Cortez, John) Checked 9 rooms, storage and hallways logged findings in spreadsheet Swapped pumps with PMs due (TC 07/17/2023 10:07:07 AM, Cortez, John) Checked 9 rooms, storage and hallways logged findings in spreadsheet Swapped pumps with PMs due
| 196192 | RDS | CMPLT | CMPLT | 2023-07-21 10:11 | 2023-07-21 10:14 | Weekly check for Imaging - July 17 | (TC 07/21/2023 10:11:03 AM, Cortez, John) Didn't check imaging suites. Checked hallways (TC 07/17/2023 10:11:03 AM, Cortez, John) Didn't check imaging suites. Checked hallways
1056055 | 196193 | CM | CMPLT | CMPLT | 2023-07-21 10:25 | 2023-07-21 10:25 | MSH-5-506 IR1 AZURION ; TCS CONTROL CONSOLE CAME UNSCREWED FROM BED RAIL AND FELL. | (TC 07/21/2023 10:24:57 AM, Whittington, Scott) MSH-5-506 IR1 AZURION ; TCS CONTROL CONSOLE CAME UNSCREWED FROM BED RAIL AND FELL. | MSH-5-506 IR1 AZURION ; TCS CONTROL CONSOLE CAME UNSCREWED FROM BED RAIL AND FELL.
1065418 | 196194 | CM | CMPLT | CMPLT | 2023-07-21 12:42 | 2023-08-03 11:40 | issue with Hyland on JULY 14- opened ticket | (TC 07/21/2023 12:42:25 PM, Samanian, Kamran) checked the connections and they are fine. Contacted Hyland to open the ticket and sent log file , ticket# 01612189 (TC 08/03/2023 11:39:30 AM, Samanian, Kamran) requested to close the ticket. No news from Anita.
1061207 | 196195 | CM | CMPLT | CMPLT | 2023-07-21 12:55 | 2023-07-21 12:55 | Setup the Camera
1058172 | 196196 | CM | CMPLT | CMPLT | 2023-07-21 14:04 | 2023-07-21 14:04 | APS EPIDURAL 2023 July Library Downloaded | (TC 07/21/2023 02:03:43 PM, Reid, Emmerson ) complete
67763 | 196197 | CM | CMPLT | CMPLT | 2023-07-21 14:06 | 2023-07-21 14:06 | APS EPIDURAL 2023 July Library Downloaded
1058174 | 196198 | CM | CMPLT | CMPLT | 2023-07-21 14:08 | 2023-07-21 14:08 | APS EPIDURAL 2023 July Library Downloaded
1058169 | 196199 | CM | CMPLT | CMPLT | 2023-07-21 14:10 | 2023-07-21 14:10 | APS EPIDURAL 2023 July Library Downloaded
67761 | 196200 | CM | CMPLT | CMPLT | 2023-07-21 14:10 | 2023-07-21 14:10 | APS EPIDURAL 2023 July Library Downloaded
67756 | 196201 | PM | CMPLT | CMPLT | 2023-07-21 14:20 | 2023-07-21 14:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/21/2023 02:36:39 PM, Reid, Emmerson ) APS EPIDURAL 2023 July Library Downloaded PM complete, see proceedure for details.
| 196202 | RDS | CMPLT | CMPLT | 2023-07-24 08:03 | 2023-07-24 10:05 | Weekly check for ED
| 196203 | RDS | CMPLT | CMPLT | 2023-07-24 08:08 | 2023-07-24 13:01 | Weekly check for ICU
63438 | 196204 | CM | CMPLT | CMPLT | 2023-07-24 08:10 | 2023-07-24 08:11 | Damaged power cord | (TC 07/24/2023 08:10:20 AM, Baghirzade, Nara) Replaced power cord. Functional tests passed. Ran a feed at various rates; no issues/errors with the feed. Occluded the line; got an occlusion alarm and feed stopped. Occlusion alarm is functional.
| 196205 | CM | CMPLT | CMPLT | 2023-07-24 08:37 | 2023-07-25 08:21 | Weekly Check for CCU | (TC 07/24/2023 08:20:55 AM, Scero, Tomlyn) Weekly checks complete
68985 | 196206 | CM | CMPLT | CMPLT | 2023-07-24 08:47 | 2023-07-24 10:28 | Not charging when plugged in | (TC 07/24/2023 10:27:09 AM, Canjimavadivel, Kamallesh ) Damage found in power cord on plug end. Replaced with a spare cord. Device charging normally. Returned to ED.
| 196207 | RDS | CMPLT | CMPLT | 2023-07-24 09:10 | 2023-10-10 10:26 | Weekly Checks For PACU | (TC 10/10/2023 10:25:58 AM, Reid, Emmerson ) Complete
| 196208 | RDS | CMPLT | CMPLT | 2023-07-24 09:11 | 2023-07-24 09:19 | Weekly Checks For PACU | (TC 07/24/2023 09:18:22 AM, Reid, Emmerson ) Complete, see proceedure.
501620 | 196209 | CM | CMPLT | CMPLT | 2023-07-24 09:54 | 2023-07-26 12:21 | MSH OPG SPP - C51 Transducer - Distortion at strain gauge | (TC 07/24/2023 09:53:37 AM, Chau, Ken) case #0122027704 -ordered probe as it was causing distortion at strain gauge (TC 07/26/2023 12:20:01 PM, Chau, Ken) Probe retired: replaced with MIE#1067419 | case #0122027704
68269 | 196210 | CM | CMPLT | CMPLT | 2023-07-24 10:07 | 2023-07-28 09:13 | Msmt Ext Malfunction | (TC 07/24/2023 10:07:00 AM, Canjimavadivel, Kamallesh ) Isolated issue down to FMS module. Took spare FMS module (1051394) and upgraded to rev N. Swapped out with defective FMS module. (TC 07/24/2023 01:41:53 PM, Canjimavadivel, Kamallesh ) Found corrosion and oxidation on MSL connecter. Created a case with Philips for quote (TC 07/24/2023 02:26:04 PM, Canjimavadivel, Kamallesh ) Spoke with Philips over the phone and received quote for part. Placed order (TC 07/28/2023 09:12:32 AM, Canjimavadivel, Kamallesh ) Replaced MSL down board, tested ok in biomed shop. Swapped out with spare in unit.
1051394 | 196211 | CM | CMPLT | CMPLT | 2023-07-24 10:10 | 2023-07-24 10:11 | Software upgrade | (TC 07/24/2023 10:10:30 AM, Canjimavadivel, Kamallesh ) Upgraded from Rev J to Rev N. Installed as spare in ED Y41
| 196212 | CM | CMPLT | CMPLT | 2023-07-24 10:16 | 2023-07-26 09:26 | Schedule maintenance contract for L&D OR tables | (TC 07/24/2023 10:19:30 AM, Rendon Lopez , Elizabeth) 1 year schedule maintenance contract for L&D OR tables (TC 07/26/2023 09:25:05 AM, Rendon Lopez , Elizabeth) Sent PO to company, contract confirmed, will schedule PMs
| 196213 | CM | CMPLT | CMPLT | 2023-07-24 10:21 | 2023-08-02 15:03 | Spare X2 batteries | (TC 08/02/2023 03:02:39 PM, Rendon Lopez , Elizabeth) Received batteries, labelled them
| 196214 | CM | CMPLT | CMPLT | 2023-07-24 10:28 | 2023-07-24 10:28 | checked the n-care for IP address browser | (TC 07/24/2023 10:28:21 AM, Samanian, Kamran) done!
| 196215 | CM | CMPLT | CMPLT | 2023-07-24 10:29 | 2023-07-24 10:30 | PACU2S not showing up in FP | (TC 07/24/2023 10:29:22 AM, Samanian, Kamran) verified that the FP not showing up PACU2S. checked few things and NFF. Asked Emmerson to renboot the PACU2S and resolved issue.
1047469 | 196216 | PM | CMPLT | CMPLT | 2023-07-24 10:47 | 2023-07-24 10:47 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 6th year PM after 5th year PM last year. No GS500. Replaced PS500 batteries, V battery, O2 filter, air filter and O2 pressure regulator. calibrated. Tested. Functional. SW.2.51 Leakage: 81uA and 80 ohm.
63436 | 196217 | CM | CMPLT | CMPLT | 2023-07-24 10:55 | 2023-07-24 10:58 | Door off hinge and damaged power cord | (TC 07/24/2023 10:56:57 AM, Baghirzade, Nara) Cleaned hinges and replaced the door back; sits properly and covers the loading area. Replaced the power cord. Ran a feed at various rates; no issues encountered. Clamped the feed line and occlusion alarm was functional. Buzzer at good volume level.
67919 | 196218 | CM | CMPLT | CMPLT | 2023-07-24 10:58 | 2023-07-24 11:01 | System failure 25 | (TC 07/24/2023 10:58:40 AM, Baghirzade, Nara) Replaced the 9V battery with a rechargeable 9V battery. Fan was not seated properly, causing system failure 25. Re-seated the fan and incubator boots up with no errors or failures. During functional check, the RH sensor was inaccurately displaying the humidity; off by 6%. Calibrated the RH sensor. Set 75% RH and 35C temp; measured 75.5% RH and 35C temp.
| 196219 | CM | CMPLT | CMPLT | 2023-07-24 11:37 | 2023-07-24 11:36 | Correction of PICiX servers serial numbers in TMS | (TC 07/24/2023 11:36:40 AM, Samanian, Kamran) done!
| 196220 | CM | CMPLT | CMPLT | 2023-07-24 12:48 | 2023-07-24 12:52 | Wheelchair drop off/coordination | (TC 07/24/2023 12:48:28 PM, Baghirzade, Nara) Henna requested a wheelchair for their clinic as their elderly patients struggling walking from the waiting room area to the exam rooms. Contacted Michael Badour a week ago and he had connected with Ula to dispatch a wheelchair for Ophthalmology. Met with Ula today to pick up the wheelchair and dropped it off with Henna at the ophthalmology clinic.
1066462 | 196221 | FC | CMPLT | CMPLT | 2023-07-24 12:54 | 2023-07-24 12:55 | QC checklist | (TC 07/24/2023 12:54:20 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
67103 | 196222 | CM | CMPLT | CMPLT | 2023-07-24 13:32 | 2023-07-24 13:33 | TEMP problem | (TC 07/24/2023 01:31:59 PM, Samanian, Kamran) verified the problem, replaced the probe, configured the monitor. tested and passed
| 196223 | CM | CMPLT | CMPLT | 2023-07-24 13:37 | 2023-07-31 12:51 | huddle meeting
| 196224 | CM | CMPLT | CMPLT | 2023-07-24 13:37 | 2023-07-24 13:37 | Biomed meeting | Biomed meeting
| 196225 | CM | CMPLT | CMPLT | 2023-07-24 13:38 | 2023-07-24 13:38 | PM huddle | (TC 07/24/2023 01:37:40 PM, Baghirzade, Nara) PM huddle
67144 | 196226 | CM | CMPLT | CMPLT | 2023-07-24 13:38 | 2023-07-24 13:39 | Temperature probe broken, not measuring temperature | (TC 07/24/2023 01:38:17 PM, Baghirzade, Nara) Replaced temperature probe. Functional tests passed for: NIBP Temperature SpO2 Handed the system back to Vimy.
1047862 | 196227 | CM | CMPLT | CMPLT | 2023-07-24 13:40 | 2023-07-24 13:42 | Plungerhead replacement | (TC 07/24/2023 01:40:13 PM, Baghirzade, Nara) Replaced the plunger head assembly; there were cracks on both sides of the plunger head. Was able to salvage spare parts as not all parts required replacement. Claws move properly and not too much tension on the lever.
| 196228 | CM | CMPLT | CMPLT | 2023-07-24 13:42 | 2023-07-24 13:42 | Developing a laser checklist | (TC 07/24/2023 01:42:20 PM, Baghirzade, Nara) Worked on developing a laser checklist; after last laser committee meeting, I added what was discussed from the feedback. Working on determining the NHZ and LTCA for lasers, starting with the CO2 laser.
| 196229 | CM | CMPLT | CMPLT | 2023-07-24 13:44 | 2023-07-24 13:46 | Meeting regarding oxygen tank and crash cart | (TC 07/24/2023 01:44:11 PM, Baghirzade, Nara) Met with Henna to discuss the ARC's response regarding a crash cart and AED; crash cart is the clinic's responsibility as they are a satellite site. As for the AED, Steve from RT will give an update. Henna also inquired about the oxygen tank and who replenishes it; relayed this information to Adeel for next steps.
1047469 | 196230 | CM | CMPLT | CMPLT | 2023-07-24 14:07 | 2023-07-28 08:09 | ordering VN500 parts for RT | received parts for VN500 vents (TC 07/28/2023 08:09:49 AM, Arslan, Olcay) received parts for VN500 vents. PO#7202253983
| 196231 | CM | CMPLT | CMPLT | 2023-07-24 17:22 | 2023-07-24 17:22 | Site coverage.
| 196232 | CM | CMPLT | CMPLT | 2023-07-25 07:48 | 2023-07-24 11:36 | Huddle meeting
1061209 | 196233 | CM | CMPLT | CMPLT | 2023-07-25 08:08 | 2023-07-25 11:37 | light number 3 arm is out of tension | (TC 07/25/2023 08:07:43 AM, Samanian, Kamran) received a call from OR8 that exam light 3 arm is loosen. verified the problem. not able to fix due to starting a case soon. surgeon was ok with it . it will be fixed after case is done (TC 07/25/2023 11:37:03 AM, Samanian, Kamran) back to the room and adjsuted the tension, tested and passed.
1063260 | 196234 | PM | CMPLT | CMPLT | 2023-07-25 08:37 | 2023-07-25 09:49 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/25/2023 09:45:02 AM, Rhuland, Courtney) Came down stating "not turning on" Device was missing cord retainer and power cord was not connected properly Cord retainer replaced Pad that came down with system was defective Changed pad and completed PM Irradiance - 49.9 uW/cm^2/nm (large pad) Hours - 289.6 Ground resistance - 11.0 uA Resistance - 0.132 Ohms
67960 | 196235 | CM | CMPLT | CMPLT | 2023-07-25 08:44 | 2023-07-25 08:45 | System failure 25 | (TC 07/25/2023 08:44:06 AM, Baghirzade, Nara) Fan was not properly seated causing error 25 upon boot up. Re-seated the fan. Replaced 9V with 9V rechargeable. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 76% Incubator is fully functional.
1054976 | 196236 | CM | CMPLT | CMPLT | 2023-07-25 08:51 | 2023-07-25 08:52 | SpO2 not functioning | (TC 07/25/2023 08:51:09 AM, Baghirzade, Nara) Red LNC-10 cable was damaged; replaced cable. Functional tests passed for: -NIBP -SpO2 -Temperature
1055947 | 196237 | CM | CMPLT | CMPLT | 2023-07-25 08:53 | 2023-07-25 08:54 | BP hose leaking | (TC 07/25/2023 08:52:50 AM, Baghirzade, Nara) BP hose had a leak; replaced the BP hose. During functional check, temperature probe was not measuring. Replaced the temperature probe and well. All functional tests passed for: NIBP SpO2 Temperature
1056443 | 196238 | CM | CMPLT | NFF | 2023-07-25 08:55 | 2023-07-25 08:56 | BP did not measure | (TC 07/25/2023 08:55:13 AM, Baghirzade, Nara) Could not find an issue with measuring BP; took multiple measurements successfully. Powered the root down and rebooted, took more measurements and no issues/errors. All functional tests passed for: NIBP SpO2 Temperature No fault found.
| 196239 | RDS | CMPLT | CMPLT | 2023-07-25 09:17 | 2023-07-25 09:22 | Weekly Check - NICU/Resus | (TC 07/25/2023 09:21:06 AM, Ferreira, Qwynn) Sweeped the unit and checked 3 rooms - all were good Checked defib and noticed pads were expired Will email May Everything else was good
| 196240 | RDS | CMPLT | CMPLT | 2023-07-25 09:23 | 2023-07-25 09:25 | Weekly Check - 12S | (TC 07/25/2023 09:25:08 AM, Ferreira, Qwynn) Weekly Check
| 196241 | RDS | CMPLT | CMPLT | 2023-07-25 09:26 | 2023-07-25 09:32 | Weekly Check - Oncology | (TC 07/25/2023 09:27:42 AM, Ferreira, Qwynn) Weekly Check
| 196242 | CM | CMPLT | CMPLT | 2023-07-25 09:46 | 2023-07-25 09:47 | EMG server upgrade plan meeting | (TC 07/25/2023 09:46:26 AM, Baghirzade, Nara) EMG server upgrade plan meeting with Kamran, Nima, and Andrew
1059180 | 196243 | CM | CMPLT | CMPLT | 2023-07-25 10:01 | 2023-07-25 10:06 | Battery Depleted | (TC 07/25/2023 10:05:12 AM, Reid, Emmerson ) Tried conditioning the old battery more than 5x Replaced with a new battery.
1051357 | 196244 | CM | CMPLT | CMPLT | 2023-07-25 10:07 | 2023-07-25 10:08 | Not turning on | (TC 07/25/2023 10:06:57 AM, Rhuland, Courtney) Power cord was not fully connected Strain relief was missing, replaced with spare Functional check - OK
67958 | 196245 | CM | CANCL | CANCL | 2023-07-25 10:11 |  | Ordering incubator hood
| 196246 | CM | CMPLT | CMPLT | 2023-07-25 10:20 | 2023-07-25 10:20 | meeting for EMG-Internal | (TC 07/25/2023 10:20:27 AM, Samanian, Kamran) meeting and some works around
1058059 | 196247 | CM | CMPLT | CMPLT | 2023-07-25 10:27 | 2023-07-25 12:51 | RT/AA reported in between cases self test failed. | (TC 07/25/2023 11:12:01 AM, Scero, Tomlyn) Pulled device into OR12 for testing. Faheem called and said between cases yesterday they performed a selftest and everything failed and came up red. He didn't troubleshoot, but just replaced the AGM. I performed the self test and eerything was green and passed. I ran the AGM in different vent modes with a test lung to confirm that there were no errors/alarms. Told Peter the device is fine.
1061207 | 196248 | CM | CMPLT | CMPLT | 2023-07-25 12:31 | 2023-07-25 12:32 | OR 6 camera requested to be removed | (TC 07/25/2023 12:31:36 PM, Samanian, Kamran) done and returned to the drawer in the nursing desk
65919 | 196249 | CM | CMPLT | CMPLT | 2023-07-25 12:59 | 2023-07-25 13:00 | Clutch and rear case replacement | (TC 07/25/2023 12:59:32 PM, Baghirzade, Nara) During functional check, the plunger was not going down properly and was repeatedly getting stuck. Clutch was worn out. Replaced clutch. Rear case was damaged; replaced rear case. Attached new asset tag. Transfered the lock latch over to the new case.
1054275 | 196250 | PM | CMPLT | CMPLT | 2023-07-25 13:00 | 2023-07-25 13:02 | PM-BATTERY-Annual PM for: Bed Echo-Flex 5002 | (TC 07/25/2023 01:00:30 PM, Scero, Tomlyn) PM complete. LC 19uA GWR .18 ohms
66009 | 196251 | PM | CMPLT | CMPLT | 2023-07-25 13:12 | 2023-07-25 13:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 01:12:17 PM, Rhuland, Courtney) Rear and front cases replaced PM completed All tests passed Refer to ASM for details
| 196252 | CM | CMPLT | CMPLT | 2023-07-25 13:17 | 2023-07-25 13:19 | Meeting with Julie re: wound vacs inventory | (TC 07/25/2023 01:16:45 PM, Baghirzade, Nara) Met with Julie regarding the wound vacs on 14th floor. Julie had traded in 3 VACs earlier this year so had to update the information on TMS as 2/3 VACs were no longer on site. A loaner VAC was also retired as we had not seen it since 2022. Need to meet with Nima or Adeel to discuss communications regarding the VACs so that we can keep our database up to date.
| 196253 | CM | CMPLT | CMPLT | 2023-07-25 14:25 | 2023-07-25 14:26 | PM huddle | (TC 07/25/2023 02:25:23 PM, Baghirzade, Nara) PM huddle
| 196254 | CM | CMPLT | CMPLT | 2023-07-25 14:32 | 2023-07-25 14:32 | biomed meeting | biomed meeting
66474 | 196255 | CM | CMPLT | CMPLT | 2023-07-25 14:33 | 2023-07-25 14:33 | left front caster is off the floor | adjusted left front caster
69189 | 196256 | FC | CMPLT | CMPLT | 2023-07-25 15:42 | 2023-07-25 15:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/25/2023 03:41:49 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
| 196257 | CM | CMPLT | CMPLT | 2023-07-25 16:49 | 2023-07-25 16:49 | Site Coverage
1051140 | 196258 | CM | CMPLT | CMPLT | 2023-07-26 07:47 | 2023-07-26 09:37 | Arm needs to be tightened | (TC 07/26/2023 09:36:26 AM, Canjimavadivel, Kamallesh ) Opened up spring assembly and reassembled. Arm is staying up in all positions. Light tested ok. Plugged in to charge before sending back to unit.
67936 | 196259 | CM | CMPLT | CMPLT | 2023-07-26 08:04 | 2023-07-26 08:05 | System failure 25 | (TC 07/26/2023 08:04:02 AM, Baghirzade, Nara) System failure 25 was caused by the fan not seated properly; reseated the fan. Replaced 9V battery with 9V rechargeable battery. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75.5% RH Incubator is fully functional.
65919 | 196260 | CM | CMPLT | CMPLT | 2023-07-26 08:33 | 2023-08-02 14:47 | Unable to calibrate plunger force | (TC 07/26/2023 08:32:46 AM, Baghirzade, Nara) Tried to calibrate the plunger force but was unable to and calibration resulted in a fail. Opened the device and ensured that all connections are seated properly. The clutch was replaced but did not resolve the issue. Will have to order a new plunger assembly. (TC 08/02/2023 02:46:13 PM, Baghirzade, Nara) Replaced plunger assembly. Calibration passed.
65919 | 196261 | CM | CMPLT | CMPLT | 2023-07-26 08:38 | 2023-08-02 07:51 | Ordering plunger assembly | (TC 07/26/2023 08:38:42 AM, Baghirzade, Nara) SC 1003011390 (TC 08/02/2023 07:50:00 AM, Baghirzade, Nara) PO 7202255644 (TC 02/14/2024 07:54:23 AM, Baghirzade, Nara) Finally located the paperwork for the plunger as it was not located; plunger was used for another NICU syringe pump by another colleague. Will reorder a new one.
63876 | 196262 | CM | CMPLT | CMPLT | 2023-07-26 08:45 | 2023-08-02 13:03 | Ordering hood | (TC 07/26/2023 08:49:08 AM, Baghirzade, Nara) SC 1003011415 (TC 08/02/2023 12:41:08 PM, Baghirzade, Nara) PO 7202255641
1054158 | 196263 | CM | CMPLT | CMPLT | 2023-07-26 08:55 | 2023-07-26 08:56 | Button broken | (TC 07/24/2023 08:55:45 AM, Rendon Lopez , Elizabeth) Button broken. Replaced with spare, tested function-OK
67821 | 196264 | PM | CMPLT | CMPLT | 2023-07-26 08:57 | 2023-07-26 08:59 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/24/2023 08:57:49 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed. PM changed to 3yrs
1067395 | 196265 | CM | CMPLT | CMPLT | 2023-07-26 08:59 | 2023-07-26 09:00 | QC checklist | (TC 07/26/2023 08:59:48 AM, Baghirzade, Nara) QC checklist completed. Returned to 18 South
66563 | 196266 | CM | CMPLT | CMPLT | 2023-07-26 08:59 | 2023-07-26 09:01 | Battery replaced | (TC 07/24/2023 08:59:50 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed. PM changed to 3yrs. Placed in L&D
66550 | 196267 | PM | CMPLT | CMPLT | 2023-07-26 09:06 | 2023-07-26 09:07 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/25/2023 09:06:51 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed. PM changed to 3yrs
66557 | 196268 | PM | CMPLT | CMPLT | 2023-07-26 09:08 | 2023-07-26 09:10 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/24/2023 09:09:16 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed. PM changed to 3yrs
1051389 | 196269 | CM | CMPLT | NFF | 2023-07-26 09:15 | 2023-07-26 09:16 | not picking up on monitor | (TC 07/25/2023 09:15:13 AM, Rendon Lopez , Elizabeth) Tested, no fault found
67224 | 196270 | CM | CMPLT | CMPLT | 2023-07-26 09:20 | 2023-08-14 09:34 | Warmer beeping | (TC 07/25/2023 09:20:12 AM, Rendon Lopez , Elizabeth) Found back switch off, turned on. Testing. (TC 07/26/2023 01:26:20 PM, Rendon Lopez , Elizabeth) Replaced blender, tested Resus system-OK. Warmer is now not initializing and power alarm on, testing test points in Power PCA (TC 07/27/2023 01:26:20 PM, Rendon Lopez , Elizabeth) Device is receiving AC power but not DC. Replaced power PCA, not resolved. Same issue. needs new power supply, inquired for part, ordered (TC 08/02/2023 08:42:09 AM, Rendon Lopez , Elizabeth) Replaced power supply following the manual. Warmer is working, however it shut down after some time, intermittent problem. Troubleshooted, replaced CPU, same issue, no errors on logs. Replaced standby switch, seems to be working. Ordered standby switch, testing in progress (TC 08/04/2023 02:56:19 PM, Rendon Lopez , Elizabeth) Replaced standby switch and wire connector, working. No audio indication of alarms, speaker connector damaged, ordered speaker. (TC 08/10/2023 09:33:23 AM, Rendon Lopez , Elizabeth) Replaced speaked, tested all functions-OK. Returned to L&D Room 15
1061215 | 196271 | CM | CMPLT | CMPLT | 2023-07-26 10:28 | 2023-07-26 10:31 | Emergency sound coming from room | (TC 07/26/2023 10:28:03 AM, Reid, Emmerson ) Reeba called about a souind coming from the boom in OR14, however it was the Pyxis, after a Reboot the sound went away. I did leave a ticket with the vendor to note if this happens again.
1067419 | 196272 | CM | CMPLT | CMPLT | 2023-07-26 10:45 | 2023-07-26 12:23 | NEW C5-1 PROBE INCOMING OPG SPP
66020 | 196273 | CM | CMPLT | CMPLT | 2023-07-26 10:47 | 2023-07-26 10:48 | Broken cases | (TC 07/26/2023 10:47:32 AM, Rhuland, Courtney) Front and rear cases replaced PM complete All tests passed Refer to ASM for details
1049940 | 196274 | CM | CMPLT | CMPLT | 2023-07-26 10:50 | 2023-08-28 15:28 | Distorted image on touch screen display | (TC 07/26/2023 10:50:37 AM, Fan, Tak) Distorted image on start up. Restarted, persisted. Checked cabling. Swapped with touch screen on other machine (1063026, 416586VE95A ). Distortion on the touch screen. Ordered new touch screen. GC200111, Operating Panel Upper Aurora 164L0025, Cable ECG MARQ. AHA/America 164L0027, Leadwires ECG MARQ. AHA / AMERICA GA307239, Gel Cup, Quanitity: 2 (TC 07/31/2023 10:29:42 AM, Fan, Tak) Follow up. All parts but touch screen came in. Touch screen on backorder. (TC 08/28/2023 03:27:14 PM, Fan, Tak) Screen on backorder indefinitely. Ordered refurbished monitor. Replaced. Returned bad display to GE.
1061316 | 196275 | CM | CMPLT | CMPLT | 2023-07-26 10:56 | 2023-09-18 08:34 | Cannot burn QR study to burn | (TC 07/26/2023 10:56:09 AM, Fan, Tak) Tested on multiple systems, cannot copy from Viewpoint (Images could not be found). Study is viewable on system and on a viewpoint client. Contacted Viewpoint support (Ref 432494) Could not find an issue. Other studies can be imported. They will look into it. Exported onto encrypted USB from client, OK. Imported from USB to system, OK. Burned to non-dicom CD (DICOM too large), OK. (TC 07/31/2023 01:26:20 PM, Fan, Tak) Still a problem on random studies. Copied study on encrypted USB. Burned it onto a CD with laptop. OK. Request was for images not dicom files, burned another CD. Contacted Sinai IT to allow burner rights to the PC, REF 034833656. (TC 08/08/2023 11:58:00 AM, Fan, Tak) Request approved. Robin from IT needs more network info. Plugged DVD drive into PC and leaving in for now. IT will change settings again in near future. No Reply from GE.
1060000 | 196276 | CM | CMPLT | CMPLT | 2023-07-26 11:17 | 2023-07-26 11:20 | Segment screen was frozen | (TC 07/26/2023 11:17:17 AM, Baghirzade, Nara) I was able to get the task manager screen up and from there I shut down the computer. After booting it up, it was stuck on the command window and would not load the segment screen. I was unable to get to the task manager this time so I had to do a hard reboot by unplugging the monitor from AC and plugging it back in. The command window took longer than expected to load but after a couple of minutes, it loaded the segment screen. Central is fully functional.
| 196277 | CM | CMPLT | CMPLT | 2023-07-26 11:36 | 2023-07-26 11:37 | going to IVF to check the exam room 2 & 3 | (TC 07/26/2023 11:36:34 AM, Samanian, Kamran) checked the BPs, BAT is out and needs to be replaced.
68287 | 196278 | CM | CMPLT | CMPLT | 2023-07-26 12:09 | 2023-12-21 09:26 | Needle EMG freezing | (TC 07/26/2023 12:09:16 PM, Baghirzade, Nara) Needle EMG was freezing during patient test; swapped the amplifiers around to try to rule out whether the cause is the amplifier or the main control box. Vinod will report back about whether the issue has resolved. (TC 07/28/2023 01:33:18 PM, Baghirzade, Nara) Issue is persisting in the same room (1643), will need to change the base unit as this is the only thing that the amplifier connects to. Checked that all connections are secure. The second EMG did not have this issue which is why I've determined that the base unit is the issue. Emailed Rob Duga for a quote. (TC 10/31/2023 12:47:36 PM, Baghirzade, Nara) Tested amplifier and electrodes for noise at 10 uV scale on data input connectors 1 and 2. Both DINs are good, no excessive noise. Signal is less than 10 uV when shorting all 3 leads. Ran the HW programmer however it didn't affect the amplifier. Will contact Natus for remote support. (TC 11/06/2023 08:35:29 AM, Baghirzade, Nara) Placed amplifier back on the system. Ran EDX HW programmer successfully. Waiting on the next Natus remote troubleshooting steps. (TC 12/01/2023 09:19:05 AM, Baghirzade, Nara) Dr. Lazarou reported on Wednesday (Nov 29) that the during needle EMG, the waveforms were falling flat and no signal getting picked up from the muscle. After placing a new needle and pressing the connector in firmly, it worked however this is not the optimal way to conduct the work. Took the amplifier down to biomed, Samir and I opened the device, took off the old solder joints completely from the connector, took the connector off itself for inspection (all looked good except for the metal contacts inside the connector were spreading further apart, Samir pushed them back together). Cleaned all the through holes for the connector, resoldered it, cleaned the soldering area with flux-off and wd40. Placed amplifier back on the EMG system, ran the HW programmer - completed successfully. Ran the HW diagnostics and auto test for the amplifier, all passed. The test log was created and saved locally to the EMG. Will wait for feedback from Dr. Lazarou.
501071 | 196279 | CM | CMPLT | CMPLT | 2023-07-26 12:19 | 2023-08-16 14:53 | MSH CEOU OPG - IU22 - System Shut Down Itself | (TC 07/26/2023 12:18:39 PM, Chau, Ken) user said it shutdown on them during use 2x -ran system test, passed -provide logs to FSE -waiting to see if PSU requires replacement (TC 07/31/2023 12:58:57 PM, Chau, Ken) Replaced PSU but system still restarting after boot up -put old PSU back -tested unit with FSE Cody and cannot find what can cause it -ran system test -Cody will be replacing UMB tomorrow (TC 08/01/2023 01:12:09 PM, Chau, Ken) New MAC registered with MSH IT, JDMIPACs tested okay UMB replaced and software reloaded by FSE -not restarting atm but will keep WO open (TC 08/16/2023 02:51:13 PM, Whittington, Scott) msh ceou rm 12 reuires measurement pkg loaded after umb replacement | Case#0122041358
| 196280 | CM | CMPLT | CMPLT | 2023-07-26 12:41 | 2023-07-31 12:53 | Battery for IVF - BPPro 3400 VSM | (TC 07/26/2023 12:41:16 PM, Samanian, Kamran) contacted Setvens and asked for the price with providing PN (TC 07/26/2023 12:41:16 PM, Samanian, Kamran) contacted Setvens and asked for the price with providing PN. S/C#1003011894 (TC 07/31/2023 12:51:28 PM, Samanian, Kamran) Received the battery and used it PO# 7202255639
1061214 | 196281 | CM | CMPLT | CMPLT | 2023-07-26 14:24 | 2023-07-26 14:25 | iPAD issue ( light controller ) - dead- black screen | (TC 07/26/2023 02:24:05 PM, Samanian, Kamran) verified the problem. very very low battery. removed the ipad from wall. disconnected and reconnected USB cable, ipad came back up with 1% charge and started charging. noticed error message wifi with host 1 however all lights were being controlled with ipad and no issue.
66762 | 196282 | CM | CMPLT | NFF | 2023-07-26 14:26 | 2023-07-26 14:57 | "broken" | (TC 07/26/2023 02:26:46 PM, Cortez, John) Completed functional test/PM successfully, Unable to fnd any errors or anything odd in the event logs
1053491 | 196283 | CM | CMPLT | CMPLT | 2023-07-27 07:58 | 2024-01-08 11:05 | Error code 20005 | (TC 07/27/2023 12:41:29 PM, Canjimavadivel, Kamallesh ) Got email from Stan regarding issue with vent. Went up to take a look and downloaded logs. Stan also reached out to vendor as it may be related to previous issue on another vent. Jason came in and downloaded logs to send to head office in Sweden. (TC 01/08/2024 11:03:41 AM, Canjimavadivel, Kamallesh ) Vent was upgraded with latest software that was supposed to address issue. Returned to circulation
65944 | 196284 | CM | CMPLT | CMPLT | 2023-07-27 08:40 | 2023-07-27 09:02 | Broken front and rear case | (TC 07/27/2023 08:57:31 AM, Baghirzade, Nara) The front and rear case were very damaged and when I opened the device, the internal housing was also damaged. Replaced: -front case -rear case -latch/lock -mylar gasket -pressure disc holder -internal housing --added the NICU barcode and asset tag on the new cases. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1062093 | 196285 | CM | CMPLT | CMPLT | 2023-07-27 09:12 | 2023-07-27 09:12 | mouse cursor movement not smooth
1049215 | 196286 | CM | CMPLT | CMPLT | 2023-07-27 09:15 | 2023-07-27 09:22 | Patch update | (TC 07/27/2023 09:15:16 AM, Canjimavadivel, Kamallesh ) Rebooted PC and successfully installed patches
1061206 | 196287 | CM | CMPLT | NFF | 2023-07-27 09:29 | 2023-07-27 09:33 | Monitors showing blue screens | (TC 07/27/2023 09:30:38 AM, Reid, Emmerson ) Went to check the issue but they resolved it before I got there
1054537 | 196288 | PM | CMPLT | CMPLT | 2023-07-27 09:32 | 2023-07-27 09:33 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/27/2023 09:32:14 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
69436 | 196289 | PM | CMPLT | CMPLT | 2023-07-27 10:29 | 2023-07-27 10:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/27/2023 10:28:40 AM, Rhuland, Courtney) Pressure and rate calibrated PM complete All tests passed post calibration Refer to ASM for details
1057743 | 196290 | CM | CMPLT | CMPLT | 2023-07-27 10:53 | 2023-07-27 10:53 | Mouse pointer lagging |  | Mouse pointer lagging
1059798 | 196291 | CM | CMPLT | CMPLT | 2023-07-27 11:00 | 2023-07-27 11:00 | trackball sticky
1065409 | 196292 | CM | CMPLT | CMPLT | 2023-07-27 11:08 | 2023-07-27 11:08 | 1 | (TC 07/27/2023 11:07:58 AM, Samanian, Kamran) received a call that VSM not functional, and showing leak. arrived in the clinic and verified the problem, replaced the cuff, tested and passed
1062999 | 196293 | PM | CMPLT | CMPLT | 2023-07-27 11:10 | 2023-07-27 11:11 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/27/2023 11:10:06 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1067051 | 196294 | CM | CMPLT | CMPLT | 2023-07-27 11:11 | 2023-10-13 08:38 | "makes odd noise", error code 1150 | (TC 07/27/2023 11:11:00 AM, Cortez, John) Confirmed issue, sent service request to hillrom (TC 10/12/2023 08:37:53 AM, Cortez, John) Vince repaired bed and provided service report
64582 | 196295 | PM | CMPLT | CMPLT | 2023-07-27 12:35 | 2023-07-27 12:36 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/27/2023 12:35:21 PM, Baghirzade, Nara) Battery replaced. Device tests passed. PM completed.
1056544 | 196296 | CM | CMPLT | CMPLT | 2023-07-27 12:37 | 2023-07-27 12:38 | Has not checked in on the defib dashboard | (TC 07/27/2023 12:36:47 PM, Baghirzade, Nara) AED did not check in on the dashboard. Manually checked it in. At first it would not check in and had to get in the supervisor menu, and did a check in from there which went through. Check in was taking longer than expected to be completed.
1067426 | 196297 | ACC | CMPLT | CMPLT | 2023-07-27 13:24 | 2023-08-04 10:24 | Incoming inspection | (TC 07/27/2023 02:15:10 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged Assembled light and put batteries on charging dock (TC 08/03/2023 10:22:52 AM, Canjimavadivel, Kamallesh ) Dropped it off with Dr. Dushenski
1067427 | 196298 | ACC | CMPLT | CMPLT | 2023-07-27 13:26 | 2023-08-04 10:24 | Incoming inspection | (TC 07/27/2023 02:16:29 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged Assembled light and put batteries on charging dock
65889 | 196299 | CM | CMPLT | CMPLT | 2023-07-27 14:00 | 2023-07-27 14:00 | Channel error | (TC 07/27/2023 01:59:48 PM, Rhuland, Courtney) Checked error log but could not replicate error Functional check - OK
1063000 | 196300 | PM | CMPLT | CMPLT | 2023-07-27 14:07 | 2023-07-27 14:07 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/27/2023 02:06:41 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1059798 | 196301 | CM | CMPLT | CMPLT | 2023-07-27 14:23 | 2023-07-27 14:23 | ordering part
| 196302 | CM | CMPLT | CMPLT | 2023-07-27 14:25 | 2023-07-27 14:25 | site coverage
| 196303 | CM | CMPLT | CMPLT | 2023-07-27 14:38 | 2023-08-02 13:35 | order for biomed stock- sterile water for injection | (TC 07/27/2023 02:47:57 PM, Cortez, John) placed order (TC 07/31/2023 01:34:51 PM, Cortez, John) order arrived
68991 | 196304 | CM | CMPLT | CMPLT | 2023-07-27 14:45 | 2023-07-27 14:45 | Bent power cord | (TC 07/27/2023 02:44:41 PM, Rhuland, Courtney) Power cord replaced Functional check - OK
| 196305 | CM | CMPLT | CMPLT | 2023-07-27 15:00 | 2023-09-15 09:27 | LVP pressure sensor order- for biomed stock | (TC 07/27/2023 03:03:59 PM, Cortez, John) Placed order for 20x sensors (TC 09/14/2023 09:26:47 AM, Cortez, John) received
| 196306 | CM | CMPLT | CMPLT | 2023-07-27 15:43 | 2023-07-27 15:43 | site coverage
| 196307 | CM | CMPLT | CMPLT | 2023-07-27 15:46 | 2023-07-27 15:46 | site coverage
| 196308 | CM | CMPLT | CMPLT | 2023-07-27 15:48 | 2023-07-27 15:48 | site coverage
| 196309 | CM | CMPLT | CMPLT | 2023-07-27 16:30 | 2023-07-27 16:31 | boom "not working" in ICU pod B room 16 | (TC 07/24/2023 04:29:37 PM, Cortez, John) Noticed that the power switch was in off position on secondary boom. Unable to find any other issues
1046787 | 196310 | CM | CMPLT | CMPLT | 2023-07-27 16:32 | 2023-08-10 09:31 | "Won't inflate" | (TC 07/24/2023 04:32:35 PM, Cortez, John) Noticed power cords were loosely connected to bed. Secured cords and added more slack. Bed inflates properly Noticed footboard was damaged (mounting rod detached from footboard) Sent service request to Hillrom (TC 08/08/2023 09:30:19 AM, Cortez, John) Bed was repaired by Vince (Hillrom)
1061420 | 196311 | CM | CMPLT | CMPLT | 2023-07-27 16:38 | 2023-07-27 16:39 | "movement functions not working" | (TC 07/24/2023 04:38:12 PM, Cortez, John) Confirmed issue Power cycled bed Tested ok
1061412 | 196312 | CM | CMPLT | CMPLT | 2023-07-28 07:08 | 2023-07-28 06:30 | Display unresponsive | (TC 07/28/2023 06:30:00 AM, Baghirzade, Nara) Display was unresponsive; reset the bed by lifting the footboard. Display is functional; checked all icons/functions. Bed is fully functional.
67917 | 196313 | CM | CMPLT | CMPLT | 2023-07-28 08:54 | 2023-07-28 09:00 | Loud noise when powering the incubator | (TC 07/28/2023 08:56:58 AM, Baghirzade, Nara) Fan was not seated properly; re-seated the fan. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
1047844 | 196314 | CM | CMPLT | NFF | 2023-07-28 09:02 | 2023-07-28 09:03 | "Not working" | (TC 07/28/2023 09:01:38 AM, Rhuland, Courtney) Checked error log and found error related to speaker Could not replicate error and all startup tests passed multiple times Ran test infusions and no fault found
1062235 | 196315 | CM | CMPLT | CMPLT | 2023-07-28 09:33 | 2023-07-28 09:34 | No function on thebed, cannot move anything | (TC 07/28/2023 09:32:59 AM, Ferreira, Qwynn) tested the bed and no function was present. Checked the head board and noticed two pins were bent. Used a flat head and repositioned the pins. Reset the bed and tested - OK All functions were restored. Bed working
| 196316 | RDS | CMPLT | CMPLT | 2023-07-28 09:49 | 2023-07-28 09:52 | Weekly Rounds: 18S and PACU PODA | (TC 07/28/2023 09:49:09 AM, Baghirzade, Nara) Weekly round completed at 18S and PACU PODA
| 196317 | RDS | CMPLT | CMPLT | 2023-07-28 09:52 | 2023-07-28 09:56 | Weekly Rounds: 11S | (TC 07/28/2023 09:52:25 AM, Baghirzade, Nara) Weekly round at 11s completed.
| 196318 | CM | CMPLT | NFF | 2023-07-28 10:37 | 2023-07-28 10:38 | SpO2 cable check | (TC 07/28/2023 10:37:35 AM, Baghirzade, Nara) SpO2 cable is functional. Cable and connectors are in good condition. Tested on different systems and is functional.
66072 | 196319 | CM | CMPLT | CMPLT | 2023-07-28 12:21 | 2023-07-28 12:21 | clip is missing | Emmerson printed new clips for the lift. tested. functional. returned.
1047469 | 196320 | CM | CMPLT | CMPLT | 2023-07-28 12:23 | 2023-07-28 12:23 | defective PBA OLED controller | Replaced defective PBA OLED controller. Tested. functional.
| 196321 | CM | CMPLT | CMPLT | 2023-07-28 12:26 | 2023-07-28 12:26 | Stryker bed scale calibration is OFF | Stryker bed scale calibration is OFF. Talked to user and explained them to remove the patient and zero to sclae. They will wait.
| 196322 | CM | CMPLT | CMPLT | 2023-07-28 12:32 | 2023-08-01 08:31 | Troubleshooting with Masimo | (TC 07/28/2023 12:31:49 PM, Baghirzade, Nara) Troubleshooting remotely with Maismo tech and Naz regarding the key lock authentication failure error. Swapped the ethernet cable for MyUTN but unable to verify whether connectivity with UTN is good. Masimo tech was able to ping and connect to UTN but seems like the server and the UTN are having connection issues. Samuel (masimo tech) will look into it and inquire wiht his team if we can bypass the UTN.
| 196323 | RDS | CMPLT | CMPLT | 2023-07-28 13:36 | 2023-07-28 13:39 | Weekly Rounds: NICU | (TC 07/28/2023 01:36:20 PM, Baghirzade, Nara) Weekly rounds in NICU completed.
| 196324 | RDS | CMPLT | CMPLT | 2023-07-28 13:39 | 2023-07-28 13:41 | Weekly rounds: RESUS | (TC 07/28/2023 01:39:25 PM, Baghirzade, Nara) Weekly rounds in RESUS is completed.
| 196325 | CM | CMPLT | CMPLT | 2023-07-28 13:44 | 2023-07-28 13:49 | 10N and 10S Masimo viewstations off network | (TC 07/28/2023 01:43:46 PM, Cortez, John) Worked with Kamran to troubleshoot issue. Apparently IT swapped the switch (#5) on the 10th floor closet yesterday and didn't reconfigure the port for 10S. We were unable to trace cable for 10N from patch panel to switch, so Mirza from IT reconfigured another port on switch #4. Issue was resolved
1047581 | 196326 | CM | CMPLT | CMPLT | 2023-07-28 13:46 | 2023-07-28 13:47 | Functional check and power cord change | (TC 07/28/2023 01:45:43 PM, Baghirzade, Nara) Replaced power cord. Functional tests passed. Ran a feed at various rates and no issues/errors during the feed. Occluded the line and the occlusion alarm was functional.
63928 | 196327 | CM | CMPLT | CMPLT | 2023-07-28 13:55 | 2023-07-28 14:00 | "not working good" | (TC 07/24/2023 01:54:57 PM, Cortez, John) Tested with multiple simulators, tested ok. Leads, connectors are all good At times, printer gets jammed and device halts printing. Says "printer door opened". Reinstalled paper. No fault found. Probably an electrode issue, some electrodes were outside the pack and dried out. Tested new electrodes on myself and ECG signal was stable
| 196328 | RDS | CMPLT | CMPLT | 2023-07-28 14:04 | 2023-07-28 14:13 | Weekly check for ICU, July 24 | (TC 07/28/2023 02:04:25 PM, Cortez, John) Checked rooms 7,8,14,16 Checked hallways, equipment room Swapped pumps due for PM, did functional test on aspirators Found overview station between 14/15 on login screen, logged in Boom on 16 was switched off (TC 07/24/2023 02:04:25 PM, Cortez, John) Checked rooms 7,8,14,16 Checked hallways, equipment room Swapped pumps due for PM, did functional test on aspirators Found overview station between 14/15 on login screen, logged in Boom on 16 was switched off
68224 | 196329 | CM | CMPLT | CMPLT | 2023-07-28 14:32 | 2023-07-31 10:17 | No ECG tracing | (TC 07/31/2023 10:15:43 AM, Scero, Tomlyn) In the room nurse says there's no ECG tracing despite new leads. I borught the unit downstairs and tested ECG with a simulator and it's working fine. Possible electrode issues unknown. Put the X2 on the shelf for use as a spare.
63031 | 196330 | CM | CMPLT | CMPLT | 2023-07-28 14:33 | 2023-09-05 14:32 | Battery doesn't charge | (TC 07/24/2023 02:32:49 PM, Cortez, John) Swapped batteries, still doesn't charge Swapped aspirator with spare (TC 08/14/2023 03:45:26 PM, Cortez, John) Swapped power supply board with one from spare aspirator. Seems to now charge/work under battery power. However, swapped original board back and it still functions normally (TC 09/05/2023 02:30:19 PM, Cortez, John) Functional test passed
1065004 | 196331 | FC | CMPLT | CMPLT | 2023-07-28 14:54 | 2023-07-28 14:55 | Functional test | (TC 07/24/2023 02:54:21 PM, Cortez, John) Functional test passed
1065021 | 196332 | FC | CMPLT | CMPLT | 2023-07-28 15:04 | 2023-07-28 15:05 | Functional check | (TC 07/24/2023 03:04:06 PM, Cortez, John) Functional check passed
1057054 | 196333 | CM | CMPLT | CMPLT | 2023-07-28 15:08 | 2023-07-28 15:08 | All 4 workstations are not recieving pre-fetches |  | All 4 Mamo workstations are not recieving pre-fetches 12th floor called Hologic case # 01329893 I talked to Paul in Hologic IT He asked me to get A numbers for 3 patients. The numbers I gave him are 858238902 858064190 858015528
32510 | 196334 | CM | CMPLT | CMPLT | 2023-07-28 15:10 | 2023-07-28 15:10 | detector cover does not latch | fixed latches on both sides
| 196335 | CM | CMPLT | CMPLT | 2023-07-28 15:12 | 2023-07-28 15:12 | Site coverage
1065016 | 196336 | FC | CMPLT | CMPLT | 2023-07-28 15:49 | 2023-07-28 15:50 | Functional check | (TC 07/24/2023 03:49:28 PM, Cortez, John) Functional check passed
1065017 | 196337 | FC | CMPLT | CMPLT | 2023-07-28 15:52 | 2023-07-28 15:52 | Functional check | (TC 07/24/2023 03:51:43 PM, Cortez, John) Functional check passed
1065023 | 196338 | FC | CMPLT | CMPLT | 2023-07-28 15:54 | 2023-07-28 15:55 | Functional check | (TC 07/24/2023 03:54:19 PM, Cortez, John) Functional check passed
1065007 | 196339 | FC | CMPLT | CMPLT | 2023-07-28 15:56 | 2023-07-28 15:57 | Functional check | (TC 07/24/2023 03:56:18 PM, Cortez, John) Functional check passed
32560 | 196340 | CM | CMPLT | CMPLT | 2023-07-30 16:39 | 2023-07-30 16:37 | AC unit flood to CT area | AC unit flood to CT area CT shutdown July 28 CT powered up/tested July 30 CT phantom tested OK
| 196341 | RDS | CMPLT | CMPLT | 2023-07-31 08:28 | 2023-07-31 09:06 | Weekly check for ED
| 196342 | RDS | CMPLT | CMPLT | 2023-07-31 08:29 | 2023-08-01 13:38 | Weekly check for ICU
| 196343 | RDS | CMPLT | CMPLT | 2023-07-31 08:56 | 2023-07-31 09:21 | Weekly Checks For PACU | (TC 07/31/2023 09:20:03 AM, Reid, Emmerson ) Done.
1054521 | 196344 | CM | CMPLT | CMPLT | 2023-07-31 08:57 | 2023-07-31 10:26 | Broken PIM holder | (TC 07/31/2023 10:20:02 AM, Canjimavadivel, Kamallesh ) Replaced arm assembly with spare from storage Replaced damaged keyboard cover Tested ok
62989 | 196345 | CM | CMPLT | CMPLT | 2023-07-31 09:04 | 2023-07-31 09:05 | Missing adjustment knob | (TC 07/31/2023 09:04:00 AM, Baghirzade, Nara) Replaced the adjustment knob. Tightened the screw. All functional tests passed.
1067395 | 196346 | CM | CMPLT | CMPLT | 2023-07-31 09:14 | 2023-08-01 08:30 | Organizing pickup of loaner | (TC 07/31/2023 09:17:36 AM, Baghirzade, Nara) Called 3M to organize a pickup of the loaner from biomed. Hospital owned VAC already received. Confirmation #: 12772934 (TC 08/01/2023 08:27:39 AM, Baghirzade, Nara) Loaner was picked up. Attached pick up confirmation.
1062147 | 196347 | CM | CMPLT | CMPLT | 2023-07-31 09:34 | 2023-07-31 09:41 | No power to bed | (TC 07/31/2023 09:38:04 AM, Ferreira, Qwynn) received call from 12S for bed not powering on Went to check and confirmed bed had no power despite being plugged in Checked head board and the pins were bent slightly - maybe issue? Fixed the pins and reset headboard as well as power switch Bed powered on and checked all functions - Working Let the front clerk know that the bed is ready to go.
66432 | 196348 | CM | CMPLT | CMPLT | 2023-07-31 10:47 | 2023-07-31 10:56 | Not turning on | (TC 07/31/2023 10:55:14 AM, Scero, Tomlyn) Replaced battery adaptor and tested ECG with my ECG simulator. Telepack working fine.
| 196349 | CM | CMPLT | CMPLT | 2023-07-31 11:37 | 2023-07-31 11:39 | Weekly Check in CCU | (TC 07/31/2023 11:37:02 AM, Scero, Tomlyn) Weekly checks complete.
| 196350 | CM | CMPLT | CMPLT | 2023-07-31 12:46 | 2023-07-31 12:47 | ENDO Emrg Cart monitor | (TC 07/31/2023 12:46:33 PM, Samanian, Kamran) unit mentioned the monitor tilted. verified it , it is heavy . a bit adjustment applied on the arm. Porvided informaytion to ENDO manager
| 196351 | CM | CMPLT | CMPLT | 2023-07-31 13:35 | 2023-07-31 13:36 | Meeting | (TC 07/31/2023 01:35:09 PM, Rhuland, Courtney) Huddle + bi-weekly meeting
| 196352 | CM | CMPLT | CMPLT | 2023-07-31 14:24 | 2023-07-31 14:24 | meeting with Olympus for VPN access | (TC 07/31/2023 02:24:23 PM, Samanian, Kamran) done!
1067436 | 196353 | ACC | CMPLT | CMPLT | 2023-07-31 14:37 | 2023-08-01 10:39 | Incoming inspection | (TC 07/31/2023 03:03:20 PM, Canjimavadivel, Kamallesh ) Unboxed and added to database Asset tagged (TC 08/01/2023 10:38:29 AM, Canjimavadivel, Kamallesh ) Labelled power cord Dropped to unit for labelling
1067437 | 196354 | ACC | CMPLT | CMPLT | 2023-07-31 14:40 | 2023-08-01 10:38 | Incoming inspection | (TC 08/01/2023 10:37:43 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged Labelled power cord Dropped to unit for labelling
1067438 | 196355 | ACC | CMPLT | CMPLT | 2023-07-31 14:42 | 2023-08-01 10:37 | Incoming inspection | (TC 08/01/2023 10:36:36 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged Labelled power cord Dropped to unit for labelling
| 196356 | CM | CMPLT | CMPLT | 2023-07-31 14:55 | 2023-08-02 10:31 | NAS order | (TC 07/31/2023 03:01:36 PM, Samanian, Kamran) S/C#1003015342 (TC 07/31/2023 03:01:36 PM, Samanian, Kamran) S/C#1003015342 PO#7202258262 (TC 08/02/2023 10:31:12 AM, Samanian, Kamran) order received
26151 | 196357 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26795 | 196358 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-17 15:18 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 08/17/2023 03:17:51 PM, Reid, Emmerson ) PM Complete, see attached PM Checklist for details
51581 | 196359 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52114 | 196360 | PM | CMPLT | NA | 2023-08-01 03:21 | 2023-08-01 09:34 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/01/2023 09:31:41 AM, Rhuland, Courtney) Repaired bed tilt mechanism Began recall and noticed humidifier needed replacing When removing humidifier, chassis broke where attached Unit to be retired and used for spare parts
52234 | 196361 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
60088 | 196362 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-19 09:37 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/19/2023 09:37:08 AM, Samanian, Kamran) No longer owned by ENDO. Moved from ENDO account to Biomed Account- No more X2 in ENDO
62600 | 196363 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-19 09:39 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/19/2023 09:38:56 AM, Samanian, Kamran) No longer owned by ENDO. Moved from ENDO account to Biomed Account- No more X2 in ENDO
63348 | 196364 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-17 15:12 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 08/17/2023 03:11:27 PM, Reid, Emmerson ) PM Compelete, see checklist for results.
63349 | 196365 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Smoke Evacuation Systems, Surgical
63484 | 196366 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
64075 | 196367 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-02-09 08:21 | PM ANNUAL: Power Systems, Uninterruptible | (TC 02/09/2024 08:19:54 AM, Canjimavadivel, Kamallesh ) Was kept in biomed. Batteries replaced and installed in 12th floor closet. Updated asset info
64760 | 196368 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-06-24 14:53 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 06/24/2024 02:52:35 PM, Cortez, John) Replaced battery
64954 | 196369 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Smoke Evacuation Systems, Surgical
64973 | 196370 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Smoke Evacuation Systems, Surgical
65262 | 196371 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-11-09 07:35 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 73uA and 88 ohm. RT need the unit could not finish the PM! Finished PM. Functional.
65885 | 196372 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65929 | 196373 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-13 14:29 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 02:28:25 PM, Rhuland, Courtney) PM completed by BD
65941 | 196374 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-27 10:28 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 10/27/2023 10:27:06 AM, Rhuland, Courtney) Replaced rear case (recall) PM complete All tests passed Refer to ASM for details
65942 | 196375 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
66007 | 196376 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-27 14:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/27/2023 01:42:04 PM, Rhuland, Courtney) PM complete All tests passed Refer to ASM for details
66046 | 196377 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-27 13:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2024 02:59:20 PM, Reid, Emmerson ) Replaced the kaypad due to damage
66049 | 196378 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66062 | 196379 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-14 10:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 10:27:18 AM, Huang, Jessica) replaced the power cord and completed PM
66063 | 196380 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66349 | 196381 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Smoke Evacuation Systems, Surgical
66467 | 196382 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-12 10:34 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 73uA and 88 ohm.
66578 | 196383 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: RACKS, MODULE
66659 | 196384 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-06-24 14:46 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 06/24/2024 02:45:57 PM, Cortez, John) Replaced batteries
66706 | 196385 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-12-11 12:56 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 12/11/2023 12:54:02 PM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. Replaced RH sensor as it failed calibration 3 times. SE door latch was replaced (found broken). VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. No errors found in error log. Time was updated as it was off by a few minutes. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Heater is fully functional. Functional tests passed. Electrical safety results - Leakage current = 48.0 uA Power cord resistance = 77.0 mohms
66772 | 196386 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | PM schedule changed to 3 years.
66773 | 196387 | PM | CNP | EIU | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66774 | 196388 | PM | CNP | EIU | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66776 | 196389 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66782 | 196390 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66788 | 196391 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66868 | 196392 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-14 09:38 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/11/2023 09:38:04 AM, Rendon Lopez , Elizabeth) Battery replaced, PM completed, all tests passed
66871 | 196393 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66873 | 196394 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-19 10:31 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 10/19/2023 10:29:56 AM, Scero, Tomlyn) Battery replaced. PM tests complete and passed. No leaks
67175 | 196395 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67755 | 196396 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67823 | 196397 | PM | CMPLT | UNL | 2023-08-01 03:21 | 2023-09-28 10:52 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/27/2023 10:51:25 AM, Rendon Lopez , Elizabeth) Unable to locate
67892 | 196398 | PM | CMPLT | UNL | 2023-08-01 03:21 | 2024-03-13 13:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 03/11/2024 01:21:16 PM, Rendon Lopez , Elizabeth) Unable to locate for 4+ years. Retire from service
67934 | 196399 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-12-18 14:52 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 12/18/2023 02:37:31 PM, Rhuland, Courtney) PM complete Calibrated: Line voltage Ref voltage Main voltage Scale Functional check - 35C, 75% humidity Read - 35C, 74.6% humidity Electrical safety: leakage - 79uA Resistance - 219 mOhms
67952 | 196400 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-11-15 10:58 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/15/2023 10:44:43 AM, Rhuland, Courtney) Repaired broken bed tilt Calibratd LV and main voltage Humidity failed repeatedly during PM after calibration Replaced RH sensor Completed latch recall PM passed Calibrated and verified scale Functional check - OK Electrical safety - OK Ground leakage - 85uA Resistance - 22mOhms
67960 | 196401 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-12 08:45 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 10/12/2023 08:39:52 AM, Rhuland, Courtney) Repaired hood hinge parts Replaced worn side door hinge LV and main voltage calibrated Scale calibrated Functional check - OK Temp set to 35 - read 35 Humidity set to 75% - read 75.5% Electrical safety - OK
68072 | 196402 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-01-23 14:24 | 5029 | (TC 01/23/2024 02:23:46 PM, Baghirzade, Nara) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 19 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 77 cmH2O Less than 80 = 73 cmH2O (TC 01/23/2024 02:24:09 PM, Rhuland, Courtney) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 19 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 77 cmH2O Less than 80 = 73 cmH2O
68288 | 196403 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68598 | 196404 | PM | CMPLT | PMFAIL | 2023-08-01 03:21 | 2023-09-26 12:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/26/2023 12:47:00 PM, Campbell, Connor) Plunger position failed in PM Plunger position cannot be calibrated The plunger is much too sticky and makes it impossible to use Ordering a new plunger, will perform the PM and functional check when it arrives
68648 | 196405 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-21 18:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68674 | 196406 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68682 | 196407 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-04-24 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68697 | 196408 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-02 14:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 02:57:25 PM, Canjimavadivel, Kamallesh ) Cleaned PC Replaced battery PM tests pass
68701 | 196409 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-26 13:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/21/2024 02:58:50 PM, Reid, Emmerson ) Replaced the power cord
68714 | 196410 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-04 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/03/2023 09:21:17 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
68729 | 196411 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68762 | 196412 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68765 | 196413 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68774 | 196414 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68779 | 196415 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-05 11:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 11:37:21 AM, Reid, Emmerson ) Replaced the battery. Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
68791 | 196416 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68849 | 196417 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-04-30 13:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/30/2024 01:28:21 PM, Rhuland, Courtney) Sent out for PM - completed by BD Replaced battery Loaded network and data set
68862 | 196418 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-25 14:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 08:39:27 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 45 mohms Leakage current: 61.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/25/2023 02:09:32 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3784 mamphr New battery capacity: 4008 mamphr PM completed.
68896 | 196419 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-05-08 15:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 02/13/2024 09:46:54 AM, Cortez, John) Replaced battery. Conditioned overnight. Batt capacity: 4135 mAhrs
68906 | 196420 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-02-02 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 01/31/2024 02:59:20 PM, Scero, Tomlyn) PM complete and tests passed. Battery replaced. (TC 02/02/2024 08:50:54 AM, Scero, Tomlyn) electrical safety tests complete.
68917 | 196421 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68929 | 196422 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68943 | 196423 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-30 11:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/30/2023 11:43:55 AM, Campbell, Connor) battery replaced refer to procedure for ESA PM complete/passed, refer to ASM
68944 | 196424 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68952 | 196425 | PM | CANCL | CANCL | 2023-08-01 03:21 | 2024-03-26 13:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68984 | 196426 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68990 | 196427 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68993 | 196428 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69008 | 196429 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-18 07:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/17/2023 12:44:06 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 40 mohms Leakage current: 66.3 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/18/2023 07:40:48 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3393 mamphr New battery capacity: 4048 mamphr PM completed.
69094 | 196430 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-05 11:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 11:05:11 AM, Campbell, Connor) PM complete/passed, refer to ASM
69100 | 196431 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-20 13:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/20/2023 01:24:09 PM, Cortez, John) Cleaned, Downloaded error/event logs, LVP display dim, replaced Calibrated infusion rate, PM completed successfully, results logged in ASM
69125 | 196432 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-27 10:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/21/2024 03:16:20 PM, Reid, Emmerson ) Replaced the front casing
69126 | 196433 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-15 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/15/2023 03:10:10 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69160 | 196434 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69171 | 196435 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69183 | 196436 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-19 20:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69201 | 196437 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-24 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/24/2023 10:22:12 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69202 | 196438 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69205 | 196439 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69247 | 196440 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-30 07:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 07:41:19 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69251 | 196441 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-07 14:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 02:32:24 PM, Campbell, Connor) Pressure and rate calibrated PM complete, refer to ASM supervised by nara
69257 | 196442 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-31 09:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/31/2023 09:15:10 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69260 | 196443 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-14 14:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:41:16 PM, Huang, Jessica) replaced the female iui connector and completed PM
69283 | 196444 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69289 | 196445 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-13 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/13/2023 01:44:35 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69428 | 196446 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-20 16:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/20/2023 04:10:21 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69438 | 196447 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-25 11:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69443 | 196448 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-26 16:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/26/2024 04:58:10 PM, Cortez, John) Replaced patient side pressure sensor, calibrated PM complete (TC 03/26/2024 04:59:33 PM, Cortez, John) Forgot to check bezel
69485 | 196449 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-18 08:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/18/2023 08:54:35 AM, Campbell, Connor) PM complete/passed, refer to ASM
69528 | 196450 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-13 10:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/13/2023 10:07:23 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69584 | 196451 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-14 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 01:59:18 PM, Huang, Jessica) replaced the female iui connector and completed PM
69617 | 196452 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-04-02 16:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69647 | 196453 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-18 10:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/18/2023 10:42:33 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69671 | 196454 | PM | CANCL | CANCL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
530501 | 196455 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Integrated Power Console
67359 | 196456 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-21 08:31 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 09/21/2023 08:29:46 AM, Campbell, Connor) Drained and cleaned O-rings replaced All alarm tests passed Temp check passed Refilled Ground leakage - 13.4 Ground resistance - 315mOhms PM complete/passed
67360 | 196457 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-28 12:51 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 09/21/2023 02:54:37 PM, Campbell, Connor) Drained, cleaned, and filled with new cleaning fluid Replaced o-rings General Alarm Test - OK Low reservoir Alarm Test - OK Interlock Alarm Test - OK Over Temp Alarm Test - OK Electrical Safety -Current Leakage - 13.6uA -Ground Resistance - 80 mOhm Temp check needs calibration, will refill solution and close the PM once temp check is calibrated and confirmed. (TC 09/22/2023 08:38:12 AM, Campbell, Connor) confirmed TEMP is ok refilled solution PM complete/passed
32540 | 196458 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-11 14:46 | PM - Monthly | PM Major done NRA/AIR cal L FOV done Slice count - 621850 TOS phantom a(-997.7, 10.3) b(335.4, 14.1) c(134.2, 14.1) d(98.1, 14.7) e(-105.1, 13.4) f(1.7, 14.3) -System passed all tests
32560 | 196459 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-05 14:23 | PM - Monthly | PM completed
32580 | 196460 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-06 20:45 | PM - Monthly | PM cancelled due to service issues at TGH
500230 | 196461 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-11-28 10:57 | PM - Annual
1046059 | 196462 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1046087 | 196463 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-26 08:19 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/26/2023 08:18:02 AM, Campbell, Connor) replaced water chamber front plate PM complete/passed, refer to procedure
1046207 | 196464 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-01-02 12:16 | PM - Electrical Safety | (TC 12/29/2023 12:16:18 PM, Fan, Tak) Max A = 0.5 / 10.9 uA
1046273 | 196465 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-15 13:12 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/30/2023 03:24:33 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL pre (TC 09/15/2023 01:12:39 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046274 | 196466 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-14 21:28 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/30/2023 03:28:18 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 09/14/2023 09:28:01 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046275 | 196467 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-14 21:29 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/31/2023 11:44:46 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL pre (TC 09/14/2023 09:28:50 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046276 | 196468 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-25 09:29 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/31/2023 11:45:43 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL pre (TC 09/25/2023 09:29:21 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046286 | 196469 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-18 10:16 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 08/18/2023 10:15:17 AM, Reid, Emmerson ) PM Complete, see attached PM Checklist for details.
1046397 | 196470 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1046399 | 196471 | PM | ACTIV | ISSUE | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1046451 | 196472 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-08-18 10:18 | Annual PM | (TC 08/18/2023 10:17:10 AM, Reid, Emmerson ) PM Complete, see attached PM Checklist for details.
1046453 | 196473 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-11-03 10:38 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 11/02/2023 10:37:18 AM, Reid, Emmerson ) PM Complete, see attached report
1047045 | 196474 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-11 19:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047096 | 196475 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-10-27 13:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/27/2023 01:42:32 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 62 mohms Leakage current: 56.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
1047097 | 196476 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-03-20 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047104 | 196477 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-01 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 02:21:26 PM, Canjimavadivel, Kamallesh ) Cleaned PCU Replaced battery PM complete
1047144 | 196478 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-05-08 14:26 | Annual PM | (TC 03/04/2024 02:24:18 PM, Scero, Tomlyn) PM completed by GE Healthcare March 2024. Service report is attached.
1047159 | 196479 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2024-05-08 14:29 | Annual PM | (TC 03/01/2024 02:27:44 PM, Scero, Tomlyn) PM completed by GE Healthcare. Service report attached.
1047472 | 196480 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-11-08 12:39 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 6th year PM after 5th year PM last year. No GS500. Replaced PS500 batteries, V battery, air filter, O2 filter and O2 pressure regulator. Calibrated. Tested. Functional. SW.2.51 Leakage: 82uA and 90 ohm.
1047829 | 196481 | PM | CANCL | CANCL | 2023-08-01 03:21 | 2024-09-12 14:59 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
1047862 | 196482 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-28 14:27 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 09/28/2023 01:42:53 PM, Baghirzade, Nara) Samir did some testing with the board but as a result the configuration was erased. Assembled the pump back together but there were 2 parts missing: the holder for the plunger wire and the screw that keeps the holder in place. Samir used a copper terminal end instead to help keep the wire in place. Reloaded the configuration to the latest: March 2019. First load of the configuration failed as it kept timing out. Second load of the configuration was successful. Confirmed that the configuration was uploaded successfully. Updated date and time. (TC 09/28/2023 02:25:10 PM, Baghirzade, Nara) PM passed and completed. Refer to procedure for results. Ran a basic infusion multiple times at varying rates; all infusions were completed successfully. Ran the flow delivery accuracy test twice to ensure that the results are within the limits; both times the flow delivery accuracy was well within the limits. Pump is fully functional.
1048472 | 196483 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-15 16:47 | HARP - Semi Annual
1048472 | 196484 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-15 16:48 | PM |  | PM Minor
1048537 | 196485 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-22 16:51 | HARP - Semi-Annual | HARP
1048537 | 196486 | PM | CMPLT | CMPLT | 2023-08-01 03:21 | 2023-09-21 22:52 | PM - Semi-Annual | PM
1048646 | 196487 | PM | CNP | ENL | 2023-08-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048653 | 196488 | PM | CNP | ENL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048674 | 196489 | PM | CNP | ENL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048678 | 196490 | PM | CNP | ENL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048695 | 196491 | PM | CNP | ENL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049057 | 196492 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-08-09 10:17 | Planned Maintenance | (TC 08/09/2023 10:10:03 AM, Canjimavadivel, Kamallesh ) Set points: 43.5 38.2 32.4 Low temp alarm: 32.8 Primary over temp: 48.5 Secondary over temp: 52.9 PM complete
1050088 | 196493 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-10-16 16:06 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 10/10/2023 04:02:50 PM, Rendon Lopez , Elizabeth) 4 year PM parts replaced (TC 10/11/2023 04:02:50 PM, Rendon Lopez , Elizabeth) PM tests in progress (TC 10/13/2023 04:02:50 PM, Rendon Lopez , Elizabeth) Running ventilation mode tests, O2 reading unstable, intermitently reading 21% and not reaching higher than 94%. Troubleshooting (TC 10/16/2023 04:02:50 PM, Rendon Lopez , Elizabeth) Connected AGM to Visia, calibrated PGA O2. Tested all ventilation functions, readings Okay. PM completed, see attached report. Returned device to OR 5
1050249 | 196494 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-08-04 08:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/04/2023 08:07:14 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 37 mohms Leakage current: 50.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
1051358 | 196495 | PM | CANCL | CANCL | 2023-08-01 03:22 | 2024-10-02 09:08 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
1052940 | 196496 | PM | CANCL | CANCL | 2023-08-01 03:22 |  | PM - Electrical Safety | Probe damaged.
1053348 | 196497 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-08-30 20:21 | MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG | (TC 08/30/2023 03:18:34 PM, Whittington, Scott) MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG (TC 08/30/2023 03:19:35 PM, Whittington, Scott) MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG (TC 08/30/2023 03:19:46 PM, Dolinar, John) MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG (TC 08/30/2023 06:19:00 PM, Whittington, Scott) MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG | MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG
1055097 | 196498 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-03-15 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055164 | 196499 | PM | CNP | ENL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055168 | 196500 | PM | CNP | ENL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055773 | 196501 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-02-27 10:30 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 02/26/2024 12:57:39 PM, Rhuland, Courtney) Transmitted logs stored on defib Deleted all logs off defib Pushed software and CP upgrade Enabled automatic log transfer and daylight savings time change (TC 02/27/2024 10:19:57 AM, Rhuland, Courtney) PM complete See attached checklist for details
1055776 | 196502 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-18 09:20 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 09/14/2023 09:19:21 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached checklist
1055787 | 196503 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-18 09:15 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 09/11/2023 09:14:14 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached checklist
1056665 | 196504 | PM | CANCL | CANCL | 2023-08-01 03:22 | 2024-08-23 13:17 | PM-MAJOR_ Annual PM for: Microscope -Surgical
1056667 | 196505 | PM | ACTIV | PV | 2023-08-01 03:22 |  | PM-MAJOR-Annual PM for : Microscope- Opthalmology
1056911 | 196506 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | PM-MAJOR_ Annual PM for: Microscope -Surgical
1057568 | 196507 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-08-18 10:11 | PM MAJOR - ANNUAL PM for: Electrosurgical Unit | (TC 08/18/2023 10:10:53 AM, Reid, Emmerson ) PM Complete, see attached PM Checklist for details.
1058267 | 196508 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-18 10:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 09:46:05 AM, Rendon Lopez , Elizabeth) Battery replaced, charging (TC 09/15/2023 10:51:22 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
1058272 | 196509 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-11-09 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/09/2023 01:28:56 PM, Campbell, Connor) PM passed/complete refer to procedure for ESA
1059795 | 196510 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-15 16:19 | HARP | HARP
1059795 | 196511 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-14 16:21 | PM - Semi-Annual | PM
1059796 | 196512 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-08 16:16 | HARP | HARP
1059796 | 196513 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-15 16:17 | PM - Semi-Annual | PM
1059810 | 196514 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-08 16:22 | HARP | HARP
1059810 | 196515 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-07 16:23 | PM - Semi-Annual | PM
1061335 | 196516 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-20 08:59 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 09/19/2023 08:59:01 AM, Rendon Lopez , Elizabeth) PM completed, passed
1061348 | 196517 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2023-09-06 08:14 | Planned Maintenance | (TC 09/05/2023 07:50:03 AM, Canjimavadivel, Kamallesh ) PM complete, see attached checklist
1064666 | 196518 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-02-09 09:08 | PM-MAJOR - Annual PM For: Electrosurgical Units
1064692 | 196519 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064693 | 196520 | PM | CANCL | CANCL | 2023-08-01 03:22 | 2024-10-04 09:34 | Annual PM
1064694 | 196521 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064695 | 196522 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064696 | 196523 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064700 | 196524 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064701 | 196525 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064702 | 196526 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064703 | 196527 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064704 | 196528 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064705 | 196529 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064706 | 196530 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064713 | 196531 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064714 | 196532 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064715 | 196533 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064716 | 196534 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064717 | 196535 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064718 | 196536 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064719 | 196537 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064720 | 196538 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064721 | 196539 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064722 | 196540 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064723 | 196541 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064724 | 196542 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064725 | 196543 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064726 | 196544 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064727 | 196545 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064728 | 196546 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064729 | 196547 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064730 | 196548 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
1064733 | 196549 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1064734 | 196550 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual PM
1065041 | 196551 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:03:21 PM, Rhuland, Courtney) PM completed by BD
1065059 | 196552 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065060 | 196553 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065061 | 196554 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:20:40 PM, Rhuland, Courtney) PM completed by BD
1065062 | 196555 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065063 | 196556 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-26 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065064 | 196557 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065065 | 196558 | PM | CANCL | CANCL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065066 | 196559 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-26 10:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065067 | 196560 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065068 | 196561 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065069 | 196562 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-26 10:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065070 | 196563 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-26 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065071 | 196564 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:13:20 PM, Rhuland, Courtney) PM completed by BD
1065072 | 196565 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065073 | 196566 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065074 | 196567 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:23:39 PM, Rhuland, Courtney) PM completed by BD
1065075 | 196568 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065076 | 196569 | PM | CNP | ENL | 2023-08-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065077 | 196570 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:26:55 PM, Rhuland, Courtney) PM completed by BD
1065078 | 196571 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:08:07 PM, Rhuland, Courtney) PM completed by BD
1065079 | 196572 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065080 | 196573 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065081 | 196574 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065082 | 196575 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065083 | 196576 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:22:15 PM, Rhuland, Courtney) PM completed by BD
1065084 | 196577 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065085 | 196578 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065086 | 196579 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:04:28 PM, Rhuland, Courtney) PM completed by BD
1065087 | 196580 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:06:57 PM, Rhuland, Courtney) PM completed by BD
1065088 | 196581 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:16:50 PM, Rhuland, Courtney) PM completed by BD
1065089 | 196582 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065090 | 196583 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065091 | 196584 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-24 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065092 | 196585 | PM | CMPLT | CMPLT | 2023-08-01 03:22 | 2024-04-04 13:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:25:13 PM, Rhuland, Courtney) PM completed by BD
1066122 | 196586 | PM | ACTIV | ISSUE | 2023-08-01 03:22 |  | Annual Maintenance
| 196587 | CM | CMPLT | CMPLT | 2023-08-01 07:34 | 2023-07-31 07:34 | huddle meeting | (TC 07/31/2023 07:34:33 AM, Samanian, Kamran) done!
1066114 | 196588 | FC | CMPLT | CMPLT | 2023-08-01 07:43 | 2023-08-01 07:56 | QC check | (TC 08/01/2023 07:55:54 AM, Canjimavadivel, Kamallesh ) QC check complete
| 196589 | PA | CMPLT | CMPLT | 2023-08-01 08:31 | 2023-08-02 14:49 | Water bags for stock | (TC 08/01/2023 08:33:50 AM, Canjimavadivel, Kamallesh ) Placed order for water
| 196590 | CM | CMPLT | CMPLT | 2023-08-01 08:44 | 2023-08-08 11:54 | Ordering syringe pressure sensors for biomed stock | (TC 08/01/2023 08:57:32 AM, Baghirzade, Nara) SC 1003015658 (TC 08/08/2023 11:53:14 AM, Baghirzade, Nara) PO 7202258743
| 196591 | CM | CMPLT | CMPLT | 2023-08-01 08:50 | 2023-08-14 09:57 | Battery adaptors for telemetry packs | (TC 08/14/2023 09:56:44 AM, Scero, Tomlyn) received and put on stock shelf
| 196592 | CM | ACTIV | ISSUE | 2023-08-01 09:00 |  | uploaded the Alaris pump Dataset | (TC 08/01/2023 09:00:04 AM, Chukri, Samir) Uploaded the new dataset emaild confirm with Pump team and checked with pcu pump and emailed to the rest of biomed team to update the Alaris system Maintenance. (TC 08/10/2023 02:28:59 PM, Chukri, Samir) dataset transfer update and emailed to Leyla. (TC 12/07/2023 02:35:29 PM, Chukri, Samir) discussed with Jessica and elizabeth regarding the Alaris pump wi-fi connection in BH
| 196593 | PA | CMPLT | CMPLT | 2023-08-01 09:08 | 2023-08-08 14:36 | Sterile Water
1067439 | 196594 | CM | CMPLT | CMPLT | 2023-08-01 09:32 | 2023-08-01 14:25 | 6 year PM | (TC 08/01/2023 02:24:37 PM, Scero, Tomlyn) PM parts replaced and tests complete. No leaks
1067440 | 196595 | CM | CMPLT | CMPLT | 2023-08-01 09:36 | 2023-08-01 14:24 | 4 year PM | (TC 08/01/2023 02:23:52 PM, Scero, Tomlyn) PM parts replaced and tests complete. No leaks
68706 | 196596 | PM | CMPLT | CMPLT | 2023-08-01 09:42 | 2023-08-01 09:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/01/2023 09:42:04 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 77 mohms Leakage current: 69.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
67944 | 196597 | CM | CMPLT | CMPLT | 2023-08-01 09:46 | 2023-08-01 09:47 | System failure 25 | (TC 08/01/2023 09:45:53 AM, Baghirzade, Nara) Fan was improperly seated; re-seated the fan. Replaced 9V battery with rechargeable 9V battery. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
1061209 | 196598 | CM | CMPLT | CMPLT | 2023-08-01 10:24 | 2023-08-01 10:26 | Monitors 1 and 3 not holding up | (TC 08/01/2023 10:24:31 AM, Reid, Emmerson ) The monitors were tilting forward and needed to b tightened. I did this for both monitors as well as tighten the horizontal rotation. Working now.
69664 | 196599 | CM | CMPLT | CMPLT | 2023-08-01 10:26 | 2023-08-01 10:27 | Door not closing/latching alarm | (TC 08/01/2023 10:26:20 AM, Rhuland, Courtney) Downloaded error logs - no error related found Ran check in and functional - OK No fault found Refer to ASM for details
52114 | 196600 | CM | CMPLT | CMPLT | 2023-08-01 12:58 | 2023-08-01 13:01 | Dismantling and disposal of incubator hood | (TC 08/01/2023 12:59:37 PM, Baghirzade, Nara) Dismantled the hood and took off all salvageable parts such as the air/humidity probes, latches, grommets, port doors, and door parts. Disposed of the hood.
32510 | 196601 | CM | CMPLT | CMPLT | 2023-08-01 13:10 | 2023-08-01 13:09 | MSH DR1 GUARD PROTECT | (TC 08/01/2023 01:09:44 PM, Whittington, Scott) MSH DR1 GUARD PROTECT | MSH DR1 GUARD PROTECT
1056125 | 196602 | PM | CMPLT | CMPLT | 2023-08-01 13:37 | 2023-08-01 14:25 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/01/2023 02:24:21 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1053010 | 196603 | CM | CMPLT | CMPLT | 2023-08-01 13:56 | 2023-08-01 13:58 | "Failed while saving patient data" | (TC 08/01/2023 01:56:14 PM, Fan, Tak) Checked system. Was brought out of exam room. Seems OK. Used network cable from original room, not connecting. Cable appears damaged. Replaced cable. Exam not on Viewpoint. Sent exam no problem, verified on viewpoint. | Store succeeded to: Local. Store failed to: MSHECHO02.
1049904 | 196604 | CM | CMPLT | CMPLT | 2023-08-01 14:06 | 2023-08-01 14:06 | Inner drive arm cover broken. | (TC 07/31/2023 02:07:15 PM, Dolinar, John) Repaired inner drive handle cover. | Repaired arm.
63873 | 196605 | CM | CMPLT | CMPLT | 2023-08-01 14:08 | 2023-08-01 14:09 | Latch recall | (TC 08/01/2023 02:08:28 PM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
| 196606 | CM | CMPLT | CMPLT | 2023-08-01 14:10 | 2023-08-01 14:10 | Courtesy Contract Partner Care Service
1049905 | 196607 | CM | CMPLT | CMPLT | 2023-08-01 14:10 | 2023-08-01 14:10 | Left cart cover chipped. DRX corner bumper missing. | (TC 08/01/2023 02:11:16 PM, Dolinar, John) Replaced cover and bumper. (TC 07/31/2023 02:11:16 PM, Dolinar, John) Replaced cover and bumper. | Replaced cover and bumper.
63873 | 196608 | CM | CMPLT | CMPLT | 2023-08-01 14:11 | 2023-08-01 14:13 | Replaced hood | (TC 08/01/2023 02:11:30 PM, Baghirzade, Nara) Replaced the hood from a retired incubator. Placed humidity/air probe sensor probes on the hood. Replaced old corner grommets.
| 196609 | CM | CMPLT | CMPLT | 2023-08-01 14:13 | 2023-08-01 14:13 | Site coverage.
| 196610 | CM | CMPLT | CMPLT | 2023-08-01 14:16 | 2023-08-01 14:16 | Follow up for mock inspection | (TC 08/01/2023 02:17:23 PM, Dolinar, John) Follow up for mock inspection
| 196611 | CM | CMPLT | CMPLT | 2023-08-01 14:35 | 2023-08-11 07:54 | myUTN Masimo defective
1058263 | 196612 | CM | CMPLT | CMPLT | 2023-08-01 14:52 | 2023-08-01 14:53 | Battery does not reach more than 4 hours of runtime | (TC 08/01/2023 02:52:07 PM, Baghirzade, Nara) Left it charging for a few hours; battery runtime does not go above 4 hours. Replaced battery. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 46 mohms Leakage current: 66.9 uA
1061333 | 196613 | FC | CMPLT | CMPLT | 2023-08-01 15:42 | 2023-08-01 15:43 | QC checklist | (TC 08/01/2023 03:42:25 PM, Baghirzade, Nara) QC checklist completed. Returned to 18South
66509 | 196614 | CM | CMPLT | CMPLT | 2023-08-01 18:00 | 2023-08-08 13:13 | Recorder not printing | (TC 08/01/2023 06:02:30 PM, Rendon Lopez , Elizabeth) Removed fetal monitor from room, replaced with spare. Testing, intermittend "Check paper" error, cleaned sensors, not resolved. Recorder previously repaired, ordered a new replacement. Replaced recorder with new spare, calibrated, tested-OK
69427 | 196615 | CM | CMPLT | CMPLT | 2023-08-01 18:15 | 2023-08-01 18:18 | Error - Segment Failure | (TC 07/28/2023 06:16:30 PM, Rendon Lopez , Elizabeth) Error 210.6041 - Segment failure upon turn on. Replaced LCD assembly. Downloaded error logs, continuous bottle pressure sensor error, replaced bottle PS. Calibrated
1051292 | 196616 | CM | CMPLT | CMPLT | 2023-08-01 18:20 | 2023-08-01 18:21 | Not reading | (TC 08/01/2023 06:20:15 PM, Rendon Lopez , Elizabeth) Physical damage, unrepairable. Retire from service, removed battery and used in 1067441
1067441 | 196617 | CM | CMPLT | CMPLT | 2023-08-01 18:22 | 2023-08-01 18:22 | In service | (TC 08/01/2023 06:21:43 PM, Rendon Lopez , Elizabeth) Added to TMS, replaced battery with old transducer from 2019. Charged, tested, working.
67311 | 196618 | CM | CMPLT | CMPLT | 2023-08-01 18:23 | 2023-08-03 08:47 | Top warmer stock at 21C | (TC 08/01/2023 06:22:51 PM, Rendon Lopez , Elizabeth) Troubleshooting. Reset top warmer, E-98, not warming up. Removed top module from OR, contacted company for warranty. Testing module in office stand alone, seems to be working (TC 08/02/2023 08:45:53 AM, Rendon Lopez , Elizabeth) Ran module overnight, working properly. Reseted several times-OK. Returned to OR 3
1060029 | 196619 | CM | CMPLT | CMPLT | 2023-08-01 18:24 | 2023-08-01 18:25 | Blue light malfunction | (TC 08/01/2023 06:24:29 PM, Rendon Lopez , Elizabeth) Confirmed issue, unrepairable, retired from service
1062695 | 196620 | CM | CMPLT | CMPLT | 2023-08-01 18:26 | 2023-08-01 18:26 | Blue light malfunction | (TC 08/01/2023 06:25:56 PM, Rendon Lopez , Elizabeth) Confirmed issue. Replaced transducer, updated SN, charged-tested-OK
1050623 | 196621 | FC | CMPLT | NFF | 2023-08-02 06:42 | 2023-08-01 16:00 | Functional check | (TC 08/01/2023 04:00:00 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. No fault found. (TC 08/02/2023 07:47:53 AM, Baghirzade, Nara) Ran a basic infusion at varying rates; completed with no issues.
63873 | 196622 | CM | CMPLT | CMPLT | 2023-08-02 07:01 | 2023-08-02 09:13 | Created by Failed PM WO # 191605 : PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/02/2023 09:11:56 AM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. Replaced RH sensor. Calibrated RH sensor. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Greased the tilt structure. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 75%. Functional tests passed. Electrical safety results - Leakage current = 30.6 uA Power cord resistance = 120 mohms | (TC 08/01/2023 05:00:00 PM, Baghirzade, Nara) Ran the functional test for more than an hour. RH % is off by more than 5% - calibrated the RH sensor multiple times but the RH sensor is not reading the humidity correctly and will not stabilize. RH sensor needs to be replaced.
65914 | 196623 | FC | CMPLT | CMPLT | 2023-08-02 07:05 | 2023-08-02 07:05 | Functional check | (TC 08/02/2023 07:04:45 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. (TC 08/02/2023 07:47:06 AM, Baghirzade, Nara) Ran a basic infusion at varying rates; no errors during infusion. Infusion completed.
68678 | 196624 | CM | CMPLT | CMPLT | 2023-08-02 07:11 | 2023-08-02 07:12 | Battery capacity low | (TC 08/02/2023 07:10:55 AM, Rhuland, Courtney) Battery was less than 2 years old so optimal conditioning was performed Old capacity - 2274 mamphr New capacity - 4075 mamphr Functional check - OK
68957 | 196625 | CM | CMPLT | CMPLT | 2023-08-02 08:15 | 2023-08-04 10:07 | Not turning on | (TC 08/02/2023 01:48:09 PM, Canjimavadivel, Kamallesh ) Opened up device to verify fuses were intact. Tested with different keypad and working okay. Swapped keypad and reassembled. Pushed new data set. PM performed, all tests pass. Starting battery conditioning. (TC 08/04/2023 10:05:57 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete New capacity 4019 mamphr | Not turning on when plugged in but turns on when not plugged in
| 196626 | CM | CMPLT | CMPLT | 2023-08-02 08:24 | 2023-07-31 08:23 | Biomeed meeting | Biomeed meeting
1061455 | 196627 | CM | CMPLT | CMPLT | 2023-08-02 08:26 | 2023-08-01 08:25 | patient bed not working | reset the bed. tested. functional.
| 196628 | RDS | CMPLT | CMPLT | 2023-08-02 09:04 | 2023-08-02 09:11 | Weekly Check - 12S | (TC 08/02/2023 09:10:47 AM, Ferreira, Qwynn) procedure for results
| 196629 | RDS | CMPLT | CMPLT | 2023-08-02 09:12 | 2023-08-02 09:15 | Weekly Check - NICU | (TC 08/01/2023 09:13:53 AM, Ferreira, Qwynn) check procedure, may ordered pads, hopefully they come in shortly
| 196630 | RDS | CMPLT | CMPLT | 2023-08-02 09:17 | 2023-08-02 09:19 | Weekly Check - Oncology | (TC 08/01/2023 09:18:21 AM, Ferreira, Qwynn) See procedure, this unit is awesome, no problems
1065162 | 196631 | CM | CMPLT | CMPLT | 2023-08-02 09:32 | 2023-08-02 10:55 | spO2 not working | (TC 08/02/2023 10:53:41 AM, Ferreira, Qwynn) Tested the unit, and the spO2 was alarming changed probe but no result tested with new probe body cable alarm went away. The cable in the unit looked to be new, saving to test on another unit Returned device with new probe body cable
| 196632 | CM | CMPLT | CMPLT | 2023-08-02 10:29 | 2023-08-02 10:30 | PACSGear Server status | (TC 08/02/2023 10:29:00 AM, Samanian, Kamran) Services in the server were acting unresponsive in yesterday, restarted services ( by Atif ) and after a whiel they were stopped. Server rebooted and it seemed everything worked. this morning checked the status of jobs and all were completed from yesterday.
65128 | 196633 | CM | CMPLT | CMPLT | 2023-08-02 10:44 | 2023-08-02 10:45 | Not working | (TC 08/02/2023 10:43:42 AM, Reid, Emmerson ) Pickd up CADD pump from PACU. Said not working. I put batteries in and it turned on. Dose card was giving an error. I checked the pins and everything seemed fine but the error kept coming up that the button was stuck. Checked with another dose cord and it was working. Disposed of the dose cord and returned the pump.
1047862 | 196634 | PA | CMPLT | CMPLT | 2023-08-02 11:58 | 2023-08-08 11:55 | Ordering parts for repair | (TC 08/02/2023 12:01:35 PM, Baghirzade, Nara) SC 1003017344 (TC 08/08/2023 11:54:53 AM, Baghirzade, Nara) PO 7202259572
67123 | 196635 | CM | CMPLT | NFF | 2023-08-02 12:15 | 2023-08-02 12:16 | Reported that review option is not working | (TC 08/02/2023 12:15:12 PM, Baghirzade, Nara) Could not find an issue with the review button. Took multiple measurements and was able to cycle through them by using the review button. Left the device for some time and retook BP, SpO2 and temp measurements; I was able to cycle through them. All functional tests passed. Device is functional. No fault found.
1059795 | 196636 | CM | CMPLT | CMPLT | 2023-08-02 12:45 | 2023-08-01 23:44 | Broken monitor replacement and C-Arm break replacement. | Monitor was replaced, breaks were repaired and replaced.
1059828 | 196637 | PM | CMPLT | CMPLT | 2023-08-02 13:16 | 2023-08-02 14:05 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/02/2023 02:03:42 PM, Rhuland, Courtney) Back right wheel nut tightened Drained, disinfected, rinsed and refilled with sterile water
1062543 | 196638 | CM | CMPLT | CMPLT | 2023-08-02 13:30 | 2023-08-02 13:33 | "keeps alarming" | (TC 07/31/2023 01:30:24 PM, Cortez, John) Moved sliding sections to correct position and acknowledged on screen. Seems to clear alarm
1065355 | 196639 | CM | CMPLT | CMPLT | 2023-08-02 13:56 | 2023-08-08 11:51 | MSH US PVT 661 CRYSTAL DAMAGE SOURCE SPARE | (TC 08/02/2023 01:56:28 PM, Whittington, Scott) MSH US PVT 661 CRYSTAL DAMAGE SOURCE SPARE | MSH US PVT 661 CRYSTAL DAMAGE SOURCE SPARE
62971 | 196640 | PA | CMPLT | CMPLT | 2023-08-02 14:52 | 2023-08-29 13:18 | Ordering parts for repair | (TC 08/02/2023 02:55:13 PM, Baghirzade, Nara) SC 1003017693 (TC 08/29/2023 01:17:07 PM, Baghirzade, Nara) PO 7202259838
62966 | 196641 | CM | CMPLT | CMPLT | 2023-08-03 07:17 | 2023-08-03 07:19 | "Broken" | (TC 08/03/2023 07:17:32 AM, Rhuland, Courtney) Tag stated the device was broken Back switch was not on Turned on switch and performed functional check
63876 | 196642 | CM | CMPLT | CMPLT | 2023-08-03 07:40 | 2023-08-03 07:41 | Dismantling and disposal of damaged hood | (TC 08/03/2023 07:40:38 AM, Baghirzade, Nara) Removed all latches, grommets, window, portholes, doors, and other attachments on the hood. Disposed of the hood.
67920 | 196643 | CM | CMPLT | CMPLT | 2023-08-03 07:46 | 2023-08-03 07:49 | Broken inside door | (TC 08/03/2023 07:48:24 AM, Rhuland, Courtney) Chipped inside door replaced with spare Functional check - OK
| 196644 | PA | CMPLT | CMPLT | 2023-08-03 08:13 | 2023-08-14 10:24 | Orings for Zeus | (TC 08/03/2023 08:15:53 AM, Scero, Tomlyn) Ordered on ereq (TC 08/14/2023 10:23:44 AM, Scero, Tomlyn) received.
68714 | 196645 | CM | CMPLT | CMPLT | 2023-08-03 08:49 | 2023-08-03 08:50 | Will not turn on | (TC 08/02/2023 08:49:21 AM, Rendon Lopez , Elizabeth) Found battery missing screws. Replaced screws, tested-OK. Run battery conditioning-OK
69458 | 196646 | PM | CMPLT | CMPLT | 2023-08-03 08:51 | 2023-08-03 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/02/2023 08:50:52 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69559 | 196647 | CM | CMPLT | CMPLT | 2023-08-03 08:53 | 2023-09-18 10:40 | Created by Failed PM WO # 189648 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/15/2023 10:39:51 AM, Rendon Lopez , Elizabeth) Replaced pressure sensor. PM completed, all tests passed | (TC 08/02/2023 08:51:40 AM, Rendon Lopez , Elizabeth) Found Bottle pressure sensor voltage stocked on high. Sensor to be replaced, waiting on part
63876 | 196648 | CM | CMPLT | CMPLT | 2023-08-03 09:35 | 2023-08-03 09:37 | Replaced hood | (TC 08/03/2023 09:36:06 AM, Baghirzade, Nara) Replaced hood. Placed all latches, grommets, flip door, porthole door, hinges, and doors back on the new hood.
1058584 | 196649 | CM | CMPLT | CMPLT | 2023-08-03 10:03 | 2023-08-03 10:07 | Not working | (TC 08/03/2023 10:03:04 AM, Reid, Emmerson ) Came without batteries so i added new ones. Cleaned the lense and tested on the checker. Gave an error the first couple times right away so I let it sit overnight. Today it worked first try.
1063746 | 196650 | PM | CMPLT | CMPLT | 2023-08-03 10:13 | 2023-08-03 11:13 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/03/2023 11:12:30 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 196651 | CM | CMPLT | CMPLT | 2023-08-03 10:33 | 2023-09-22 14:58 | OR 10&11 decant | (TC 08/02/2023 02:27:00 PM, Chukri, Samir) meeting
1066475 | 196652 | CM | CMPLT | CMPLT | 2023-08-03 11:31 | 2023-08-03 11:31 | Filter Replacement | (TC 08/03/2023 11:31:16 AM, Samanian, Kamran) Chris ( Stryker ) came in and replace the filter
1066476 | 196653 | CM | CMPLT | CMPLT | 2023-08-03 11:32 | 2023-08-03 11:32 | Filter Replacement | (TC 08/03/2023 11:32:41 AM, Samanian, Kamran) Chris ( Stryker ) came in and replace the filter
1058606 | 196654 | CM | CMPLT | CMPLT | 2023-08-03 11:34 | 2023-08-03 11:34 | filter replacement | (TC 08/03/2023 11:34:22 AM, Samanian, Kamran) Chris ( Stryker ) came in and replace the filter
63876 | 196655 | CM | CMPLT | CMPLT | 2023-08-03 11:50 | 2023-08-04 11:11 | Created by Failed PM WO # 194334 : PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/04/2023 11:10:18 AM, Baghirzade, Nara) Replaced RH sensor. PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temp to 36C, measured 36C on the analyzer. Functional tests passed. | (TC 08/03/2023 11:48:04 AM, Baghirzade, Nara) RH sensor failed during PM. Calibrated the RH sensor and measurement was not accurate. Re-calibrated the sensor and letting incubator rest and stabilize before I start the functional check again. If RH sensor fails again, will be replacing the RH sensor.
1067483 | 196656 | CM | ACTIV | ISSUE | 2023-08-03 11:55 |  | NAS PREP | (TC 08/02/2023 11:55:12 AM, Samanian, Kamran) started prep (TC 08/03/2023 11:55:32 AM, Samanian, Kamran) more prep- added accounts (TC 08/04/2023 10:37:02 AM, Samanian, Kamran) mapped the T drive to my laptop and moved old folders to it. after that old folders were deleted from T drive in domain
67958 | 196657 | CM | CMPLT | CMPLT | 2023-08-03 12:15 | 2023-08-03 12:16 | Loud noise when switching on | (TC 08/03/2023 12:15:36 PM, Baghirzade, Nara) Fan was not seated properly; re-seated the fan. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% humidity. Incubator is fully functional.
65919 | 196658 | CM | CMPLT | CMPLT | 2023-08-03 12:21 | 2024-02-14 07:54 | Possible defective new plunger assembly | (TC 08/03/2023 12:20:37 PM, Baghirzade, Nara) Replaced with new plunger assembly. During functional check, the split nut kept failing the test. Opened the syringe again and checked all connections, and the plunger assembly. Everything was seated in properly. Had Elizabeth also take a look and she couldn't find a problem with it either. Re-seated the plunger assembly, ran a functional and now it is intermittently working; when plunger is at the highest point, the split nut fails. Will reach out to BD for troubleshooting or replacement of plunger as it is suspected to be defective from manufacturing. (TC 08/04/2023 09:39:58 AM, Baghirzade, Nara) Spoke to Justin Seybert from BD; shipped the plunger to him for assessment as possible defective plunger out of box. Waybill # 383841 (TC 10/25/2023 01:54:13 PM, Baghirzade, Nara) Have not heard back from BD or Justin Seybert for 2 months. Sent a follow up e-mail to Justin. He has been busy and will get back to me this week regarding the state of the plunger assembly. (TC 10/26/2023 01:38:45 PM, Baghirzade, Nara) Response from Justin Seybert (BD): Good morning Nara, I installed the plunger assembly into a known good pump and ran through testing and was unable to duplicate the stated problem, part appears to be good. I will have the part returned to inventory for repair stock, and I am requesting a new replacement to be sent out to you. Regards, Justin (TC 02/14/2024 07:52:03 AM, Baghirzade, Nara) Finally located the paperwork for the plunger; plunger was used for another NICU syringe pump. Will reorder a new one. (TC 02/14/2024 07:52:03 AM, Baghirzade, Nara) Finally located the paperwork for the plunger as it was not located; plunger was used for another NICU syringe pump by another colleague. Will reorder a new one.
1065355 | 196659 | ACC | CMPLT | CMPLT | 2023-08-03 12:30 | 2023-08-08 11:51 | MSH US CANON PROBE PVT 661 ACCEPTANCE SN# 99B07Z2305 MFG DECEMBER 2007 | (TC 08/03/2023 12:33:21 PM, Whittington, Scott) MSH US CANON PROBE PVT 661 ACCEPTANCE SN# 99B07Z2305 MFG DECEMBER 2007 | MSH US CANON PROBE PVT 661 ACCEPTANCE SN# 99B07Z2305 MFG DECEMBER 2007
1066454 | 196660 | FC | CMPLT | CMPLT | 2023-08-03 14:13 | 2023-08-03 14:13 | QC checklist | (TC 08/03/2023 02:12:55 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
| 196661 | CM | CMPLT | CMPLT | 2023-08-03 14:17 | 2023-08-03 14:18 | Biomed cleanup | (TC 08/03/2023 02:17:18 PM, Baghirzade, Nara) Biomed cleanup
| 196662 | CM | CMPLT | CMPLT | 2023-08-03 14:36 | 2023-08-03 14:36 | Biomed meeting | Biomed meeting
1051879 | 196663 | PA | CMPLT | CMPLT | 2023-08-03 14:50 | 2023-08-09 13:01 | Ordering parts for repair | (TC 08/03/2023 02:51:56 PM, Baghirzade, Nara) SC 1003018700 (TC 08/09/2023 01:00:21 PM, Baghirzade, Nara) PO 7202260614
| 196664 | CM | CMPLT | CMPLT | 2023-08-03 15:34 | 2023-09-01 11:52 | misc biomed - aug 2023 | (TC 08/03/2023 03:33:49 PM, Cortez, John) biomed dept cleanup (TC 08/03/2023 03:36:48 PM, Cortez, John) Went through spare radicals+docking stations and retired faulty units. Kept the rest (TC 08/03/2023 02:23:49 PM, Cortez, John) Philips meeting (TC 08/08/2023 03:56:57 PM, Cortez, John) Showed Barb (PCM 10N) how to turn off audible alarm from bariatric bed (TC 08/09/2023 09:54:57 AM, Cortez, John) Replaced AED cabinet battery in 14 microbiology (TC 08/10/2023 09:55:53 AM, Cortez, John) Lent training AED to Vincy (10S CNS) (TC 08/09/2023 02:17:17 PM, Cortez, John) Helped Nima test new linings for suction regulators Initially found that linings dont fit current containers Steve Arai found an alternative container that fits/works (TC 08/09/2023 02:18:50 PM, Cortez, John) 14S called about a procuity bed in 009-A, they want to know if it can be used I checked bed and informed them that it needs service, and they need to find a spare While there was also notified about the AUx port on 009-B not working, verified issue Olcay is already putting in a service request for both beds (TC 08/16/2023 11:32:41 AM, Cortez, John) Picked up training AED from Vincy (10S) (TC 08/16/2023 03:40:17 PM, Cortez, John) systems committee meeting (TC 08/16/2023 03:41:00 PM, Cortez, John) masimo "key authentication error" meeting (TC 08/16/2023 03:40:17 PM, Cortez, John) systems committee meeting (TC 08/14/2023 03:42:37 PM, Cortez, John) huddle (TC 08/16/2023 03:43:37 PM, Cortez, John) New canister/liner testing for regulators (TC 08/17/2023 04:30:35 PM, Cortez, John) TMS cleanup (TC 08/17/2023 04:31:06 PM, Cortez, John) PM meeting (TC 08/17/2023 04:31:31 PM, Cortez, John) philips biweekly meeting (TC 08/17/2023 04:56:27 PM, Cortez, John) Becky mentioned Masimo teset server is down. Able to ping test server, seems to be a wifi issue. Becky will contact IT (TC 08/18/2023 04:02:50 PM, Cortez, John) Gathered up items/equipment for disposal on Monday August 21 (TC 08/18/2023 04:04:44 PM, Cortez, John) Inventory of ICU beds with Kam (TC 08/18/2023 04:05:16 PM, Cortez, John) Responded to Brad (plexxus?) regarding regulator liner/container testing (TC 08/18/2023 04:06:28 PM, Cortez, John) Tried to figure out Masimo Root configuration issue with Nara and Jessica (TC 08/18/2023 04:13:13 PM, Cortez, John) Activated one recall account and reponded to Hamilton recall (TC 08/21/2023 01:59:04 PM, Cortez, John) Assisted movers in disposing biomed items/equipment (TC 08/21/2023 02:00:45 PM, Cortez, John) Addressed open onerecall notices (TC 08/22/2023 02:02:25 PM, Cortez, John) Addressed Masimo 10N client down alert, but noticed there's a "device SN 2000077415 is alarming" alert Unable to physically locate SN in question (TC 08/22/2023 02:20:01 PM, Cortez, John) Registered for Draeger training (TC 08/21/2023 04:56:54 PM, Cortez, John) Walkthrough Masimo Root monitor with Maisha (TC 08/21/2023 04:57:36 PM, Cortez, John) huddle (TC 08/23/2023 03:59:02 PM, Cortez, John) Helped Amarjeet (10N) find flowmeter ordering info (TC 08/28/2023 01:00:04 PM, Cortez, John) Fluoroscopy manager called regarding ceiling light in RM 509. Issue was addressed by the time I got there (TC 08/29/2023 02:45:19 PM, Cortez, John) sent Radicals to bed for storage (TC 08/28/2023 11:47:26 AM, Cortez, John) huddle (TC 08/30/2023 11:48:07 AM, Cortez, John) Adeel one one one (TC 08/30/2023 11:51:01 AM, Cortez, John) Looked into flights/hotel for Draeger training
| 196665 | CM | CMPLT | CMPLT | 2023-08-04 08:09 | 2023-08-03 08:09 | Biomed meeting | Biomed meeting
69595 | 196666 | CM | CMPLT | CMPLT | 2023-08-04 08:41 | 2023-08-04 08:42 | Returned from UHN | (TC 08/04/2023 08:41:00 AM, Rhuland, Courtney) Device cleaned Check-in complete All tests passed Refer to ASM for details
63962 | 196667 | CM | CMPLT | NFF | 2023-08-04 09:27 | 2023-08-04 09:28 | Warmer beeping | (TC 08/03/2023 09:27:15 AM, Rendon Lopez , Elizabeth) Warmer in OR 3, testing, no issues. Checked errors, none.
1056560 | 196668 | CM | CMPLT | CMPLT | 2023-08-04 09:54 | 2023-08-04 09:56 | Battery replacement | (TC 08/04/2023 09:55:32 AM, Canjimavadivel, Kamallesh ) Replaced battery, self test pass Sent device history to server manually
1068132 | 196669 | CM | CMPLT | CMPLT | 2023-08-04 09:57 | 2023-08-28 10:54 | MSH OPG SPP - Broken 3D Probe (RIC5-9-D) | (TC 08/04/2023 10:13:03 AM, Chau, Ken) Ordering a replacement, customer service unsure how to process broken probe.... they will be paging out to FSE to reach out for a replacement possibility? (TC 08/28/2023 10:27:01 AM, Fan, Tak) Replaced probe with SN: 301446KR1, (purolator: 334261045275), RMA 50026096. | probe punctured and requires replacement -case#1-548989604163 -additional case# created by customer serivce ..... 1-548990918080 -part#5499592 -using Site#416596VE10D
| 196670 | CM | CMPLT | CMPLT | 2023-08-04 10:35 | 2023-08-04 10:36 | working on MASIMO , key lock authentication error message | (TC 08/04/2023 10:35:19 AM, Samanian, Kamran) issue appeared again. emailed masimo to inform the issue while we are still waiting to receive the UTN. rebooted 14S and 18S1, script checked by masimo. for now all looks ok.
66070 | 196671 | CM | CMPLT | CMPLT | 2023-08-04 11:12 | 2023-08-04 11:13 | Broken cases | (TC 08/04/2023 11:12:27 AM, Rhuland, Courtney) Front and rear cases replaced Check in PM performed All tests passed Refer to ASM for details
| 196672 | RDS | CMPLT | CMPLT | 2023-08-04 11:13 | 2023-08-04 11:16 | Weekly Round NICU | (TC 08/04/2023 11:13:04 AM, Baghirzade, Nara) Weekly round in NICU completed. See procedure for results.
| 196673 | RDS | CMPLT | CMPLT | 2023-08-04 11:16 | 2023-08-04 11:18 | Weekly round in RESUS | (TC 08/04/2023 11:16:15 AM, Baghirzade, Nara) Weekly round in RESUS completed. See procedure for results.
| 196674 | RDS | CMPLT | CMPLT | 2023-08-04 11:18 | 2023-08-04 11:20 | Weekly round in 11S | (TC 08/04/2023 11:18:17 AM, Baghirzade, Nara) Weekly round in 11S completed. See procedure for results.
| 196675 | RDS | CMPLT | CMPLT | 2023-08-04 11:21 | 2023-08-04 11:25 | Weekly round in 18S and PACU PODA | (TC 08/04/2023 11:21:12 AM, Baghirzade, Nara) Weekly round in 18S/PACU PODA completed. See procedure for results.
53011 | 196676 | FC | CMPLT | CMPLT | 2023-08-04 11:46 | 2023-08-04 11:49 | Functional check | (TC 08/04/2023 11:46:18 AM, Cortez, John) Functional check passed, battery not charging, replaced
1056544 | 196677 | CM | CMPLT | CMPLT | 2023-08-04 12:05 | 2023-08-04 12:06 | Has not checked in more than 7 days | (TC 08/04/2023 12:05:35 PM, Baghirzade, Nara) Manually checked in.
66692 | 196678 | CM | CMPLT | CMPLT | 2023-08-04 12:53 | 2023-08-04 12:54 | System failure 25 | (TC 08/04/2023 12:53:02 PM, Baghirzade, Nara) Fan was improperly seated; re-seated the fan. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
1061662 | 196679 | CM | CMPLT | CMPLT | 2023-08-04 13:03 | 2023-08-04 13:04 | Hose connector damaged | (TC 08/04/2023 01:02:56 PM, Cortez, John) Replaced hose with spare
1061211 | 196680 | CM | ACTIV | ISSUE | 2023-08-04 13:04 |  | 3D Camera issues | (TC 08/04/2023 01:04:20 PM, Reid, Emmerson ) spoke with Rose. She had concerns about the 3D scope being old and mentioned an incident with it freezing. It turns out that it was the VMC monitor that froze due to a loose cable. She still mentioned having a backup system and wanted it tested. I sent an email to Andrew about this and quickly tested the system. Everything seemed fine.
1065013 | 196681 | FC | CMPLT | CMPLT | 2023-08-04 13:11 | 2023-08-04 13:14 | Functional check | (TC 08/04/2023 01:11:31 PM, Cortez, John) Functional check passed
1065006 | 196682 | FC | CMPLT | CMPLT | 2023-08-04 13:21 | 2023-08-04 13:22 | Functional check | (TC 08/04/2023 01:21:23 PM, Cortez, John) Functional check passed
1065009 | 196683 | CM | CMPLT | CMPLT | 2023-08-04 13:26 | 2023-08-04 13:27 | Functional check | (TC 08/04/2023 01:26:41 PM, Cortez, John) Functional check passed
| 196684 | RDS | CMPLT | CMPLT | 2023-08-04 13:41 | 2023-08-04 13:45 | Weekly check for ICU - August 4 | (TC 08/04/2023 01:41:22 PM, Cortez, John) Checked rooms 31, 9, 11, 12 Arms slip in 31 and 12 9 missing mount for thermometer
| 196685 | RDS | CMPLT | CMPLT | 2023-08-04 13:46 | 2023-08-04 13:50 | Weekly check for 10N - August 4 | (TC 08/04/2023 01:46:23 PM, Cortez, John) Rooms occupied, checked hallways and equipment storage room Talked with Natalia (PCM) about discarding a few items in storage, including welch allyn vitals tower, nurse call cables, gas cylinder holders Will look into KCI pump, asset# 62766, 10N doesn't need it
| 196686 | RDS | CMPLT | CMPLT | 2023-08-04 13:51 | 2023-08-04 13:58 | Weekly check for 10S - August 4 | (TC 08/04/2023 01:51:30 PM, Cortez, John) Rooms occupied, checked hallways and equipment storage room. Vitals monitors need cuffs+spo2 sensors. Will remind unit to order (TC 08/04/2023 01:51:30 PM, Cortez, John) Rooms occupied, checked hallways and equipment storage room. Vitals monitors need cuffs+spo2 sensors. Will remind unit to order Isotour mattress pump has broken hose connector, will swap swapped pumps due for maintenance
502044 | 196687 | RQ | CMPLT | CMPLT | 2023-08-04 15:27 | 2023-09-13 19:42 | collimator lock/latch mechanism failing |  | collimatotor lock/latch mechanism failing on ERGO camera
66548 | 196688 | PM | CMPLT | CMPLT | 2023-08-04 15:46 | 2023-08-04 15:48 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/04/2023 03:47:17 PM, Rendon Lopez , Elizabeth) Battery rep[laced. PM completed, all tests passed
66547 | 196689 | PM | CMPLT | CMPLT | 2023-08-04 15:49 | 2023-08-04 15:51 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/04/2023 03:49:32 PM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
67914 | 196690 | CM | CMPLT | NFF | 2023-08-08 07:22 | 2023-08-08 09:08 | Humidity and Temperature not working correctly | (TC 08/08/2023 09:07:25 AM, Ferreira, Qwynn) Came down for issue with temp and humidity Tested both at 35degC and 75% RH Read 35.2 and 76.3 % RH Returned to unit NFF
65963 | 196691 | CM | CMPLT | CMPLT | 2023-08-08 07:51 | 2023-08-08 07:52 | Channel error; no error specified on biomed tag | (TC 08/08/2023 07:51:14 AM, Baghirzade, Nara) No error specified; downloaded error logs, did not note any significant errors. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. Ran a basic infusion at various rates; no errors or issues noted. Infusion completed successfully.
67919 | 196692 | CM | CMPLT | CMPLT | 2023-08-08 07:56 | 2023-08-08 07:57 | System failure 25 | (TC 08/04/2023 07:56:04 AM, Ferreira, Qwynn) System failure 25 was present when powering on the incubator Reseated the fan - Working correctly now Replaced 2 port doors Repalced 2 hinge springs Functional check - OK
| 196693 | RDS | CMPLT | CMPLT | 2023-08-08 08:02 | 2023-08-08 09:23 | Weekly check for ED
| 196694 | RDS | CMPLT | CMPLT | 2023-08-08 08:04 | 2023-08-08 14:40 | Weekly check for ICU
| 196695 | OC | CMPLT | CMPLT | 2023-08-08 08:06 | 2023-08-14 07:48 | Kam On-call; August 4 - 11, 2023 | (TC 08/06/2023 12:20:00 PM, Canjimavadivel, Kamallesh ) Received called from ICU regarding issue with lift. Verified with nurse the lift in question was a ceiling lift in one of the patient rooms. Advised them to contact building services. Did not need to come in. (TC 08/09/2023 04:10:00 PM, Canjimavadivel, Kamallesh ) Received call from locating at 4:07 PM asking what time on-call hours start because 14S couldn't reach anyone on biomed line. Asked them to put me through to 14S. Ward clerk was asking the status of a bed as they reported someone came earlier to look at the bed and didn't provide info on whether the bed can be used or not. I informed them I will have to follow up the next morning as I'm no longer on site. Called John who was still on site and asked if he can take a look and provide instructions to the unit whether the bed is usable or not.
65976 | 196696 | CM | CMPLT | CMPLT | 2023-08-08 08:22 | 2023-08-08 08:25 | Channel error | (TC 08/08/2023 08:23:28 AM, Rhuland, Courtney) Channel error 354.6770 Reseated pressure sensor Check in performed All tests passed and channel error cleared Refer to ASM for details
67929 | 196697 | CM | CMPLT | CMPLT | 2023-08-08 08:36 | 2023-08-08 08:37 | Latch recall | (TC 08/08/2023 08:36:40 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
65889 | 196698 | CM | CMPLT | CMPLT | 2023-08-08 08:55 | 2023-08-08 09:08 | Channel error | (TC 08/08/2023 08:55:13 AM, Rhuland, Courtney) Pressure sensor reseated Missing screw replaced Check in complete All tests passed Refer to ASM for details
1067501 | 196699 | CM | CMPLT | CMPLT | 2023-08-08 09:03 | 2023-08-08 09:05 | Repurposed retired thermometer | (TC 08/08/2023 09:03:28 AM, Baghirzade, Nara) Reurposed spare thermometer for NICU in place of an old Filac thermometer that has been retired. All functional tests passed. Printed and placed new asset sticker.
| 196700 | CM | CMPLT | CMPLT | 2023-08-08 09:23 | 2023-08-22 08:14 | SpO2 malfunction reported in room 15 | (TC 08/08/2023 09:23:24 AM, Canjimavadivel, Kamallesh ) Brought up spare X3 and cable to troubleshoot. Nurse already swapped cable with another one. Brought cable down and verified issue is with cable. Submitted product complaint with Masimo for warranty replacement. (TC 08/10/2023 08:36:21 AM, Canjimavadivel, Kamallesh ) Received reply from vendor and they initiated warranty exchange. Requested no charge PO (TC 08/10/2023 10:30:21 AM, Canjimavadivel, Kamallesh ) Forwarded PO to vendor. Shipped out defective cables via KJV (TC 08/22/2023 08:12:51 AM, Canjimavadivel, Kamallesh ) Received replacement cables, dropped off to ICU
1056466 | 196701 | CM | CMPLT | CMPLT | 2023-08-08 10:35 | 2023-08-08 11:24 | NIBP hose leaking | (TC 08/08/2023 11:23:38 AM, Baghirzade, Nara) Replaced NIBP hose. All functional checks passed for: NIBP Temp SpO2
| 196702 | CM | CANCL | CANCL | 2023-08-08 10:42 |  | database test | (TC 08/08/2023 10:59:08 AM, Chukri, Samir) test
| 196703 | CM | CMPLT | CMPLT | 2023-08-08 10:51 | 2023-08-09 07:55 | Unlocking Masimo Kite tablets | (TC 08/08/2023 10:51:03 AM, Canjimavadivel, Kamallesh ) Called help desk to create tickets to get tablets unlocked. (TC 08/08/2023 02:36:46 PM, Canjimavadivel, Kamallesh ) Reset 2 tablets to factory and installed updates
67076 | 196704 | CM | CANCL | CANCL | 2023-08-08 11:02 |  | test | (TC 08/08/2023 11:03:33 AM, Chukri, Samir) test
1054775 | 196705 | CM | CMPLT | CMPLT | 2023-08-08 12:48 | 2023-08-08 12:49 | BP Not inflating | (TC 08/08/2023 12:48:10 PM, Ferreira, Qwynn) tested the device and BP was not working but the cuff seemed to be the issue Infomed the unit that a new cuff was required and that the machine was working once a new cuff is installed Returned to unit
1059778 | 196706 | FC | CMPLT | CMPLT | 2023-08-08 13:20 | 2023-08-08 13:21 | QC checklist | (TC 08/08/2023 01:20:17 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1067500 | 196707 | CM | CMPLT | CMPLT | 2023-08-08 13:43 | 2023-08-08 13:48 | Incoming inspection and asset tagging | (TC 08/08/2023 01:43:37 PM, Baghirzade, Nara) Incoming inspection completed; microscope is fully functional. Assembled by Philip from Leica. Asset tagged. Electrical safety - Leakage current = 60.2 uA Power cord resistance = 39 mohms
| 196708 | PA | CMPLT | CMPLT | 2023-08-08 13:59 | 2023-09-08 07:49 | Supplies
1061316 | 196709 | CM | CMPLT | CMPLT | 2023-08-08 14:01 | 2023-08-08 14:02 | Pediatrician PCs upgrade for better hardware | (TC 08/08/2023 09:01:00 AM, Fan, Tak) Coordinated with Jose for SW install.
1067505 | 196710 | CM | CMPLT | CMPLT | 2023-08-08 14:07 | 2023-08-08 14:11 | Broken probe | (TC 08/08/2023 02:07:28 PM, Baghirzade, Nara) Probe was physically broken; replaced probe with a spare one in biomed. Thermometer was not on TMS; added to database and asset tagged. Functional tests passed.
65969 | 196711 | PM | CMPLT | CMPLT | 2023-08-08 14:34 | 2023-08-08 14:35 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/08/2023 02:34:35 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1061211 | 196712 | CM | CMPLT | CMPLT | 2023-08-08 15:07 | 2023-08-08 15:14 | moving monitors to OR 9 | (TC 08/08/2023 03:07:26 PM, Samanian, Kamran) Moved monitors to room 9, worked on VMC3 and TvONE set up to make 3D scope working. (TC 08/08/2023 03:12:08 PM, Reid, Emmerson ) all AV 3D system moved, set up and tested, passed
| 196713 | CM | ACTIV | ISSUE | 2023-08-08 15:15 |  | FetaLink- FM in OPG ( John Chen ) for fetalink intgeration testing | (TC 08/08/2023 03:15:12 PM, Samanian, Kamran) moved and left in OPG (TC 08/09/2023 09:38:16 AM, Samanian, Kamran) tested connectivity and passed. There are issues with FetaLink and Vital Signs in the chart. (TC 08/10/2023 01:19:55 PM, Samanian, Kamran) discussed internally for iBUS new domain integration (TC 09/19/2023 08:26:27 AM, Samanian, Kamran) met with Isabel for checking MHR in FM
1061959 | 196714 | CM | CMPLT | NFF | 2023-08-08 15:49 | 2023-08-08 15:50 | "side rail controls not working" | (TC 08/08/2023 03:48:42 PM, Cortez, John) No fault found, controls working fine Power cycled bed
1062027 | 196715 | CM | CMPLT | NFF | 2023-08-08 15:51 | 2023-08-08 15:53 | "unable to use scale" | (TC 08/08/2023 03:51:12 PM, Cortez, John) No fault found, scale is displaying weight values No error messages Spoke with nurse on restarting bed and how to power cycle it
1062001 | 196716 | CM | CMPLT | CMPLT | 2023-08-08 15:54 | 2023-08-29 14:44 | "head section not moving up or down" | (TC 08/08/2023 03:53:40 PM, Cortez, John) Confirmed issue. One end of actuator for head section detached from bed Called stryker and requested service (TC 08/29/2023 02:43:06 PM, Cortez, John) Bed was repaired by Mark
1061333 | 196717 | FC | CMPLT | CMPLT | 2023-08-09 06:56 | 2023-08-09 06:56 | QC checklist | (TC 08/09/2023 06:55:56 AM, Baghirzade, Nara) QC checklist completed. Returned to 18 South
1054983 | 196718 | CM | CMPLT | CMPLT | 2023-08-09 07:21 | 2023-08-09 07:23 | BP hose leaking | (TC 08/09/2023 07:21:24 AM, Baghirzade, Nara) BP hose was leaking at Root connector end; replaced BP hose. Noticed that no configuration was loaded to the root; config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1061203 | 196719 | CM | CMPLT | CMPLT | 2023-08-09 07:42 | 2023-08-09 07:43 | Monitor Balancing | (TC 08/08/2023 07:42:14 AM, Reid, Emmerson ) Balanced the monitor tilt. Working now.
1061209 | 196720 | CM | CMPLT | CMPLT | 2023-08-09 07:43 | 2023-08-09 07:44 | Monitor Balancing | (TC 08/08/2023 07:43:38 AM, Reid, Emmerson ) Balanced the monitor tilt. Working now.
| 196721 | PA | CMPLT | CMPLT | 2023-08-09 07:58 | 2023-08-22 14:25 | Formlabs Print Supplies
| 196722 | RDS | CMPLT | CMPLT | 2023-08-09 08:10 | 2023-08-09 09:13 | Weekly Checks For PACU | (TC 08/09/2023 09:11:37 AM, Reid, Emmerson ) Everything was good except PHI was being left on monitors and some MX800s are being turned off. I emailed Wendy about resolving this.
68294 | 196723 | FC | CMPLT | CMPLT | 2023-08-09 09:05 | 2023-08-09 09:10 | PM-MAJOR - Annual PM for: Infusion Pumps | (TC 08/09/2023 09:06:00 AM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1056911 | 196724 | CM | CMPLT | CMPLT | 2023-08-09 09:14 | 2023-08-09 09:15 | Image too Bright | (TC 08/09/2023 09:14:00 AM, Reid, Emmerson ) Saturation was too high on the video but good through the lense. I went through the pre-set colour settings until there was one they liked. They als requested to put video on the monitors but did not have a ScaleOR. I setup one for them.
69520 | 196725 | PM | CMPLT | CMPLT | 2023-08-09 09:45 | 2023-08-09 09:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 09:45:52 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 196726 | CM | CMPLT | CMPLT | 2023-08-09 10:27 | 2023-08-09 10:34 | Validation of plasma thawer | (TC 08/09/2023 10:33:37 AM, Baghirzade, Nara) Marietta requested to validate the plasma thawer post-installation. Measured leakage current and voltage. Filled out the validation form.
69593 | 196727 | PM | CMPLT | CMPLT | 2023-08-09 11:07 | 2023-08-09 11:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2023 11:07:34 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1047862 | 196728 | CM | CMPLT | CMPLT | 2023-08-09 13:09 | 2023-09-28 13:41 | Motor not running error | (TC 08/09/2023 01:09:29 PM, Baghirzade, Nara) Motor will not run when trying to prime or run a feed; replaced the motor but issue is still persistent. Placed the old motor back as it is functional and checked with DMM. Replaced the worm gear; motor still doesn't run. According to the manual, the sensor could have failed on the main PCB but the service manual does not have any PCB drawings and do not allow board-level repairs. Contacted Alla to request the repair quote from ICU medical. (TC 09/28/2023 01:39:45 PM, Baghirzade, Nara) The plunger was getting stuck. The clutch was cleaned and lubricated. The plunger/driver was also cleaned and lubricated. Motor is now running.
1049058 | 196729 | PM | CMPLT | CMPLT | 2023-08-09 13:23 | 2023-08-09 13:29 | Planned Maintenance | (TC 08/09/2023 01:23:17 PM, Canjimavadivel, Kamallesh ) Set points: 41.0 36.3 32.6 Low temp: 34.2 Primary over temp: 46.5 Secondary over temp: 50.1 PM complete
1061207 | 196730 | CM | CMPLT | CMPLT | 2023-08-09 15:02 | 2023-09-01 13:22 | Move OR 11 to 6 | (TC 08/09/2023 03:02:34 PM, Reid, Emmerson ) Moved Equipment (TC 09/01/2023 01:21:23 PM, Reid, Emmerson ) Equipment moved back into OR 11 and OR 6 and checked.
| 196731 | CM | CMPLT | CMPLT | 2023-08-10 09:02 | 2023-08-24 11:23 | Cell Washer- helmer, SN# 0002426 | (TC 08/10/2023 09:02:27 AM, Samanian, Kamran) Received the request from core lab ( shelley Solomon EXT. 4502 ) about the Helmer Cell washer. Lid is not staying when it is open. verified the issue. Contacted company and was given the PN, price and how to replace it. Sent out email to Shelley and Marietta for ordering the part. (TC 08/09/2023 09:02:27 AM, Samanian, Kamran) Received the request from core lab ( shelley Solomon EXT. 4502 ) about the Helmer Cell washer. Lid is not staying when it is open. verified the issue. Contacted company and was given the PN, price and how to replace it. Sent out email to Shelley and Marietta for ordering the part. (TC 08/24/2023 11:20:52 AM, Samanian, Kamran) part was ordered by lab. recived the parts. brought the device down and replaced the spring ( brass sleve needs to be replaced in future ) . tested and passed. returned the washer back to the lab.
1061212 | 196732 | CM | CMPLT | CMPLT | 2023-08-10 09:05 | 2023-08-09 09:05 | moving Olympus stuffs to OR 6 due to OR door repair. | (TC 08/09/2023 09:05:06 AM, Samanian, Kamran) done, tested and passed
| 196733 | CM | CANCL | CANCL | 2023-08-10 09:06 | 2023-08-09 09:37 | FetaLink Integration testing with CCE in OPG location | (TC 08/09/2023 09:06:07 AM, Samanian, Kamran) tested connectivity and passed. There are issues with FetaLink and Vital Signs in the chart.
| 196734 | RDS | CMPLT | CMPLT | 2023-08-10 09:18 | 2023-08-09 09:19 | Weekley check for ENDO | (TC 08/09/2023 09:18:27 AM, Samanian, Kamran) no issue or concern. shelves in ENDO 1 will need to be brough lower. will schedule it
| 196735 | RDS | CMPLT | CMPLT | 2023-08-10 09:21 | 2023-08-09 09:21 | Weekly check for 9S | (TC 08/09/2023 09:20:43 AM, Samanian, Kamran) no issue, updated them about ECG machine which will be ready tomorrow.
1061210 | 196736 | CM | CMPLT | NFF | 2023-08-10 09:39 | 2023-08-10 09:40 | Call for recoording not working | (TC 08/10/2023 09:39:28 AM, Reid, Emmerson ) When I got there it was resolved NFF
1061216 | 196737 | CM | CMPLT | CMPLT | 2023-08-10 09:40 | 2023-08-10 09:41 | TS7000 remote not working | (TC 08/10/2023 09:40:42 AM, Reid, Emmerson ) Changed remote with another from the hallway Working now,
1054543 | 196738 | CM | CMPLT | CMPLT | 2023-08-10 09:42 | 2023-08-10 09:44 | BAT low, not charging, | (TC 08/09/2023 09:41:48 AM, Samanian, Kamran) started working on the machine on AUG. 8. verified the problem. BATs are 4 years old, Replaced both. reseted ECG machine. Synced with the iECG server. Services print head. Serviced ECG adapter clips and reconnected. checked leadwires and they are in good shape. Fully charged BAT. tested ECG machine and passed. (TC 08/10/2023 09:44:53 AM, Samanian, Kamran) double checked the function and passed
1061210 | 196739 | CM | CMPLT | CMPLT | 2023-08-10 12:36 | 2023-08-10 12:40 | Dr needed to download video recording | (TC 08/10/2023 12:36:13 PM, Reid, Emmerson ) Spoke with Jaron about how to do this. Ended up downloaded by USB and forwarding instruction.
1049902 | 196740 | CM | CMPLT | CMPLT | 2023-08-10 13:19 | 2023-08-23 12:11 | MSH R62 WILL NOT MOVE, FRONT BUMP DETECTION ACTIVE; JEFF HAD SAME PROBLEM LAST WEEK | (TC 08/10/2023 01:18:44 PM, Whittington, Scott) MSH R62 WILL NOT MOVE, FRONT BUMP DETECTION ACTIVE; JEFF HAD SAME PROBLEM LAST WEEK PUSHING IN AND RELEASING FRONT BUMPER DID NOT CLEAR ERROR FULL REBOOOT ALLOWED SYSTEM TO MOVE | MSH R62 WILL NOT MOVE, FRONT BUMP DETECTION ACTIVE; JEFF HAD SAME PROBLEM LAST WEEK PUSHING IN AND RELEASING FRONT BUMPER DID NOT CLEAR ERROR FULL REBOOOT ALLOWED SYSTEM TO MOVE
1049941 | 196741 | CM | CMPLT | CMPLT | 2023-08-10 15:20 | 2023-08-10 15:21 | TEE failed to recognize yesterday but worked on other system | (TC 08/10/2023 03:20:05 PM, Fan, Tak) Tested with spare TEE, no problems. Checked connections but could not see any issues. TTE probe worked yesterday too but not TEE. Retrieved logs, but they don't show anything around the time frame. Probe may not have been picked up by the system.
| 196742 | CM | CMPLT | CMPLT | 2023-08-10 15:22 | 2023-08-23 12:17 | MSH SPP >HSC STUDIES IMAGES NOT TXFD | (TC 08/10/2023 03:25:53 PM, Whittington, Scott) MSH SPP >HSC STUDIES IMAGES NOT TXFD From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: August 10, 2023 2:43 PM To: Al-Jalam, Adnan - The Hospital for Sick Children (SickKids) <adnan.al-jalam@sickkids.ca> Cc: Proulx, Amy - Sinai Health System <Amy.Proulx@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca>; Riaz, Mirza - Sinai Health System <Mirza.Riaz@sinaihealth.ca>; Hollingshead, Andrea - Sinai Health System <Andrea.Hollingshead@sinaihealth.ca>; Thakur, Varsha - The Hospital for Sick Children (SickKids) <varsha.thakur@sickkids.ca>; O'Connor, Anne - Sinai Health System <Anne.O'Connor@sinaihealth.ca> Subject: Re: [External] RE: MSH SPP>HSC FETAL STUDY TRANSFER Hi Varsha Do you have details on studies missing from MSH SPP>HSC Thanks Scott 416-402-2462 | MSH SPP >HSC STUDIES IMAGES NOT TXFD
| 196743 | CM | CMPLT | CMPLT | 2023-08-10 18:09 | 2023-08-10 18:09 | ENDO Room # 1 , shelves on the Boom lowered | (TC 08/10/2023 06:08:49 PM, Samanian, Kamran) done, tested and passed
| 196744 | CM | CMPLT | CMPLT | 2023-08-10 18:10 | 2023-08-31 07:55 | PACSGear image test for new iBUS | (TC 08/10/2023 06:10:07 PM, Samanian, Kamran) image was sent to TEST server but not able to retrive on the powerchart (TC 08/24/2023 08:13:13 AM, Samanian, Kamran) had Philips and John on the call. Pointed image capture to the TESTserver, created patinet test bit not able to pull up. it seems worklist AE, IP and port needs to be set (TC 08/31/2023 07:54:31 AM, Samanian, Kamran) checked image capture for the test patients, found patients, captured images and sent, test was passed from biomed side
| 196745 | CON | CMPLT | CMPLT | 2023-08-10 21:06 | 2023-08-23 12:12 | MSH MOL INSPECTION PREP | (TC 08/11/2023 03:49:30 PM, Whittington, Scott) MSH MOL INSPECTION PREP (TC 08/11/2023 03:49:30 PM, Whittington, Scott) MSH MOL INSPECTION PREP | MSH MOL INSPECTION PREP
| 196746 | CM | CMPLT | CMPLT | 2023-08-11 08:23 | 2023-08-11 08:24 | OR3 equipment back to the room | (TC 08/11/2023 08:22:59 AM, Samanian, Kamran) returned all equipment to OR3 as OR3 door repair is done.
| 196747 | CM | CMPLT | CMPLT | 2023-08-11 08:25 | 2023-08-17 16:26 | Daily check | (TC 08/11/2023 08:24:44 AM, Samanian, Kamran) done and email was sent (TC 08/14/2023 08:41:21 AM, Samanian, Kamran) done and email was sent (TC 08/16/2023 07:50:02 AM, Samanian, Kamran) done and email was sent (TC 08/15/2023 07:50:22 AM, Samanian, Kamran) done and email was sent
69130 | 196748 | PM | CMPLT | CMPLT | 2023-08-11 10:31 | 2023-08-11 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/11/2023 10:31:18 AM, Cortez, John) Cleaned, Downloaded error/event logs, Door latch damaged, replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
| 196749 | CM | CMPLT | CMPLT | 2023-08-11 11:19 | 2023-08-21 11:15 | Parts for medfusion 4000 pump | (TC 08/11/2023 11:18:57 AM, Chukri, Samir) ordered repair parts for medfusion pump Asset # 1047840 and stock (TC 08/21/2023 11:14:36 AM, Chukri, Samir) received cables waiting the Gauge. (TC 01/18/2024 07:45:52 AM, Chukri, Samir) received the 3000 calibration kit.
65909 | 196750 | CM | CMPLT | CMPLT | 2023-08-11 11:27 | 2023-08-14 10:25 | Broken cases | (TC 08/11/2023 11:26:39 AM, Rhuland, Courtney) Replaced broken front and rear cases (TC 08/14/2023 10:24:30 AM, Rhuland, Courtney) PM completed All tests passed Refer to ASM for details
1046035 | 196751 | CM | CMPLT | CMPLT | 2023-08-11 13:14 | 2023-08-11 13:15 | Det cal required. | (TC 08/11/2023 01:14:17 PM, Fan, Tak) Calibrated. Passed.
1049362 | 196752 | CM | CMPLT | CMPLT | 2023-08-11 14:03 | 2023-08-11 14:04 | MSH CT Emerg - Injector - System Does Not Latch / Unlatch | (TC 08/11/2023 02:03:27 PM, Chau, Ken) Contrast on syringe causing it to cease -removed all contrast and tested unit okay
| 196753 | RDS | CMPLT | CMPLT | 2023-08-11 14:03 | 2023-08-11 14:06 | Weekly check for ICU - August 8-11 | (TC 08/11/2023 02:05:06 PM, Cortez, John) Checked rooms 16, 14, 13, 24, 8, 9, 12 Webview between 14/16 logged out, logged back in Swapped pumps due for PM Room 13 boom sliding
| 196754 | RDS | CMPLT | CMPLT | 2023-08-11 14:07 | 2023-08-11 14:11 | Weekly check for 10N - August 8-11 | (TC 08/08/2023 02:10:32 PM, Cortez, John) Checked equipment storage room, hallways
| 196755 | RDS | CMPLT | CMPLT | 2023-08-11 14:12 | 2023-08-11 14:15 | Weekly check for 10S - August 8-11 | (TC 08/08/2023 02:14:01 PM, Cortez, John) Checked rooms 5 and 30-A, equipment storage and hallways
1046196 | 196756 | CM | CMPLT | CMPLT | 2023-08-11 15:38 | 2023-08-11 15:37 | PM | SIEMENS completed scheduled PM Mechanical PM done
39506 | 196757 | CM | CMPLT | CMPLT | 2023-08-12 07:24 | 2023-08-11 07:24 | MSH ER CT INJECTOR FAILED, INSTALL SPARE INJECTOR | (TC 08/12/2023 07:24:31 AM, Whittington, Scott) MSH ER CT INJECTOR FAILED, INSTALL SPARE INJECTOR (TC 08/08/2023 07:24:00 PM, Whittington, Scott) MSH ER CT INJECTOR FAILED, INSTALL SPARE INJECTOR (TC 08/08/2023 07:24:00 PM, Whittington, Scott) MSH ER CT INJECTOR FAILED, INSTALL SPARE INJECTOR | MSH ER CT INJECTOR FAILED, INSTALL SPARE INJECTOR
32510 | 196758 | CM | CMPLT | CMPLT | 2023-08-12 07:32 | 2023-08-31 10:20 | MSH DR1 OGP , COLLIMATOR, TABLESIDE CNTRL | (TC 08/12/2023 05:17:00 PM, Whittington, Scott) MSH DR1 OGP , COLLIMATOR, TABLESIDE CNTRL | MSH DR1 OGP , COLLIMATOR, TABLESIDE CNTRL
66692 | 196759 | CM | CMPLT | CMPLT | 2023-08-14 07:32 | 2023-08-14 08:44 | Broken latch | (TC 08/14/2023 08:42:30 AM, Rhuland, Courtney) Replaced broken side door latch Calibrated RH during functional check Functional check - OK
1062999 | 196760 | PM | CMPLT | CMPLT | 2023-08-14 07:34 | 2023-08-14 08:56 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/14/2023 08:55:34 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 196761 | RDS | CMPLT | CMPLT | 2023-08-14 08:01 | 2023-08-14 08:32 | Weekly Checks For PACU | (TC 08/14/2023 08:30:47 AM, Reid, Emmerson ) Checks all complete. They still need to work on turning the monitors into Standby and removing PHI
| 196762 | RDS | CMPLT | CMPLT | 2023-08-14 08:01 | 2023-08-14 08:41 | Weekly check for ED
| 196763 | RDS | CMPLT | CMPLT | 2023-08-14 08:08 | 2023-08-17 10:16 | Weekly check for ICU
1061447 | 196764 | CM | CMPLT | CMPLT | 2023-08-14 08:10 | 2023-08-11 08:09 | right side controls not working | reset the unit. tested. functional.
1055535 | 196765 | CM | CMPLT | CMPLT | 2023-08-14 08:43 | 2023-08-14 08:46 | Loud feedback | (TC 08/11/2023 08:43:58 AM, Rendon Lopez , Elizabeth) Confirmed problem, previously repaired. Retired
1056193 | 196766 | CM | CMPLT | CMPLT | 2023-08-14 08:47 | 2023-08-14 08:48 | Equipment malfunction | (TC 08/11/2023 08:47:24 AM, Rendon Lopez , Elizabeth) Confirmed problem, previously repaired. Retired
1051427 | 196767 | CM | CMPLT | CMPLT | 2023-08-14 08:48 | 2023-08-14 08:49 | Equipment malfunction | (TC 08/11/2023 08:48:27 AM, Rendon Lopez , Elizabeth) Confirmed problem, previously repaired. Retired
1064341 | 196768 | CM | CMPLT | CMPLT | 2023-08-14 09:07 | 2023-08-14 09:20 | Table going down by itself | (TC 08/09/2023 09:07:57 AM, Rendon Lopez , Elizabeth) Confirmed issue, defective remote control. Kim will order replacement and spare
65967 | 196769 | CM | CMPLT | CMPLT | 2023-08-14 09:27 | 2023-08-14 09:33 | Created by Failed PM WO # 194346 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/10/2023 09:27:27 AM, Rendon Lopez , Elizabeth) Replaced clamp, front case and LCD under recall. PM completed, all tests passed | (TC 08/09/2023 09:23:26 AM, Rendon Lopez , Elizabeth) Rear case and handle cracked, replaced under recall. Fails clamp close position test
63303 | 196770 | CM | CMPLT | CMPLT | 2023-08-14 09:35 | 2023-08-28 14:22 | Intermittently doesn't turn on | (TC 08/10/2023 09:34:54 AM, Rendon Lopez , Elizabeth) Removed from room 15, testing, NFF. Will replace stand by switch. Damaged/cracked bed assembly and resus system knobs, inquired for pricing to order parts (TC 08/15/2023 03:08:49 PM, Rendon Lopez , Elizabeth) ordered replacement parts and spare for L&D (TC 08/23/2023 09:03:46 AM, Rendon Lopez , Elizabeth) Replaced power board fuses, standby switch, and speaker, tested, working, no issues. Replaced Resus system knobs, and PIP knob that were cracked. Whistling noise when turned on, missing cap on small tube, replaced, tested Resus-OK. (TC 08/28/2023 02:17:31 PM, Rendon Lopez , Elizabeth) Replaced bed assembly that was cracked, tested all functions-Ok. Returned to L&D
66509 | 196771 | CM | CMPLT | CMPLT | 2023-08-14 09:42 | 2023-08-15 14:31 | monitor turning off while monitoring patient | (TC 08/11/2023 09:42:02 AM, Rendon Lopez , Elizabeth) Removed from Room 5, spare monitor working in room. Testing in office, no issues found. (TC 08/14/2023 02:30:22 PM, Rendon Lopez , Elizabeth) Continuous testing with no issues. Returned monitor to room 5
1063000 | 196772 | PM | CMPLT | CMPLT | 2023-08-14 10:15 | 2023-08-14 10:15 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/14/2023 10:15:01 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 196773 | CM | CMPLT | CMPLT | 2023-08-14 10:25 | 2023-08-16 10:04 | Weekly Checks in CCU | (TC 08/16/2023 10:04:01 AM, Scero, Tomlyn) Weekly checks complete
1062236 | 196774 | CM | CMPLT | CMPLT | 2023-08-14 10:40 | 2023-08-14 10:42 | Bed Alarming - no function | (TC 08/14/2023 10:40:07 AM, Ferreira, Qwynn) Bed alarm was for the scale it was not properly setup Reset the bed and tested - working
1061204 | 196775 | CM | CMPLT | CMPLT | 2023-08-14 10:43 | 2023-08-14 10:44 | Connect foot boom again | (TC 08/14/2023 10:43:03 AM, Reid, Emmerson ) someone removed the boom from OR3, so i reconnected it between cases
1061216 | 196776 | CM | CMPLT | CMPLT | 2023-08-14 10:44 | 2023-08-14 10:45 | Moved ESU smoke evac and insufflator into room | (TC 08/14/2023 10:44:19 AM, Reid, Emmerson ) Complete.
1061207 | 196777 | CM | CMPLT | CMPLT | 2023-08-14 10:46 | 2023-09-01 13:23 | Rattling from boom rack | (TC 08/14/2023 10:46:27 AM, Reid, Emmerson ) Boom rack was ratting pretty loadly. I adjusted the position of the equipment a bit and it made it a lot more quiet. Will look for a more permanent solution
1054537 | 196778 | PM | CMPLT | CMPLT | 2023-08-14 12:02 | 2023-08-14 12:07 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/14/2023 12:02:34 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1056125 | 196779 | PM | CMPLT | CMPLT | 2023-08-14 12:52 | 2023-08-14 14:31 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/14/2023 02:30:46 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
65899 | 196780 | CM | CMPLT | CMPLT | 2023-08-14 14:24 | 2023-08-14 14:25 | Channel error | (TC 08/14/2023 02:24:20 PM, Rhuland, Courtney) Pressure sensor reseated Rear case replaced Left side handle replaced Check-in complete All tests passed Refer to ASM for details
1061217 | 196781 | CM | CMPLT | NFF | 2023-08-14 14:26 | 2023-08-14 14:26 | Olympus monitor not logging on | (TC 08/14/2023 02:25:43 PM, Reid, Emmerson ) Went up to see and staff had not pressed "logon" After that they were fine. NFF
| 196782 | CM | CMPLT | CMPLT | 2023-08-14 14:32 | 2023-08-14 14:32 | Huddle | (TC 08/14/2023 02:31:55 PM, Rhuland, Courtney) Huddle
1054383 | 196783 | CM | CMPLT | CMPLT | 2023-08-14 14:46 | 2023-08-15 07:58 | Battery failed | (TC 08/14/2023 02:51:37 PM, Reid, Emmerson ) Would not turn on when unplugged. When I plugged it in and turned it on there was an error that the battery needed to be replaced. I replaced the battery and I am letting it charge overnight. (TC 08/15/2023 07:57:23 AM, Reid, Emmerson ) Tested, its working now.
1047841 | 196784 | CM | CMPLT | CMPLT | 2023-08-15 07:21 | 2023-08-15 07:22 | Error "maintenance required" upon bootup | (TC 08/15/2023 07:20:58 AM, Baghirzade, Nara) Checked maintenance date on device and it was set to 2024, triggering maintenance alarm too early. Connected medfusion to pharmguard toolbox; in device setup, changed maintenance year to 2030. Functional tests passed. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
1065715 | 196785 | CM | CMPLT | CMPLT | 2023-08-15 07:23 | 2023-08-15 07:25 | Broken prongs on power cord | (TC 08/15/2023 07:24:21 AM, Baghirzade, Nara) Replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 41 mohms Leakage current: 60.6 uA
65963 | 196786 | CM | CMPLT | CMPLT | 2023-08-15 08:16 | 2023-08-15 08:18 | Channel error | (TC 08/15/2023 08:16:10 AM, Rhuland, Courtney) Channel error 354.6770 Reseated pressure sensor Replaced cracked rear case Check-in complete All tests passed Refer to ASM for details
65970 | 196787 | CM | CMPLT | CMPLT | 2023-08-15 08:23 | 2023-08-15 08:24 | Channel error | (TC 08/15/2023 08:22:58 AM, Rhuland, Courtney) Channel error 354.6770 Reseated pressure sensor Check-in complete All tests passed Refer to ASM for details
1055788 | 196788 | CM | CMPLT | CMPLT | 2023-08-15 08:46 | 2023-08-15 08:47 | Unplugged | (TC 08/15/2023 08:45:47 AM, Reid, Emmerson ) Defib alert that it was not plugged in. I plugged it back in to the wall and charging light appeared
| 196789 | CM | CMPLT | CMPLT | 2023-08-15 09:06 | 2023-08-18 12:30 | EMG server upgrade, data migration, and testing
62979 | 196790 | CM | CMPLT | CMPLT | 2023-08-15 10:15 | 2023-08-15 10:16 | Wobbly stand | (TC 08/15/2023 10:15:26 AM, Rhuland, Courtney) Attempted to tighten screws holding post Could not access screws, replaced stand with spare Functional check - OK
1056502 | 196791 | CM | CMPLT | CMPLT | 2023-08-15 13:07 | 2023-09-11 08:30 | Service watchdog error and constant "powering off" state | (TC 08/15/2023 01:06:39 PM, Baghirzade, Nara) Rebooted the radical multiple times, but cannot bypass the service watchdog timer. After a couple of seconds, the radical goes into "powering off" mode until it is physically switched off. Could not find any information in the service manuals. Contacted Masimo tech service for further support. (TC 08/16/2023 07:31:51 AM, Baghirzade, Nara) Could not troubleshoot/mitigate this issue. Masimo tech recommended power cycling; issue still persists. Will have to send it in for service to Masimo. (TC 08/17/2023 12:38:42 PM, Baghirzade, Nara) SC 1003028134 Sending for repair (TC 08/21/2023 02:17:36 PM, Baghirzade, Nara) Sending radical for repair. RMA: 1038548 KJV courier tracking: 386545 (TC 09/01/2023 07:13:44 AM, Baghirzade, Nara) Received Radical 7. PO 7202268132 (TC 09/11/2023 08:29:01 AM, Baghirzade, Nara) Radical 7 replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Removed the radical battery. Printed asset tag and device number.
1066454 | 196792 | FC | CMPLT | CMPLT | 2023-08-15 13:13 | 2023-08-15 13:14 | QC checklist | (TC 08/15/2023 01:13:27 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
1054475 | 196793 | CM | CMPLT | CMPLT | 2023-08-15 13:28 | 2023-09-26 11:10 | Created by Failed PM WO # 193453 : PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 08/15/2023 11:34:04 AM, Cortez, John) Requested service with Hillrom (TC 08/16/2023 03:48:34 PM, Cortez, John) Talked with Vince, and he mentioned that the rail needs to be replaced, and is on order (TC 09/19/2023 03:58:45 PM, Cortez, John) Completed rest of PM (TC 09/19/2023 11:08:12 AM, Cortez, John) Bed repair completed by Vince | (TC 08/15/2023 01:23:51 PM, Cortez, John) Replaced bed+intellidrive batteries Right head rail, bent Pendant physical damage 60 mOhms, 25 uAmps Bed unable to lower completely due to bent siderail. Will place service request
67816 | 196794 | CM | CMPLT | NFF | 2023-08-15 13:35 | 2023-08-15 13:51 | ECG noise | (TC 08/15/2023 01:49:28 PM, Rhuland, Courtney) Tag stated "monitor picks up reading from ECG cable despite nothing plugged in" Performed functional check and ran various ECG tests and no fault found No noise was present when disconnected from simulator
1056405 | 196795 | CM | CMPLT | NFF | 2023-08-15 13:50 | 2023-08-15 13:53 | Temperature always measuring 37.6C | (TC 08/15/2023 01:51:43 PM, Baghirzade, Nara) Temperature probe in good condition and properly seated; measured temperature multiple times in different modes (aux, oral); could not replicate the error that was reported. Temperature probe is fully functional and no fault found. Functional tests passed for NIBP, temp, and SpO2.
69531 | 196796 | PM | CMPLT | CMPLT | 2023-08-15 14:10 | 2023-08-15 14:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/15/2023 02:10:15 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1055776 | 196797 | CM | CMPLT | CMPLT | 2023-08-15 14:13 | 2023-08-15 14:14 | AC unplugged | (TC 08/15/2023 02:12:47 PM, Baghirzade, Nara) Defib dashboard alerted that defib was unplugged from AC. Power cord was loose at the defib end but still plugged in the wall. Ensured that power cord was properly seated in the defib.
| 196798 | CM | CMPLT | CMPLT | 2023-08-15 14:23 | 2023-08-18 09:58 | panda warmer spare control board | (TC 08/15/2023 02:24:50 PM, Rendon Lopez , Elizabeth) Spare control board will be charged when used
67227 | 196799 | CM | CMPLT | CMPLT | 2023-08-15 14:26 | 2023-08-15 14:28 | Error 11 displayed on screen | (TC 08/14/2023 02:25:40 PM, Rendon Lopez , Elizabeth) Warmer removed from OR2, error showed up during use, no harm. Troubleshooting confirmed issue, found control board defective (TC 08/15/2023 02:25:40 PM, Rendon Lopez , Elizabeth) Replaced control board and power board fuses. calibrated scale, tested for hours on different temps, working properly. Returned to L&d Room 17
66518 | 196800 | CM | CMPLT | CMPLT | 2023-08-15 14:32 | 2023-08-15 14:34 | BP keeps inflating | (TC 08/14/2023 02:33:00 PM, Rendon Lopez , Elizabeth) Removed monitor from ODU I, confirmed issue, replaced BP assembly. Tested all functions-OK
66500 | 196801 | CM | CMPLT | CMPLT | 2023-08-15 14:35 | 2023-08-15 14:36 | Loud printer | (TC 08/15/2023 02:34:54 PM, Rendon Lopez , Elizabeth) Checked in room, paper issue. Cleaned recorder head, replaced paper, printing with no issues
1050479 | 196802 | CM | CMPLT | CMPLT | 2023-08-15 14:47 | 2023-08-28 14:17 | battery replacement needed | (TC 08/15/2023 02:47:34 PM, Rendon Lopez , Elizabeth) Servo n notification, battery due for replacement. ordered 2 batteries (TC 08/28/2023 02:16:52 PM, Rendon Lopez , Elizabeth) Received batteries, delivered to Rex
1050482 | 196803 | CM | CMPLT | CMPLT | 2023-08-15 14:48 | 2023-08-28 14:17 | Battery replacement needed | (TC 08/15/2023 02:48:46 PM, Rendon Lopez , Elizabeth) Notification battery repalcement needed. ordered battery (TC 08/28/2023 02:16:14 PM, Rendon Lopez , Elizabeth) Received batteries, delivered to Rex
1067515 | 196804 | CM | CMPLT | CMPLT | 2023-08-16 07:56 | 2023-08-14 11:38 | Installation and set up | (TC 08/14/2023 07:55:48 AM, Samanian, Kamran) Installed Device, given remote access to Samuel to check the device, it was supposed to come in Pre-Configured but not and he had to work on it. finalized set up, licensed system, tested all and passed. email attached. ticket#256D6CG
| 196805 | CM | CMPLT | CMPLT | 2023-08-16 08:03 | 2023-08-24 16:59 | EMG Server Upgrade - Started | (TC 08/15/2023 08:03:21 AM, Samanian, Kamran) Chris from Natus joined the call and remotely connected to the clients and server. He backed up 4 individual files ( results, log, config, ,,) and moved them to new server database ( exiting new physical new EEG server ) . then started moving patients results into the new DB server. left the job as it was taking time (TC 08/18/2023 04:58:32 PM, Samanian, Kamran) connected with Christina. one of the client was syncing with delay. she deleted corrupted test patient and it resolved the issue.
1061209 | 196806 | CM | CMPLT | NFF | 2023-08-16 08:31 | 2023-08-16 08:32 | Light's tablet not working | (TC 08/16/2023 08:30:59 AM, Reid, Emmerson ) Reported that it was not working but everything seemed fine, all controls worked and the battery was charged. NFF
1061210 | 196807 | CM | CMPLT | NFF | 2023-08-16 08:33 | 2023-08-16 08:35 | Boom can not move | (TC 08/16/2023 08:33:49 AM, Reid, Emmerson ) Told they could not move the boom, but actually they were not happy with how the room was rearranged overnight. I helped them move the equipment. NFF
| 196808 | PA | CMPLT | CMPLT | 2023-08-16 08:58 | 2023-08-21 09:54 | Sterile water
69357 | 196809 | PM | CMPLT | CMPLT | 2023-08-16 09:17 | 2023-08-16 09:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 09:17:40 AM, Rhuland, Courtney) Rate calibrated PM complete All tests passed post calibration Refer to ASM for details
1051085 | 196810 | CM | CMPLT | CMPLT | 2023-08-16 09:27 | 2023-08-16 10:03 | SPO2 intermittent | (TC 08/16/2023 09:53:31 AM, Scero, Tomlyn) Confirmed that readings were intermittent. Replaced the Cable and sensor
63968 | 196811 | CM | CMPLT | CMPLT | 2023-08-16 10:12 | 2023-08-17 08:44 | Not working with Instrument | (TC 08/16/2023 10:20:34 AM, Reid, Emmerson ) Called tech support who made a ticket and connected to the Sales Rep. It sounds like the instrument and the device are working on incompatable sofeware versions. In the room they brought in two devices and two instruments and both had the same issue. Spoke with the vendor and they plan on coming in today to help resolve this issue. Copying this one the other asset as well. (TC 08/16/2023 02:27:08 PM, Reid, Emmerson ) Met with the Vendor and set them up with the three Harmonic devices. They are working on upgrades now. (TC 08/17/2023 08:43:09 AM, Reid, Emmerson ) Software was upgraded to 2018-1 and TMS was cleaned up for this model.
1062949 | 196812 | CM | CMPLT | CMPLT | 2023-08-16 10:25 | 2023-08-17 08:44 | Device not working with instrument | (TC 08/16/2023 10:25:13 AM, Reid, Emmerson ) See WO on asset# 63968 (TC 08/17/2023 08:44:11 AM, Reid, Emmerson ) Software was upgraded to 2018-1 and TMS was cleaned up for this model.
1059778 | 196813 | FC | CMPLT | CMPLT | 2023-08-16 11:11 | 2023-08-16 11:12 | QC checklist | (TC 08/16/2023 11:11:17 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
| 196814 | CM | CMPLT | CMPLT | 2023-08-16 11:17 | 2023-08-14 11:17 | Biomed meeting | Biomed meeting
66474 | 196815 | CM | CMPLT | CMPLT | 2023-08-16 11:18 | 2023-08-16 11:18 | needs wheel alignment | needs wheel alignment/ Adjusted wheell. Returned back to RT
63814 | 196816 | CM | CMPLT | CMPLT | 2023-08-16 11:20 | 2023-08-16 11:20 | broken bracket | replaced broken bracet. tested. functional.
1047834 | 196817 | CM | CMPLT | CMPLT | 2023-08-16 11:21 | 2023-08-16 11:21 | alarming? | tested unit. functional. returned back to RT
65855 | 196818 | CM | CMPLT | CMPLT | 2023-08-16 11:22 | 2023-08-15 11:21 | Functional check | Tested. Functional.
53006 | 196819 | CM | CMPLT | CMPLT | 2023-08-16 11:23 | 2023-08-15 11:22 | Functional check | Tested. Functional.
63138 | 196820 | CM | CMPLT | CMPLT | 2023-08-16 11:25 | 2023-08-15 11:24 | Functional check | Tested. Functional.
1066462 | 196821 | FC | CMPLT | CMPLT | 2023-08-16 12:19 | 2023-08-16 12:20 | Eject button not working properly +QC checklist | (TC 08/16/2023 12:19:06 PM, Baghirzade, Nara) Could not find an issue with the eject button; working properly. QC checklist completed. Returned to 14 South
1050039 | 196822 | PM | CMPLT | CMPLT | 2023-08-16 12:27 | 2023-08-16 12:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/16/2023 12:27:27 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
501063 | 196823 | CM | CMPLT | CMPLT | 2023-08-16 13:18 | 2023-08-16 13:18 | No display on main monitor | (TC 08/16/2023 01:17:53 PM, Fan, Tak) Confirmed issue. Checked cabling. Used PC monitor, which was fine. Placed call for monitor. (TC 08/17/2023 03:36:43 PM, Fan, Tak) Retrieved from receiving, replaced monitor.
1051886 | 196824 | CM | CMPLT | CMPLT | 2023-08-16 13:48 | 2023-08-28 08:53 | replace valve
1059980 | 196825 | PM | CMPLT | CMPLT | 2023-08-16 15:30 | 2023-08-16 15:39 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 08/16/2023 03:30:20 PM, Cortez, John) Bed+Intellidrive batteries were replaced 110 mOhms, 22 uAmps Pendant missing clip PM passed
| 196826 | CM | CMPLT | CMPLT | 2023-08-17 08:10 | 2023-09-01 10:06 | Neptune waste left behind after repair issue. | (TC 08/16/2023 08:09:46 AM, Chukri, Samir) Meeting with Jessica,Adeel, Emmerson and Kamran . (TC 09/01/2023 10:02:18 AM, Chukri, Samir) MDRD tooke clean the defective and put it in e-waste box
67963 | 196827 | CM | CMPLT | CMPLT | 2023-08-17 08:28 | 2023-08-17 10:17 | System failure error | (TC 08/17/2023 08:28:00 AM, Baghirzade, Nara) Levelled the incubator with Courtney. (TC 08/17/2023 10:15:28 AM, Rhuland, Courtney) Reseated fan to fix system error 25 Noticed incubator was not level Levelled incubator base with Nara Calibrated RH Functional check - OK | Originally sent down for a system failure error but Courtney noted that the incubator was not levelled
67971 | 196828 | CM | CMPLT | CMPLT | 2023-08-17 08:29 | 2023-08-17 08:30 | Broken latch | (TC 08/17/2023 08:28:41 AM, Baghirzade, Nara) Replaced broken latch. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
1061220 | 196829 | CM | CMPLT | CMPLT | 2023-08-17 10:09 | 2023-08-17 10:11 | Microscope image shaking | (TC 08/17/2023 10:08:54 AM, Reid, Emmerson ) Called up to OR19. Image was shacking but iit was not due to the microscope, spoke with vendor who confirmed it. I belive the patient was moving a bit. They were fine to proceed with proceedure.
1053981 | 196830 | CM | CMPLT | CMPLT | 2023-08-17 11:28 | 2023-08-17 11:29 | Video cutting in/out during use | (TC 08/17/2023 11:27:40 AM, Baghirzade, Nara) Cable is damaged internally on the side closest to the scope. Replaced the cable. Tested the video output and video was in good resolution. Played with the cable and video was stable. Dropped off at Resus.
32520 | 196831 | CM | CMPLT | CMPLT | 2023-08-17 11:49 | 2023-08-31 11:08 | MSH DR2 COLLIMATOR FIELD ILLUMINATION LED INTERMETTENT | (TC 08/17/2023 01:32:54 PM, Whittington, Scott) MSH DR2 COLLIMATOR FIELD ILLUMINATION LED INTERMETTENT | MSH DR2 COLLIMATOR FIELD ILLUMINATION LED INTERMETTENT
501133 | 196832 | CM | CMPLT | CMPLT | 2023-08-17 11:53 | 2023-08-21 12:05 | MSH 5 US IU22 SITE 553814 SW LOAD STALLS PAGE 17 | (TC 08/17/2023 12:03:48 PM, Whittington, Scott) US IU22 SITE 553814 SW LOAD STALLS PAGE 17 SW VERSION 6.3.7.745 REF # 0122148223 ERAN (TC 08/18/2023 11:40:28 AM, Whittington, Scott) MSH 5 US IU22 SITE 553814 SW LOAD STALLS PAGE 17 JILLIAN THOMPSON ON SITE (TC 08/21/2023 12:03:48 PM, Whittington, Scott) MSH 5 US IU22 SITE 553814 SW LOAD STALLS PAGE 17 SW VERSION 6.3.7.745 REF # 0122148223 ERAN JILLIAN RELOADED SW, BUT THE BACKUP DISC NOT CURRENT, WAS FOR TWH INFO RELOADED SETTINGS FROM USB, BURNED A DVD AND LABELLED SYSTEM WITH CORRECTMSH AE (TC 08/21/2023 12:03:48 PM, Whittington, Scott) MSH 5 US IU22 SITE 553814 SW LOAD STALLS PAGE 17 SW VERSION 6.3.7.745 REF # 0122148223 ERAN JILLIAN RELOADED SW, BUT THE BACKUP DISC NOT CURRENT, WAS FOR TWH INFO RELOADED SETTINGS FROM USB, BURNED A DVD AND LABELLED SYSTEM WITH CORRECTMSH AE (TC 09/08/2023 11:30:20 AM, Fan, Tak) SW issue reoccurred. UMB and HDDs were replaced. Applied for MAC filtering. Setup network. Tested, returned to dept. (TC 09/07/2023 11:46:27 AM, Fan, Tak) Checked system with Sara. | MSH 5 US IU22 SITE 553814 SW LOAD STALLS PAGE 17 SW VERSION 6.3.7.745 REF # 0122148223 ERAN
| 196833 | CM | CMPLT | CMPLT | 2023-08-17 12:22 | 2023-08-17 12:23 | Incubator replacement meeting | (TC 08/17/2023 12:22:29 PM, Baghirzade, Nara) Incubator replacement meeting
1053347 | 196834 | CM | CMPLT | CMPLT | 2023-08-17 13:31 | 2023-09-20 16:47 | MSH BI 1270 H1 Face Shield releasing and sliding inward without intentional release activation. | (TC 08/17/2023 01:30:48 PM, Whittington, Scott) MSH BI 1270 H1 Face Shield releasing and sliding inward without intentional release activation. (TC 08/24/2023 12:05:26 PM, Whittington, Scott) MSH BI 1270 H1 Face Shield releasing and sliding inward without intentional release activation. RECEIVED AND REPLACED FACESHIELD PN# ASY-10641 | From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Thursday, August 17, 2023 10:41 AM To: Breast Health Support <BreastHealth.Support@Hologic.com> Cc: Veeneman, Jon <Jon.Veeneman@hologic.com>; Moe - Dynamic Imaging <moe@dynamicimaging.ca>; Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Murphy, Eimear - Sinai Health System <Eimear.Murphy@sinaihealth.ca>; Vaillancourt, Melanie - Sinai Health System <Melanie.Vaillancourt@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: MSH H1 sn# 3DM160700412 Face Shield releasing and sliding inward without release activation Good morning MSH H1 Face Shield releasing and sliding inward without intentional release activation. Thanks Scott 416-402-2462 Get Outlook for iOS [aka.?ms] This e-mail may contain confidential and/or privileged information for the sole use Good morning MSH H1 Face Shield releasing and sliding inward without intentional release activation. Thanks Scott
| 196835 | CM | CMPLT | CMPLT | 2023-08-17 13:35 | 2023-08-17 13:35 | Electrical safety testing of lab equipment in blood bank | (TC 08/17/2023 01:34:57 PM, Baghirzade, Nara) Tested voltage and amperage for lab equipment. Noted the value and what type of test equipment on their validation sheet.
1059828 | 196836 | PM | CMPLT | CMPLT | 2023-08-17 13:52 | 2023-08-17 14:25 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/17/2023 02:23:27 PM, Rhuland, Courtney) Worked with Nara to fix loose back caster Drained, disinfected, rinsed and refilled with sterile water
| 196837 | CM | CMPLT | CMPLT | 2023-08-17 14:27 | 2023-08-17 14:27 | Meetings | (TC 08/17/2023 02:26:38 PM, Rhuland, Courtney) PM update & incubator replacement meetings
67825 | 196838 | PM | CMPLT | CMPLT | 2023-08-17 14:39 | 2023-08-17 14:40 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/16/2023 02:39:23 PM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
69415 | 196839 | PM | CMPLT | CMPLT | 2023-08-17 15:11 | 2023-08-17 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/17/2023 03:11:01 PM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
63349 | 196840 | CM | CMPLT | CMPLT | 2023-08-17 15:14 | 2023-08-17 15:15 | Not working | (TC 08/17/2023 03:14:16 PM, Reid, Emmerson ) Was told it was not working. Found that it was not turning on. Checked the fuses and one was blown. I replaced both and it is working now.
1061967 | 196841 | CM | CMPLT | CMPLT | 2023-08-17 16:34 | 2023-09-05 15:09 | "Difficult to extend foot section" | (TC 08/17/2023 04:34:26 PM, Cortez, John) Confirmed issue Was able to depress handle, but foot section seems to be stuck Patient still in bed sleeping, will try again later (TC 09/05/2023 03:01:52 PM, Cortez, John) Pneumatic tubing dislodged on left side, reattached to frame Tested ok
1062137 | 196842 | CM | CMPLT | CMPLT | 2023-08-17 16:39 | 2023-08-17 16:56 | Ground pin on ower cord missing | (TC 08/17/2023 04:39:33 PM, Cortez, John) Replaced cord
1054962 | 196843 | CM | CMPLT | CMPLT | 2023-08-17 17:00 | 2023-08-18 11:29 | "Not functional" | (TC 08/17/2023 05:00:14 PM, Cortez, John) Tested all parameters BP hose connector leaking, swapped connector from spare Probe well has debris inside, swapped SPO2 flatlines randomly (TC 08/18/2023 11:23:17 AM, Cortez, John) Swapped radical and SPO2 extension. Appears more stable
501066 | 196844 | CM | CMPLT | CMPLT | 2023-08-17 19:22 | 2023-08-24 08:53 | MSH CEOU IU22 - RM#4 - System Red Screen and Error 561 | (TC 08/17/2023 07:21:42 PM, Chau, Ken) Ran system test and all passed Control panel LCD screen displaying half red Once restarted generates error FSE inform it will require a UMB replacement case #0122147289 (TC 08/18/2023 11:42:20 AM, Whittington, Scott) MSH CEOU RM 4 Ran system test and all passed Control panel LCD screen displaying half red Once restarted generates error FSE inform it will require a UMB replacement case #0122147289 JILLIAN THOMPSON ON-SITE PROVIDE DOLLY TO MOVE UMB INTO OPG MAC WHITELISTING REQUIRED WHEN ADVISED OF MAC | case #0122147289
1061214 | 196845 | CM | CMPLT | CMPLT | 2023-08-18 05:56 | 2023-08-18 06:05 | Annual Inspection | (TC 08/18/2023 06:03:15 AM, Reid, Emmerson ) Checked the monitors and tightened two of them (1 and 3) Checked the booms all were working and moving well without drift. Checked iPad and was fully charged (95% and charging. Inspection Complete.
1061217 | 196846 | CM | CMPLT | CMPLT | 2023-08-18 07:13 | 2023-08-18 07:17 | Annual Inspection | (TC 08/18/2023 07:14:05 AM, Reid, Emmerson ) Checked all the monitors, good. Checked the iPad, fully charged. Checked the booms, all good no drifting. Inspection Complete.
1049062 | 196847 | PM | CMPLT | CMPLT | 2023-08-18 07:25 | 2023-08-18 07:58 | Planned Maintenance | (TC 08/18/2023 07:48:08 AM, Reid, Emmerson ) PM Complete, see proceedure.
63873 | 196848 | CM | CMPLT | CMPLT | 2023-08-18 07:34 | 2023-08-18 07:35 | Door hinge, latch, and hood clip was missing/damaged | (TC 08/18/2023 07:34:20 AM, Baghirzade, Nara) Replaced door latch, hinge, and hood clip. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
1061220 | 196849 | CM | CMPLT | CMPLT | 2023-08-18 08:00 | 2023-08-18 08:02 | Annual Inspection | (TC 08/18/2023 08:00:45 AM, Reid, Emmerson ) Checked all the monitors, good. Checked the booms and needed to fix the drifting in 3 spots (2 on one boom). iPad fully charged. Inspection complete.
| 196850 | RDS | CMPLT | CMPLT | 2023-08-18 08:49 | 2023-08-18 08:51 | Rounds: 18South and PACU PODA | (TC 08/18/2023 08:50:20 AM, Baghirzade, Nara) Completed rounds in 18S and PACU PODA. See procedure document.
| 196851 | RDS | CMPLT | CMPLT | 2023-08-18 08:51 | 2023-08-18 08:53 | Rounds: 11 South | (TC 08/18/2023 08:52:40 AM, Baghirzade, Nara) Rounds in 11 South completed. See procedure document.
| 196852 | RDS | CMPLT | CMPLT | 2023-08-18 08:54 | 2023-08-18 08:59 | Rounds: NICU and RESUS | (TC 08/18/2023 08:57:51 AM, Baghirzade, Nara) Rounds in NICU and RESUS completed. See procedure.
62699 | 196853 | CM | CMPLT | CMPLT | 2023-08-18 09:41 | 2023-08-18 09:55 | reprogrammed for oncology | (TC 08/18/2023 09:53:15 AM, Scero, Tomlyn) This cart was repurposed to outpatient oncology for temporary use until they receive their own cart. I deleted the access to iECG and database because they won't use ECG orders. They will only be printing ECGs. Funcitonal checks complete and passed.
1065115 | 196854 | CM | CMPLT | CMPLT | 2023-08-18 09:50 | 2023-09-27 14:33 | UroCap not connecting | (TC 08/18/2023 09:50:46 AM, Reid, Emmerson ) Laborie Ticket# 58096 UroCap was not connecting (SN UC5-12-30009074) Tried troubleshooting with vendor but we could not resolve. They said it was under warranty and would send a new one by next week. (TC 09/08/2023 03:02:27 PM, Reid, Emmerson ) Went over with the new urocap and troubleshooted on the phone with tech suport. Still was not connecting. They believe it is an issue still with the (new) urocap. (TC 09/27/2023 02:32:28 PM, Reid, Emmerson ) Connected the new Urocap and it worked! Then calibrated. Device is working now.
1061215 | 196855 | CM | CMPLT | CMPLT | 2023-08-18 10:04 | 2023-08-18 10:04 | Monitors needed to be tightened | (TC 08/18/2023 10:04:02 AM, Reid, Emmerson ) Complete
69329 | 196856 | PM | CMPLT | CMPLT | 2023-08-18 10:19 | 2023-08-18 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/18/2023 10:19:19 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1060147 | 196857 | CM | CMPLT | CMPLT | 2023-08-18 10:56 | 2023-08-25 13:19 | Error code E0.6 | (TC 08/18/2023 12:45:54 PM, Canjimavadivel, Kamallesh ) Tried swapping batteries and probe as per service manual, error still present (TC 08/22/2023 11:07:13 AM, Canjimavadivel, Kamallesh ) Called tech support to further troubleshooting. Attempted to factory reset and issue did not resolve. Checked warranty is still valid. Initiated replacement return. (TC 08/25/2023 01:17:28 PM, Canjimavadivel, Kamallesh ) Received replacement from vendor. Updated serial number in TMS. Packed up defective thermometer as per instructions from vendor for shipping.
1061216 | 196858 | CM | CMPLT | CMPLT | 2023-08-18 12:54 | 2023-08-18 12:54 | Olympus VMC monitor froze | (TC 08/18/2023 12:53:41 PM, Reid, Emmerson ) Reseated connectors, working now.
67429 | 196859 | CM | CMPLT | CMPLT | 2023-08-18 13:35 | 2023-08-18 13:37 | Equipment assessment for disposal | (TC 08/18/2023 01:35:28 PM, Baghirzade, Nara) Sri reached out to have a scale assessed for disposal. Scale is very old, not calibrated, and no parts available for it. Scale has not been used in a long time and has been in storage. Have spoken to Samir and agreed to retire this device. Sri will arrange with Michael Badour for the disposal.
| 196860 | CM | CMPLT | CMPLT | 2023-08-18 14:38 | 2023-08-17 14:37 | Biomed meeting | Biomed meeting
52170 | 196861 | CM | CMPLT | NFF | 2023-08-21 08:03 | 2023-08-21 08:04 | Humidifier not working | (TC 08/21/2023 08:03:23 AM, Rhuland, Courtney) Ran functional check for 30 minutes Humidity was accurate No fault found
1063746 | 196862 | PM | CMPLT | CMPLT | 2023-08-21 08:20 | 2023-08-21 08:45 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/21/2023 08:44:29 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
66683 | 196863 | CM | CMPLT | CMPLT | 2023-08-21 09:02 | 2023-08-21 09:09 | Latch recall | (TC 08/21/2023 09:08:35 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
1066454 | 196864 | CM | CMPLT | CMPLT | 2023-08-21 09:29 | 2023-08-30 09:01 | Canister pops off without being ejected | (TC 08/21/2023 09:28:48 AM, Baghirzade, Nara) Contacted 3M to send the VAC in for service. Unit covered under warranty. Confirmation #: 12812228 Requested loaner while VAC is out for service. Loaner confirmation: 13033556 (TC 08/22/2023 06:23:20 AM, Baghirzade, Nara) Attached pickup report for 14S VAC (TC 08/30/2023 09:00:39 AM, Baghirzade, Nara) VAC was repaired by 3M. VAC was received and directly dropped off at 14S with Julie and Monica. (TC 08/30/2023 09:00:39 AM, Baghirzade, Nara) VAC was repaired by 3M. Attached QC checklist.
65983 | 196865 | CM | CMPLT | CMPLT | 2023-08-21 09:47 | 2023-10-13 12:54 | Created by Failed PM WO # 194350 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/21/2023 02:00:23 PM, Baghirzade, Nara) Replaced rear case. Replaced left handle. Reprinted asset tag. Transferred lock latch on to rear case. Noticed that the mylar gasket was also damaged and there was some kind of a leak (sugar feed) all over the pressure sensor holder and mylar gasket; cleaned the pressure sensor holder and need to replace the mylar gasket. Do not have biomed stock for mylar gasket, will have to place an order. Removed and disposed of old mylar gasket. (TC 10/13/2023 12:51:52 PM, Baghirzade, Nara) Replaced mylar gasket. Replaced pressure sensor disc. Asssembled pump. Pressure sensor disc was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 08/21/2023 09:45:35 AM, Baghirzade, Nara) Pressure disc accuracy verification failed during PM; checked pressure disc with analog sensor test and pressure is not being detected. Will replace the pressure sensor disc.
1049608 | 196866 | CM | CMPLT | CMPLT | 2023-08-21 11:46 | 2023-08-21 11:46 | MSH SPP - RM6 Hall D - GE Voluson - No Network | (TC 08/21/2023 11:45:48 AM, Chau, Ken) Network wont connect -connected with IT at MSH (Mirza) and was able to get connectivity back -Mirza had to add/remove network
| 196867 | CM | CMPLT | CMPLT | 2023-08-21 12:12 | 2023-08-21 12:13 | Disassembling devices for retirement/disposal | (TC 08/21/2023 12:12:01 PM, Baghirzade, Nara) Removed old blood pressure devices from Family Medicine Clinic to retire and dispose of. (TC 08/21/2023 12:12:33 PM, Rhuland, Courtney) Removed old blood pressure devices from Family Medicine Clinic to retire and dispose of.
| 196868 | CM | CMPLT | CMPLT | 2023-08-21 13:50 | 2023-08-21 13:50 | Monday's PM Huddle | (TC 08/21/2023 01:49:56 PM, Baghirzade, Nara) Monday's PM Huddle
| 196869 | CM | CMPLT | CMPLT | 2023-08-21 13:52 | 2023-08-21 13:53 | IsoTour pumps missing power cords on 11S | (TC 08/21/2023 01:51:34 PM, Baghirzade, Nara) Only had 1 spare power cord that is specifically for IsoTour pumps. Gave the power cord to Sri. Had to place a medical grade power cord on the second IsoTour pump. Requested more power cords from Melissa at Stryker; confirmed with Nima that power cords are covered under service contract. Once the power cords come in, I will replace the medical grade power cord with the IsoTour power cord.
| 196870 | CM | CMPLT | CMPLT | 2023-08-21 14:19 | 2023-08-21 14:19 | Biomed meeting | Biomed meeting
| 196871 | CM | CMPLT | CMPLT | 2023-08-21 14:28 | 2023-09-05 15:13 | Spare fetal monitor BP assy | (TC 08/21/2023 02:28:36 PM, Rendon Lopez , Elizabeth) Ordering spare parts for fetal monitor. Will be charged when used Received
| 196872 | CM | CMPLT | CMPLT | 2023-08-21 14:37 | 2023-09-14 13:22 | Spare pump parts | (TC 08/21/2023 03:11:10 PM, Rendon Lopez , Elizabeth) Spare pump parts ordering (TC 08/28/2023 04:02:26 PM, Rendon Lopez , Elizabeth) Received batteries, front cases, and bezels (TC 09/14/2023 01:21:22 PM, Rendon Lopez , Elizabeth) Received all parts
1047840 | 196873 | CM | CMPLT | CMPLT | 2023-08-21 14:40 | 2023-08-21 14:41 | Update date and time | (TC 08/21/2023 02:40:15 PM, Baghirzade, Nara) Updated date and time. Updated PM schedule. Confirmed that the correct configuration was loaded.
| 196874 | CM | CMPLT | CMPLT | 2023-08-21 14:43 | 2023-08-21 14:44 | huddle | (TC 08/21/2023 02:43:32 PM, Rhuland, Courtney) Huddle/presentation
1064418 | 196875 | CM | CMPLT | CMPLT | 2023-08-21 15:08 | 2023-08-21 15:09 | Port Door Fell off | (TC 08/21/2023 03:07:58 PM, Ferreira, Qwynn) Port Door was seen hanging off the incubator Replaced the door and inspected the rest of the incubator A previous repalced door was starting to tear at the top of the seem - Notified Adeel/Samir about this.
66500 | 196876 | CM | CMPLT | CMPLT | 2023-08-21 15:18 | 2023-08-21 17:27 | Monitor not working | (TC 08/21/2023 03:56:12 PM, Rendon Lopez , Elizabeth) Tested in room, found NIBP not inflating. Removed and placed spare. (TC 08/21/2023 03:56:12 PM, Rendon Lopez , Elizabeth) Replaced NIBP assy, calibrated, tested-OK. Al functions OK. Returned to Room 8
1049362 | 196877 | CM | CMPLT | CMPLT | 2023-08-21 15:30 | 2023-08-21 15:31 | MSH ER Injector - Defective Set / Contrast Side Fluid | (TC 08/21/2023 03:30:13 PM, Chau, Ken) Contrast on injector causing plunger to freeze up Cleaned contrast and informed supervisor of issue To inspect new sets
1057112 | 196878 | CM | CMPLT | CMPLT | 2023-08-21 17:28 | 2023-08-21 17:33 | Needs to be calibrated | (TC 08/21/2023 05:31:49 PM, Rendon Lopez , Elizabeth) Tested all functions, no fault found. Calibrated NIBP, tested-OK.
1067554 | 196879 | CM | CMPLT | CMPLT | 2023-08-22 06:28 | 2023-08-22 06:31 | Received loaner | (TC 08/22/2023 06:28:34 AM, Baghirzade, Nara) Attached drop off receipt and inspection card. Printed asset tag. QC completed by 3M. Returned to 14S.
62727 | 196880 | CM | CMPLT | CMPLT | 2023-08-22 06:35 | 2023-08-22 06:36 | Removing device from service | (TC 08/22/2023 06:34:48 AM, Baghirzade, Nara) Device not supported, no parts/repair available. Removed device from service and will be retired/disposed. Clinic was provided with Masimos as a replacement.
60142 | 196881 | CM | CMPLT | CMPLT | 2023-08-22 06:37 | 2023-08-22 06:37 | Removing device from service | (TC 08/22/2023 06:36:35 AM, Baghirzade, Nara) Device not supported, no parts/repair available. Removed device from service and will be retired/disposed. Clinic was provided with Masimos as a replacement.
67145 | 196882 | CM | CMPLT | CMPLT | 2023-08-22 06:39 | 2023-08-22 06:40 | Removing device from service | (TC 08/22/2023 06:39:20 AM, Baghirzade, Nara) Device not supported, no parts/repair available. Removed device from service and will be retired/disposed. Clinic was provided with Masimos as a replacement.
67151 | 196883 | CM | CMPLT | CMPLT | 2023-08-22 06:40 | 2023-08-22 06:41 | Removing device from service | (TC 08/22/2023 06:40:16 AM, Baghirzade, Nara) Device not supported, no parts/repair available. Removed device from service and will be retired/disposed. Clinic was provided with Masimos as a replacement.
1055126 | 196884 | CM | CMPLT | NFF | 2023-08-22 08:13 | 2023-08-22 08:14 | No issue described on tag | (TC 08/22/2023 08:13:11 AM, Baghirzade, Nara) Unable to find the issue with the PCU. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 175 mohms Leakage current: 61.1 uA No fault found.
| 196885 | RDS | CMPLT | CMPLT | 2023-08-22 08:15 | 2023-08-22 08:54 | Weekly check for ED
| 196886 | RDS | CMPLT | CMPLT | 2023-08-22 08:16 | 2023-08-23 13:59 | Weekly check for ICU
1055346 | 196887 | CM | CMPLT | CMPLT | 2023-08-22 08:20 | 2023-08-22 08:21 | Removing device from service | (TC 08/22/2023 08:19:55 AM, Baghirzade, Nara) Removed scale from service; defective and no parts available. Nelia is replacing all infant scales at the clinic (capital approval). Removed measuring rod as it is in good condition. Disposed of the scale.
66069 | 196888 | CM | CMPLT | CMPLT | 2023-08-22 08:46 | 2023-08-23 14:39 | Not turning on | (TC 08/22/2023 08:46:23 AM, Rhuland, Courtney) Replaced front case and display with spare AC power not registering Replaced power cord Charging battery before check-in (TC 08/23/2023 02:39:03 PM, Rhuland, Courtney) Check-in complete All tests passed
| 196889 | CM | CMPLT | CMPLT | 2023-08-22 10:10 | 2023-08-29 08:02 | MOL Inspection Biomed. | (TC 08/22/2023 10:09:59 AM, Chukri, Samir) inspection the Biomed workshop safety. with Nicol from occupation health
| 196890 | PA | CMPLT | CMPLT | 2023-08-22 10:12 | 2023-08-29 13:19 | Ordering for LTRI | (TC 08/22/2023 10:48:45 AM, Baghirzade, Nara) SC 1003031344 (TC 08/29/2023 01:18:40 PM, Baghirzade, Nara) PO 7202269884
65970 | 196891 | CM | CMPLT | CMPLT | 2023-08-22 11:44 | 2023-08-22 11:45 | Channel error | (TC 08/22/2023 11:43:52 AM, Rhuland, Courtney) Channel error 354.6690 Reseated pressure sensor Calibrated and performed check-in Error cleared All tests passed
1057033 | 196892 | CM | CMPLT | CMPLT | 2023-08-22 11:46 | 2023-09-13 10:24 | No charge PO for root replacement | (TC 08/22/2023 11:49:45 AM, Baghirzade, Nara) Requested no-charge PO for a root replacement. The Root that was received a week ago has incorrect software version and unable to push the MSH configuration to it. Masimo to provide a new Root with the correct software free of charge. (TC 08/22/2023 11:50:48 AM, Baghirzade, Nara) SC 1003031671 (TC 08/23/2023 07:34:49 AM, Baghirzade, Nara) RMA 1038765 KJV 386883 (TC 08/23/2023 07:34:49 AM, Baghirzade, Nara) RMA 1038765 KJV 386883 No charge PO 7202269891 (TC 09/13/2023 10:23:13 AM, Baghirzade, Nara) Received replacement Root under NO CHARGE PO: 7202269891
68812 | 196893 | PM | CMPLT | CMPLT | 2023-08-22 12:34 | 2023-08-23 07:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/22/2023 12:34:14 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 73 mohms Leakage current: 62.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/23/2023 07:39:54 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4110 mamphr PM completed.
52411 | 196894 | CM | CMPLT | CMPLT | 2023-08-22 13:33 | 2023-08-22 13:57 | Binocular head eye pieces not articulating | (TC 08/21/2023 01:33:52 PM, Cortez, John) Confirmed issue. Uninstalled binocular head and disassembled. Noticed that there is a buildup of dirt/debris, which is preventing eyepieces from articulating. Cleaned with alcohol and reinstalled on slit lamp. Dr. Dushenski verified that lamp is functioning normally
| 196895 | CM | CMPLT | CMPLT | 2023-08-22 14:00 | 2023-08-25 16:46 | ordering spare force triad foot pedal
| 196896 | CM | CMPLT | CMPLT | 2023-08-22 14:08 | 2023-08-22 14:08 | Biomed meeting
1047851 | 196897 | CM | CMPLT | CMPLT | 2023-08-22 14:49 | 2023-08-23 14:06 | System failure | (TC 08/23/2023 01:57:54 PM, Rhuland, Courtney) When powering on, error present "Primary Audible Alarm Post" Opened device and reseated primary speaker Noted there was a message in OneRecall from ICU medical stating software problems for this error Ran multiple test infusions over the course of 2 days Powered on and off device multiple times and error could not be replicated after reseating speaker
1063000 | 196898 | PM | CMPLT | CMPLT | 2023-08-22 14:49 | 2023-08-22 14:50 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/22/2023 02:49:14 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
66692 | 196899 | CM | CMPLT | CMPLT | 2023-08-23 07:33 | 2023-08-23 08:57 | Broken door | (TC 08/23/2023 08:54:56 AM, Rhuland, Courtney) Broken door that came down was the wrong door (door has not been recalled) Replaced door with correct door Replaced porthole door Calibrated RH Functional check - OK
| 196900 | CM | CMPLT | CMPLT | 2023-08-23 08:09 | 2023-08-22 08:10 | ERBE training | (TC 08/21/2023 08:10:13 AM, Samanian, Kamran) ERBE VIO3 training (TC 08/22/2023 08:10:36 AM, Samanian, Kamran) APC training
66155 | 196901 | CM | CMPLT | CMPLT | 2023-08-23 08:16 | 2023-08-23 08:17 | Display does not work | (TC 08/23/2023 08:16:20 AM, Baghirzade, Nara) When removing probe from probe holder, the screen remains blank but the thermometer beeps. Replaced the unit entirely with a spare biomed unit. Functional test passed. Changed the SN on the asset. Printed asset tag.
66043 | 196902 | PM | CMPLT | CMPLT | 2023-08-23 08:27 | 2023-08-24 08:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 08:27:40 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 35 mohms Leakage current: 86.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/24/2023 08:34:07 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4083 mamphr PM completed.
66043 | 196903 | PM | CANCL | DUP | 2023-08-23 08:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66037 | 196904 | PM | CMPLT | CMPLT | 2023-08-23 08:38 | 2023-08-24 10:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 08:37:52 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 180 mohms Leakage current: 81.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 08/24/2023 10:28:51 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4097 mamphr PM completed.
1054537 | 196905 | PM | CMPLT | CMPLT | 2023-08-23 09:09 | 2023-08-23 09:18 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/23/2023 09:11:10 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1067564 | 196906 | CM | CMPLT | CMPLT | 2023-08-23 09:45 | 2023-08-23 11:31 | Incoming inspection | (TC 08/23/2023 11:10:23 AM, Canjimavadivel, Kamallesh ) Unboxed and assembled onto stand. Got previously delivered accessories from storage. Installed battery pack. Delivered device to ICU RT room
1062999 | 196907 | PM | CMPLT | CMPLT | 2023-08-23 10:02 | 2023-08-23 10:37 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/23/2023 10:36:38 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1058228 | 196908 | PM | CMPLT | CMPLT | 2023-08-23 10:27 | 2023-08-23 10:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 10:27:01 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69389 | 196909 | PM | CMPLT | CMPLT | 2023-08-23 10:32 | 2023-08-23 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 10:32:18 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 196910 | CM | ACTIV | ISSUE | 2023-08-23 12:12 |  | meeting for IBE testing with new iBUS | (TC 08/23/2023 12:12:16 PM, Samanian, Kamran) still in progress, Cerner was not able to see TEST1 bed label (TC 09/12/2023 09:55:07 AM, Samanian, Kamran) Tested IBE-AGM with cerner and all passed (TC 09/19/2023 08:25:45 AM, Samanian, Kamran) preparing monitor and X2 for NICU UAT testing (TC 10/13/2023 11:37:07 AM, Samanian, Kamran) done (TC 10/13/2023 11:37:07 AM, Samanian, Kamran) done
| 196911 | CM | CMPLT | CMPLT | 2023-08-23 12:13 | 2023-08-23 12:13 | OR exam table meeting - space | (TC 08/23/2023 12:13:38 PM, Samanian, Kamran) done
66027 | 196912 | CM | CMPLT | CMPLT | 2023-08-23 13:43 | 2023-08-23 14:22 | Broken cases | (TC 08/23/2023 02:21:45 PM, Rhuland, Courtney) Replaced front and rear cases Performed check-in All tests passed Refer to ASM for details
69266 | 196913 | PM | CMPLT | CMPLT | 2023-08-23 14:05 | 2023-08-23 14:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 02:05:33 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69117 | 196914 | PM | CMPLT | CMPLT | 2023-08-23 14:17 | 2023-08-23 14:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/23/2023 02:17:22 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1061209 | 196915 | CM | CMPLT | CMPLT | 2023-08-23 14:31 | 2023-08-23 14:32 | Light iPad blank | (TC 08/23/2023 02:31:20 PM, Reid, Emmerson ) iPad was asleep. I changed the settings so it would stay on.
| 196916 | CM | CMPLT | CMPLT | 2023-08-23 14:35 | 2023-08-23 14:37 | Othalmoscope not working | (TC 08/23/2023 02:35:01 PM, Reid, Emmerson ) Went over to chekc the opthalmoscope and the light was not working. I brought it back and fixed it by replacing the headpiece becuase we did not have bullbs that fit. emailed with them to create account on TMS
1049362 | 196917 | CM | CMPLT | CMPLT | 2023-08-23 15:04 | 2023-08-31 08:29 | MSH - Injector Emerg - Plunger on Side A Not Latching | (TC 08/23/2023 03:03:54 PM, Chau, Ken) Cleaned plungers Cleaned Latch Calibrated system -still doesnt always latch with fluid attached, without fluid no issue -may require a motor replacement or piston tips (TC 08/28/2023 07:07:32 PM, Chau, Ken) Replaced all plunger hooks and tested both A and B side without issues -tested with fluid and without -tested with techs and no issue -informed if issue comes back to report issue -leaving WO open for another week
39506 | 196918 | CM | CMPLT | CMPLT | 2023-08-23 15:06 | 2023-08-28 19:07 | MSH Spare Injector - Error on startup | (TC 08/23/2023 03:06:30 PM, Chau, Ken) System generating error 3026 -pointing towards 48V power supply -tested voltage on base and boards all generating 48V and 15V -replaced base as it occured several times intermittently -system has been on and no issues for over 4+ hours (TC 08/28/2023 07:06:44 PM, Chau, Ken) Tested spare injector after several days when replaced base PSU and no issue -setup in room and no issue | Error 3026 - Pointing towardss 48V PSU
68784 | 196919 | FC | CMPLT | CMPLT | 2023-08-23 15:41 | 2023-08-24 08:37 | PCU sent down to biomed | (TC 08/23/2023 03:41:32 PM, Cortez, John) No errors found in log PM competed successfully Conditioning battery overnight (TC 08/23/2023 08:35:56 AM, Cortez, John) 4107 mAhs
1067437 | 196920 | FC | CMPLT | CMPLT | 2023-08-24 08:11 | 2023-08-24 08:11 | QC checklist | (TC 08/24/2023 08:11:05 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1059553 | 196921 | CM | CMPLT | CMPLT | 2023-08-24 09:56 | 2023-08-24 09:56 | Windows update | (TC 08/24/2023 09:55:43 AM, Baghirzade, Nara) Update completed for central Zone C
1059552 | 196922 | CM | CMPLT | CMPLT | 2023-08-24 09:57 | 2023-08-24 09:57 | Windows update | (TC 08/24/2023 09:56:40 AM, Baghirzade, Nara) Windows update completed for central Zone B
1059543 | 196923 | CM | CMPLT | CMPLT | 2023-08-24 09:57 | 2023-08-24 09:58 | Windows update | (TC 08/24/2023 09:57:18 AM, Baghirzade, Nara) Update completed for central Zone A.
| 196924 | CM | CMPLT | CMPLT | 2023-08-24 11:01 | 2023-08-24 11:01 | PFT/CPET meeting | (TC 08/24/2023 11:00:52 AM, Baghirzade, Nara) PFT/CPET meeting
1061214 | 196925 | CM | CMPLT | CMPLT | 2023-08-24 11:33 | 2023-08-25 08:25 | Vitals DVI connector fell into Boom | (TC 08/24/2023 11:33:01 AM, Reid, Emmerson ) DVI connector fell into the boom when I removed the cable. Need to fix this between cases. (TC 08/25/2023 08:24:53 AM, Reid, Emmerson ) Fixed
1061216 | 196926 | CM | CMPLT | CMPLT | 2023-08-24 11:34 | 2023-08-24 11:35 | Vitals not showing on monitor | (TC 08/24/2023 11:34:08 AM, Reid, Emmerson ) Reattached cable to boom. Working now
1061215 | 196927 | CM | ACTIV | ISSUE | 2023-08-24 11:36 |  | Monitors shutting down | (TC 08/24/2023 11:35:41 AM, Reid, Emmerson ) Monitors shut down multiple times when the FT10 was used. I tried moving the cables apart but this did not help. I moved the ESU plug into another outlet. I also noticed the foot peddal had some tears in the insulation and needs to be replaced. (TC 09/07/2023 11:02:53 AM, Reid, Emmerson ) Issue was happenning again. Power cables were moved back to the boom, so I moved them to the wall outlet and it fixed the issue. Spoke with olympus to see if they can investigate. (TC 09/18/2023 11:06:08 AM, Reid, Emmerson ) Met with Derreck. Jessica, Adeel, and another Olympus tech. We could not recreate the issue. We switched the camera processor with another one and will see if this helps.
1067567 | 196928 | CM | CMPLT | CMPLT | 2023-08-24 16:34 | 2023-08-24 16:34 | Initial Accept | (TC 08/24/2023 04:34:22 PM, Samanian, Kamran) performed safet test and passed
1067568 | 196929 | CM | CMPLT | CMPLT | 2023-08-24 16:35 | 2023-08-24 16:35 | Initial Accept | (TC 08/24/2023 04:35:11 PM, Samanian, Kamran) performed safet test and passed
1067569 | 196930 | CM | CMPLT | CMPLT | 2023-08-24 16:36 | 2023-08-24 16:36 | Initial Accept | (TC 08/24/2023 04:35:47 PM, Samanian, Kamran) performed safet test and passed
1067570 | 196931 | CM | CMPLT | CMPLT | 2023-08-24 16:36 | 2023-08-24 16:36 | Initial Accept | (TC 08/24/2023 04:36:39 PM, Samanian, Kamran) performed safet test and passed
68552 | 196932 | CM | CMPLT | CMPLT | 2023-08-24 17:18 | 2023-08-24 22:08 | Flow sensor failure followed by O2 sensor failure, not reading | (TC 08/24/2023 05:19:14 PM, Rendon Lopez , Elizabeth) Device was pulled from OR 1, failure during surgery. Testing, flow sensor failure alarm, sensors wet, replaced both flow sensors, ran self tests-OK. Downloaded logs, several PATO sensor failures during time of complaint, but unable to replicate issue. Tested O2 readings, takes a long time to change and reach set value. ordering PATO sensor.
68377 | 196933 | CM | CMPLT | CMPLT | 2023-08-24 17:23 | 2023-08-24 22:10 | Device fails self test | (TC 08/24/2023 05:25:48 PM, Rendon Lopez , Elizabeth) large leak in breathing block, fails self test, pressure sensor failure. Ordered pressure sensor
1059828 | 196934 | CM | CMPLT | NFF | 2023-08-25 08:29 | 2023-08-25 08:32 | Patient probe 1 not lighting up | (TC 08/25/2023 08:29:18 AM, Baghirzade, Nara) Plugged in probe to port 1; lit up green. Plugged in probe to port 2; also functional. Went in to service mode, verified that ports are functional. There was a computer cable that was being used, which is probably what was being used to plug in. Removed this cable from the Altrix entirely. Could not find fault with the device. Patient cable is fully functional and in good condition. Returned to Zone A NICU
| 196935 | OC | CMPLT | CMPLT | 2023-08-25 09:15 | 2023-08-25 09:22 | Elizabeth Rendon Lopez On call Aug 18- Aug 24, 2023 | (TC 08/25/2023 03:15:00 AM, Rendon Lopez , Elizabeth) LIM monitor in main OR 1 alarming. Came to the hospital, faulty power bar in room, removed and replaced with new one.
62672 | 196936 | CM | CMPLT | CMPLT | 2023-08-25 10:18 | 2023-08-25 11:06 | Not working | (TC 08/24/2023 10:18:12 AM, Canjimavadivel, Kamallesh ) Multiple otoscope/ophthalmoscope were reported not working. Went up to take a look. Was able to repair 2 devices, brought back 2 other devices to biomed to further troubleshoot. Both fuses were blown. The other device had damaged handles. Swapped handles from this device with the other damaged handles. (TC 08/25/2023 10:28:02 AM, Canjimavadivel, Kamallesh ) Spoke to CNS to discuss next steps. ED has a few parts that might be able to use. Need heads for both handles. Also mentioned parts are not available as this model is discontinued. They are okay with 3 working devices until they move into new space. (TC 08/25/2023 11:04:08 AM, Canjimavadivel, Kamallesh ) Repaired the 3rd device using heads found in ED storage room. Installed on wall in AC and informed unit they have 3 working devices. This 4th unit to be retired for parts.
67965 | 196937 | CM | CMPLT | CMPLT | 2023-08-25 10:20 | 2023-09-12 12:18 | Wrong door placed on the incubator | (TC 08/25/2023 10:20:28 AM, Baghirzade, Nara) The opposite door was placed on the incubator, which means another incubator has the opposite door as well. Tried to find the original door; checked RESUS, RESUS storage, L&D ORs and hallway, NICU storage, and NICU rooms. Could not locate the correct door as well as the second incubator with the wrong door. Informed the nurse at the NICU central to let the staff know as well as inform me if they find the omnibeds with the wrong doors. (TC 08/25/2023 10:32:31 AM, Baghirzade, Nara) Second door had missing recall latch and broken latch on south end. Replaced all latches. Replaced missing grommets. (TC 09/12/2023 12:16:52 PM, Baghirzade, Nara) Connor (coop student) swapped the latches around and used a new NE latch for the door under my supervision.
| 196938 | CM | CMPLT | CMPLT | 2023-08-25 11:09 | 2023-08-25 11:09 | iBUS-Cerner weekly meeting | (TC 08/25/2023 11:08:49 AM, Samanian, Kamran) discussed progress on FetaLink, VSM, BD, IBE
25379 | 196939 | CM | CMPLT | CMPLT | 2023-08-25 11:10 | 2023-08-25 11:22 | Incident Report: SAFER# 62467 | (TC 08/25/2023 11:21:22 AM, Reid, Emmerson ) Investigated and completed report. Report is attached to WO and saved on the T-drive. Unit will be retired.
| 196940 | CM | CMPLT | CMPLT | 2023-08-25 11:37 | 2023-08-25 11:39 | Ophthalmoscope bulb replacements | (TC 08/25/2023 11:37:47 AM, Baghirzade, Nara) Replaced ophthalmoscope bulbs. Met with Sally and Ludmila to discuss the BP installation they want in the procedure room and they requested a quote for an adult scale for their triage area.
| 196941 | CM | CMPLT | CMPLT | 2023-08-25 12:00 | 2023-08-25 12:00 | Masimo UTN SN# 27120200500109 | (TC 08/25/2023 11:59:57 AM, Samanian, Kamran) done! PO# 7202259039
69178 | 196942 | PM | CMPLT | CMPLT | 2023-08-25 12:05 | 2023-08-25 12:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/25/2023 12:05:23 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 196943 | PA | CMPLT | CMPLT | 2023-08-25 13:31 | 2023-08-29 08:41 | Ordering sterile water for altrix cleaning - NICU | (TC 08/25/2023 01:35:42 PM, Baghirzade, Nara) SC 1003035313 (TC 08/29/2023 08:39:40 AM, Baghirzade, Nara) Received water Order # 415261472
1061210 | 196944 | CM | CMPLT | CMPLT | 2023-08-25 13:31 | 2023-08-25 13:35 | Error message on the scope monitor blocking screen | (TC 08/25/2023 01:31:58 PM, Reid, Emmerson ) There was a message on all the monitors that a printer was not connected. They wanted the message gone, so i waited for a good time incase i interupted the monitors. I pressed the ESC key on the keyboard and it cleared the message. Working now.
1064743 | 196945 | CM | CMPLT | CMPLT | 2023-08-25 14:56 | 2023-08-25 14:56 | new incoming inspection | (TC 06/06/2023 02:55:59 PM, Huang, Jessica) assembled the device, new incoming inspection
1064928 | 196946 | CM | CMPLT | CMPLT | 2023-08-25 14:59 | 2023-08-25 14:59 | new incoming inspection | (TC 06/06/2023 02:59:01 PM, Huang, Jessica) assembled the device, new incoming inspection
1064929 | 196947 | CM | CMPLT | CMPLT | 2023-08-25 15:00 | 2023-08-25 15:00 | new incoming inspection | (TC 06/06/2023 02:59:58 PM, Huang, Jessica) assembled the device, new incoming inspection
1066964 | 196948 | CM | CMPLT | CMPLT | 2023-08-25 15:01 | 2023-08-25 15:01 | assembled the device, new incoming inspection | (TC 06/06/2023 03:00:48 PM, Huang, Jessica) assembled the device, new incoming inspection
1066965 | 196949 | CM | CMPLT | CMPLT | 2023-08-25 15:02 | 2023-08-25 15:02 | assembled the device, new incoming inspection | (TC 06/06/2023 03:02:19 PM, Huang, Jessica) assembled the device, new incoming inspection
1066966 | 196950 | CM | CMPLT | CMPLT | 2023-08-25 15:03 | 2023-08-25 15:03 | assembled the device, new incoming inspection | (TC 06/06/2023 03:03:12 PM, Huang, Jessica) assembled the device, new incoming inspection
1066967 | 196951 | CM | CMPLT | CMPLT | 2023-08-25 15:04 | 2023-08-25 15:04 | assembled the device, new incoming inspection | (TC 06/06/2023 03:04:10 PM, Huang, Jessica) assembled the device, new incoming inspection
1066968 | 196952 | CM | CMPLT | CMPLT | 2023-08-25 15:05 | 2023-08-25 15:05 | assembled the device, new incoming inspection | (TC 06/06/2023 03:04:55 PM, Huang, Jessica) assembled the device, new incoming inspection
1066969 | 196953 | CM | CMPLT | CMPLT | 2023-08-25 15:10 | 2023-08-25 15:10 | assembled the device, new incoming inspection | (TC 06/06/2023 03:10:33 PM, Huang, Jessica) assembled the device, new incoming inspection
1066970 | 196954 | CM | CMPLT | CMPLT | 2023-08-25 15:13 | 2023-08-25 15:13 | assembled the device, new incoming inspection | (TC 06/06/2023 03:13:11 PM, Huang, Jessica) assembled the device, new incoming inspection
1066971 | 196955 | CM | CMPLT | CMPLT | 2023-08-25 15:15 | 2023-08-25 15:15 | assembled the device, new incoming inspection | (TC 06/13/2023 03:15:38 PM, Huang, Jessica) assembled the device, new incoming inspection
1066972 | 196956 | CM | CMPLT | CMPLT | 2023-08-25 15:16 | 2023-08-25 15:16 | assembled the device, new incoming inspection | (TC 06/13/2023 03:16:24 PM, Huang, Jessica) assembled the device, new incoming inspection
66464 | 196957 | CM | CMPLT | CMPLT | 2023-08-25 15:17 | 2023-10-11 08:29 | Ordering parts for VN500 | Received parts. Charge to RT . PO#7202272601
1066973 | 196958 | CM | CMPLT | CMPLT | 2023-08-25 15:19 | 2023-08-25 15:19 | assembled the device, new incoming inspection | (TC 06/13/2023 03:18:46 PM, Huang, Jessica) assembled the device, new incoming inspection
1066974 | 196959 | CM | CMPLT | CMPLT | 2023-08-25 15:19 | 2023-08-25 15:19 | assembled the device, new incoming inspection | (TC 06/13/2023 03:19:28 PM, Huang, Jessica) assembled the device, new incoming inspection
1066975 | 196960 | CM | CMPLT | CMPLT | 2023-08-25 15:20 | 2023-08-25 15:20 | assembled the device, new incoming inspection | (TC 06/13/2023 03:20:20 PM, Huang, Jessica) assembled the device, new incoming inspection
1066976 | 196961 | CM | CMPLT | CMPLT | 2023-08-25 15:22 | 2023-08-25 15:22 | assembled the device, new incoming inspection | (TC 06/13/2023 03:21:57 PM, Huang, Jessica) assembled the device, new incoming inspection
1066977 | 196962 | CM | CMPLT | CMPLT | 2023-08-25 15:23 | 2023-08-25 15:23 | assembled the device, new incoming inspection | (TC 06/13/2023 03:22:50 PM, Huang, Jessica) assembled the device, new incoming inspection
1066978 | 196963 | CM | CMPLT | CMPLT | 2023-08-25 15:24 | 2023-08-25 15:24 | assembled the device, new incoming inspection | (TC 06/13/2023 03:23:50 PM, Huang, Jessica) assembled the device, new incoming inspection
1066979 | 196964 | CM | CMPLT | CMPLT | 2023-08-25 15:25 | 2023-08-25 15:25 | assembled the device, new incoming inspection | (TC 06/13/2023 03:24:44 PM, Huang, Jessica) assembled the device, new incoming inspection
1066980 | 196965 | CM | CMPLT | CMPLT | 2023-08-25 15:25 | 2023-08-25 15:25 | assembled the device, new incoming inspection | (TC 06/13/2023 03:25:26 PM, Huang, Jessica) assembled the device, new incoming inspection
1066981 | 196966 | CM | CMPLT | CMPLT | 2023-08-25 15:27 | 2023-08-25 15:27 | assembled the device, new incoming inspection
1066982 | 196967 | CM | CMPLT | CMPLT | 2023-08-25 15:28 | 2023-08-25 15:28 | assembled the device, new incoming inspection
1066983 | 196968 | CM | CMPLT | CMPLT | 2023-08-25 15:29 | 2023-08-25 15:29 | assembled the device, new incoming inspection | (TC 06/20/2023 03:29:26 PM, Huang, Jessica) assembled the device, new incoming inspection
1066984 | 196969 | CM | CMPLT | CMPLT | 2023-08-25 15:30 | 2023-08-25 15:30 | assembled the device, new incoming inspection | (TC 06/20/2023 03:30:20 PM, Huang, Jessica) assembled the device, new incoming inspection
1066985 | 196970 | CM | CMPLT | CMPLT | 2023-08-25 15:31 | 2023-08-25 15:31 | assembled the device, new incoming inspection | (TC 06/20/2023 03:31:26 PM, Huang, Jessica) assembled the device, new incoming inspection
1066986 | 196971 | CM | CMPLT | CMPLT | 2023-08-25 15:32 | 2023-08-25 15:32 | assembled the device, new incoming inspection | (TC 06/20/2023 03:32:40 PM, Huang, Jessica) assembled the device, new incoming inspection
1067157 | 196972 | CM | CMPLT | CMPLT | 2023-08-25 15:33 | 2023-08-25 15:33 | assembled the device, new incoming inspection | (TC 06/20/2023 03:33:39 PM, Huang, Jessica) assembled the device, new incoming inspection
1067158 | 196973 | CM | CMPLT | CMPLT | 2023-08-25 15:34 | 2023-08-25 15:34 | assembled the device, new incoming inspection | (TC 06/20/2023 03:34:29 PM, Huang, Jessica) assembled the device, new incoming inspection
1067160 | 196974 | CM | CMPLT | CMPLT | 2023-08-25 15:35 | 2023-08-25 15:35 | assembled the device, new incoming inspection | (TC 06/20/2023 03:35:23 PM, Huang, Jessica) assembled the device, new incoming inspection
1067261 | 196975 | CM | CMPLT | CMPLT | 2023-08-25 15:36 | 2023-08-25 15:36 | assembled the device, new incoming inspection | (TC 06/20/2023 03:36:29 PM, Huang, Jessica) assembled the device, new incoming inspection
1067262 | 196976 | CM | CMPLT | CMPLT | 2023-08-25 15:37 | 2023-08-25 15:37 | assembled the device, new incoming inspection | (TC 06/27/2023 03:37:27 PM, Huang, Jessica) assembled the device, new incoming inspection
1067480 | 196977 | CM | CMPLT | CMPLT | 2023-08-25 15:38 | 2023-08-25 15:38 | assembled the device, new incoming inspection | (TC 06/27/2023 03:38:15 PM, Huang, Jessica) assembled the device, new incoming inspection
1067481 | 196978 | CM | CMPLT | CMPLT | 2023-08-25 15:40 | 2023-08-25 15:40 | assembled the device, new incoming inspection | (TC 06/27/2023 03:39:54 PM, Huang, Jessica) assembled the device, new incoming inspection
1067482 | 196979 | CM | CMPLT | CMPLT | 2023-08-25 15:41 | 2023-08-25 15:41 | assembled the device, new incoming inspection | (TC 08/25/2023 03:40:40 PM, Huang, Jessica) assembled the device, new incoming inspection
| 196980 | CM | CMPLT | CMPLT | 2023-08-25 15:42 | 2023-08-25 15:42 | sent masimo monitors from HBH to MSH
| 196981 | CM | CMPLT | CMPLT | 2023-08-25 15:45 | 2023-08-25 15:45 | screen protector | (TC 07/10/2023 03:44:56 PM, Huang, Jessica) took all the protectors out of the box and stored them in OR 12 (TC 07/17/2023 03:46:39 PM, Huang, Jessica) installed the protectors in the OR rooms
| 196982 | CM | ACTIV | ISSUE | 2023-08-25 15:49 |  | department meeting
| 196983 | CM | ACTIV | ISSUE | 2023-08-25 16:05 |  | on call daily checking
| 196984 | CM | CMPLT | CMPLT | 2023-08-25 16:11 | 2023-08-25 16:11 | OR 13 ipad checking
1061215 | 196985 | CM | CMPLT | CMPLT | 2023-08-25 16:15 | 2023-08-25 16:15 | ipad was off | (TC 08/21/2023 04:14:49 PM, Huang, Jessica) disconnected the power cord and reconnected it a few times, it started charging
| 196986 | CM | CMPLT | CMPLT | 2023-08-25 16:18 | 2023-08-25 16:18 | the table isnt working in OR 5 | (TC 08/21/2023 04:18:06 PM, Huang, Jessica) received a call from OR 5 and saying the table wasnt working. went to the room and was told the table worked
| 196987 | CM | CMPLT | CMPLT | 2023-08-25 16:47 | 2023-08-25 16:47 | OR 1 lIM system
1061214 | 196988 | CM | CMPLT | CMPLT | 2023-08-25 16:49 | 2023-08-25 16:49 | vitals didnt display on the wall monitor | (TC 08/25/2023 04:49:15 PM, Huang, Jessica) got approached by Camilla from OR 13 saying vital couldnt display on the wall monitor. went to the room and checked it. working fine
| 196989 | CM | CMPLT | CMPLT | 2023-08-25 17:09 | 2023-08-25 17:09 | OR 1 service request | (TC 08/22/2023 05:10:04 PM, Huang, Jessica) received the service request online regarding Tourniquets in OR 1. checked both machines and working fine
1051886 | 196990 | CM | CMPLT | CMPLT | 2023-08-25 17:14 | 2023-08-25 17:14 | missing screws from the power connector | (TC 08/16/2023 05:14:05 PM, Huang, Jessica) replaced the missing screws and functional checked
120013 | 196991 | CM | CMPLT | CMPLT | 2023-08-25 17:15 | 2023-08-25 17:15 | removed from service
541131 | 196992 | CM | CMPLT | CMPLT | 2023-08-25 17:20 | 2023-08-25 17:20 | not working | (TC 08/15/2023 05:20:55 PM, Huang, Jessica) replaced the SPO2 sensor
1046467 | 196993 | CM | CMPLT | CMPLT | 2023-08-25 17:22 | 2023-08-25 17:22 | not working | (TC 08/25/2023 05:21:59 PM, Huang, Jessica) replaced the SPO2 sensor
1048045 | 196994 | CM | CMPLT | CMPLT | 2023-08-25 17:23 | 2023-08-25 17:23 | not working | (TC 08/15/2023 05:22:51 PM, Huang, Jessica) sent the device to KCI for service repair
| 196995 | CM | CMPLT | CMPLT | 2023-08-25 17:24 | 2023-08-25 17:24 | checking the battery on the ECG cart
1063936 | 196996 | CM | CMPLT | CMPLT | 2023-08-25 17:27 | 2023-08-25 17:27 | PM | (TC 08/14/2023 05:27:10 PM, Huang, Jessica) PM completed
531048 | 196997 | CM | CMPLT | CMPLT | 2023-08-25 17:28 | 2023-08-25 17:28 | lower suction level | (TC 08/11/2023 05:28:25 PM, Huang, Jessica) replaced the valve and functional checked
| 196998 | CM | CMPLT | CMPLT | 2023-08-25 17:38 | 2023-08-25 17:38 | checked suction tree in OR | (TC 08/11/2023 05:37:56 PM, Huang, Jessica) picked up the broken suction device from OR repaired the stand and functional checked
1063791 | 196999 | CM | CMPLT | CMPLT | 2023-08-25 17:40 | 2023-08-25 17:40 | not working | (TC 08/10/2023 05:39:39 PM, Huang, Jessica) sent the headlight to the vendor for warranty repair
531093 | 197000 | CM | CMPLT | CMPLT | 2023-08-25 17:44 | 2023-08-25 17:44 | lower sunction level | (TC 08/09/2023 05:43:55 PM, Huang, Jessica) replaced the valve, functional checked
| 197001 | CM | CMPLT | CMPLT | 2023-08-25 17:48 | 2023-08-25 17:48 | active the port to connect the medical device in OR 6 | (TC 08/09/2023 05:48:29 PM, Huang, Jessica) robotic needed to be moved from OR 11 to OR 6 contacted IT / Helpdesk to active the port
1051897 | 197002 | CM | CMPLT | CMPLT | 2023-08-25 17:51 | 2023-08-25 17:51 | lower suction level | (TC 08/09/2023 05:50:55 PM, Huang, Jessica) replaced the valve and functional checked
541203 | 197003 | CM | CMPLT | CMPLT | 2023-08-25 17:53 | 2023-08-25 17:53 | removed from service
541042 | 197004 | CM | CMPLT | CMPLT | 2023-08-25 17:55 | 2023-08-25 17:55 | not working | (TC 08/08/2023 05:54:57 PM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
541029 | 197005 | CM | CMPLT | CMPLT | 2023-08-25 17:56 | 2023-08-25 17:56 | not working | (TC 08/08/2023 05:56:00 PM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
1056475 | 197006 | CM | CMPLT | CMPLT | 2023-08-28 08:04 | 2023-08-28 08:04 | BP needs calibration | Calibrated BP. tested. Functional.
68069 | 197007 | CM | CMPLT | CMPLT | 2023-08-28 08:05 | 2023-08-28 08:05 | weight is not accurate | installed missing foot. Fixed loose column. Calibrated. Tested. Functional.
| 197008 | RDS | CMPLT | CMPLT | 2023-08-28 08:57 | 2023-08-28 09:01 | Weekly Checks For PACU | (TC 08/28/2023 09:00:29 AM, Reid, Emmerson ) Checks complete
| 197009 | CM | CMPLT | CMPLT | 2023-08-28 09:52 | 2023-09-05 13:17 | Ordering distilled water for biomed stock | (TC 08/28/2023 10:09:30 AM, Cortez, John) Ordered for biomed stock
1057546 | 197010 | CM | CMPLT | CMPLT | 2023-08-28 09:52 | 2023-08-28 10:02 | missing face plate | (TC 08/28/2023 09:52:37 AM, Scero, Tomlyn) Replaced missing bezel (faceplate). no other indications of what it was sent down for. Test BP, SPO2, ECG, temp, and IBP. All functions passed.
| 197011 | RDS | CMPLT | CMPLT | 2023-08-28 10:03 | 2023-08-28 11:03 | Weekly checks in CCU | (TC 08/28/2023 11:02:08 AM, Scero, Tomlyn) weekly check complete. spoke with dianne regarding ECG trucks and how to order replacements. Telemetery questions came up as well that I had to answer.
63809 | 197012 | CM | CMPLT | NFF | 2023-08-28 10:34 | 2023-08-28 10:35 | BP not inflating passed 70 | (TC 08/28/2023 10:33:58 AM, Canjimavadivel, Kamallesh ) Checked for leaks in cuff, did not find any. BP function working as intended, was not able to replicate issue reported.
1059938 | 197013 | CM | CMPLT | CMPLT | 2023-08-28 10:50 | 2023-08-28 11:28 | Not charging | (TC 08/28/2023 11:14:02 AM, Canjimavadivel, Kamallesh ) Power cord was damaged. Replaced cord with spare cord and tested charging. Keyboard cover was also ripped. Removed and replaced with new one.
| 197014 | CM | ACTIV | ISSUE | 2023-08-28 11:26 |  | Dual Guage Regulator order for APC test
1058584 | 197015 | CM | CMPLT | CMPLT | 2023-08-28 12:17 | 2023-08-28 12:20 | Not working | (TC 08/28/2023 12:17:08 PM, Reid, Emmerson ) Checked on the calibrator and passed first try. Working now.
1065273 | 197016 | CM | CMPLT | CMPLT | 2023-08-28 12:52 | 2023-08-28 12:54 | Wand cable connector head came apart | (TC 08/28/2023 12:52:29 PM, Reid, Emmerson ) Wand cable connector head came apart. I was able to put it back together and it was working. The clinician wanted to have the vendor validate it.
1067333 | 197017 | CM | CMPLT | CMPLT | 2023-08-28 14:12 | 2023-08-28 14:12 | demo device
120021 | 197018 | CM | CMPLT | CMPLT | 2023-08-28 14:13 | 2023-08-28 14:13 | not working | (TC 07/04/2023 02:15:32 PM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
120020 | 197019 | CM | CMPLT | CMPLT | 2023-08-28 14:17 | 2023-08-28 14:17 | not working | (TC 07/04/2023 02:16:46 PM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
| 197020 | CM | CMPLT | CMPLT | 2023-08-28 14:18 | 2023-08-28 14:18 | assisting Chris from Medtronic to upgrade software on FT 10
1048272 | 197021 | CM | CMPLT | CMPLT | 2023-08-28 14:19 | 2023-08-28 14:19 | not working | (TC 07/05/2023 02:19:25 PM, Huang, Jessica) fixed the stand, replaced the temperature probe
1054071 | 197022 | CM | CMPLT | CMPLT | 2023-08-28 14:21 | 2023-08-28 14:21 | not working | (TC 07/05/2023 02:20:54 PM, Huang, Jessica) fixed the stand, replaced the temperature probe
1054073 | 197023 | CM | CMPLT | CMPLT | 2023-08-28 14:22 | 2023-08-28 14:22 | not working | (TC 07/05/2023 02:21:50 PM, Huang, Jessica) fixed the stand, replaced the temperature probe
531022 | 197024 | CM | CMPLT | CMPLT | 2023-08-28 14:23 | 2023-08-28 14:23 | not working | (TC 07/10/2023 02:23:13 PM, Huang, Jessica) troubleshooting the device, replaced the motor. working fine so sent it back to the unit. The device was sent it again with the same issue. replaced the valve, tested it, working fine
1046035 | 197025 | CM | CMPLT | CMPLT | 2023-08-28 15:33 | 2023-09-18 08:31 | MSH ER DRX EVO E5-02(ER) Image artifact | (TC 08/28/2023 10:16:00 AM, Fan, Tak) Replicated artifact on table. Dent in the center causing a tear-drop shaped artifact. Contacted Carestream for quote. (TC 08/29/2023 09:44:00 AM, Fan, Tak) Tested images with Eimear. Artifact is actually caused by the dented surface dispersing the x-rays causing apparent radiopaqueness. Smoothing out the dent causes the dent to look more radiotransparent, much less non-uniformity. Eimear will check with rads on image quality acceptability. | MSH ER DRX EVO E5-02(ER) Image artifact Site: Customer #: 936886 Address: Service Zone: 1 Contact: Phone: email: Phone / email Material No Quantity Description Unit Price Unit Discount Line Total SPAC8671 1 TABLE TOP WHITE $14,322.64 $14,322.64
| 197026 | CM | CMPLT | CMPLT | 2023-08-28 15:59 | 2023-08-28 15:58 | weekly Huddle | (TC 08/28/2023 03:58:44 PM, Samanian, Kamran) done!
1049227 | 197027 | CM | ACTIV | ISSUE | 2023-08-28 16:20 |  | Changing the server role | (TC 08/17/2023 04:20:32 PM, Samanian, Kamran) reloaded Rev C from scratch and rebuild the server (TC 08/28/2023 04:21:09 PM, Samanian, Kamran) reconfigure the server as Warm Standby (TC 08/28/2023 04:21:34 PM, Samanian, Kamran) tried to patch the Physio 2 and crashed , so Warm Standby Server renamed and IP reconfigured to act as Physio 2
32560 | 197028 | CM | CMPLT | CMPLT | 2023-08-28 19:54 | 2023-08-28 19:52 | CT scan /display not booting/R sys errors | CT scan /display not booting/R sys errors Service completed troubleshooting Service completed Raw data/image drive calibrations CT phantom tested OK Service tested scan/dispaly computer for possible faults/tested OK
1059778 | 197029 | FC | CMPLT | CMPLT | 2023-08-29 07:56 | 2023-08-29 07:57 | QC checklist | (TC 08/29/2023 07:56:32 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1066462 | 197030 | FC | CMPLT | CMPLT | 2023-08-29 07:57 | 2023-08-29 07:57 | QC checklist | (TC 08/29/2023 07:57:05 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
| 197031 | OC | ACTIV | ISSUE | 2023-08-29 09:04 |  | Samir Chukri On-Call | (TC 08/28/2023 11:15:00 PM, Chukri, Samir) Got a call regarding the sample tube not working. explain that biomed not looking after this issue and they should call Building Service (TC 12/15/2023 10:03:01 PM, Chukri, Samir) Received call From L&D O.R #3 surgical table not working, checked found that Biomed not looking after this unit , Medaris medical comany lookingter this uint called back Halina Zenew (charge nurse) to explain to her and she required me to call tomorrow to give them the comany name and tel. (TC 12/19/2023 09:27:46 AM, Chukri, Samir) Received a call from Triage regarding three fetal monitor not connecting to Cerner. I advise her to contact IT team. (TC 02/28/2024 11:03:22 PM, Chukri, Samir) Received call around 10:07 from 14S that Masimo view not working . remotly access to Masimo webside and found computer was turned off . re-called 14S department nurse and guide it to turn on their computer. promblem solved . (TC 03/05/2024 10:00:00 PM, Chukri, Samir) Received call around 10:00 pm from O.R room # 9 three display monitors not working . checked by phone with Kamran, and returnthe call back to O.R nurse, they said the problem solved and no need more help. informed Emmerson next day to check O.R #9 .
65982 | 197032 | CM | CMPLT | CMPLT | 2023-08-29 09:31 | 2023-08-29 09:56 | Channel error 352.6770 | (TC 08/29/2023 09:54:49 AM, Baghirzade, Nara) Rear case and left handle were broken; replaced rear case and left handle. Re-seated the pressure sensor as it was only replaced the year prior. Re-seated the pressure sensor. Calibrated the pressure sensor. Check-in tests passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1054421 | 197033 | CM | CMPLT | CMPLT | 2023-08-29 09:46 | 2023-08-29 09:46 | not working | (TC 08/22/2023 09:46:20 AM, Huang, Jessica) replaced the SPO2 sensor
1046531 | 197034 | CM | CMPLT | CMPLT | 2023-08-29 09:47 | 2023-08-29 09:47 | not working | (TC 08/22/2023 09:47:01 AM, Huang, Jessica) replaced the SPO2 sensor
541265 | 197035 | CM | CMPLT | CMPLT | 2023-08-29 09:48 | 2023-08-29 09:48 | not working | (TC 08/22/2023 09:47:38 AM, Huang, Jessica) problem not found
541274 | 197036 | CM | CMPLT | CMPLT | 2023-08-29 09:48 | 2023-08-29 09:48 | not working | (TC 08/22/2023 09:50:28 AM, Huang, Jessica) replaced the SPO2 senor and the temperature probe
1061337 | 197037 | CM | CMPLT | CMPLT | 2023-08-29 09:51 | 2023-08-29 09:51 | not working | (TC 08/22/2023 09:51:13 AM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
541029 | 197038 | CM | CMPLT | CMPLT | 2023-08-29 09:52 | 2023-08-29 09:52 | not working | (TC 08/22/2023 09:52:04 AM, Huang, Jessica) problem not found
1059828 | 197039 | CM | CMPLT | CMPLT | 2023-08-29 09:53 | 2023-08-29 09:53 | Cleaning | (TC 08/29/2023 09:53:16 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
120016 | 197040 | CM | CMPLT | CMPLT | 2023-08-29 09:54 | 2023-08-29 09:54 | not working | (TC 08/22/2023 09:53:38 AM, Huang, Jessica) replaced the SPO2 sensor
541037 | 197041 | CM | CMPLT | CMPLT | 2023-08-29 09:54 | 2023-08-29 09:54 | not working | (TC 08/22/2023 09:54:23 AM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
541258 | 197042 | CM | CMPLT | CMPLT | 2023-08-29 09:55 | 2023-08-29 09:55 | bad battery | (TC 08/22/2023 09:55:26 AM, Huang, Jessica) replaced the battery
1067589 | 197043 | ACC | CMPLT | CMPLT | 2023-08-29 11:15 | 2023-08-29 11:20 | Incoming inspection and asset tagging | (TC 08/29/2023 11:19:26 AM, Baghirzade, Nara) Incoming inspection completed. Asset tagged. Unit is fully functional. Electrical safety results - Current leakage: 47.3 uA Power cord resistance: 82 mohms
68377 | 197044 | CM | CMPLT | CMPLT | 2023-08-29 11:28 | 2023-08-30 11:58 | Bag pressure sensor failure | (TC 08/29/2023 11:31:49 AM, Scero, Tomlyn) Called to OR 12 for broken device. Self test fails with pressure sensor failure. Downloaded DSR to confirm which pressure sensor is failing. DSR says its the Bag pressure sensor. Bag pressure sensor test is passing in the IPM-L, but error is still occuring. Called Draeger to get a quote for a pressure sensor and ordered a new one on ereq. (TC 08/29/2023 10:27:04 AM, Scero, Tomlyn) Found a spare pressure sensor in stock. Replaced bag pressure sensor and passed self tests. Continue with more tests to confirm funtional tomorrow. (TC 08/30/2023 10:28:16 AM, Scero, Tomlyn) Tests as per IPM-L complete and passed. Put unit back together and self test complete.
1067591 | 197045 | ACC | CMPLT | CMPLT | 2023-08-29 11:50 | 2023-08-29 11:54 | Incoming inspection and asset tagging | (TC 08/29/2023 11:52:49 AM, Baghirzade, Nara) Incoming inspection completed. Asset tagged. Unit is fully functional. Electrical safety results - Leakage current: 47.2 uA Power cord resistance: 98 mohms
68755 | 197046 | CM | CMPLT | CMPLT | 2023-08-29 11:59 | 2023-09-15 09:01 | System error 120.4610 | (TC 08/29/2023 11:59:23 AM, Canjimavadivel, Kamallesh ) Downloaded error logs. Error code associated to low battery level. Started battery conditioning test. (TC 09/01/2023 10:12:39 AM, Canjimavadivel, Kamallesh ) Battery conditioning test aborted multiple times as it did not complete. Tried swapping battery and power cord, same error displaying. Manual suggested replacing power supply board. Open PC and found connector in back of power supply board not connected. Also found broken connector on logic board. Placed order for logic board. (TC 09/14/2023 12:59:03 PM, Canjimavadivel, Kamallesh ) Replaced logic board and flashed PC (TC 09/15/2023 08:58:03 AM, Canjimavadivel, Kamallesh ) Performed check-in tests, all pass Battery conditioning complete, new capacity 3968 mamphr.
1050256 | 197047 | CM | CMPLT | CMPLT | 2023-08-29 12:01 | 2023-08-30 08:18 | Reported does not turn on | (TC 08/29/2023 03:19:22 PM, Canjimavadivel, Kamallesh ) Checked the logs. Cleaned pump and performed check in tests. Discovered some keys on keypad do not work. Need to replace keypad. (TC 08/30/2023 08:17:24 AM, Canjimavadivel, Kamallesh ) Replaced keypad Check-in tests pass
1047913 | 197048 | CM | CMPLT | CMPLT | 2023-08-29 12:48 | 2023-08-29 13:00 | Water needs to be renewed | (TC 08/28/2023 12:48:10 PM, Cortez, John) Scheduled with Gabby (Dovigi manager) a time to come up and change water (TC 08/29/2023 12:55:09 PM, Cortez, John) Drained and refilled circuit with distilled water Rooms was busy at one point so had to take multiple trips Will need to send one 4L distilled water jug to Dovigi to replace the one I borrowed
| 197049 | CM | CMPLT | CMPLT | 2023-08-29 13:15 | 2023-09-14 07:39 | 3D print tryout for CNS camera | (TC 08/29/2023 01:14:34 PM, Baghirzade, Nara) Emmerson printed a stand/mount for the CNS camera so they can properly mount it on the IV pole - 3D print did not work for the camera. Emmerson to redesign and print a new stand/mount for the camera. Followed up with Carol regarding this as well as to get a sense of how many studies they have coming up so that Emmerson has the opportunity to try out his prints as well. (TC 09/12/2023 12:02:06 PM, Baghirzade, Nara) 3D print did not work out as the camera needs to be able to rotate. Qwynn instead mounted the camera and its mount on to a pole clamp. This worked out better but needs to be approved by Carol or her team. Coordinated to meet with Carol to show her the camera setup and if it works for them.
1064443 | 197050 | CM | CMPLT | CMPLT | 2023-08-29 13:26 | 2023-08-29 13:26 | not working | (TC 07/13/2023 01:25:45 PM, Huang, Jessica) cleaned the sensor and passed calibration
1061974 | 197051 | CM | CMPLT | CMPLT | 2023-08-29 13:28 | 2023-08-29 13:28 | not working | (TC 06/13/2023 01:28:08 PM, Huang, Jessica) cleaned the sensor and passed calibration
1063117 | 197052 | CM | CMPLT | CMPLT | 2023-08-29 13:34 | 2023-08-29 13:34 | not working | (TC 06/13/2023 01:35:15 PM, Huang, Jessica) cleaned the sensor and passed calibration
1063746 | 197053 | CM | CMPLT | CMPLT | 2023-08-29 13:34 | 2023-08-29 13:34 | Cleaning | (TC 08/29/2023 01:34:17 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
500846 | 197054 | CM | CMPLT | CMPLT | 2023-08-29 14:14 | 2023-08-31 10:47 | MSH 5 US RM 3 SYSTEM FREEZES, REQUIRES BREAKER ON/OFF TO RESTART | (TC 08/29/2023 02:13:49 PM, Whittington, Scott) MSH 5 US RM 3 SYSTEM FREEZES, REQUIRES BREAKER ON/OFF TO RESTART (TC 08/29/2023 08:46:00 PM, Whittington, Scott) MSH 5 US RM 3 SYSTEM FREEZES, REQUIRES BREAKER ON/OFF TO RESTART | MSH 5 US RM 3 SYSTEM FREEZES, REQUIRES BREAKER ON/OFF TO RESTART
62696 | 197055 | CM | CMPLT | CMPLT | 2023-08-29 14:14 | 2023-08-29 14:35 | Lead #3 not working | (TC 08/29/2023 02:30:32 PM, Canjimavadivel, Kamallesh ) RL lead was damaged. Replaced lead and tested with simulator. Tested ok
1046035 | 197056 | CM | CMPLT | CMPLT | 2023-08-29 14:49 | 2023-08-29 14:50 | Bottom handle two button OTC movement button not responsive. | (TC 08/29/2023 02:49:39 PM, Fan, Tak) Replaced two button OTC PCB.
502366 | 197057 | CM | CMPLT | NA | 2023-08-29 14:54 | 2023-09-11 11:01 | XY Brakes not working and ECG freezes every 10-20 seconds | (TC 08/29/2023 02:53:53 PM, Fan, Tak) XY brake assembly not working. Checked button, OK. Ran test, doesn't do anything. Checked voltages on motor control module and up to brakes. All the brakes are open somehow. (TC 08/30/2023 12:21:31 PM, Fan, Tak) Confirmed issue. Only ECG freezing. Reseated all the physio and front panel cables. Problem still persists. Ordered Physio module. (TC 08/31/2023 02:56:04 PM, Fan, Tak) Replaced ECG module but the problem was still persisting. Checked cabling, performed software reload (without erasure). Voltages from PS go through BEP to the physio/front end modules. Checked voltages, OK. Removed modules in case of interference. ECG started working normally after physio module USB cable was removed and plugged back in, but problem comes back every restart. Reloaded software again on spare old HDD, problem resolved. (Other HDD is only half the size (500mb)). Will reload again on original HDD in near future. (TC 09/06/2023 11:43:00 AM, Fan, Tak) Replaced XY assembly with Toni and Deo from GE. Lower panel damaged, ordered replacement. User OK with system being down for now. (TC 09/07/2023 05:00:00 PM, Fan, Tak) Replaced the lower panel. Reassembled system. (TC 09/08/2023 10:00:00 AM, Fan, Tak) Reloaded software for ECG issue on old HDD.
1063119 | 197058 | CM | CMPLT | CMPLT | 2023-08-29 15:02 | 2023-08-29 15:02 | not working | (TC 06/13/2023 03:02:32 PM, Huang, Jessica) cleaned the sensor and passed calibration
1063120 | 197059 | CM | CMPLT | CMPLT | 2023-08-29 15:03 | 2023-08-29 15:03 | not working | (TC 06/13/2023 03:04:33 PM, Huang, Jessica) cleaned the sensor and passed calibration
1053010 | 197060 | CM | CMPLT | CMPLT | 2023-08-29 15:05 | 2023-08-29 15:08 | System crashed | (TC 08/29/2023 03:04:41 PM, Fan, Tak) Crashed during end of study. System restarted OK. Checked logs, nothing pops out. Crashed while processing image "89". Could have been a system memory issue.
1063121 | 197061 | CM | CMPLT | CMPLT | 2023-08-29 15:05 | 2023-08-29 15:05 | not working | (TC 06/13/2023 03:05:25 PM, Huang, Jessica) cleaned the sensor and passed calibration
1063366 | 197062 | CM | CMPLT | CMPLT | 2023-08-29 15:06 | 2023-08-29 15:06 | not working | (TC 06/13/2023 03:06:05 PM, Huang, Jessica) cleaned the sensor and passed calibration
1063367 | 197063 | CM | CMPLT | CMPLT | 2023-08-29 15:14 | 2023-08-29 15:14 | not working | (TC 06/13/2023 03:13:52 PM, Huang, Jessica) cleaned the sensor and passed calibration
1063370 | 197064 | CM | CMPLT | CMPLT | 2023-08-29 15:14 | 2023-08-29 15:14 | not working | (TC 06/13/2023 03:14:38 PM, Huang, Jessica) cleaned the sensor and passed calibration
1063365 | 197065 | CM | CMPLT | CMPLT | 2023-08-29 15:15 | 2023-08-29 15:15 | not working | (TC 06/13/2023 03:15:27 PM, Huang, Jessica) cleaned the sensor and passed calibration
502231 | 197066 | CM | CMPLT | CMPLT | 2023-08-29 15:48 | 2023-08-28 19:46 | Problem with table motion. | Table motion adjusted/calibrated, also the cable hose for the powr cable was adjusted.
1066115 | 197067 | FC | CMPLT | CMPLT | 2023-08-29 15:56 | 2023-08-29 15:56 | QC check | (TC 08/29/2023 03:55:52 PM, Canjimavadivel, Kamallesh ) QC check complete
1058279 | 197068 | PM | CMPLT | CMPLT | 2023-08-30 06:55 | 2023-08-30 07:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 07:02:31 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 91 mohms Leakage current: 65.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
| 197069 | RDS | CMPLT | CMPLT | 2023-08-30 07:16 | 2023-08-30 09:11 | Weekly checks for ED
| 197070 | RDS | CMPLT | CMPLT | 2023-08-30 07:17 | 2023-08-31 14:19 | Weekly checks for ICU
67969 | 197071 | CM | CMPLT | CMPLT | 2023-08-30 07:18 | 2023-08-30 07:19 | Opposite doors were attached | (TC 08/30/2023 07:18:42 AM, Baghirzade, Nara) Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
1056573 | 197072 | CM | CMPLT | CMPLT | 2023-08-30 07:45 | 2023-08-30 07:46 | Updating asset info | (TC 08/30/2023 07:45:19 AM, Canjimavadivel, Kamallesh ) Updated asset to reflect training in dashboard and TMS
52252 | 197073 | CM | CMPLT | CMPLT | 2023-08-30 08:02 | 2023-08-30 08:03 | Opposite doors attached | (TC 08/30/2023 08:02:25 AM, Baghirzade, Nara) Swapped doors around. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
| 197074 | CM | CMPLT | CMPLT | 2023-08-30 08:49 | 2023-08-30 08:50 | BP install in procedure room | (TC 08/30/2023 08:48:48 AM, Baghirzade, Nara) Installed the BP on to the wall mount in procedure room. BP device is fully functional. Functional checks of BPs in exam rooms 11, 12, and 14 are good.a
1047144 | 197075 | CM | CMPLT | CMPLT | 2023-08-30 08:59 | 2024-01-25 09:54 | techs unable to login | (TC 08/30/2023 01:00:00 PM, Scero, Tomlyn) GE tech was onsite to service the CASE system and was unable to log in. Kamran and I spent sometime with him in the stress lab trying to troubleshoot ways to login. A help tickets was created and eventually Justin and Westley were called. (TC 01/25/2024 09:46:53 AM, Scero, Tomlyn) I followed up with GE Tech 'Krunal.JAGODANA@ge.com' for a follow up or service report. Unfortunately he never returned my email. The issues was resolved though.
1067554 | 197076 | CM | CMPLT | CMPLT | 2023-08-30 09:02 | 2023-08-30 13:40 | Coordinated pick up with 3M | (TC 08/30/2023 09:02:11 AM, Baghirzade, Nara) Coordinated pickup of loaner VAC with 3M. Pick-up confirmation #: 12827881 (TC 08/30/2023 01:40:10 PM, Baghirzade, Nara) Picked up by 3M.
1067598 | 197077 | CM | CMPLT | CMPLT | 2023-08-30 09:30 | 2023-09-12 09:01 | 6 year PM | (TC 09/12/2023 09:00:07 AM, Scero, Tomlyn) 6 year PM parts replaced. Tests complete and passed.
68392 | 197078 | CM | CMPLT | CMPLT | 2023-08-30 09:36 | 2023-10-30 10:09 | Failing self test | (TC 09/08/2023 09:29:52 AM, Scero, Tomlyn) Tested DIVA and failed self test. Replaced 3 year PM parts and still fails. Also fails self tests; 5054 Safety function 5051 Diva leak valve DUMA voltage 2607 (red) Will need to send to Draeger Germany for flat rate repair. (TC 09/12/2023 02:43:18 PM, Scero, Tomlyn) PO requested from plexxus (TC 09/13/2023 10:50:48 AM, Scero, Tomlyn) PO received from plexxus. RMA received from Draeger online. 308567626 KJV shipping scheduled a pick up. (TC 10/30/2023 10:06:02 AM, Scero, Tomlyn) Device returned. Service report uploaded. Tested unit in AGM complete
1054758 | 197079 | CM | CMPLT | CMPLT | 2023-08-30 10:05 | 2023-08-30 10:05 | not working | (TC 06/06/2023 10:05:14 AM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
1046528 | 197080 | CM | CMPLT | CMPLT | 2023-08-30 10:06 | 2023-08-30 10:06 | not working | (TC 06/06/2023 10:06:34 AM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
1056125 | 197081 | PM | CMPLT | CMPLT | 2023-08-30 10:11 | 2023-08-30 10:16 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/30/2023 10:11:12 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
67922 | 197082 | CM | CMPLT | CMPLT | 2023-08-30 10:28 | 2023-08-30 10:28 | Latch recall | (TC 08/30/2023 10:28:21 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
68552 | 197083 | CM | CMPLT | CMPLT | 2023-08-30 10:30 | 2023-08-30 11:57 | O2 reading issue | (TC 08/30/2023 11:54:42 AM, Scero, Tomlyn) While I was away last week the device was pulled and said it gave an O2 issue. When tested the machine I found that the sevo was giving an error along with a leak in the breathing system. I removed the diva and replaced the breathing system and restarted the selft test. It passed and is fully functional. I ran the device in different vent modes to test the O2 concentration. PATO sensor is working fine and O2 measuring correctly. No issue with O2 sensor at the moment.
| 197084 | PA | CMPLT | CMPLT | 2023-08-30 10:42 | 2023-09-01 13:15 | Virox wipes | (TC 08/30/2023 10:47:06 AM, Canjimavadivel, Kamallesh ) Placed an order for wipes (TC 09/01/2023 01:14:16 PM, Canjimavadivel, Kamallesh ) Received wipes and put away in cabinet
67922 | 197085 | FC | CMPLT | CMPLT | 2023-08-30 11:14 | 2023-08-30 11:15 | Functional checks | (TC 08/30/2023 11:14:19 AM, Baghirzade, Nara) Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 76% RH. Incubator is fully functional.
1066460 | 197086 | FC | CMPLT | CMPLT | 2023-08-30 11:19 | 2023-08-30 11:19 | QC checklist | (TC 08/30/2023 11:19:11 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
1067604 | 197087 | MEP | CMPLT | CMPLT | 2023-08-30 11:55 | 2024-05-01 08:19 | MSH 5 US DEPLOY SPRESSO APLIO 500 ACCEPTANCE | (TC 08/30/2023 11:55:08 AM, Whittington, Scott) MSH 5 US DEPLOY SPRESSO APLIO 500 ACCEPTANCE (TC 08/30/2023 11:55:28 AM, Dolinar, John) MSH 5 US DEPLOY SPRESSO APLIO 500 ACCEPTANCE | MSH 5 US DEPLOY SPRESSO APLIO 500 ACCEPTANCE
69484 | 197088 | CM | ACTIV | ISSUE | 2023-08-30 13:39 |  | Not working | (TC 08/30/2023 01:38:44 PM, Reid, Emmerson ) Downloaded, cleared and viewed log history. I did not see any recent errors to understand what was not working. Found that the AIL sensor is damaged and needs to be replaced. Waiting on parts.
1059982 | 197089 | PM | CMPLT | PMFAIL | 2023-08-30 14:26 | 2023-08-30 14:29 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 08/30/2023 02:26:41 PM, Cortez, John) Right side head rail bent Will request service from hillrom
1059982 | 197090 | CM | CMPLT | CMPLT | 2023-08-30 14:29 | 2023-10-19 15:47 | Created by Failed PM WO # 197089 : PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/17/2023 03:43:02 PM, Cortez, John) Bed repaired by Vince (Hillrom) | (TC 08/30/2023 02:26:41 PM, Cortez, John) Right side head rail bent Will request service from hillrom
| 197091 | CM | CMPLT | CMPLT | 2023-08-31 07:51 | 2023-09-01 09:41 | multiple tasks | (TC 08/29/2023 07:51:12 AM, Samanian, Kamran) 1- looking into GE CASE stress test system with GE and Dianne due to logging issue and loosing the path 2 - CCU physio data checking for CCU bed 8 3 - Looking into scanner for VSM at OPG
68034 | 197092 | PA | CMPLT | CMPLT | 2023-08-31 07:52 | 2023-09-05 10:40 | Ordering incubator hood | (TC 08/31/2023 07:54:57 AM, Baghirzade, Nara) SC 1003039388 (TC 09/05/2023 10:37:51 AM, Baghirzade, Nara) PO 7202275099
64269 | 197093 | CM | CMPLT | CMPLT | 2023-08-31 10:01 | 2023-08-31 10:29 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 08/31/2023 10:00:45 AM, Baghirzade, Nara) Took monitor down from room 88. Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. Kamran and I connected it to the support tool however we could not make any changes using the tool. Brought monitor to NICU procedure room, connected it to the CNS machine and demo mode was successfully showing all parameters on the CNS display. Put label RS232 on the monitor. Put monitor back in room 88. (TC 08/31/2023 10:29:02 AM, Baghirzade, Nara) RS232 board SN: DE48609359
| 197094 | CM | CMPLT | CMPLT | 2023-08-31 10:07 | 2023-08-31 10:07 | meeting with katrina and John for Image capture | (TC 08/31/2023 10:07:13 AM, Samanian, Kamran) discussed about the results and arranged for next meeting for more testings
68673 | 197095 | CM | CMPLT | NFF | 2023-08-31 12:13 | 2023-08-31 12:14 | Reported that battery is low even when plugged in to AC | (TC 08/31/2023 12:12:36 PM, Baghirzade, Nara) No fault found with the device. Battery was fully depleted; left it for a day to charge. Checked battery status on ASM and battery is within capacity. Charged fully and is operational on battery. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 49 mohms Leakage current: 76.9 uA
67932 | 197096 | CM | CMPLT | CMPLT | 2023-08-31 13:21 | 2023-08-31 13:21 | Latch recall | (TC 08/31/2023 01:21:18 PM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
| 197097 | CM | CMPLT | CMPLT | 2023-08-31 13:24 | 2023-08-31 13:25 | Meeting with Julie and Monica | (TC 08/31/2023 01:24:15 PM, Baghirzade, Nara) Met with Julie and Monica (managing the wound VACs). They had some confusion on which VACs they sent out for repair vs loaner. Also noted which label is for which VAC for their records and will input the labels in TMS as well.
1056206 | 197098 | CM | CMPLT | CMPLT | 2023-08-31 14:06 | 2023-08-31 14:07 | Patched and crashed - Rebuilt | (TC 08/31/2023 02:05:58 PM, Samanian, Kamran) tried to patch the server through FP server and it was crahsed. removed the USB from motherboard and rebuilt the server . Reconfigure it as STANDBY server until it will be patched soon and will be renamed again
50599 | 197099 | PM | ACTIV | ISSUE | 2023-09-01 03:26 |  | PM-MAJOR - Annual PM For: Scales, Infant
60165 | 197100 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Radiofrequency Generators
62380 | 197101 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-01-19 14:00 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 01/19/2024 01:57:59 PM, Rhuland, Courtney) PM complete RF leakage: 40mA Se attached checklist for details
62982 | 197102 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-10-02 09:57 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
63722 | 197103 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
64867 | 197104 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-12-06 11:03 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 12/06/2023 11:02:24 AM, Campbell, Connor) Pm complete/passed, refer to procedure
64881 | 197105 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Smoke Evacuation Systems, Surgical
64910 | 197106 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: High Flow Insufflator
65173 | 197107 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-10-31 10:29 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 10/31/2023 10:26:35 AM, Reid, Emmerson ) PM Complete, see proceedure for details. Note that one of the pins on the dose corrd had broken off in the CADD. I removed the pin. The Dose cord does still work to deliver a dose but the light does not work. I will send it out like this.
65749 | 197108 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM Major Annual: Microstream CO2 extension
65898 | 197109 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65903 | 197110 | PM | CMPLT | PMFAIL | 2023-09-01 03:27 | 2023-09-20 08:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/14/2023 10:45:42 AM, Rendon Lopez , Elizabeth) Front case cracked.
65943 | 197111 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-20 12:38 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/20/2023 12:37:46 PM, Campbell, Connor) PM complete/passed, refer to ASM
66024 | 197112 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66040 | 197113 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-05 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/05/2023 10:32:17 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 47 mohms Leakage current: 92.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66054 | 197114 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-05 11:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 10:59:36 AM, Cortez, John) see wo 189466
66060 | 197115 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-08 08:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 11:09:35 AM, Campbell, Connor) 2023/09/07 PM complete, refer to ASM Battery conditioning started 2023/09/08 Battery Condition complete - old capacity 3400 mAh - new capacity 4109 mAh
66460 | 197116 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-10-16 12:54 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 73uA and 89 ohm.
66669 | 197117 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66670 | 197118 | PM | CNP | ENL | 2023-09-01 03:27 |  | BP Check
66738 | 197119 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-01-24 11:59 | PM-MAJOR - Annual PM For: Giraffe Shuttle | (TC 01/24/2024 11:58:13 AM, Baghirzade, Nara) PM passed and completed. Fan is functional. Alarms work when foot pedal is activated. Battery health is good; green LED illuminated. Electrical safety: power cord resistance = 245 mohms leakage current = 85.4 uA (TC 01/24/2024 11:58:22 AM, Rhuland, Courtney) PM passed and completed. Fan is functional. Alarms work when foot pedal is activated. Battery health is good; green LED illuminated. Electrical safety: power cord resistance = 245 mohms leakage current = 85.4 uA
66771 | 197120 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-28 10:20 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | PM completed Sept 2021, not due for 3 years or Sept 2024. pm rescheduled
66793 | 197121 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 07:53 | BP Check
66794 | 197122 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66795 | 197123 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-06-18 12:44 | BP Check
66796 | 197124 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66797 | 197125 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-06-03 11:16 | PM | (TC 06/03/2024 11:15:53 AM, DeGuzman, Leila) PM completed. Go to procedure to see details.
66798 | 197126 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 07:55 | BP Check
66799 | 197127 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-05-31 13:31 | BP Check
66800 | 197128 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 07:56 | BP Check
66801 | 197129 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 07:56 | BP Check
66802 | 197130 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 08:17 | BP Check
66807 | 197131 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 08:30 | BP Check
66808 | 197132 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 08:32 | BP Check
66809 | 197133 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66810 | 197134 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66811 | 197135 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 08:35 | BP Check
66812 | 197136 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66813 | 197137 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 11:53 | BP Check
66814 | 197138 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66818 | 197139 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66822 | 197140 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66823 | 197141 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 11:56 | BP Check
66826 | 197142 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 11:57 | BP Check
66827 | 197143 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66828 | 197144 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66830 | 197145 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 11:58 | BP Check
66831 | 197146 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:00 | BP Check
66833 | 197147 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:01 | BP Check
66834 | 197148 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66835 | 197149 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66838 | 197150 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:07 | BP Check
66841 | 197151 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:08 | BP Check
66842 | 197152 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-18 11:56 | BP Check | No BP check, PM only.
66844 | 197153 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
66845 | 197154 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:10 | BP Check
66846 | 197155 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:12 | BP Check
66847 | 197156 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:13 | BP Check
66851 | 197157 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-16 09:10 | BP Check
66852 | 197158 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:15 | BP Check
66853 | 197159 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:16 | BP Check
66854 | 197160 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-09-12 13:50 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/12/2024 01:50:16 PM, Rhuland, Courtney) Supervised (TC 09/12/2024 01:49:28 PM, Alvarado, Karla) Performed touch screen calibration PM Completed All tests passed See procedure for details
66854 | 197161 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:20 | BP Check
66856 | 197162 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-06-03 12:14 | PM-MAJOR - Annual PM For: PATIENT MONITORS, PORTABLE | (TC 06/03/2024 12:13:33 PM, DeGuzman, Leila) PM completed. Go to procedure to see details.
66857 | 197163 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-06-03 07:11 | BP Check
66859 | 197164 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:24 | BP Check
66860 | 197165 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:26 | BP Check
66863 | 197166 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:28 | BP Check
67200 | 197167 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:30 | BP Check
67201 | 197168 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | BP Check
67203 | 197169 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:31 | BP Check
67205 | 197170 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:33 | BP Check
67206 | 197171 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:34 | BP Check
67207 | 197172 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67207 | 197173 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-25 12:35 | BP Check
67208 | 197174 | PM | CANCL | CANCL | 2023-09-01 03:27 | 2024-07-16 09:11 | BP Check
67236 | 197175 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-01-24 11:57 | PM-MAJOR - Annual PM For: Infant Resuscitation System | (TC 01/24/2024 11:56:03 AM, Baghirzade, Nara) PM completed. All values are within manufacturer specifications. With flow adjusted to 0 L/min, checked that no flow is going through. Passed. Set flow to 7 L/min - @0 O2% = 0 @20 O2% = 20 @40 O2% = 37.8 @60 O2% = 58 @80 O2% = 79 @100 O2% = 98.8 Checked flow @ 5 L/min = 5.2 L/min @ 7 L/min = 7.4 L/min @ 10 L/min = 10 L/min (TC 01/24/2024 11:56:54 AM, Rhuland, Courtney) PM completed. All values are within manufacturer specifications. With flow adjusted to 0 L/min, checked that no flow is going through. Passed. Set flow to 7 L/min - @0 O2% = 0 @20 O2% = 20 @40 O2% = 37.8 @60 O2% = 58 @80 O2% = 79 @100 O2% = 98.8 Checked flow @ 5 L/min = 5.2 L/min @ 7 L/min = 7.4 L/min @ 10 L/min = 10 L/min
67766 | 197176 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-10-25 10:02 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 10/25/2023 09:08:42 AM, Reid, Emmerson ) Upgraded Firmware to V040201 and pushed APS Epidural - Juily 2023 (TC 10/25/2023 10:00:36 AM, Reid, Emmerson ) PM complete see attached checklist. Had to run the volume test twice.
67815 | 197177 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67944 | 197178 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-01-29 09:38 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 01/29/2024 09:34:49 AM, Rhuland, Courtney) PM completed REF & Main voltage calibrated Scale calibrated & verified Functional check - OK Set 35C, 75% humidity Read 35C, 74.6% humidity Electrical safety - OK
67948 | 197179 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-11-09 09:33 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/09/2023 09:30:53 AM, Campbell, Connor) replaced SW door latch because it was broken; originally came down for this issue PM Passed Calibrated System - OK Calibrated Line Voltage - OK Calibrated Scale - OK Leakage Current Off - 138.6 uA Leakage Current ON - 86.6 uA Ground Resistance Off - 153 mOhms Ground Resistance ON - 153 mOhms Functional Test - Pass - Temp = 33.0 deg - Humidity = 75%
67961 | 197180 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-11-09 09:13 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/03/2023 02:04:54 PM, Rhuland, Courtney) Began chassis replacement - broken screw hold at humidifier (TC 11/08/2023 02:40:11 PM, Rhuland, Courtney) Completed chassis replacement Replaced humidifier PM will follow (TC 11/09/2023 09:10:55 AM, Rhuland, Courtney) Rebuild completed PM completed Calibrated LV and main voltage Calibrated and verified scale Changed time Functional check - OK Tested 35C, 75% humidity Verified 35C, 75% humidity Electical safety - OK Leakage current - 77.7 uA Resistance - 60 mOhms
67969 | 197181 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
68291 | 197182 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68295 | 197183 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68374 | 197184 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68378 | 197185 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68554 | 197186 | PM | ACTIV | ISSUE | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68640 | 197187 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-10-19 17:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/19/2023 05:29:58 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
68641 | 197188 | PM | CMPLT | PMFAIL | 2023-09-01 03:27 | 2023-09-18 10:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 09:58:17 AM, Campbell, Connor) suspected the pressure sensor was fauly, confirmed bottle pressure sensor fatal error in the logs
68642 | 197189 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68651 | 197190 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-18 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/15/2023 10:38:31 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
68656 | 197191 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68657 | 197192 | PM | CMPLT | PMFAIL | 2023-09-01 03:27 | 2023-09-18 11:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 11:32:44 AM, Campbell, Connor) Pm failed due to "check IV set" error and pump sensor fatal errors on both pressure sensors
68666 | 197193 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68669 | 197194 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68713 | 197195 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68715 | 197196 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68735 | 197197 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68746 | 197198 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68754 | 197199 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-04-03 13:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:13:07 PM, Rhuland, Courtney) PM completed by BD
68764 | 197200 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-04-30 13:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/30/2024 01:25:36 PM, Rhuland, Courtney) Sent out for PM - completed by BD Replaced battery Loaded network and data set
68776 | 197201 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-03-20 14:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68788 | 197202 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-06 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 02:48:28 PM, Campbell, Connor) pm completed, all tests passed. refer to ASM battery replaced supervised by nara
68844 | 197203 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68858 | 197204 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-10-20 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/20/2023 09:54:12 AM, Campbell, Connor) Replaced power cord as it was broken replaced battery cleaned iui connectors and bottom pump latch pm passed/complete
68905 | 197205 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68935 | 197206 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68946 | 197207 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68948 | 197208 | PM | CNP | ENL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68951 | 197209 | PM | CANCL | CANCL | 2023-09-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68968 | 197210 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2023-09-05 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/30/2023 09:48:35 AM, Cortez, John) see WO 189525
68983 | 197211 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-03-13 10:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:56:39 AM, Cortez, John) Female IUI replaced PM complete
68999 | 197212 | PM | CMPLT | CMPLT | 2023-09-01 03:27 | 2024-04-15 09:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/11/2024 09:04:16 AM, Huang, Jessica) replacedt the female IUI connector and completed PM
69032 | 197213 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69043 | 197214 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69046 | 197215 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-26 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/26/2023 09:42:38 AM, Baghirzade, Nara) Both pressure sensors were defective; ran the analog sensor test on ASM and found both sensors stay at 4.9V. Replaced bottle side sensor and patient side sensor. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69050 | 197216 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69051 | 197217 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69052 | 197218 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69065 | 197219 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-11-28 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/27/2023 09:55:46 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69068 | 197220 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69073 | 197221 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-03-21 18:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69086 | 197222 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69103 | 197223 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69104 | 197224 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69111 | 197225 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69115 | 197226 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69128 | 197227 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69133 | 197228 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-12 15:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/12/2023 03:07:17 PM, Cortez, John) Cleaned, Downloaded error/event logs, LVP display dim, replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
69135 | 197229 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-12-29 10:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 12/29/2023 10:40:05 AM, Campbell, Connor) PM complete/passed, refer to ASM
69146 | 197230 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69159 | 197231 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69174 | 197232 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69190 | 197233 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69191 | 197234 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-03 11:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 11:15:06 AM, Reid, Emmerson ) Replaced the bezel because it was not calibrating. Ran Calibration and ASM PM, all tests passed. PM Complete.
69192 | 197235 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-06 08:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 08:03:33 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69197 | 197236 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69203 | 197237 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-24 09:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69207 | 197238 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-11 08:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/11/2023 08:02:30 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69216 | 197239 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69224 | 197240 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69246 | 197241 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69252 | 197242 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-03-22 12:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2024 12:30:31 PM, Reid, Emmerson ) Replaced the front case, ran calibration and ASM PM. PM Complete.
69253 | 197243 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69270 | 197244 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69279 | 197245 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-11 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/11/2023 09:21:10 AM, Cortez, John) Cleaned, Downloaded error/event logs, LVP display dim, replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
69281 | 197246 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-03-15 08:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69299 | 197247 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69302 | 197248 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-03 14:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:30:13 PM, Rhuland, Courtney) PM completed by BD
69307 | 197249 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69309 | 197250 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69317 | 197251 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-15 15:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/15/2023 03:45:36 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
69318 | 197252 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69319 | 197253 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69332 | 197254 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69337 | 197255 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-03-14 14:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69338 | 197256 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69347 | 197257 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69351 | 197258 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-03-19 21:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69354 | 197259 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-04 13:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:02:19 PM, Rhuland, Courtney) PM completed by BD
69361 | 197260 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69381 | 197261 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69402 | 197262 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69408 | 197263 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-03 14:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:38:37 PM, Rhuland, Courtney) PM completed by BD
69411 | 197264 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-25 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/25/2023 11:23:33 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69423 | 197265 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-19 10:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 10:06:51 AM, Campbell, Connor) PM complete/passed, refer to ASM
69426 | 197266 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69435 | 197267 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-25 14:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69437 | 197268 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69456 | 197269 | PM | CMPLT | PMFAIL | 2023-09-01 03:28 | 2023-09-18 11:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 11:54:58 AM, Campbell, Connor) rate test failed
69462 | 197270 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69464 | 197271 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-19 15:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 03:21:32 PM, Cortez, John) Cleaned, Downloaded error/event logs, LVP display dim, replaced, Calibrated infusion rate, PM completed successfully, results logged in ASM
69482 | 197272 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-03-11 16:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69486 | 197273 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69491 | 197274 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-23 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69499 | 197275 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69509 | 197276 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-01 11:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 11:37:36 AM, Rendon Lopez , Elizabeth) Calibrated rate accuracy. PM completed, all tests passed
69511 | 197277 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-06 09:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 08:59:10 AM, Reid, Emmerson ) Downloaded and cleared log history. Ran ASM PM, all tests passed. PM Complete.
69513 | 197278 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-30 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69527 | 197279 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-04-02 16:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69543 | 197280 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69558 | 197281 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69559 | 197282 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-18 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Refer to WO 196647
69580 | 197283 | PM | CMPLT | PMFAIL | 2023-09-01 03:28 | 2023-09-25 09:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/25/2023 09:07:46 AM, Campbell, Connor) Display board top right segment is not working, need display replaced Front case has cracked plastic from being dropped, needs replacement to prevent liquid ingress
69587 | 197284 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69639 | 197285 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69646 | 197286 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-03-14 16:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69657 | 197287 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69665 | 197288 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-18 15:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 03:40:34 PM, Cortez, John) Cleaned, Downloaded error/event logs, Replaced LVP display and alarm LED cover Calibrated infusion rate, PM completed successfully, results logged in ASM
69667 | 197289 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69679 | 197290 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65553 | 197291 | PM | CNP | ENL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
67362 | 197292 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-21 12:07 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 09/21/2023 12:05:16 PM, Campbell, Connor) PM complete/passed Drained and cleaned the reservoir Refilled with new recirculating fluid Replaced o-rings Replaced Drain tube cap General Alarm Test - OK Low recirculating fluid alarm test - OK Interlock Alarm Test - OK Over Temp Alarm Test - OK Electrical Safety - Current Leakage - 14.0 uA - Ground Resistance - 141Ohms
63957 | 197293 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-01-23 14:23 | PM-MAJOR - Annual PM For: Infant T-Piece Resuscitator | (TC 01/23/2024 02:22:25 PM, Baghirzade, Nara) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 20 cmH2O @40 = 40 cmH2O Greater than 60 = passed - 75 cmH2O Less than 80 = 77 cmH2O (TC 01/23/2024 02:23:00 PM, Rhuland, Courtney) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 20 cmH2O @40 = 40 cmH2O Greater than 60 = passed - 75 cmH2O Less than 80 = 77 cmH2O
65614 | 197294 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-01-23 14:20 | PM-MAJOR - Annual PM For: Infant T-Piece Resuscitator | (TC 01/23/2024 02:19:24 PM, Baghirzade, Nara) Performance assurance passed and completed. All values within range. @10 = 9 cmH2O @20 = 20 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 75 cmH2O Less than 80 = 76 cmH2O (TC 01/23/2024 02:20:01 PM, Rhuland, Courtney) Performance assurance passed and completed. All values within range. @10 = 9 cmH2O @20 = 20 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 75 cmH2O Less than 80 = 76 cmH2O
65615 | 197295 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2024-01-23 14:22 | PM-MAJOR - Annual PM For: Infant T-Piece Resuscitator | (TC 01/23/2024 02:20:59 PM, Baghirzade, Nara) Performance assurance passed and completed. All values within range. @10 = 9 cmH2O @20 = 19 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 77 cmH2O Less than 80 = 75 cmH2O
500846 | 197296 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-09-29 16:59 | PM - Semi-Annual
500847 | 197297 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-02 17:12 | msh us rm 1 PM - Semi-Annual | (TC 10/02/2023 05:12:26 PM, Whittington, Scott) msh us rm 1 PM - Semi-Annual | msh us rm 1 PM - Semi-Annual
500850 | 197298 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-02 17:04 | MSH IRUS "CLYDE" PM - Semi-Annual | (TC 10/02/2023 05:03:54 PM, Whittington, Scott) MSH IRUS "CLYDE" PM - Semi-Annual | MSH IRUS "CLYDE" PM - Semi-Annual
32510 | 197299 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-20 22:32 | MSH DR1 PM - Semi-Annual | (TC 10/23/2023 11:08:00 AM, Whittington, Scott) MSH DR1 PM - Semi-Annual | MSH DR1 PM - Semi-Annual
32510 | 197300 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-20 22:32 | MSH DR1 HARP OCT 21 TRGT - Semi Annual | (TC 10/21/2023 07:06:00 PM, Whittington, Scott) MSH DR1 HARP OCT 21 TRGT - Semi Annual (TC 10/23/2023 11:07:00 AM, Whittington, Scott) MSH DR1 HARP OCT 21 TRGT - Semi Annual HARP FOLLOWUP | MSH DR1 HARP OCT 21 TRGT - Semi Annual
32540 | 197301 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-13 14:49 | PM - Monthly | PM completed Annual CT HARP/QA testing completed sc-628197 TOS phantom a(-992.1,9.6) b(333.6,13.0) c(133.4,13.4) d(102.1,15.2) e(-105.9,12.6) f(1.5,14.1) -passed all phantom tests
32560 | 197302 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-03 15:28 | PM - Monthly | PM completed Annual HARP/QA testing completed
32580 | 197303 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-04 15:45 | PM - Monthly | PM completed Annual CT/QA testing completed
500732 | 197304 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-12 16:53 | MSH-20-470 (Sports Clinic) PM- SCHEDULED SEMI-ANNUAL | (TC 10/12/2023 03:38:49 PM, Whittington, Scott) MSH-20-470 (Sports Clinic) PM- SCHEDULED SEMI-ANNUAL | MSH-20-470 (Sports Clinic) PM- SCHEDULED SEMI-ANNUAL
500800 | 197305 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-12 16:52 | msh bius PM - Semi-Annual | (TC 10/12/2023 03:37:33 PM, Whittington, Scott) msh bius PM - Semi-Annual | msh bius PM - Semi-Annual
500801 | 197306 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-12 16:52 | msh bius PM - Semi-Annual | (TC 10/12/2023 03:37:58 PM, Whittington, Scott) msh bius PM - Semi-Annual | msh bius PM - Semi-Annual
500802 | 197307 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-12 16:52 | msh bius PM - Semi-Annual | (TC 10/12/2023 03:38:25 PM, Whittington, Scott) msh bius PM - Semi-Annual | msh bius PM - Semi-Annual
500857 | 197308 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-02 16:59 | MSH US RM 8 PM - Semi-Annual | (TC 10/02/2023 04:57:56 PM, Whittington, Scott) MSH US RM 8 PM - Semi-Annual | MSH US RM 8 PM - Semi-Annual
500858 | 197309 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-02 14:51 | MSH-5-US rm 7 PM - Semi-AnnualPM - Semi-Annual | (TC 10/02/2023 02:50:55 PM, Whittington, Scott) MSH-5-US rm 7 PM - Semi-AnnualPM - Semi-Annual | MSH-5-US rm 7 PM - Semi-AnnualPM - Semi-Annual
500868 | 197310 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-02 14:42 | MSH US RM 4 PM - Semi-Annual | (TC 10/02/2023 02:42:20 PM, Whittington, Scott) MSH US RM 4 PM - Semi-Annual | MSH US RM 4 PM - Semi-Annual
500869 | 197311 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-02 14:56 | MSH US RM 6 PM - Semi-Annual | (TC 10/02/2023 02:55:39 PM, Whittington, Scott) MSH US RM 6 PM - Semi-Annual DRIVE D UNABLE TO DFRG, REQ HDD REPLACE | MSH US RM 6 PM - Semi-Annual DRIVE D UNABLE TO DFRG, REQ HDD REPLACE
501621 | 197312 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-12-28 15:27 | PM - Annual | (TC 11/30/2023 12:54:18 PM, Green, Paul) Removed injector and installed working spare. Performed PM on injector. Pot cal ok. Pressure calibration A side ok. When performing B side calibration, getting a SW -101 error on CRU. On laptop, Critical Error. "Error in Function frmSSEPPressureCal:writeStatusMessage Note this Error information to Technical Support Error number 5. Description: Invalid procedure call or argument. (TC 12/08/2023 03:50:34 PM, Green, Paul) Called TAC. Ordered mechanical assembly. Received assembly and stored with injector. Calibration gear at TWH so will return when time permits. They are ok with spare for now. (TC 12/20/2023 09:12:38 AM, Green, Paul) Installed mechanical assembly. Calibrated positon and pressure. Upon final test, piston drove through limit and physically damaged. Reported this to Bayer. Bayer will send a replacement as this one was defective. (TC 12/19/2023 09:12:38 AM, Green, Paul) Installed mechanical assembly. Calibrated positon and pressure. Upon final test, piston drove through limit and physically damaged. Reported this to Bayer. Bayer will send a replacement as this one was defective. (TC 12/28/2023 03:26:39 PM, Green, Paul) Completed pressure calibration with new cal syringes, ok. Performed final test, ok. Will install back in room when time permits.
501888 | 197313 | PM | CMPLT | CMPLT | 2023-09-01 03:28 | 2023-10-02 15:01 | msh us rm 2pm | (TC 10/02/2023 03:01:22 PM, Whittington, Scott) msh us rm 2pm | msh us rm 2pm
501935 | 197314 | PM | CANCL | CANCL | 2023-09-01 03:28 |  | PM-MAJOR - Annual PM For: Scanning Systems, Ultrasonic
1046035 | 197315 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-10-30 11:40 | MSH ER DRX EVO HARP TRGT APR 30 - Semi Annual | (TC 10/30/2023 11:39:51 AM, Fan, Tak) HARP Completed. kVp calibrated. | MSH ER DRX EVO HARP TRGT APR 30 - Semi Annual
1046035 | 197316 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-10-30 11:43 | PM Major | (TC 10/30/2023 11:40:36 AM, Fan, Tak) PM completed. Generator calibrated. (TC 11/01/2023 12:52:00 PM, Fan, Tak) Replaced wall stand face cover.
1047003 | 197317 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047004 | 197318 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047006 | 197319 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047007 | 197320 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047010 | 197321 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047011 | 197322 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-03-26 12:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047014 | 197323 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047046 | 197324 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-10-05 11:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 11:34:47 AM, Campbell, Connor) PM passed/complete, refer to ASM
1047047 | 197325 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-09-06 13:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 01:49:15 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1047049 | 197326 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047050 | 197327 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047062 | 197328 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-03-20 14:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047099 | 197329 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-03-20 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047100 | 197330 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-03-20 15:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047102 | 197331 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-03-20 15:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047362 | 197332 | PM | ACTIV | PV | 2023-09-01 03:29 |  | PM-MAJOR- Annual PM for: Testers
1048634 | 197333 | PM | CNP | ENL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048664 | 197334 | PM | CNP | ENL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049028 | 197335 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-10-24 09:52 | Planned Maintenance | (TC 10/24/2024 09:52:27 AM, Alvarado, Karla) PM completed All tests passed See procedure for details Updated Asset number Reset the hour meter.
1049046 | 197336 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-11-16 18:02 | Planned Maintenance | (TC 11/16/2023 06:01:58 PM, Huang, Jessica) defective device, sent the PO to the vendor and received the replacement unit. sent it back to OR
1049048 | 197337 | PM | ACTIV | ISSUE | 2023-09-01 03:29 |  | Planned Maintenance
1049677 | 197338 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-10-24 14:28 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/24/2023 02:27:35 PM, Campbell, Connor) Adjusted the LCD, overtemp alarm and circulation temperature pots PM complete/passed, refer to procedure
1050269 | 197339 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-09-18 16:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 04:25:43 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1050773 | 197340 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-09-08 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 09:49:25 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1053349 | 197341 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-11-01 17:49 | MSH BI 1274 H2 HARP OCT 20 TRGT SEMI-ANNUAL | (TC 11/01/2023 05:48:32 PM, Dolinar, John) MSH BI 1274 HARP OCT 20 TRGT SEMI-ANNUAL (TC 11/01/2023 05:48:53 PM, Whittington, Scott) MSH BI 1274 HARP OCT 20 TRGT SEMI-ANNUAL (TC 11/01/2023 05:48:53 PM, Whittington, Scott) MSH BI 1274 HARP OCT 20 TRGT SEMI-ANNUAL | MSH BI 1274 HARP OCT 20 TRGT SEMI-ANNUAL
1055114 | 197342 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-04-15 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/11/2024 09:06:06 AM, Huang, Jessica) replaced the power cord and completed PM
1055116 | 197343 | PM | CNP | ENL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055157 | 197344 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-09-19 15:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 03:47:45 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
1055744 | 197345 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-01-23 14:18 | PM-MAJOR - Annual PM For: Infant T-Piece Resuscitator | (TC 01/23/2024 02:15:30 PM, Baghirzade, Nara) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 20 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 78 cmH2O Less than 80 = 70 cmH2O (TC 01/23/2024 02:18:08 PM, Rhuland, Courtney) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 20 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 78 cmH2O Less than 80 = 70 cmH2O
1056125 | 197346 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-11-20 14:46 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 09/13/2023 08:38:26 AM, Canjimavadivel, Kamallesh ) Placed order for PM kit (TC 10/03/2023 01:36:04 PM, Canjimavadivel, Kamallesh ) Followed up with Plexxus regarding status of PO (TC 11/20/2023 02:27:32 PM, Canjimavadivel, Kamallesh ) Worked with Connor and showed him replacement of PM parts. Connor completed cleaning cycle and device was returned to unit.
1056934 | 197347 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-09-01 13:21 | PM Major Annual: Physiologic Monitor Modules, Exhaled Carbon Dioxide | Refer to WO#189715, PM recently completed
1056935 | 197348 | PM | CMPLT | PMFAIL | 2023-09-01 03:29 | 2023-10-31 09:09 | PM Major Annual: Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 10/25/2023 10:59:51 AM, Canjimavadivel, Kamallesh ) Started PM, final calibration did not complete (TC 10/31/2023 09:07:35 AM, Canjimavadivel, Kamallesh ) Attempted to calibrate again, fail calibration
1056943 | 197349 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056944 | 197350 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056945 | 197351 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056946 | 197352 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056947 | 197353 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056948 | 197354 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056949 | 197355 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056950 | 197356 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056951 | 197357 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056952 | 197358 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056953 | 197359 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056954 | 197360 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056955 | 197361 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056956 | 197362 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056957 | 197363 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056958 | 197364 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056959 | 197365 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056960 | 197366 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056961 | 197367 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056962 | 197368 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056963 | 197369 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056964 | 197370 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056965 | 197371 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056966 | 197372 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056967 | 197373 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056968 | 197374 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056969 | 197375 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056970 | 197376 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056971 | 197377 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056972 | 197378 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056973 | 197379 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056974 | 197380 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1056975 | 197381 | PM | ACTIV | SHP | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM for Bed electric
1057221 | 197382 | PM | ACTIV | ISSUE | 2023-09-01 03:29 |  | PM Major - Annual: Physiologic Monitor Modules, Exhaled Carbon Dioxide
1057393 | 197383 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1057574 | 197384 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-11-03 10:29 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 11/03/2023 10:29:02 AM, Reid, Emmerson ) PM Complete, see report.
1058117 | 197385 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-09-27 08:02 | PM-MAJOR-Annual PM for : Table, Operating | (TC 09/27/2023 07:58:57 AM, Reid, Emmerson ) PM Completed by Hillrom, service report attached.
1058249 | 197386 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-03-15 15:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1061411 | 197387 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-09-03 08:58 | PM-MAJOR-Annual PM For: Med Surg Bed
1062031 | 197388 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062095 | 197389 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-10-02 07:25 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 2nd year PM after 1st year PM last year. No GS500. Replaced PS500 batteries, V battery and air filter. Calibrated. Tested. Functional. SW.2.51 Leakage: 82uA and 90 ohm.
1062096 | 197390 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-09-22 11:32 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 2nd year PM after 1st year PM last year. No GS500. Replaced PS500 batteries, V battery and air filter. Calibrated. Tested. Functional. SW.2.51 Leakage: 83uA and 87 ohm.
1062097 | 197391 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-10-04 14:44 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 2nd year PM after 1st year PM last year. No GS500. Replaced PS500 batteries, V battery and air filter. Calibrated. Tested. Functional. SW.2.51 Leakage: 81uA and 93 ohm.
1062098 | 197392 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-10-27 12:50 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 2nd year PM after 1st year PM last year. No GS500. Replaced PS500 batteries, V battery and air filter. Calibrated. Tested. Functional. SW.2.51 Leakage: 81uA and 95 ohm.
1062102 | 197393 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Pressure-Vacuum Digital Meter
1062118 | 197394 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2023-12-12 09:05 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 12/07/2023 09:05:01 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1062120 | 197395 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-03-15 15:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 03/15/2024 03:36:52 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1062122 | 197396 | PM | CMPLT | CMPLT | 2023-09-01 03:29 | 2024-01-19 09:54 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/18/2024 09:53:40 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1062124 | 197397 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062125 | 197398 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062126 | 197399 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-09-16 09:54 | PM-MAJOR-Annual PM For: Med Surg Bed
1062127 | 197400 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062128 | 197401 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-11-07 10:37 | PM-MAJOR-Annual PM For: Med Surg Bed
1062129 | 197402 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062130 | 197403 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062131 | 197404 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062132 | 197405 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062133 | 197406 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062134 | 197407 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062135 | 197408 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062136 | 197409 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-11-07 11:08 | PM-MAJOR-Annual PM For: Med Surg Bed
1062137 | 197410 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062138 | 197411 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062139 | 197412 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062140 | 197413 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062141 | 197414 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-09-03 09:45 | PM-MAJOR-Annual PM For: Med Surg Bed
1062142 | 197415 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062143 | 197416 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062144 | 197417 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-11-21 14:00 | PM-MAJOR-Annual PM For: Med Surg Bed
1062145 | 197418 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062146 | 197419 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062147 | 197420 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062168 | 197421 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-11-07 11:25 | PM-MAJOR-Annual PM For: Med Surg Bed
1062169 | 197422 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-09-13 09:10 | PM-MAJOR-Annual PM For: Med Surg Bed
1062170 | 197423 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-11-07 10:33 | PM-MAJOR-Annual PM For: Med Surg Bed
1062171 | 197424 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-09-27 08:37 | PM-MAJOR-Annual PM For: Med Surg Bed
1062172 | 197425 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062173 | 197426 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062174 | 197427 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-11-07 13:27 | PM-MAJOR-Annual PM For: Med Surg Bed
1062175 | 197428 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062176 | 197429 | PM | ACTIV | READY | 2023-09-01 03:29 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062177 | 197430 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-11-07 11:11 | PM-MAJOR-Annual PM For: Med Surg Bed
1062178 | 197431 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-09-27 08:56 | PM-MAJOR-Annual PM For: Med Surg Bed
1062179 | 197432 | PM | CANCL | CANCL | 2023-09-01 03:29 | 2024-09-25 08:24 | PM-MAJOR-Annual PM For: Med Surg Bed
1062191 | 197433 | PM | CANCL | CANCL | 2023-09-01 03:29 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062192 | 197434 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062193 | 197435 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062194 | 197436 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062195 | 197437 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062196 | 197438 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062197 | 197439 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062198 | 197440 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062199 | 197441 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062200 | 197442 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062201 | 197443 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062202 | 197444 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062203 | 197445 | PM | CANCL | CANCL | 2023-09-01 03:30 |  | PM-MAJOR - ANNUAL PM FOR: STRETCHER
1062232 | 197446 | PM | ACTIV | ISSUE | 2023-09-01 03:30 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1062233 | 197447 | PM | ACTIV | ISSUE | 2023-09-01 03:30 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1062235 | 197448 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 11:23 | PM-MAJOR-Annual PM For: Med Surg Bed
1062236 | 197449 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062237 | 197450 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-27 11:34 | PM-MAJOR-Annual PM For: Med Surg Bed
1062238 | 197451 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-27 10:55 | PM-MAJOR-Annual PM For: Med Surg Bed
1062239 | 197452 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-18 13:18 | PM-MAJOR-Annual PM For: Med Surg Bed
1062240 | 197453 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-25 08:12 | PM-MAJOR-Annual PM For: Med Surg Bed
1062241 | 197454 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062242 | 197455 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 11:06 | PM-MAJOR-Annual PM For: Med Surg Bed
1062243 | 197456 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-25 08:26 | PM-MAJOR-Annual PM For: Med Surg Bed
1062244 | 197457 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 09:00 | PM-MAJOR-Annual PM For: Med Surg Bed
1062245 | 197458 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 10:36 | PM-MAJOR-Annual PM For: Med Surg Bed
1062246 | 197459 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062247 | 197460 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062248 | 197461 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-18 13:25 | PM-MAJOR-Annual PM For: Med Surg Bed
1062253 | 197462 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062254 | 197463 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062255 | 197464 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-13 09:07 | PM-MAJOR-Annual PM For: Med Surg Bed
1062256 | 197465 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-28 15:55 | PM-MAJOR-Annual PM For: Med Surg Bed
1062257 | 197466 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-27 11:55 | PM-MAJOR-Annual PM For: Med Surg Bed
1062258 | 197467 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-18 14:11 | PM-MAJOR-Annual PM For: Med Surg Bed
1062259 | 197468 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-25 08:33 | PM-MAJOR-Annual PM For: Med Surg Bed
1062260 | 197469 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-25 08:16 | PM-MAJOR-Annual PM For: Med Surg Bed
1062261 | 197470 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 09:06 | PM-MAJOR-Annual PM For: Med Surg Bed
1062262 | 197471 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-27 11:32 | PM-MAJOR-Annual PM For: Med Surg Bed
1062263 | 197472 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 09:59 | PM-MAJOR-Annual PM For: Med Surg Bed
1062264 | 197473 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-27 11:44 | PM-MAJOR-Annual PM For: Med Surg Bed
1062265 | 197474 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-13 14:46 | PM-MAJOR-Annual PM For: Med Surg Bed
1062266 | 197475 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062269 | 197476 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-21 12:20 | PM-MAJOR-Annual PM For: Med Surg Bed
1062270 | 197477 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062271 | 197478 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-22 08:53 | PM-MAJOR-Annual PM For: Med Surg Bed
1062272 | 197479 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 11:31 | PM-MAJOR-Annual PM For: Med Surg Bed
1062273 | 197480 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062274 | 197481 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-18 13:11 | PM-MAJOR-Annual PM For: Med Surg Bed
1062275 | 197482 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062276 | 197483 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062277 | 197484 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 11:05 | PM-MAJOR-Annual PM For: Med Surg Bed
1062278 | 197485 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062279 | 197486 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062280 | 197487 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-25 08:18 | PM-MAJOR-Annual PM For: Med Surg Bed
1062281 | 197488 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062282 | 197489 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 09:35 | PM-MAJOR-Annual PM For: Med Surg Bed
1062309 | 197490 | PM | CMPLT | CMPLT | 2023-09-01 03:30 | 2023-11-29 10:39 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 11/21/2023 03:06:58 PM, Canjimavadivel, Kamallesh ) Placed order for PM Kit (TC 11/28/2023 02:08:27 PM, Canjimavadivel, Kamallesh ) Replaced PM kit parts Performed all calibration and system tests Pre-op check pass Reconditioning battery (TC 11/28/2023 02:20:30 PM, Cortez, John) Helped with PM Electrical Safety: 22 uA (TC 11/29/2023 10:37:00 AM, Cortez, John) Swapped in conditioned battery: Manufactured date 10/05/20 6579 mAh 22 cycles 1% error 42.8% SOH
1062312 | 197491 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062313 | 197492 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 08:16 | PM-MAJOR-Annual PM For: Med Surg Bed
1062314 | 197493 | PM | CMPLT | CMPLT | 2023-09-01 03:30 | 2024-08-27 11:47 | PM-MAJOR-Annual PM For: Med Surg Bed
1062315 | 197494 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 13:57 | PM-MAJOR-Annual PM For: Med Surg Bed
1062316 | 197495 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-13 14:11 | PM-MAJOR-Annual PM For: Med Surg Bed
1062317 | 197496 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062318 | 197497 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062319 | 197498 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062320 | 197499 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-16 08:39 | PM-MAJOR-Annual PM For: Med Surg Bed
1062321 | 197500 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-16 08:43 | PM-MAJOR-Annual PM For: Med Surg Bed
1062322 | 197501 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-27 10:13 | PM-MAJOR-Annual PM For: Med Surg Bed
1062323 | 197502 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-08-27 10:44 | PM-MAJOR-Annual PM For: Med Surg Bed
1062324 | 197503 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 10:48 | PM-MAJOR-Annual PM For: Med Surg Bed
1062325 | 197504 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 09:10 | PM-MAJOR-Annual PM For: Med Surg Bed
1062388 | 197505 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 10:44 | PM-MAJOR-Annual PM For: Med Surg Bed
1062389 | 197506 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-18 13:22 | PM-MAJOR-Annual PM For: Med Surg Bed
1062390 | 197507 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-22 09:18 | PM-MAJOR-Annual PM For: Med Surg Bed
1062391 | 197508 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062392 | 197509 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 08:48 | PM-MAJOR-Annual PM For: Med Surg Bed
1062393 | 197510 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 10:30 | PM-MAJOR-Annual PM For: Med Surg Bed
1062394 | 197511 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 09:31 | PM-MAJOR-Annual PM For: Med Surg Bed
1062395 | 197512 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 13:55 | PM-MAJOR-Annual PM For: Med Surg Bed
1062396 | 197513 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-25 08:28 | PM-MAJOR-Annual PM For: Med Surg Bed
1062397 | 197514 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062398 | 197515 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062399 | 197516 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062400 | 197517 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062401 | 197518 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062435 | 197519 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062436 | 197520 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062437 | 197521 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062438 | 197522 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062439 | 197523 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062440 | 197524 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062441 | 197525 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062442 | 197526 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062443 | 197527 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062444 | 197528 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 11:27 | PM-MAJOR-Annual PM For: Med Surg Bed
1062445 | 197529 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062446 | 197530 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-03 08:36 | PM-MAJOR-Annual PM For: Med Surg Bed
1062447 | 197531 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062448 | 197532 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062543 | 197533 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062544 | 197534 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062545 | 197535 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062546 | 197536 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-11-07 14:17 | PM-MAJOR-Annual PM For: Med Surg Bed
1062547 | 197537 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062548 | 197538 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-13 09:04 | PM-MAJOR-Annual PM For: Med Surg Bed
1062549 | 197539 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-25 08:31 | PM-MAJOR-Annual PM For: Med Surg Bed
1062550 | 197540 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062551 | 197541 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062552 | 197542 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062553 | 197543 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062554 | 197544 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062555 | 197545 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062556 | 197546 | PM | CMPLT | CMPLT | 2023-09-01 03:30 | 2024-08-27 09:20 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062557 | 197547 | PM | CANCL | CANCL | 2023-09-01 03:30 | 2024-09-05 13:29 | PM-MAJOR-Annual PM For: Med Surg Bed
1062558 | 197548 | PM | ACTIV | READY | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062604 | 197549 | PM | ACTIV | SHP | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM for Bed electric
1062605 | 197550 | PM | ACTIV | SHP | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM for Bed electric
1062606 | 197551 | PM | ACTIV | SHP | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM for Bed electric
1062607 | 197552 | PM | ACTIV | SHP | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM for Bed electric
1062608 | 197553 | PM | ACTIV | SHP | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM for Bed electric
1062609 | 197554 | PM | ACTIV | SHP | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM for Bed electric
1062610 | 197555 | PM | ACTIV | SHP | 2023-09-01 03:30 |  | PM-MAJOR-Annual PM for Bed electric
1062655 | 197556 | PM | ACTIV | SHP | 2023-09-01 03:31 |  | PM-MAJOR-Annual PM for Bed electric
1062657 | 197557 | PM | ACTIV | SHP | 2023-09-01 03:31 |  | PM-MAJOR-Annual PM for Bed electric
1062658 | 197558 | PM | ACTIV | SHP | 2023-09-01 03:31 |  | PM-MAJOR-Annual PM for Bed electric
1063193 | 197559 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2023-10-02 17:06 | MSH 5 RPCU BONNIE PM - Semi-Annual | (TC 10/02/2023 05:06:23 PM, Whittington, Scott) MSH 5 RPCU BONNIE PM - Semi-Annual | MSH 5 RPCU BONNIE PM - Semi-Annual
1064964 | 197560 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2023-10-13 16:24 | HARP | HARP
1064964 | 197561 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2023-10-12 16:25 | PM - Semi-Annual | PM
1064973 | 197562 | PM | ACTIV | ISSUE | 2023-09-01 03:31 |  | Annual PM
1065040 | 197563 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-23 10:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/23/2024 10:10:29 AM, Rhuland, Courtney) PM completed by BD
1065045 | 197564 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-23 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/23/2024 10:08:05 AM, Rhuland, Courtney) PM completed by BD
1065046 | 197565 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-03 14:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:25:24 PM, Rhuland, Courtney) PM completed by BD
1065047 | 197566 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-05 14:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:52:39 PM, Rhuland, Courtney) PM completed by BD
1065048 | 197567 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-03 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:23:48 PM, Rhuland, Courtney) PM completed by BD
1065049 | 197568 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-26 08:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065050 | 197569 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-26 08:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065051 | 197570 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-26 08:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065052 | 197571 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-05 14:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:51:30 PM, Rhuland, Courtney) PM completed by BD
1065053 | 197572 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-03 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:21:25 PM, Rhuland, Courtney) PM completed by BD
1065054 | 197573 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-03 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:22:45 PM, Rhuland, Courtney) PM completed by BD
1065055 | 197574 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-23 10:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/23/2024 10:11:55 AM, Rhuland, Courtney) PM completed by BD
1065056 | 197575 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-05 14:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:50:15 PM, Rhuland, Courtney) PM completed by BD
1065057 | 197576 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2024-04-05 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:49:10 PM, Rhuland, Courtney) PM completed by BD
1065058 | 197577 | PM | CMPLT | CMPLT | 2023-09-01 03:31 | 2023-11-14 12:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/14/2023 12:27:57 PM, Campbell, Connor) PM passed/complete, refer to ASM refer to procedure for electrical safety battery conditioned
1046035 | 197578 | CM | CMPLT | CMPLT | 2023-09-01 07:19 | 2023-09-03 18:01 | MSH ER DRX EVO E5-02(ER) PT APPEARS ON JDMI PC SCHEDULED , DOES NOT SHOW UP ON DRX EVO WORKLIST, AND NOT ON TUBE HEAD,ENDS UP AS UNASSIGNED IMAGE | (TC 09/01/2023 07:18:57 AM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) PT APPEARS O JDMI PC SCHEDULED , DOES NOT SHOW UP ON DRX EVO WORKLIST, AND NOT ON TUBE HEAD,ENDS UP AS UNASSIGNED IMAGE (TC 09/03/2023 05:18:00 PM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) PT APPEARS O JDMI PC SCHEDULED , DOES NOT SHOW UP ON DRX EVO WORKLIST, AND NOT ON TUBE HEAD,ENDS UP AS UNASSIGNED IMAGE CALLED IN SEPT 1 AND SEPT 3 FOLLOWED SEANS INSTRUCTIONS TO CLEAR ERRORS FOUND IT WAS EXCESSIVE NUMBER OF ERROR LOG FILES ON DRIVE C, WHICH SEAN WAS ABLE TO REMOVE | MSH ER DRX EVO E5-02(ER) PT APPEARS O JDMI PC SCHEDULED , DOES NOT SHOW UP ON DRX EVO WORKLIST, AND NOT ON TUBE HEAD,ENDS UP AS UNASSIGNED IMAGE FIRST TIME IT HAS SHOWN UP ON DRX EVO WORKLIST, BUT NOT ON TUBEHEAD DISPLAY FREQUENT ISSUES WITH STUDY ON JDMIPACS PC MWL, BUT NOT ON DRX EVO SCHEDULED PT SCREEN
| 197579 | CM | CMPLT | CMPLT | 2023-09-01 07:57 | 2023-09-06 14:45 | Worked on the GE CASE to see why database not connected | (TC 08/31/2023 07:57:15 AM, Samanian, Kamran) resolved the supervisor issue and looking into depth files and configuration (TC 09/07/2023 02:44:17 PM, Samanian, Kamran) stayed late with GE tech and checked everything. finally resolved the issue ( downtime-start batch file had to run under administrator ) , Cardiosoft conncted to the server and was able to pull out all patient data (TC 09/06/2023 02:44:17 PM, Samanian, Kamran) stayed late with GE tech and checked everything. finally resolved the issue ( downtime-start batch file had to run under administrator ) , Cardiosoft conncted to the server and was able to pull out all patient data
69471 | 197580 | PM | CMPLT | CMPLT | 2023-09-01 08:27 | 2023-09-01 08:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/01/2023 08:26:49 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
62988 | 197581 | CM | CMPLT | CMPLT | 2023-09-01 09:37 | 2023-09-01 09:39 | Timer replacement and FC | (TC 09/01/2023 09:38:35 AM, Baghirzade, Nara) Replaced timer. Replaced all 6 rubber feet. All functional tests passed.
1063808 | 197582 | CM | CMPLT | CMPLT | 2023-09-01 10:00 | 2023-09-01 10:00 | MSH 15 SPP PROCEDURE ROOMUltrasound GE VOLUSON 10 HDD FILLING UP , DELETE CASES | (TC 09/01/2023 10:00:12 AM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOMUltrasound GE VOLUSON 10 HDD FILLING UP , DELETE CASES | MSH 15 SPP PROCEDURE ROOMUltrasound GE VOLUSON 10 HDD FILLING UP , DELETE CASES
68651 | 197583 | CM | CMPLT | CMPLT | 2023-09-01 10:40 | 2023-09-18 10:38 | Not working properly | (TC 08/29/2023 10:40:22 AM, Rendon Lopez , Elizabeth) Downloaded error logs, bottle pressure sensor errors. Waiting for part (TC 09/15/2023 10:37:49 AM, Rendon Lopez , Elizabeth) Replaced pressure sensor, calibrate rate-OK Door latch not being held. Replaced front door and LCD.
1057656 | 197584 | CM | ACTIV | READY | 2023-09-01 10:50 |  | Rebooted FP server and crashed | (TC 08/31/2023 10:50:14 AM, Samanian, Kamran) verified few things. it seems there was a patch and after reboot it made the server crashed. Opened the case to philips# 0122216352. Needs media to be sent and reloaded. (TC 09/12/2023 05:31:12 PM, Samanian, Kamran) Received the media with Windows 2019 server. removed BIOS USB from mother board, started installation (TC 09/14/2023 01:58:09 PM, Samanian, Kamran) recievd the media windows server 2019 and tried to install it but kept failing actication, waiting to receive the license (TC 10/12/2023 11:34:48 AM, Samanian, Kamran) Received the new media with license. Installed windows server 2019 ( philips USB is still in motherboard) and activated windows and set up proxy. did let Anthony know to do the rest, install FP, certifictes and config (TC 10/13/2023 11:36:22 AM, Samanian, Kamran) worked with Anthony for certificate installation issue
69494 | 197585 | CM | CMPLT | CMPLT | 2023-09-01 10:55 | 2023-09-01 10:57 | Alarming error | (TC 08/29/2023 10:54:47 AM, Rendon Lopez , Elizabeth) Error 242.4030 - Motor encoder error, faulty AIL assembly. Replaced AIL assy, calibrated pump, tested-OK
1055179 | 197586 | CM | CMPLT | CMPLT | 2023-09-01 10:58 | 2023-09-01 11:00 | Door latch broken | (TC 08/29/2023 10:58:52 AM, Rendon Lopez , Elizabeth) Latch not holding up. Replaced front door and LCD under recall.
69530 | 197587 | CM | CMPLT | CMPLT | 2023-09-01 11:02 | 2023-09-01 11:03 | Broken | (TC 08/29/2023 11:02:25 AM, Rendon Lopez , Elizabeth) Downloaded error logs, error 232.6040- Display failure. Replaced LCD under recall. Run on user mode-OK
1059984 | 197588 | PM | CMPLT | CMPLT | 2023-09-01 11:39 | 2023-09-01 11:44 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 08/30/2023 11:42:43 AM, Canjimavadivel, Kamallesh ) Bed+Intellidrive batteries were replaced 60 mOhms, 22 uAmps calibrated position sensors attached ground to frame PM passed
| 197589 | CM | CMPLT | CMPLT | 2023-09-01 11:53 | 2023-10-03 15:47 | misc biomed - sep 2023 | (TC 09/01/2023 11:52:45 AM, Cortez, John) Booked flight/hotel for Draeger training (TC 09/06/2023 01:51:28 PM, Cortez, John) Philips competency for techs prep (TC 09/06/2023 01:53:20 PM, Cortez, John) Tested ECG leads sent down to biomed from ICU. No fault found, will keep set as a spare (TC 09/07/2023 08:15:42 AM, Cortez, John) Met with tech from Laerdal who's onsite to install Sim Man in ICU (TC 09/07/2023 08:18:29 AM, Cortez, John) Helped Kamran go to TGH and salvage parts from treadmill (TC 09/08/2023 01:26:32 PM, Cortez, John) 1 of the 4 data jacks meant for a private Laerdal VLAN in the ICU sim room, isn't configured correctly. Reached out to Mirza from IS who eventually resolved issue (TC 09/11/2023 03:44:25 PM, Cortez, John) Reset power adapter that charges wall mounted ipad(controls OR lights) in OR 14 (TC 09/11/2023 03:47:00 PM, Cortez, John) huddle (TC 09/11/2023 03:47:16 PM, Cortez, John) Moved shipment of vital signs monitors to ED east (TC 09/12/2023 03:49:29 PM, Cortez, John) Biomed Masimo viewstation not displaying sectors. Had to change layout so that the application can view up to 40 devices (TC 09/13/2023 03:51:04 PM, Cortez, John) Helped Nara with configuring network settings on new Masimo Root (TC 09/14/2023 01:43:59 PM, Cortez, John) Prepared and tested scenario for philips competency test 2023 (TC 09/15/2023 03:56:46 PM, Cortez, John) IS meeting (TC 09/18/2023 10:09:07 AM, Cortez, John) huddle (TC 09/18/2023 10:09:41 AM, Cortez, John) Masimo has an alarm "client down 16N1" 16N1 refers to the 15 Murray viewstation (TC 09/19/2023 03:55:26 PM, Cortez, John) Helped setup and oversee Philips competency test with colleagues (TC 09/20/2023 02:57:02 PM, Cortez, John) Amanda from ICU asked about status of Philips web overview in sim room observation room. Becky mentioned theyre still waiting for a quote. Also needed to swap ethernet cable in observation room with a longer one (TC 09/19/2023 03:55:26 PM, Cortez, John) Helped setup and oversee Philips competency test with colleagues (TC 09/20/2023 03:00:34 PM, Cortez, John) Philips competency training day 2 (TC 09/20/2023 03:01:24 PM, Cortez, John) CE meeting (TC 09/20/2023 04:05:55 PM, Cortez, John) Sim Man walkthrough with Laerdal rep (TC 09/21/2023 03:51:35 PM, Cortez, John) Philips competency training day 3 (TC 09/21/2023 03:52:45 PM, Cortez, John) Discussed Masimo Radical battery recall notice with Samir (TC 09/21/2023 03:53:30 PM, Cortez, John) Looked at Stryker dashboard with Samir, went on 10th floor to power cycle beds, exported document of procuity errors (TC 09/26/2023 03:53:48 PM, Cortez, John) systems committee meeting (TC 09/26/2023 04:07:37 PM, Cortez, John) create systems committee update slides for general staff meeting (TC 09/27/2023 03:59:29 PM, Cortez, John) dept meeting (TC 09/27/2023 03:59:56 PM, Cortez, John) dropped off training aed to Vincy (10S) (TC 09/27/2023 04:04:49 PM, Cortez, John) Philips competency testing day 4 (TC 09/29/2023 03:38:11 PM, Cortez, John) pm meeting (TC 09/29/2023 03:38:37 PM, Cortez, John) philips competency test day 5 (TC 09/28/2023 04:01:40 PM, Cortez, John) Healped Kamran find missing switch A3U10NE on 10th floor. Turns out switch was renamed, which was why we couldn't locate it at first
1054927 | 197590 | CM | CMPLT | CMPLT | 2023-09-01 11:59 | 2023-09-01 12:08 | "Thermometer not working" | (TC 08/29/2023 11:59:22 AM, Cortez, John) Software upgraded to 2.1.49 External fan filter installed Probe well filled with debris, replaced Roll stand screws secured with loctite All parameters tested ok
| 197591 | RDS | CMPLT | CMPLT | 2023-09-01 12:07 | 2023-09-01 12:10 | Weekly Rounds: NICU | (TC 09/01/2023 12:07:15 PM, Baghirzade, Nara) Completed weekly rounds in NICU
| 197592 | RDS | CMPLT | CMPLT | 2023-09-01 12:11 | 2023-09-01 12:12 | Weekly Rounds: RESUS | (TC 09/01/2023 12:11:42 PM, Baghirzade, Nara) Completed weekly rounds in RESUS. See procedure.
| 197593 | RDS | CMPLT | CMPLT | 2023-09-01 12:13 | 2023-09-01 12:15 | Weekly Rounds: 18S and PACU POD A | (TC 09/01/2023 12:13:50 PM, Baghirzade, Nara) Weekly rounds in 18S and PACU POD A completed. See procedure.
| 197594 | RDS | CMPLT | CMPLT | 2023-09-01 12:15 | 2023-09-01 12:16 | Weekly Rounds: 11 South | (TC 09/01/2023 12:15:31 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure.
1061211 | 197595 | FC | CMPLT | CMPLT | 2023-09-01 13:18 | 2023-09-01 13:19 | Annual Inspection | (TC 09/01/2023 01:18:09 PM, Reid, Emmerson ) Checked the monitors and tightened one of them. Checked the booms and adjusted the brakes to reduce drift on side boom. iPad was fully charged. Inspection complete.
1059756 | 197596 | CM | CMPLT | CMPLT | 2023-09-01 13:24 | 2023-09-01 13:25 | Printer not working | (TC 09/01/2023 01:24:15 PM, Reid, Emmerson ) Went to OPG and connected the lap top to the printer. The wrong printer was selected, so i removed that one from the printers.
501106 | 197597 | CM | CMPLT | CMPLT | 2023-09-01 16:50 | 2023-09-11 11:13 | MSH 5 US IU22 TCS NOT RESPONDING, SITE 553758 | (TC 09/05/2023 11:38:38 AM, Fan, Tak) Cody wanted to performed calibration. Removed cover, inspected touchscreen, damaged apparent on multiple areas off th screen. Sent PN to Cody, he'll order a replacement. (TC 09/06/2023 03:45:00 PM, Fan, Tak) TSC sent to sinai by accident. Replaced TSC, tested, OK. | MSH 5 US IU22 TCS NOT RESPONDING, SITE 553758 REF# 0122223730 RHONDA
1056602 | 197598 | CM | CMPLT | CMPLT | 2023-09-01 22:08 | 2023-09-01 22:08 | Probe PLT1204BX transferred to Womens College Hospital. | (TC 09/01/2023 10:08:37 PM, Nambiar, Chandra) Probe was taken from TGH probe inventory, and tested and transferred to Womens College Hospital.
| 197599 | RDS | CMPLT | CMPLT | 2023-09-05 07:49 | 2023-09-05 08:30 | Weekly check for ED
| 197600 | RDS | CMPLT | CMPLT | 2023-09-05 07:50 | 2023-09-07 13:50 | Weekly check for ICU
1058135 | 197601 | CM | CMPLT | CMPLT | 2023-09-05 08:24 | 2023-09-05 09:07 | Not turning on | (TC 09/05/2023 09:04:31 AM, Canjimavadivel, Kamallesh ) Replaced batteries and cleared logs
| 197602 | RDS | CMPLT | CMPLT | 2023-09-05 09:09 | 2023-09-05 09:14 | Weekly Checks For PACU | (TC 09/05/2023 09:12:53 AM, Reid, Emmerson ) Weekly check complete.
1061219 | 197603 | CM | CMPLT | CMPLT | 2023-09-05 09:17 | 2023-09-05 09:18 | Fluro monitor not working | (TC 09/05/2023 09:16:51 AM, Reid, Emmerson ) Went to check it out but it was a JDMI issue. I connected with them to resolve it.
| 197604 | TEDU | CMPLT | CMPLT | 2023-09-05 09:25 | 2023-09-12 09:02 | stock parts for OR X3s | (TC 09/05/2023 09:26:15 AM, Scero, Tomlyn) Ordered new X3 bezels for OR X3s for stock
| 197605 | CM | CMPLT | CMPLT | 2023-09-05 10:03 | 2023-11-27 11:06 | Order for Alaris PCU straps, biomed stock | (TC 09/05/2023 10:03:30 AM, Cortez, John) Ordered straps (TC 11/27/2023 10:56:02 AM, Cortez, John) received order
65890 | 197606 | CM | CMPLT | NFF | 2023-09-05 10:29 | 2023-09-05 10:30 | Channel error; not specified | (TC 09/05/2023 10:29:21 AM, Baghirzade, Nara) Downloaded error log and noted one fatal error; error was pertaining to the syringe nut not engaging. Check-in tests passed and completed. IUI connectors and device cleaned. Device is fully functional. No fault found. Ran a basic infusion; no errors or issues noted. Infusion completed successfully.
68968 | 197607 | CM | CMPLT | CMPLT | 2023-09-05 10:30 | 2023-09-05 10:48 | needs strap | (TC 09/05/2023 10:47:03 AM, Cortez, John) added strap
66692 | 197608 | CM | CMPLT | CMPLT | 2023-09-05 10:31 | 2023-09-05 10:32 | SE latch broken | (TC 09/05/2023 10:31:06 AM, Baghirzade, Nara) Replaced SE latch. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
68034 | 197609 | CM | CMPLT | CMPLT | 2023-09-05 10:35 | 2023-09-05 10:38 | Disassembly/disposal of old hood | (TC 09/05/2023 10:36:58 AM, Baghirzade, Nara) Disassembled parts off the old hood. Disposed of old hood. Placed new hood. Placed all parts back on the new hood. Replaced 2 grommets.
68034 | 197610 | CM | CMPLT | CMPLT | 2023-09-05 10:35 | 2023-09-05 10:36 | Latch recall | (TC 09/05/2023 10:35:36 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall. NW/NE latches part of new hood.
66433 | 197611 | CM | CMPLT | CMPLT | 2023-09-05 11:14 | 2023-09-05 13:13 | switch label assignments and lock sector | (TC 09/05/2023 01:12:07 PM, Scero, Tomlyn) I switched the tele3 label assignment back to the proper tele3 segment on the central station. I also locked the tele3 bed to the segment because it kept being removed.
66431 | 197612 | CM | CMPLT | CMPLT | 2023-09-05 11:14 | 2023-09-05 13:16 | Switch label assignment | (TC 09/05/2023 01:15:03 PM, Scero, Tomlyn) Changed the assigned segment label back to tele1 to avoid confusion
1053728 | 197613 | FC | CMPLT | CMPLT | 2023-09-05 11:34 | 2023-09-05 11:41 | Functional test after use | (TC 09/05/2023 11:40:12 AM, Scero, Tomlyn) functional check complete and passed.
| 197614 | CM | CMPLT | CMPLT | 2023-09-05 13:08 | 2023-09-01 13:07 | O.R # 16 no video display | (TC 09/01/2023 01:08:01 PM, Chukri, Samir) reset the VMC and tested working fine, the room camera not working informed Emmerson to check.
| 197615 | RDS | CMPLT | CMPLT | 2023-09-05 13:18 | 2023-09-05 13:20 | Weekly checks for CCU | (TC 09/05/2023 01:19:44 PM, Scero, Tomlyn) Weekly checks complete
1049608 | 197616 | CM | CMPLT | CMPLT | 2023-09-05 13:19 | 2023-09-05 13:20 | OPG SPP - No Network Connection | (TC 09/05/2023 01:19:24 PM, Chau, Ken) Contact Mirza to check network and refresh -once IT completed network refresh system connected and tested okay
52411 | 197617 | CM | CMPLT | CMPLT | 2023-09-05 14:25 | 2023-09-07 13:51 | Knobs are sticking | (TC 09/05/2023 08:31:50 AM, Canjimavadivel, Kamallesh ) Cleaned out some of the knobs and parts moving smoothly for now. Looked for manual to see disassembly. (TC 09/06/2023 08:38:09 AM, Canjimavadivel, Kamallesh ) Emailed unit detailing work that was done yesterdat | Reported multiple knobs were not working properly. Red LED not working
1066462 | 197618 | FC | CMPLT | CMPLT | 2023-09-05 14:46 | 2023-09-05 14:46 | QC checklist | (TC 09/05/2023 02:45:43 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
| 197619 | CM | CMPLT | CMPLT | 2023-09-05 17:38 | 2023-09-05 17:38 | Site Coverage
69006 | 197620 | CM | CMPLT | CMPLT | 2023-09-06 07:03 | 2023-09-06 07:04 | Reported that the unit keeps alarming | (TC 09/06/2023 07:02:51 AM, Baghirzade, Nara) Downloaded error log - noted error pertaining to low battery levels. Checked battery status - battery is charging and within capacity range. No fault with device, needs to be charged. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 65 mohms Leakage current: 74.4 uA
1047844 | 197621 | CM | CMPLT | CMPLT | 2023-09-06 07:58 | 2023-09-06 08:01 | Error audible alarm bgnd test failed | (TC 09/06/2023 07:58:19 AM, Baghirzade, Nara) Checked speaker; speaker is functional from level 1 to 5. Checked the wiring; connectors and wiring is in tact and not loose. Ran 2 separate feeds and both completed successfully with no errors or issues. Occlusion test and alarm passed. Switched the device off and on multiple times but the error message did not reappear. All functional tests passed.
1051031 | 197622 | CM | CMPLT | CMPLT | 2023-09-06 08:32 | 2023-09-06 08:34 | Cannot view images on workstation from system | (TC 09/05/2023 04:32:00 PM, Fan, Tak) Checked workstation. System's C drive is completely full. Checked for temporary files or anything that could be cleared up. A large portion of the Windows folder is using up 35% of memory. Cannot check without admin rights. Faith Zhu will contact IT.
65612 | 197623 | CM | CMPLT | CMPLT | 2023-09-06 09:07 | 2023-09-06 09:50 | Not holding charge | (TC 09/06/2023 09:48:17 AM, Canjimavadivel, Kamallesh ) Replaced battery Replaced mounting plate All functions pass
1067763 | 197624 | ACC | CMPLT | CMPLT | 2023-09-06 09:36 | 2023-09-06 11:49 | Incoming Inspection | (TC 09/06/2023 11:45:59 AM, Scero, Tomlyn) Unpacked all of the boxes and assembled the ECG cart. Configured hospital network settings and loaded configuration settings. Functional tests complete and passed. Gave unit to 12 Oncology
| 197625 | CM | CMPLT | CMPLT | 2023-09-06 09:39 | 2023-09-01 09:40 | moving equipment from OR 9 back to Or10 | (TC 09/01/2023 09:40:34 AM, Samanian, Kamran) done!
67930 | 197626 | CM | CMPLT | CMPLT | 2023-09-06 10:20 | 2023-09-06 10:25 | Noisy fan | (TC 09/06/2023 10:23:14 AM, Campbell, Connor) noisy fan replaced fan motor greased tilt mechanism functional test humidity set 75, read 77 set temp 35, read 35 all passed worked with nara's supervision
1056665 | 197627 | CM | CMPLT | CMPLT | 2023-09-06 10:28 | 2023-09-06 10:29 | Handle came lose | (TC 09/06/2023 10:28:03 AM, Reid, Emmerson ) Handle came lose and I tightened it back up
| 197628 | CM | CMPLT | CMPLT | 2023-09-06 11:14 | 2023-09-06 11:14 | reviewed the PICiX Rev 4 Recall and responded | (TC 09/06/2023 11:14:07 AM, Samanian, Kamran) done!
1062414 | 197629 | CM | CMPLT | CMPLT | 2023-09-06 13:16 | 2023-09-06 13:16 | right side handle is rough and scratched | (TC 09/06/2023 01:16:01 PM, Samanian, Kamran) it seems patient's hand was hurt. checked the surface and filed it. much better than it looked like before,
| 197630 | CM | CMPLT | CMPLT | 2023-09-06 13:19 | 2023-09-06 13:19 | performed safety inspection on 3 + 1 Olympus Irrigation pump | (TC 09/06/2023 01:19:22 PM, Samanian, Kamran) all three passed and labeled. the last one did NOT pass and needs to be replaced.
66528 | 197631 | CM | CMPLT | CMPLT | 2023-09-06 13:32 | 2023-10-10 11:01 | Not associating | (TC 09/06/2023 09:33:53 AM, Rendon Lopez , Elizabeth) Checked device, calbe from dongle unplugged from CE, restarted, connecting no issues. Testing monitor, recorder issue, not feeding. Removed monitor from room and placed spare. (TC 09/06/2023 09:33:53 AM, Rendon Lopez , Elizabeth) Replaced recorder with new spare, ordered replacement (TC 10/05/2023 11:00:52 AM, Rendon Lopez , Elizabeth) Received replacement
| 197632 | CM | CMPLT | CMPLT | 2023-09-06 14:30 | 2024-02-07 17:28 | ordering three remotes for OR tables | (TC 12/01/2023 05:27:44 PM, Huang, Jessica) received them and installed them on the tables
1056470 | 197633 | CM | CMPLT | CMPLT | 2023-09-06 14:34 | 2023-09-06 14:34 | do not work | charged unit overnight. Tested BP, SpO2 and temp. Functional.
1057938 | 197634 | CM | CMPLT | CMPLT | 2023-09-06 14:38 | 2023-11-16 13:45 | ordering remotes for the OR table
1054419 | 197635 | CM | CMPLT | CMPLT | 2023-09-06 14:59 | 2023-10-25 10:47 | Not reading, physical damage | (TC 09/06/2023 02:58:58 PM, Rendon Lopez , Elizabeth) Confirmed issue, to be sent out for repair (TC 10/17/2023 10:46:24 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1053432 | 197636 | CM | CMPLT | CMPLT | 2023-09-06 15:05 | 2023-10-25 10:48 | Noisy | (TC 09/06/2023 03:05:56 PM, Rendon Lopez , Elizabeth) Confirmed issue, noisy cable. Sending out under warranty (TC 10/17/2023 10:47:52 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1056128 | 197637 | CM | CMPLT | CMPLT | 2023-09-06 15:07 | 2023-10-25 10:48 | Not working | (TC 09/06/2023 03:08:19 PM, Rendon Lopez , Elizabeth) Confirmed issue, static/ noisy cable Sendin out under warranty (TC 10/17/2023 10:47:17 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1057452 | 197638 | CM | CMPLT | CMPLT | 2023-09-06 15:09 | 2023-10-25 10:46 | Not working | (TC 09/06/2023 03:09:14 PM, Rendon Lopez , Elizabeth) Confirmed issue, static/ noisy cable Sendin out under warranty (TC 10/17/2023 10:45:34 AM, Rendon Lopez , Elizabeth) Repair was done, not under warranty. Incoming inspection-OK
1050752 | 197639 | CM | CMPLT | CMPLT | 2023-09-06 15:11 | 2023-09-06 15:12 | Not working | (TC 09/06/2023 03:11:20 PM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
67957 | 197640 | CM | CMPLT | CMPLT | 2023-09-06 15:11 | 2023-09-06 15:12 | System failure 25 | (TC 09/06/2023 03:11:22 PM, Baghirzade, Nara) Fan was not properly seated. Latch on SW was broken. Repaired latch and re-seated the fan. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C abd 76% RH. Incubator is fully functional.
1059742 | 197641 | CM | CMPLT | CMPLT | 2023-09-06 15:13 | 2023-09-06 15:14 | Blue light alarm | (TC 09/06/2023 03:13:02 PM, Rendon Lopez , Elizabeth) Changed device with spare. updated SN, charged, tested-OK
69064 | 197642 | PM | CMPLT | CMPLT | 2023-09-06 15:14 | 2023-09-06 15:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/06/2023 03:13:57 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1062279 | 197643 | CM | CMPLT | CMPLT | 2023-09-06 15:15 | 2023-09-26 09:10 | Battery not being detected by bed and scale requires service/calibration | (TC 09/06/2023 03:15:21 PM, Baghirzade, Nara) Battery is not being detected by the bed. Scale requires service. Initiated a service request with Stryker. (TC 09/26/2023 09:08:56 AM, Baghirzade, Nara) Bed was repaired by Stryker. Attached the service report to this WO. | Unable to use bed alarm as the scale requires service/calibration. Battery is not being detected by the bed. Tried to power cycle and turn the switch on/off but did not resolve the issue.
| 197644 | CM | CMPLT | CMPLT | 2023-09-06 15:55 | 2023-09-06 15:55 | NICU Omnibed lights meeting | (TC 09/06/2023 03:54:48 PM, Baghirzade, Nara) NICU meeting re: omnibed lights melting
1063369 | 197645 | CM | CMPLT | CMPLT | 2023-09-06 15:57 | 2023-09-06 15:58 | Tray fell to the bottom of the pole | (TC 09/06/2023 03:57:00 PM, Baghirzade, Nara) Screws were loose on the tray mount/holder. Tightened screws and lifted the tray to where Iliana wanted it.
1046035 | 197646 | CM | CMPLT | CMPLT | 2023-09-06 17:02 | 2023-09-06 17:02 | Images do not linking to patient name | (TC 09/06/2023 05:04:42 PM, Dolinar, John) Images do not linking to patient name. Drive C was full. Removed logs to create space. | Images do not linking to patient name. Drive C was full. Removed logs to create space.
| 197647 | CM | CMPLT | CMPLT | 2023-09-06 17:04 | 2023-09-06 17:04 | Site coverage
68783 | 197648 | PM | CMPLT | CMPLT | 2023-09-07 07:31 | 2023-09-07 07:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 07:30:53 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. (TC 09/07/2023 07:30:53 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 42 mohms Leakage current: 72.5 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
1058534 | 197649 | CM | CMPLT | NFF | 2023-09-07 08:21 | 2023-09-07 08:34 | Does not reflect bladder volume | (TC 09/07/2023 08:21:39 AM, Reid, Emmerson ) Tested on the phantom and it was reading ~124ml. I brought in another Prime Plus and it read the same bladder volume. Ran the self test and it passed.
| 197650 | CM | CMPLT | CMPLT | 2023-09-07 08:44 | 2023-09-07 08:53 | Glidescope cable won't attach to blade | (TC 09/06/2023 08:50:33 AM, Canjimavadivel, Kamallesh ) Cable and scope were brought down by Nigel. Found something obstructing the connection on the blade side connector. (TC 09/07/2023 08:52:21 AM, Canjimavadivel, Kamallesh ) Extracted broken plastic from connector. Gave cable and scope back to Nigel. Informed him to keep and eye out for cable with broken connector.
1065707 | 197651 | CM | CMPLT | CMPLT | 2023-09-07 08:53 | 2023-10-12 09:57 | keypad not responsive, keeps alarming | (TC 09/07/2023 08:57:06 AM, Cortez, John) Amanda sent it down yesterday and I confirmed issue. Requested RMA from BD (TC 09/20/2023 03:03:31 PM, Cortez, John) shipped device RMA 303116162 (TC 10/12/2023 09:56:12 AM, Cortez, John) Received unit back, completed check-in
69496 | 197652 | CM | CMPLT | CMPLT | 2023-09-07 09:00 | 2023-09-18 10:43 | Channel error | (TC 09/05/2023 09:00:24 AM, Rendon Lopez , Elizabeth) Downloaded error log, error 240.4150, faulty bottle pressure sensor. Waiting for part (TC 09/15/2023 10:42:25 AM, Rendon Lopez , Elizabeth) Replaced pressure sensor-OK
69311 | 197653 | CM | CMPLT | CMPLT | 2023-09-07 09:01 | 2023-09-18 10:45 | channel malfunction | (TC 09/05/2023 09:01:32 AM, Rendon Lopez , Elizabeth) Downloaded error log, error 240.4150, faulty bottle pressure sensor. Waiting for part (TC 09/15/2023 10:44:28 AM, Rendon Lopez , Elizabeth) Replaced pressure sensor-OK
69677 | 197654 | CM | CMPLT | CMPLT | 2023-09-07 09:24 | 2023-09-18 10:37 | Created by Failed PM WO # 193373 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/15/2023 10:36:07 AM, Rendon Lopez , Elizabeth) Replaced pressure sensor, calibrated rate. PM completed, all tests passed | (TC 09/05/2023 09:23:05 AM, Rendon Lopez , Elizabeth) Bottle pressure sensor malfunction
1062123 | 197655 | CM | CMPLT | NFF | 2023-09-07 09:32 | 2023-09-07 09:33 | No problem specified | (TC 09/06/2023 09:32:27 AM, Rendon Lopez , Elizabeth) Running on user mode, all functions okay, no alarms. Tested, NFF
| 197656 | CM | CMPLT | CMPLT | 2023-09-07 09:50 | 2023-09-14 12:20 | Working on implementing Expair on Citrix | (TC 09/07/2023 09:50:01 AM, Baghirzade, Nara) Working with Kamran to implement Expair on Citrix. Next step is to send Justin Fiege the expair folder with the executables on it to implement it on a test environment. (TC 09/14/2023 12:18:53 PM, Baghirzade, Nara) Drafted an instruction document for Justin Fiege on installing the expair software.
1061206 | 197657 | CM | CMPLT | CMPLT | 2023-09-07 10:50 | 2023-09-07 10:51 | Recorder not working | (TC 09/07/2023 10:50:35 AM, Reid, Emmerson ) Spoke with Olympus and they helped to troubleshoot. The NCARE needed to be rebooted twice before it started working, also VMC was rebooted but did not help. Working now.
1047865 | 197658 | CM | CMPLT | CMPLT | 2023-09-07 11:04 | 2023-09-07 11:04 | Date/time and config update | (TC 09/07/2023 11:03:56 AM, Baghirzade, Nara) Updated date and time. Updated configuration to the latest: March 2019. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
65305 | 197659 | PM | CMPLT | CMPLT | 2023-09-07 12:29 | 2023-09-07 12:31 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 09/07/2023 12:29:42 PM, Baghirzade, Nara) Replaced battery. Device test passed and completed. PM completed. (TC 09/07/2023 12:30:01 PM, Campbell, Connor) Replaced battery. Device test passed and completed. PM completed. (TC 09/07/2023 12:30:01 PM, Campbell, Connor) Replaced battery. Device test passed and completed. PM completed.
| 197660 | CM | CMPLT | CMPLT | 2023-09-07 12:32 | 2023-09-07 12:57 | Wall oto-ophthalmscope not functional | (TC 09/07/2023 12:32:01 PM, Baghirzade, Nara) When the handle is placed back on the wall unit, the otoscope remains on. Opened the unit and had swapped connectors around. One of the connectors and wiring is completely dead. Spoke to Sally and sent an email to Nelia regarding this and that they need to order a new unit. Left the handles and otoscope heads with Sally. Disposed of the unti in e-waste. (TC 09/07/2023 12:56:58 PM, Campbell, Connor) When the handle is placed back on the wall unit, the otoscope remains on. Opened the unit and had swapped connectors around. One of the connectors and wiring is completely dead. Disposed of the unti in e-waste. (TC 09/07/2023 12:56:58 PM, Campbell, Connor) When the handle is placed back on the wall unit, the otoscope remains on. Opened the unit and had swapped connectors around. One of the connectors and wiring is completely dead. Disposed of the unti in e-waste.
1063791 | 197661 | CM | CMPLT | CMPLT | 2023-09-07 13:23 | 2024-02-07 17:26 | needs a new battery charger holder | (TC 11/07/2023 01:43:43 PM, Huang, Jessica) got a call from OR desk and picked up the head light ordered the part (TC 09/07/2023 01:43:43 PM, Huang, Jessica) got a call from OR desk and picked up the head light ordered the part (TC 12/06/2023 05:26:16 PM, Huang, Jessica) followed up the order and received it replaced the part and sent it back to OR
1062926 | 197662 | CM | CMPLT | CMPLT | 2023-09-07 14:01 | 2023-11-16 13:43 | warranty repair
| 197663 | CM | ACTIV | ISSUE | 2023-09-07 14:06 |  | EREB - Parts for PM | (TC 09/07/2023 02:35:58 PM, Samanian, Kamran) S/C# 1003044599
1057835 | 197664 | CM | CMPLT | CMPLT | 2023-09-07 14:20 | 2024-02-07 17:26 | shipped it for warranty repair | (TC 09/08/2023 01:38:39 PM, Huang, Jessica) called to the room and saw the error code on the device. Contacted conmed and received the loaner (TC 01/15/2024 05:23:33 PM, Huang, Jessica) followed up with the vendor because it was taking to long got the unit back in Jan and sync with the smoke vac returned the loaner
| 197665 | CM | ACTIV | ISSUE | 2023-09-07 14:47 |  | Code Grey, Moved stuffs for ENDO to OR-PACU | (TC 09/07/2023 02:47:10 PM, Samanian, Kamran) done!
68547 | 197666 | CM | CMPLT | CMPLT | 2023-09-07 14:55 | 2023-09-12 09:03 | Self test failed. | (TC 09/08/2023 02:28:06 PM, Scero, Tomlyn) Tested device in OR and confirmed pressure sensor failure. Also a scavenging system error was occuring. Possibly not related to the device but to the OR boom. Downloaded DSR to look over. (TC 09/11/2023 02:29:51 PM, Scero, Tomlyn) Unit brought downstairs to RT and the Bag pressure sensor was replaced. Pressure sensor was from RT stock and already charged to RT anesthesia. After bag pressure sensor was replaced, unit was no longer giving an error. Self test passed. Tests passed per IPM-L. | pressure sensor failure and NGF valve failure
69500 | 197667 | CM | CMPLT | CMPLT | 2023-09-07 14:57 | 2023-09-07 14:58 | Functional check | (TC 09/07/2023 02:56:49 PM, Campbell, Connor) functional test based on ASM passed | Came down with a PCU, did a functional check
69398 | 197668 | PM | CMPLT | CMPLT | 2023-09-07 15:04 | 2023-09-07 15:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/07/2023 03:24:45 PM, Cortez, John) Cleaned, Downloaded error/event logs, Calibrated infusion rate, PM completed successfully, results logged in ASM
| 197669 | PA | CMPLT | CMPLT | 2023-09-08 06:36 | 2023-09-12 11:09 | Ordering sterile water | (TC 09/08/2023 06:48:22 AM, Baghirzade, Nara) SC 1003044890 (TC 09/12/2023 11:07:54 AM, Baghirzade, Nara) Emailed Murtala to cancel the shopping cart. Water was not received.
| 197670 | RDS | CMPLT | CMPLT | 2023-09-08 09:18 | 2023-09-08 10:11 | Weekly Rounds: NICU and RESUS | (TC 09/08/2023 09:18:02 AM, Baghirzade, Nara) Weekly rounds in NICU and RESUS completed. See procedure.
| 197671 | CM | CMPLT | CMPLT | 2023-09-08 09:50 | 2023-09-08 09:51 | continuing for ENDO move due to Code Grey | (TC 09/08/2023 09:50:22 AM, Samanian, Kamran) configured all CVs for storing images on the USB and internal memory
| 197672 | RDS | CMPLT | CMPLT | 2023-09-08 10:12 | 2023-09-08 10:12 | Weekly Rounds: 18S and PACU POD A | (TC 09/08/2023 10:11:51 AM, Baghirzade, Nara) Weekly rounds in PACU POD A and 18S completed. See procedure.
1059835 | 197673 | CM | CMPLT | CMPLT | 2023-09-08 10:31 | 2023-11-16 13:37 | ordering the tubings
| 197674 | CM | CMPLT | CMPLT | 2023-09-08 10:36 | 2023-09-08 10:37 | Integration meeting | (TC 09/08/2023 10:36:32 AM, Samanian, Kamran) done!
1051371 | 197675 | PM | CMPLT | CMPLT | 2023-09-08 10:41 | 2023-09-08 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/08/2023 10:42:57 AM, Canjimavadivel, Kamallesh ) Cleaned pump PM pass
| 197676 | RDS | CMPLT | CMPLT | 2023-09-08 10:59 | 2023-09-08 11:01 | Weekly Rounds: 11S | (TC 09/08/2023 10:59:20 AM, Baghirzade, Nara) Weekly rounds in 11S completed. See procedure
531113 | 197677 | FC | CMPLT | CMPLT | 2023-09-08 11:02 | 2023-09-08 11:03 | Functional check | (TC 09/08/2023 11:02:27 AM, Baghirzade, Nara) Checked connectors on tubing and cuff side to ensure it is compatible. Functional tests passed. Good physical condition. Followed up with Devina.
531115 | 197678 | FC | CMPLT | CMPLT | 2023-09-08 11:03 | 2023-09-08 11:04 | Functional check | (TC 09/08/2023 11:03:13 AM, Baghirzade, Nara) Checked connectors on tubing and cuff side to ensure it is compatible. Functional tests passed. Good physical condition. Followed up with Devina.
531117 | 197679 | FC | CMPLT | CMPLT | 2023-09-08 11:04 | 2023-09-08 11:04 | Functional check | (TC 09/08/2023 11:04:00 AM, Baghirzade, Nara) Checked connectors on tubing and cuff side to ensure it is compatible. Functional tests passed. Good physical condition. Followed up with Devina.
531119 | 197680 | FC | CMPLT | CMPLT | 2023-09-08 11:05 | 2023-09-08 11:08 | Functional check | (TC 09/08/2023 11:07:16 AM, Baghirzade, Nara) Checked connectors on tubing and cuff side to ensure it is compatible. Functional tests passed. Good physical condition. Followed up with Devina.
531120 | 197681 | FC | CMPLT | CMPLT | 2023-09-08 11:08 | 2023-09-08 11:08 | Functional check | (TC 09/08/2023 11:08:01 AM, Baghirzade, Nara) Checked connectors on tubing and cuff side to ensure it is compatible. Functional tests passed. Good physical condition. Followed up with Devina.
66445 | 197682 | CM | CMPLT | CMPLT | 2023-09-08 12:37 | 2023-09-08 12:36 | missing lock springs | installed missing springs. tested unit. functional.
| 197683 | CM | CMPLT | CMPLT | 2023-09-08 12:39 | 2023-09-08 12:39 | biomed meeting | one on one with Adeel
| 197684 | RDS | CMPLT | CMPLT | 2023-09-08 12:41 | 2023-09-08 12:41 | weekly check for 14 southand 16 north
1062021 | 197685 | CM | CMPLT | CMPLT | 2023-09-08 13:34 | 2023-09-08 13:34 | "scale needs to be calibrated" | (TC 09/08/2023 01:33:42 PM, Cortez, John) Confirmed error message Power cycling bed resolves issue
1062031 | 197686 | CM | CMPLT | CMPLT | 2023-09-08 13:37 | 2023-09-08 13:37 | "scale needs to be calibrated" | (TC 09/08/2023 01:36:47 PM, Cortez, John) Confirmed error message Power cycling bed resolves issue
| 197687 | CM | CMPLT | CMPLT | 2023-09-08 14:27 | 2023-09-08 14:27 | Suction regulator device battery dead | (TC 09/08/2023 02:27:34 PM, Campbell, Connor) replaced the battery checked functionality passed functionality test cleaned outer and internal housings | 2 suction regulators AMICO SN: VRA08089HH and AAD1769DD battery was dead
| 197688 | RDS | CMPLT | CMPLT | 2023-09-08 14:41 | 2023-09-08 14:44 | Weekly check for 10N- sep 5-8 | (TC 09/08/2023 02:42:55 PM, Cortez, John) Checked 124-A and 10-A No issues Retrieved Masimo kite tablets Informed clerk that bed in hallway can be put into service
| 197689 | RDS | CMPLT | CMPLT | 2023-09-08 14:47 | 2023-09-08 14:53 | Weekly check for 10N- sep 5-8 | (TC 09/04/2023 02:52:34 PM, Cortez, John) No issues
| 197690 | CM | CMPLT | CMPLT | 2023-09-08 14:50 | 2023-09-08 14:50 | knob does not turn on the suction regulator | (TC 09/08/2023 02:49:40 PM, Campbell, Connor) removed housing and removed the suction regulator module, reseated it and made sure it was not as tight as before everything now works in a functional test cleaned the inner and outer housings | Suction regulation device AMICO SN: 40100101652 the knob to adjust the suction had been overtightened during use and was not able to be moved
1061210 | 197691 | CM | ACTIV | ISSUE | 2023-09-08 15:06 |  | Issue with AUX 2 | (TC 09/08/2023 03:05:50 PM, Reid, Emmerson ) Met with Derrick and we determind it was an issue with the scalor. He is working on getting a replacement under warranty.
| 197692 | CM | CMPLT | CMPLT | 2023-09-08 15:13 | 2023-09-08 15:13 | suction regulation device not working at all ie. no suction | (TC 09/08/2023 03:13:14 PM, Campbell, Connor) cleaned inner and outer housing functionality test on device had no issues, the problem seems to not have been the suction regulation device and is likely another device causing this device to not work. TLDR: everything is fine, no repairs or maintenance was needed | Suction regulation device AMICO SN: 40100105697 suction regulation device was reported to not be working at all on regular or intermittent settings
531098 | 197693 | CM | CMPLT | CMPLT | 2023-09-08 15:27 | 2023-09-08 15:27 | low battery | (TC 09/08/2023 03:27:01 PM, Campbell, Connor) cleaned inner and outer housing replaced battery performed functional test, everything works as expected. | Suction regulator AMVEX SN: VRA08106HH battery was low, needed to be replaced
68673 | 197694 | CM | CMPLT | CMPLT | 2023-09-11 06:48 | 2023-09-11 06:53 | Not charging, AC LED intermittently switches on/off | (TC 09/11/2023 06:48:12 AM, Baghirzade, Nara) Power cord is defective; replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 40 mohms Leakage current: 73.5 uA
67959 | 197695 | PA | CMPLT | CMPLT | 2023-09-11 07:04 | 2023-09-13 12:00 | Ordering incubator hood | (TC 09/11/2023 07:27:17 AM, Baghirzade, Nara) SC 1003046034 (TC 09/13/2023 11:58:33 AM, Baghirzade, Nara) PO 7202279644
1056450 | 197696 | CM | CMPLT | CMPLT | 2023-09-11 08:31 | 2023-09-11 08:32 | Configuration load and functional tests | (TC 09/11/2023 08:31:06 AM, Baghirzade, Nara) New radical was received for this root. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1061215 | 197697 | CM | CMPLT | CMPLT | 2023-09-11 09:30 | 2023-09-12 07:24 | iPAD is out of battery, charge cable is being replaced | (TC 09/11/2023 09:29:50 AM, Samanian, Kamran) removed the charge cable from OR12 iPAD and waiting to get to the OR14 to replace. (TC 09/12/2023 07:23:05 AM, Samanian, Kamran) John replaced the cable yesterday afternoon ( SEP 11), checked the iPad this morninbg and battery was being charged ( 56%)
1051448 | 197698 | CM | CMPLT | CMPLT | 2023-09-11 09:38 | 2024-07-05 16:30 | "4DP-ON and 4DP-OFF message on screen" | (TC 09/11/2023 09:38:10 AM, Cortez, John) Confirmed issue Unit also displays software version (4.06)? and hangs before turning off Reached out to Cardinal Health (TC 07/05/2024 04:29:35 PM, Cortez, John) Retired, replaced with #66230
| 197699 | CM | CMPLT | CMPLT | 2023-09-11 09:40 | 2023-09-12 17:30 | Focal Point - Philips -Software | (TC 09/11/2023 09:45:48 AM, Samanian, Kamran) placed the order S/C#1003046395 (TC 09/12/2023 05:30:26 PM, Samanian, Kamran) Received PO# 7202280029
| 197700 | CM | CMPLT | CMPLT | 2023-09-11 09:48 | 2023-09-12 09:54 | Requested ENDO about recall and followed up- Water Filter | (TC 09/12/2023 09:53:56 AM, Samanian, Kamran) Anita replied. the device is no longer in use
1046588 | 197701 | CM | CMPLT | CMPLT | 2023-09-11 10:24 | 2023-09-11 10:24 | Defective probe; always reads a high temperature | (TC 09/11/2023 10:23:34 AM, Baghirzade, Nara) Replaced probe. Functional tests passed.
1048168 | 197702 | CM | CMPLT | NFF | 2023-09-11 10:28 | 2023-09-11 10:29 | Reported not working | (TC 09/11/2023 10:27:52 AM, Baghirzade, Nara) Could not find an issue with the thermometer. All functional tests passed.
1067815 | 197703 | CM | CMPLT | CMPLT | 2023-09-11 10:33 | 2023-09-12 09:00 | 6 YEAR PM | (TC 09/12/2023 08:59:26 AM, Scero, Tomlyn) 6 year PM parts replaced. Tests complete and passed
62596 | 197704 | CM | CMPLT | CMPLT | 2023-09-11 10:50 | 2023-09-11 10:51 | Intermittently switches on/off | (TC 09/11/2023 10:50:38 AM, Baghirzade, Nara) Unit is defective and end of support. No parts/repair available. Will supply NICU with a spare filac thermometer in its place.
| 197705 | RDS | CMPLT | CMPLT | 2023-09-11 10:52 | 2023-09-11 10:52 | Weekly check for 14 South and 16 North | weekly check passed
1067816 | 197706 | CM | CMPLT | CMPLT | 2023-09-11 10:54 | 2023-09-11 10:56 | Repurposed for NICU | (TC 09/11/2023 10:54:55 AM, Baghirzade, Nara) Repurposed spare thermometer for NICU in replacement of the retired defective thermometer. Functional tests passed. Printed and placed new asset sticker tag.
| 197707 | CM | CMPLT | CMPLT | 2023-09-11 14:15 | 2023-09-11 14:15 | Biomed meeting | Biome meeting
| 197708 | CM | CMPLT | CMPLT | 2023-09-11 14:36 | 2023-09-11 14:37 | meeting Huddle | (TC 09/11/2023 02:36:12 PM, Samanian, Kamran) done!
1046203 | 197709 | CM | CMPLT | CMPLT | 2023-09-11 14:43 | 2023-10-04 14:07 | Touch screen distorted, OP buttons unresponsive | (TC 09/11/2023 02:43:18 PM, Fan, Tak) Resident reported touch screen distorted. Panel no reponsive to pressing. Ordered touch screen and new operator panel. (TC 10/04/2023 02:06:11 PM, Fan, Tak) Replaced touchscreen and OP panel. Tested, OK.
1049941 | 197710 | CM | CMPLT | CMPLT | 2023-09-11 14:45 | 2023-09-18 08:33 | Blue screen during TEE | (TC 09/11/2023 02:51:09 PM, Fan, Tak) Study was in the No Archive archive somehow. Moved study over. Pulled logs, no events stand out. Probably not logged for blue screen errors.
62979 | 197711 | CM | CMPLT | CMPLT | 2023-09-11 15:36 | 2023-09-11 15:38 | broken/not turning on | (TC 09/11/2023 03:36:25 PM, Campbell, Connor) performed a functional test as per the service manual passed functional tests, no problem detected had to slightly adjust the intensity of the light to be in the middle of the acceptable range as it was reading a litle bit low (TC 09/12/2023 11:08:21 AM, Campbell, Connor) replaced the light timer replaced the rubber feet
1054537 | 197712 | PM | CMPLT | CMPLT | 2023-09-11 15:49 | 2023-09-11 15:50 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 09/11/2023 03:49:26 PM, Cortez, John) Completed cleaning cycle
| 197713 | RDS | CMPLT | CMPLT | 2023-09-11 15:55 | 2023-10-02 10:05 | Weekly rounds ICU - 9/11 | (TC 09/11/2023 03:55:15 PM, Cortez, John) Swapped pumps due for PM and plugged in all PCUs in pod C equipment room
| 197714 | CM | CMPLT | CMPLT | 2023-09-12 07:25 | 2023-09-12 07:25 | Added Bipolar foot Switch to FT10 in OR9 | (TC 09/12/2023 07:25:02 AM, Samanian, Kamran) done! (TC 09/12/2023 07:25:02 AM, Samanian, Kamran) done!
| 197715 | CM | CMPLT | CMPLT | 2023-09-12 07:56 | 2023-09-12 07:56 | PM huddle | (TC 09/11/2023 02:00:00 PM, Baghirzade, Nara) PM huddle
| 197716 | CON | CMPLT | CMPLT | 2023-09-12 07:57 | 2023-09-12 10:07 | Omnibed exam light setup recommendations document | (TC 09/12/2023 10:06:47 AM, Baghirzade, Nara) Completed the recommendations document for NICU RESUS
| 197717 | CM | CMPLT | CMPLT | 2023-09-12 08:36 | 2023-09-12 08:36 | Power bar installed on IV pole in OR13 | (TC 09/12/2023 08:35:58 AM, Samanian, Kamran) done! Maureen requested for that
66432 | 197718 | CM | CMPLT | CMPLT | 2023-09-12 09:12 | 2023-09-12 10:59 | cables not working | (TC 09/12/2023 10:04:55 AM, Scero, Tomlyn) Replaced patient leads. Tested and confirmed signal was good. returned unit to CCU
| 197719 | RDS | CMPLT | CMPLT | 2023-09-12 09:35 | 2023-09-13 08:24 | Weekly check for ED
| 197720 | RDS | CMPLT | CMPLT | 2023-09-12 09:47 | 2023-10-03 14:43 | Weekly check for ICU
1059828 | 197721 | CM | CMPLT | CMPLT | 2023-09-12 10:05 | 2023-09-12 10:07 | Cleaning | (TC 09/12/2023 10:05:15 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water. (TC 09/12/2023 10:05:20 AM, Campbell, Connor) Drained, disinfected, rinsed, and refilled with sterile water.
1049905 | 197722 | CM | CMPLT | CMPLT | 2023-09-12 10:30 | 2023-09-12 10:32 | Tube head stuck in park | (TC 09/12/2023 10:29:39 AM, Fan, Tak) Release latch works but still stuck when released. Wiggled tube head up, wear pads worn down and arm was up against the screw. Removed park sensor, replaced wear pads. Tested, OK.
| 197723 | CM | CMPLT | CMPLT | 2023-09-12 10:55 | 2023-09-12 10:55 | Philips work on FM and XDS and email | (TC 09/12/2023 10:55:04 AM, Samanian, Kamran) done, Sent email to Randy
| 197724 | PA | CMPLT | CMPLT | 2023-09-12 11:31 | 2023-09-13 12:02 | Ordering sterile water | (TC 09/12/2023 11:52:53 AM, Baghirzade, Nara) SC 1003048043 (TC 09/13/2023 12:01:46 PM, Baghirzade, Nara) Order# 4150262472
| 197725 | PA | CMPLT | CMPLT | 2023-09-12 11:55 | 2023-09-26 09:41 | Ordering for stock and repairs | (TC 09/12/2023 12:00:57 PM, Baghirzade, Nara) SC 1003048215 (TC 09/26/2023 09:40:00 AM, Baghirzade, Nara) PO 7202280942
1054543 | 197726 | CM | CMPLT | CMPLT | 2023-09-12 12:40 | 2023-09-12 12:42 | Reported that wifi connectivity was not working properly | (TC 09/12/2023 12:40:34 PM, Baghirzade, Nara) ECG tech reported that the wifi connectivity was not working properly and the patient data was not being transferred out. She left the device to sit overnight and in the morning, the patient data was finally transferred. Kamran and I checked the network configuration and connectivity. Unit is fully functional.
69009 | 197727 | CM | CMPLT | CMPLT | 2023-09-12 14:05 | 2023-09-12 14:07 | Not charging | (TC 09/12/2023 02:05:57 PM, Cortez, John) Power cord defective, replaced
1054462 | 197728 | CM | CMPLT | CMPLT | 2023-09-12 15:09 | 2023-10-13 09:43 | "Very low pressure" | (TC 09/12/2023 03:09:10 PM, Cortez, John) Ran with test set and checked alarms and performance. Tested ok. Output from compressor is good, ~300 mmhg Users might not be using correct fluid bag size, or are installing bag too high within chamber (TC 10/12/2023 09:36:10 AM, Cortez, John) A few weeks back connected Marlene with rep from ICU medical, to coordinate training/refresher Will close WO and monitor if issue comes back again
| 197729 | CM | CMPLT | CMPLT | 2023-09-12 15:53 | 2023-09-12 15:55 | NICU huddle | (TC 09/12/2023 03:54:49 PM, Baghirzade, Nara) NICU huddle. Present the omnibed exam light suggestions.
1061333 | 197730 | FC | CMPLT | CMPLT | 2023-09-12 16:00 | 2023-09-12 16:00 | QC checklist | (TC 09/12/2023 03:59:41 PM, Baghirzade, Nara) QC checklist completed. Returned to 18 South
| 197731 | CM | CMPLT | CMPLT | 2023-09-12 17:27 | 2023-10-05 11:33 | Testing image capture with Katrina | (TC 09/12/2023 05:27:10 PM, Samanian, Kamran) worklist worked, still images take time to show up in powerchart (TC 09/15/2023 01:57:10 PM, Samanian, Kamran) did some testing this morning, it takes 4-5 min to get the images in powerchart regardless of number of images (TC 09/29/2023 11:32:51 AM, Samanian, Kamran) done and passed (TC 10/05/2023 11:32:51 AM, Samanian, Kamran) done and passed
1054543 | 197732 | CM | CMPLT | CMPLT | 2023-09-12 17:28 | 2023-09-14 13:56 | WiFi connectivity | (TC 09/12/2023 05:28:13 PM, Samanian, Kamran) did the survey in 9S, romaing is not good, checked in ICU and it works better. it seems there are dead spots in 9S. need to do more testing (TC 09/13/2023 09:54:47 AM, Samanian, Kamran) The radio was changed from ABG to bg. Did the survey and signal was Ok ( Even one Bar was available ) and raoming was fine.
| 197733 | CM | CMPLT | CMPLT | 2023-09-12 17:30 | 2023-09-12 17:29 | meet with Kam to discuss about ER EAST | (TC 09/12/2023 05:29:38 PM, Samanian, Kamran) done!
| 197734 | CM | CMPLT | CMPLT | 2023-09-13 08:10 | 2023-09-20 11:34 | Requesting no charge PO for returning ERBE irrigation pumps ( consignment ) to Vantage | (TC 09/13/2023 08:09:51 AM, Samanian, Kamran) Here are the serial numbers to return: 11285114 11448550 11448548 11527600 11448549 11541332 all returned with foot pedal. there are 5 boxes. Used Purolator, (TC 09/15/2023 01:55:22 PM, Samanian, Kamran) scheduled a pick up for purolator, SEP 19 9Am-2Pm. tracking# 334303428927 (TC 09/20/2023 10:22:38 AM, Samanian, Kamran) delivered 5 boxes to Vantage
1056125 | 197735 | PM | CMPLT | CMPLT | 2023-09-13 08:25 | 2023-09-13 10:07 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 09/13/2023 10:06:37 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
1062270 | 197736 | CM | CMPLT | CMPLT | 2023-09-13 08:37 | 2023-09-13 08:38 | No functionality on the bed - nothing working | (TC 09/13/2023 08:37:18 AM, Ferreira, Qwynn) Reset the bed and tested all functionas. Bed working once again
| 197737 | PA | CMPLT | CMPLT | 2023-09-13 08:38 | 2023-09-26 09:35 | Ordering CMAC cables | (TC 09/13/2023 10:01:16 AM, Baghirzade, Nara) SC 1003049143 (TC 09/26/2023 09:34:41 AM, Baghirzade, Nara) PO 7202282284
1050098 | 197738 | CM | CMPLT | CMPLT | 2023-09-13 08:51 | 2023-09-13 08:52 | Screen not working | (TC 09/13/2023 08:51:00 AM, Canjimavadivel, Kamallesh ) Tried turning on device, not working. Found battery completely depleted. Placed on charger and tested with test battery. Device working as intended.
| 197739 | CM | CMPLT | CMPLT | 2023-09-13 09:38 | 2023-09-13 09:38 | 2x SpO2 reuseable trunk cords not working | (TC 09/13/2023 09:38:07 AM, Campbell, Connor) Tested both trunk cables with known good probes, one of them functioned normally, the other was non-functioning Both were handed off to Kam to deal with as he saw fit | Both cords came in saying they were non-functioning
| 197740 | CM | CMPLT | CMPLT | 2023-09-13 09:41 | 2023-09-13 09:41 | Disposable SpO2 trunk cord not working | (TC 09/13/2023 09:41:07 AM, Campbell, Connor) Tested the cord with known working sensor, was determined to be working normally. Handed off to Kam to deal with as he saw fit.
1059938 | 197741 | CM | CMPLT | CMPLT | 2023-09-13 09:46 | 2023-09-13 09:50 | V2 not working | (TC 09/13/2023 09:46:37 AM, Canjimavadivel, Kamallesh ) Confirmed issue, replaced V1-V3 chest lead set. V7 also damaged. Replaced lead. Tested good with simulator
| 197742 | CM | CMPLT | CMPLT | 2023-09-13 09:46 | 2023-09-13 09:46 | SpO2 probe not reading, no light on probe | (TC 09/13/2023 09:46:21 AM, Campbell, Connor) Tested both the SpO2 probe and the trunk cable with known good components The probe was determined to be faulty the Trunk cable was determined to be in working condition Both were given to Kam to deal with as he saw fit | Probe came in attached to the trunk cable
67920 | 197743 | CM | CMPLT | CMPLT | 2023-09-13 09:54 | 2023-09-13 09:55 | Broken porthole doors (2) and functional check | (TC 09/13/2023 09:54:35 AM, Baghirzade, Nara) Replaced 2 porthole doors. Both were broken off. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 75% RH. Incubator is fully functional.
67148 | 197744 | CM | CMPLT | CMPLT | 2023-09-13 10:20 | 2023-09-13 10:33 | BP not working | (TC 09/13/2023 10:31:39 AM, Canjimavadivel, Kamallesh ) Hose connector to cuff was damaged. Replaced connector and tested okay
| 197745 | CM | CMPLT | CMPLT | 2023-09-13 10:29 | 2023-09-12 10:28 | Masimo work | Masimo work
| 197746 | CM | CMPLT | CMPLT | 2023-09-13 10:30 | 2023-09-13 10:30 | TMS work | TMS work
| 197747 | CM | CMPLT | CMPLT | 2023-09-13 11:14 | 2023-09-13 11:15 | 1-1 meeting with Adeel | (TC 09/13/2023 11:14:14 AM, Baghirzade, Nara) 1-1 meeting with Adeel
1054543 | 197748 | CM | CMPLT | CMPLT | 2023-09-13 12:04 | 2023-09-13 12:05 | Amplifier was intermittently switching on/off | (TC 09/13/2023 12:03:58 PM, Baghirzade, Nara) USB patient data cable was defective. Replaced the cable. Good physical condition; battery was replaced recently. Used a simulator to check all leads; all leads are functional. Unit is functional. (TC 09/13/2023 12:04:45 PM, Campbell, Connor) USB patient data cable was defective. Replaced the cable. Good physical condition; battery was replaced recently. Used a simulator to check all leads; all leads are functional. Unit is functional.
| 197749 | CM | CMPLT | CMPLT | 2023-09-13 12:45 | 2023-12-08 08:21 | NICU Draeger Boom Case Cracked | (TC 09/13/2023 12:45:14 PM, Ferreira, Qwynn) Called Drager to obtain pricing and part number for the Boom/pendant case that was cracked. Will be submitting purchasing soon. (TC 12/07/2023 08:19:35 AM, Ferreira, Qwynn) Worked with Draeger to get new cases covers in as they shipped damaged ones. They shipped a new set to replace the broken ones. New ones arrived and boom cover will be replaced
1050065 | 197750 | PM | CMPLT | CMPLT | 2023-09-13 13:35 | 2023-09-14 14:26 | ANNUAL PM | (TC 09/14/2023 02:24:26 PM, Baghirzade, Nara) PM completed. Battery is fully functional. Calibrated CO2 with 5% CO2 gas = 37.9 mmHg NIBP passed. Replaced the BP cuff. BP hose in good condition. Set = 120/80 HR 60, Measured = 120/79 HR 60 Set = 80/40 HR 40, Measured = 79/39 HR 40 SpO2 passed. Set = 99%, Measured = 99% Set = 85%, Measured 85% ECG passed. Set 60 BPM and 20 RESP , Measured 60 BPM and 20 RESP. All leads are good. No noise or artifacts detected. All alarms are functional; tested low and high limits. All passed. Printer is functional. Power cord in good condition. Electrical safety results: Power cord resistance = 58 mohms Current leakage = 52.3 uA
1067824 | 197751 | CM | CMPLT | CMPLT | 2023-09-13 13:51 | 2023-09-14 10:31 | 6 YR PM | (TC 09/14/2023 10:30:03 AM, Scero, Tomlyn) 6 year PM parts replaced. tests complete and passed. No leaks present.
1067437 | 197752 | FC | CMPLT | CMPLT | 2023-09-13 13:58 | 2023-09-13 13:59 | QC checklist | (TC 09/13/2023 01:58:58 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1063000 | 197753 | PM | CMPLT | CMPLT | 2023-09-13 15:05 | 2023-09-13 15:05 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 09/13/2023 03:04:59 PM, Cortez, John) Completed cleaning cycle
502089 | 197754 | RQ | CMPLT | CMPLT | 2023-09-13 15:10 | 2023-09-13 19:40 | camera unable to return to default position | (TC 09/13/2023 07:36:38 PM, Paz, Ramon (Tino)) After performing a scan in stretcher mode, customer could not reposition the system to 180 degrees. Customer is saying that message pull clutch. Clutch was released. pulled back bed and moved Detector 2 trunnion to zero degrees. transformed gantry to 180 then homed. | detector 2 motion error
67959 | 197755 | CM | CMPLT | CMPLT | 2023-09-13 15:12 | 2023-09-13 15:16 | Replacement hood installation | (TC 09/13/2023 03:12:45 PM, Campbell, Connor) Took all removable parts off the old hood and disposed of it Attached new hood to the machine Performed functional check, humidity sensor was not detecting the specified humidity Removed humidity sensor and observed a wire had broken due to corrosion Replaced the humidity sensor with a new one ITEM: 6600-0784-700 ST2B000744 LOT: 04NOV2022 LPN:U01-199930 performed another functional check and determined the device now operates correctly in all aspects, humidity and temperature are within specified error ranges
65698 | 197756 | CM | CMPLT | CMPLT | 2023-09-13 15:30 | 2024-02-14 10:04 | Hose is ripped | (TC 09/13/2023 03:30:31 PM, Cortez, John) Picked up unit from OR desk Hose is ripped near connector that connects to mat Contacted Jessica and she mentioned that the hose needs to be ordered
1062242 | 197757 | CM | CMPLT | CMPLT | 2023-09-13 15:39 | 2023-09-13 15:40 | "bed exit alarm not working" | (TC 09/13/2023 03:39:22 PM, Cortez, John) Scale error present, Power cycled bed, error cleared Bed exit alarm tested ok
1061989 | 197758 | CM | CMPLT | CMPLT | 2023-09-13 15:43 | 2023-09-13 15:45 | "scale issue" | (TC 09/13/2023 03:42:49 PM, Cortez, John) Didn't see any errors, Scale seems to be ok. Informed nurse that the scale needs to zeroed when patient is off bed, then the patient can be re-weighed.
1061992 | 197759 | CM | CMPLT | NFF | 2023-09-13 15:46 | 2023-09-13 15:49 | "bed exit alarm not working" | (TC 09/13/2023 03:46:23 PM, Cortez, John) Didn't see any errors, Patient needs to be re-weighed again before bed exit alarm can be activated
32510 | 197760 | CM | CMPLT | CMPLT | 2023-09-13 17:10 | 2023-09-13 17:10 | Guard error |  | Reposition room and turned the generator on/off
| 197761 | CM | CMPLT | CMPLT | 2023-09-13 17:11 | 2023-09-13 17:11 | Site coverage
501070 | 197762 | CM | CMPLT | CMPLT | 2023-09-13 17:42 | 2023-09-13 17:44 | MSH CEOU RM 10 DID NOT HAVE BURNER OR DOWNTIME AVAILABLE AS DICOM DESTINATIONS | (TC 09/13/2023 05:41:50 PM, Whittington, Scott) MSH CEOU RM 10 DID NOT HAVE BURNER OR DOWNTIME AVAILABLE AS DICOM DESTINATIONS REMOVED THE SETTINGS FOR THESE TWO DESTINATIONS, AND RE ENTERED OK SYSTEMS HAVE VERY HIGH HDD | MSH CEOU RM 10 DID NOT HAVE BURNER OR DOWNTIME AVAILABLE AS DICOM DESTINATIONS
| 197763 | CM | CMPLT | CMPLT | 2023-09-13 17:46 | 2023-09-13 17:46 | MSH CEOU iU22 HDD HAVE VERY HIGH STORAGE LEVELS; PURGE AND DELETE PT DB ON ALL 12 SYSTEMS | (TC 09/13/2023 05:46:07 PM, Whittington, Scott) MSH CEOU iU22 HDD HAVE VERY HIGH STORAGE LEVELS; PURGE AND DELETE PT DB ON ALL 12 SYSTEMS (TC 09/13/2023 05:46:07 PM, Whittington, Scott) MSH CEOU iU22 HDD HAVE VERY HIGH STORAGE LEVELS; PURGE AND DELETE PT DB ON ALL 12 SYSTEMS | MSH CEOU iU22 HDD HAVE VERY HIGH STORAGE LEVELS; PURGE AND DELETE PT DB ON ALL 12 SYSTEMS
| 197764 | RDS | CMPLT | CMPLT | 2023-09-14 07:52 | 2023-09-18 12:43 | Weekly checks in CCU | (TC 09/11/2023 12:42:42 PM, Scero, Tomlyn) Weekly checks complete
1067836 | 197765 | ACC | CMPLT | CMPLT | 2023-09-14 08:18 | 2023-09-29 13:42 | Incoming inspection | (TC 09/29/2023 01:41:59 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067837 | 197766 | ACC | CMPLT | CMPLT | 2023-09-14 08:20 | 2023-09-29 13:35 | Incoming inspection | (TC 09/29/2023 01:34:22 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067838 | 197767 | ACC | CMPLT | CMPLT | 2023-09-14 08:22 | 2023-09-29 13:27 | Incoming inspection | (TC 09/29/2023 01:27:08 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067839 | 197768 | ACC | CANCL | CANCL | 2023-09-14 08:23 | 2024-08-23 08:46 | Incoming inspection | Unable to locate, missing after adding to TMS
1067840 | 197769 | ACC | CMPLT | CMPLT | 2023-09-14 08:24 | 2023-09-28 14:41 | Incoming inspection | (TC 09/28/2023 02:41:10 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067841 | 197770 | ACC | CMPLT | CMPLT | 2023-09-14 08:25 | 2023-09-28 14:33 | Incoming inspection | (TC 09/28/2023 02:32:34 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067842 | 197771 | ACC | CMPLT | CMPLT | 2023-09-14 08:26 | 2023-09-28 14:13 | Incoming inspection | (TC 09/28/2023 02:12:02 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067843 | 197772 | ACC | CMPLT | CMPLT | 2023-09-14 08:27 | 2023-09-28 14:09 | Incoming inspection | (TC 09/28/2023 02:08:35 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067844 | 197773 | ACC | CMPLT | CMPLT | 2023-09-14 08:28 | 2023-09-28 13:59 | Incoming inspection | (TC 09/28/2023 01:58:42 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067845 | 197774 | ACC | CMPLT | CMPLT | 2023-09-14 08:29 | 2023-09-28 13:53 | Incoming inspection | (TC 09/28/2023 01:52:30 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067846 | 197775 | ACC | CMPLT | CMPLT | 2023-09-14 08:30 | 2023-09-28 13:38 | Incoming inspection | (TC 09/28/2023 01:37:48 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067847 | 197776 | ACC | CMPLT | CMPLT | 2023-09-14 08:34 | 2023-09-28 13:24 | Incoming inspection | (TC 09/28/2023 01:23:47 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067848 | 197777 | ACC | CMPLT | CMPLT | 2023-09-14 08:36 | 2023-09-29 14:12 | Incoming inspection | (TC 09/29/2023 02:11:34 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067849 | 197778 | ACC | CMPLT | CMPLT | 2023-09-14 08:40 | 2023-09-29 14:18 | Incoming inspection | (TC 09/29/2023 02:18:05 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067850 | 197779 | ACC | CMPLT | CMPLT | 2023-09-14 08:41 | 2023-09-29 14:25 | Incoming inspection | (TC 09/29/2023 02:24:57 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067851 | 197780 | ACC | CMPLT | CMPLT | 2023-09-14 08:43 | 2023-09-29 14:31 | Incoming inspection | (TC 09/29/2023 02:30:58 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067852 | 197781 | ACC | CMPLT | CMPLT | 2023-09-14 08:45 | 2023-09-29 14:38 | Incoming inspection | (TC 09/29/2023 02:37:14 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067853 | 197782 | ACC | CANCL | CANCL | 2023-09-14 08:47 | 2024-08-23 08:45 | Incoming inspection | Unable to locate, missing after adding to TMS
66458 | 197783 | PM | CMPLT | CMPLT | 2023-09-14 08:53 | 2023-09-14 08:53 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 70uA and 85 ohm.
1067854 | 197784 | ACC | CANCL | CANCL | 2023-09-14 08:54 | 2024-08-23 08:44 | Incoming inspection | Unable to locate, missing after adding to TMS
1067855 | 197785 | ACC | CMPLT | CMPLT | 2023-09-14 08:55 | 2023-09-29 14:05 | Incoming inspection | (TC 09/29/2023 02:04:24 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067856 | 197786 | ACC | CMPLT | CMPLT | 2023-09-14 08:57 | 2023-10-23 14:37 | Incoming inspection | (TC 10/23/2023 02:36:53 PM, Canjimavadivel, Kamallesh ) Asset tagged and all tests pass
1067857 | 197787 | CM | CMPLT | CMPLT | 2023-09-14 09:08 | 2023-09-29 13:43 | Incoming inspection | (TC 09/29/2023 01:42:49 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067858 | 197788 | ACC | CMPLT | CMPLT | 2023-09-14 09:09 | 2023-09-29 13:36 | Incoming inspection | (TC 09/29/2023 01:35:20 PM, Canjimavadivel, Kamallesh ) Functional test pass
502173 | 197789 | RQ | CMPLT | CMPLT | 2023-09-14 09:09 | 2023-10-25 14:17 | Coral software not working, other software slow |  | Thinking System software slow processing, Coral software not opening
1067859 | 197790 | ACC | CMPLT | CMPLT | 2023-09-14 09:10 | 2023-09-29 13:28 | Incoming inspection | (TC 09/29/2023 01:27:43 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067860 | 197791 | ACC | CMPLT | CMPLT | 2023-09-14 09:12 | 2023-09-28 14:42 | Incoming inspection | (TC 09/28/2023 02:41:57 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067861 | 197792 | ACC | CMPLT | CMPLT | 2023-09-14 09:13 | 2023-09-28 14:32 | Incoming inspection | (TC 09/28/2023 02:31:48 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067862 | 197793 | ACC | CMPLT | CMPLT | 2023-09-14 09:16 | 2023-09-28 14:13 | Incoming inspection | (TC 09/28/2023 02:13:06 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067863 | 197794 | ACC | CMPLT | CMPLT | 2023-09-14 09:17 | 2023-09-28 14:08 | Incoming inspection | (TC 09/28/2023 02:08:03 PM, Canjimavadivel, Kamallesh ) Functional test pass
1063746 | 197795 | CM | CMPLT | CMPLT | 2023-09-14 09:19 | 2023-09-14 09:20 | Cleaning | (TC 09/14/2023 09:19:34 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water. (TC 09/14/2023 09:19:41 AM, Campbell, Connor) Drained, disinfected, rinsed, and refilled with sterile water.
1067864 | 197796 | ACC | CMPLT | CMPLT | 2023-09-14 09:19 | 2023-09-28 14:00 | Incoming inspection | (TC 09/28/2023 01:59:40 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067865 | 197797 | ACC | CMPLT | CMPLT | 2023-09-14 09:20 | 2023-09-28 13:51 | Incoming inspection | (TC 09/28/2023 01:50:56 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067866 | 197798 | ACC | CMPLT | CMPLT | 2023-09-14 09:21 | 2023-09-28 13:41 | Incoming inspection | (TC 09/28/2023 01:40:36 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067867 | 197799 | ACC | CMPLT | CMPLT | 2023-09-14 09:22 | 2023-09-28 13:36 | Incoming inspection | (TC 09/28/2023 01:36:07 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067868 | 197800 | ACC | CMPLT | CMPLT | 2023-09-14 09:23 | 2023-09-29 14:13 | Incoming inspection | (TC 09/29/2023 02:12:15 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067869 | 197801 | ACC | CMPLT | CMPLT | 2023-09-14 09:24 | 2023-09-29 14:19 | Incoming inspection | (TC 09/29/2023 02:18:45 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067870 | 197802 | ACC | CMPLT | CMPLT | 2023-09-14 09:25 | 2023-09-29 14:24 | Incoming inspection | (TC 09/29/2023 02:23:55 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067871 | 197803 | ACC | CMPLT | CMPLT | 2023-09-14 09:32 | 2023-09-29 14:32 | Incoming inspection | (TC 09/29/2023 02:31:23 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067872 | 197804 | ACC | CMPLT | CMPLT | 2023-09-14 09:33 | 2023-09-29 14:37 | Incoming inspection | (TC 09/29/2023 02:36:39 PM, Canjimavadivel, Kamallesh ) Functional test pass
1053728 | 197805 | FC | CMPLT | CMPLT | 2023-09-14 10:16 | 2023-09-14 10:17 | Functional check after use | (TC 09/14/2023 10:17:00 AM, Scero, Tomlyn) Functional checks complete and passed.
| 197806 | RDS | CMPLT | CMPLT | 2023-09-14 10:19 | 2023-09-14 10:22 | Weekly Rounds: NICU | (TC 09/14/2023 10:20:29 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure.
| 197807 | RDS | CMPLT | CMPLT | 2023-09-14 10:22 | 2023-09-14 10:24 | Weekly Round: RESUS | (TC 09/14/2023 10:22:29 AM, Baghirzade, Nara) Weekly round in RESUS completed. See procedure.
| 197808 | RDS | CMPLT | CMPLT | 2023-09-14 10:24 | 2023-09-14 10:26 | Weekly Rounds: 11S | (TC 09/14/2023 10:24:20 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure.
| 197809 | RDS | CMPLT | CMPLT | 2023-09-14 10:27 | 2023-09-14 10:28 | Weekly Rounds: PACU POD A and 18 South | (TC 09/14/2023 10:26:58 AM, Baghirzade, Nara) Weekly rounds on 18 South and PACU POD A completed. See procedure.
65956 | 197810 | CM | CMPLT | CMPLT | 2023-09-14 11:40 | 2023-09-14 11:41 | Channel error; no issue specified | (TC 09/14/2023 11:40:21 AM, Baghirzade, Nara) Downloaded and reviewed the error log. Latest errors are "log only". Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65993 | 197811 | CM | CMPLT | CMPLT | 2023-09-14 12:17 | 2023-09-14 12:17 | Latch broken and "channel error"; no issue specified | (TC 09/14/2023 12:16:38 PM, Baghirzade, Nara) Replaced lock latch. Check-in tests passed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1061215 | 197812 | CM | CMPLT | CMPLT | 2023-09-14 12:23 | 2023-11-03 10:44 | grounding checking on the boom
| 197813 | CM | CMPLT | CMPLT | 2023-09-14 12:34 | 2023-09-14 12:34 | TMS work | TMS work
| 197814 | CM | CMPLT | CMPLT | 2023-09-14 12:53 | 2024-03-25 14:27 | OneRecall Meeting | (TC 09/14/2023 12:52:50 PM, Chukri, Samir) OneRecall domain meeting
| 197815 | CM | CMPLT | CMPLT | 2023-09-14 13:07 | 2024-07-19 09:26 | Replacement parts for Nitronox | (TC 09/14/2023 01:09:10 PM, Rendon Lopez , Elizabeth) Working with Gillian, several rooms missing nitronox parts. ordered replacements (TC 02/27/2024 01:25:54 PM, Rendon Lopez , Elizabeth) More Nitronox missing valves, ordered more and follow up on previous order (TC 07/19/2024 09:26:01 AM, Rendon Lopez , Elizabeth) Received most parts, 2 demand valves on back order Complete
| 197816 | PA | CNP | ENL | 2023-09-14 13:07 |  | Realize Medical - Elucis VR Software | Software was not purchased under this WO
1067874 | 197817 | ACC | CMPLT | CMPLT | 2023-09-14 13:23 | 2023-09-29 13:29 | Incoming inspection | (TC 09/29/2023 01:29:05 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067875 | 197818 | ACC | CMPLT | CMPLT | 2023-09-14 13:31 | 2023-09-28 14:43 | Incoming inspection | (TC 09/28/2023 02:42:49 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067876 | 197819 | ACC | CMPLT | CMPLT | 2023-09-14 13:32 | 2023-09-28 14:14 | Incoming inspection | (TC 09/28/2023 02:14:03 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067878 | 197820 | ACC | CMPLT | CMPLT | 2023-09-14 13:39 | 2023-09-29 13:55 | Incoming inspection | (TC 09/29/2023 01:55:12 PM, Canjimavadivel, Kamallesh ) Functional test pass
1049885 | 197821 | PM | CMPLT | CMPLT | 2023-09-14 13:41 | 2023-09-14 14:11 | ANNUAL PM | (TC 09/14/2023 01:47:13 PM, Baghirzade, Nara) PM completed. Battery is fully functional. Calibrated CO2 with 5% CO2 gas = 37.7 mmHg NIBP passed. Replaced the BP cuff. BP hose in good condition. Set = 120/80 HR 60, Measured = 119/80 HR 60 Set = 80/40 HR 40, Measured = 78/39 HR 40 SpO2 passed. Set = 99%, Measured = 99% Set = 85%, Measured 85% ECG passed. Set 60 BPM and 20 RESP , Measured 60 BPM and 20 RESP . All leads are good. No noise or artifacts detected. All alarms are functional; tested low and high limits. All passed. Printer is functional. Power cord in good condition. Electrical safety results: Power cord resistance = 62 mohms Current leakage = 55.9 uA
1067879 | 197822 | ACC | CMPLT | CMPLT | 2023-09-14 13:42 | 2023-09-29 14:13 | Incoming inspection | (TC 09/29/2023 02:13:12 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067880 | 197823 | ACC | CMPLT | CMPLT | 2023-09-14 13:43 | 2023-09-29 14:20 | Incoming inspection | (TC 09/29/2023 02:19:55 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067881 | 197824 | CM | CMPLT | CMPLT | 2023-09-14 13:45 | 2023-09-28 13:50 | Incoming inspection | (TC 09/28/2023 01:49:40 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067882 | 197825 | ACC | CMPLT | CMPLT | 2023-09-14 13:47 | 2023-09-28 14:01 | Incoming inspection | (TC 09/28/2023 02:00:42 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067883 | 197826 | ACC | CMPLT | CMPLT | 2023-09-14 13:48 | 2023-09-28 13:36 | Incoming inspection | (TC 09/28/2023 01:35:10 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067884 | 197827 | ACC | CMPLT | CMPLT | 2023-09-14 13:50 | 2023-09-29 14:04 | Incoming inspection | (TC 09/29/2023 02:03:16 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067885 | 197828 | ACC | CMPLT | CMPLT | 2023-09-14 13:51 | 2023-09-29 14:27 | Incoming inspection | (TC 09/29/2023 02:26:16 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067886 | 197829 | ACC | CMPLT | CMPLT | 2023-09-14 13:52 | 2023-09-28 13:42 | Incoming inspection | (TC 09/28/2023 01:41:25 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067887 | 197830 | ACC | CMPLT | CMPLT | 2023-09-14 13:53 | 2023-09-29 14:39 | Incoming inspection | (TC 09/29/2023 02:38:55 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067888 | 197831 | ACC | CMPLT | CMPLT | 2023-09-14 13:54 | 2023-09-29 13:37 | Incoming inspection | (TC 09/29/2023 01:36:24 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067889 | 197832 | ACC | CMPLT | CMPLT | 2023-09-14 13:56 | 2023-10-23 14:39 | Incoming inspection | (TC 10/23/2023 02:38:46 PM, Canjimavadivel, Kamallesh ) Asset tagged and all tests pass
1067890 | 197833 | ACC | CMPLT | CMPLT | 2023-09-14 13:57 | 2023-09-29 14:33 | Incoming inspection | (TC 09/29/2023 02:32:42 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067891 | 197834 | ACC | CMPLT | CMPLT | 2023-09-14 13:59 | 2023-09-28 14:06 | Incoming inspection | (TC 09/28/2023 02:05:15 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067892 | 197835 | ACC | CMPLT | CMPLT | 2023-09-14 14:02 | 2023-09-29 14:01 | Incoming inspection | (TC 09/29/2023 02:00:40 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067893 | 197836 | ACC | CMPLT | CMPLT | 2023-09-14 14:04 | 2023-09-28 14:51 | Incoming inspection | (TC 09/28/2023 02:50:54 PM, Canjimavadivel, Kamallesh ) 1067991
1067894 | 197837 | ACC | CMPLT | CMPLT | 2023-09-14 14:06 | 2023-09-29 13:44 | Incoming inspection | (TC 09/29/2023 01:44:06 PM, Canjimavadivel, Kamallesh ) Functional test pass
1050107 | 197838 | PM | CMPLT | CMPLT | 2023-09-14 14:55 | 2023-09-14 14:57 | ANNUAL PM | (TC 09/14/2023 02:55:30 PM, Baghirzade, Nara) PM completed. Battery is fully functional. Calibrated CO2 with 5% CO2 gas = 37.7 mmHg NIBP passed. BP cuff in good condition, no leaks. BP hose in good condition. Set = 120/80 HR 60, Measured = 120/80 HR 60 Set = 80/40 HR 40, Measured = 80/38 HR 40 SpO2 passed. Set = 99%, Measured = 99% Set = 85%, Measured 85% ECG passed. Set 60 BPM and 20 RESP , Measured 60 BPM and 20 RESP . All leads are good. No noise or artifacts detected. All alarms are functional; tested low and high limits. All passed. Printer is functional. Power cord in good condition. Electrical safety results: Power cord resistance = 84 mohms Current leakage = 55.8 uA
67931 | 197839 | CM | CMPLT | CMPLT | 2023-09-14 15:13 | 2023-09-14 15:16 | Door latch recall and the hood was not level, it was sloping down | (TC 09/14/2023 03:13:39 PM, Campbell, Connor) Replaced all the old door latches with the new ones as per the recall Flipped device onto its side and tightened the base plate at the bottom of the device and added lock tite to the bolts, this made the unit level again
1046274 | 197840 | CM | CMPLT | CMPLT | 2023-09-14 21:30 | 2023-09-14 21:30 | MSH CFRH 7flr APLIO 300 OPERATOR SET BUTTON REPORTED DIFFICULT | (TC 09/14/2023 09:30:07 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 OPERATOR SET BUTTON REPORTED DIFFICULT APPEARED OK | MSH CFRH 7flr APLIO 300 OPERATOR SET BUTTON REPORTED DIFFICULT
32520 | 197841 | CM | CMPLT | CMPLT | 2023-09-14 21:33 | 2023-09-14 21:33 | MSH DR2 EXPOSURE | (TC 09/14/2023 09:33:24 PM, Whittington, Scott) MSH DR2 EXPOSURE | MSH DR2 EXPOSURE
32510 | 197842 | CM | CMPLT | CMPLT | 2023-09-14 21:35 | 2023-09-14 21:35 | MSH DR1 EXPOSURE | (TC 09/14/2023 09:34:44 PM, Whittington, Scott) MSH DR1 EXPOSURE (TC 09/14/2023 09:34:44 PM, Whittington, Scott) MSH DR1 EXPOSURE NEW EXPOSURE SWITCH SOURCED & INSTALLED | MSH DR1 EXPOSURE
| 197843 | OC | CMPLT | CMPLT | 2023-09-15 08:13 | 2023-09-22 08:26 | On Call Week Sept 15-Sept 22 2023 | (TC 09/15/2023 08:20:14 AM, Scero, Tomlyn) Daily checks (TC 09/18/2023 08:37:32 AM, Scero, Tomlyn) daily checks (TC 09/16/2023 08:37:48 AM, Scero, Tomlyn) Called at 9:13pm from CCU regarding their telemetry overview not turning on. While troubleshooting with the nurse we narrowed it down to the display monitor had lost power. I tried to get her to look for a loose connection or loss of power but she couldn't find one. She said I could call back in the morning and if it was on yet I'd have to come in. Tried to login into Citrix to check that tele1O was still running but my log in wouldn't work. Called help desk and had to get them to renew my network access. Once it was renewed, I could get into citrix but I couldn't open the Biomed desktop on my home computer. Called CCU first them in the morning Sunday to speak with Antonella. She found the loose cable and reconnected the power to the monitor. Everything is working fine. (TC 09/19/2023 09:26:36 AM, Scero, Tomlyn) daily checks complete (TC 09/20/2023 10:59:09 AM, Scero, Tomlyn) daily checks (TC 09/21/2023 08:25:57 AM, Scero, Tomlyn) daily checks complete
| 197844 | CM | CMPLT | CMPLT | 2023-09-15 08:26 | 2023-09-15 08:26 | TMS work | TMS work
67931 | 197845 | PM | CMPLT | CMPLT | 2023-09-15 08:47 | 2023-09-15 08:51 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/15/2023 08:49:15 AM, Campbell, Connor) PM complete/passed Functional test passed/complete Refer to procedure for leakage/resistance test
1055513 | 197846 | CM | CMPLT | CMPLT | 2023-09-15 10:08 | 2023-09-15 10:12 | P04 error | (TC 09/15/2023 10:07:57 AM, Campbell, Connor) Tested the thermometer to determine the issue, came up with a p06 error Tested the probe on another working unit, determined the problem was the probe Put a new probe in the unit and performed a full functional check thermometer works as expected now | determined it was not a p04 error it was a p06 error
1062479 | 197847 | CM | CMPLT | CMPLT | 2023-09-15 10:19 | 2023-09-15 10:19 | P04 error/not working | (TC 09/15/2023 10:19:25 AM, Campbell, Connor) Tested the thermometer, determined that the probe was measuring correctly but was taking much too long to obtain measurement, p04 error was not observed tested probe on another known working unit, same issue of slow measurement replaced the probe with a new one thermomter works as expected
1067342 | 197848 | CM | CMPLT | CMPLT | 2023-09-15 10:54 | 2023-09-21 11:14 | high leakage and fluctuating leak | During a case they got “high leakage and fluctuating leak. Vent would pause ventilation then resume”.n Contacted Jason from Getinge. He came to Sinai and collected logs. Started testing and running unit. No Issues. Waiting for Getinge service engineer. Getinge servic eteam came and checked the unit. It is OK to use. Attached service report. SERVICE DESCRIPTION MCC SERVO N - warranty repair SERVICE TICKET DETAILS Issue: High Leakage, Fluctuating leak value. Vent would pause ventilation then resume. Action: Extract and and evaluate the logs for trends. No notable errors reported, logs attached. Completed pre-use checks according to 3.2 Pre-use check of User's Manual SERVO-U Ventilator System v4.0, Infologic 1.93. Ran the vent with patient circuit with no errors. Verifed with several operational modes. No issues. Results: PUC tests passed without errors. Return the unit to clinical use. FAULT CODES RESOLUTION CODES Safe to operate
| 197849 | CM | CMPLT | CMPLT | 2023-09-15 10:56 | 2023-09-15 10:56 | Biomed meeting | Biomed meeting
66488 | 197850 | CM | CMPLT | CMPLT | 2023-09-15 12:51 | 2023-09-15 12:52 | No light coming from the device | (TC 09/15/2023 12:51:07 PM, Ferreira, Qwynn) tested the battery and there was power when switched on but no light source. Found the issue to be the physical switch and replaced the switch. Tested and working fine now
| 197851 | CM | CMPLT | CMPLT | 2023-09-15 13:54 | 2023-09-15 13:54 | testing AGM UAT cerner | (TC 09/15/2023 01:54:32 PM, Samanian, Kamran) done and passed
| 197852 | CM | ACTIV | ISSUE | 2023-09-15 14:01 |  | ESU interference issue with SONY monitors on the boom | (TC 09/14/2023 02:00:55 PM, Samanian, Kamran) received a call from OR14 that monitors get interrupted when ESU is activated however ESU has been plugged in to the wall. by the time arrived to the OR14 room, was told that ESU, BED and IV power bar are plugged into the wall and issue resolved. 0.5 hour later was received another call that issue is still present. Mobile cart brought to the room to use. need Olympus and Building services to look into room grounding (TC 09/15/2023 02:03:20 PM, Samanian, Kamran) olympus and buidling services checked the room and no issue found. only minor as teh olympus processor was a bit out of outlet so was pushed in. jessica was present during the check. she was told by OR that olympus integration team must come in and check the rack. until then the room stays out of service.
1067896 | 197853 | CM | CMPLT | CMPLT | 2023-09-15 14:21 | 2023-09-15 14:23 | Connector leaking | (TC 09/15/2023 02:20:54 PM, Campbell, Connor) tested known good connector and this solved the leaking problem this means the hose was the part leaking, not the device itself leaking hose was replaced perfomed a functional test function is as expected performed a electrical safety test, refer to procedure electrical safety passed
1067897 | 197854 | ACC | CMPLT | CMPLT | 2023-09-15 14:34 | 2023-09-29 14:05 | Incoming inspection | (TC 09/29/2023 02:05:01 PM, Canjimavadivel, Kamallesh ) Functional test pass
1061333 | 197855 | FC | CMPLT | CMPLT | 2023-09-15 15:09 | 2023-09-15 15:11 | QC Checklist | (TC 09/15/2023 03:09:31 PM, Rendon Lopez , Elizabeth) Functional check-OK, defaulted new patient. Pole clamp broken, replaced with spare.
| 197856 | RDS | CMPLT | CMPLT | 2023-09-18 08:12 | 2023-09-18 10:41 | Weekly check for ED
1059657 | 197857 | CM | CMPLT | CMPLT | 2023-09-18 09:11 | 2023-09-29 15:40 | Wires detached from machine | (TC 09/11/2023 09:11:42 AM, Rendon Lopez , Elizabeth) Light connector to base damaged, device is under warranty, contacted company and requested replacement (TC 09/29/2023 12:18:54 PM, Rendon Lopez , Elizabeth) Received new unit, replaced. New LOT 2331 Returned damaged unit to WA
1067900 | 197858 | ACC | CMPLT | CMPLT | 2023-09-18 09:20 | 2023-09-29 14:23 | Incoming inspection | (TC 09/29/2023 02:22:55 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067901 | 197859 | CM | CMPLT | CMPLT | 2023-09-18 09:22 | 2023-09-28 14:19 | Incoming inspection | (TC 09/28/2023 02:19:04 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067902 | 197860 | ACC | CMPLT | CMPLT | 2023-09-18 09:23 | 2023-09-29 14:10 | Incoming inspection | (TC 09/29/2023 02:09:45 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067903 | 197861 | ACC | CMPLT | CMPLT | 2023-09-18 09:24 | 2023-09-29 14:36 | Incoming inspection | (TC 09/29/2023 02:36:12 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067904 | 197862 | CM | CMPLT | CMPLT | 2023-09-18 09:25 | 2023-09-29 13:33 | Incoming inspection | (TC 09/29/2023 01:32:58 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067905 | 197863 | CM | CMPLT | CMPLT | 2023-09-18 09:26 | 2023-09-29 13:41 | Incoming inspection | (TC 09/29/2023 01:40:51 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067906 | 197864 | CM | CMPLT | CMPLT | 2023-09-18 09:27 | 2023-09-28 14:38 | Incoming inspection | (TC 09/28/2023 02:37:44 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067907 | 197865 | ACC | CMPLT | CMPLT | 2023-09-18 09:28 | 2023-09-29 14:17 | Incoming inspection | (TC 09/29/2023 02:16:26 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067908 | 197866 | ACC | CMPLT | CMPLT | 2023-09-18 09:29 | 2023-09-29 14:31 | Incoming inspection | (TC 09/29/2023 02:27:41 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067909 | 197867 | ACC | CMPLT | CMPLT | 2023-09-18 09:31 | 2023-10-23 14:38 | Incoming inspection | (TC 10/23/2023 02:37:55 PM, Canjimavadivel, Kamallesh ) Asset tagged and all tests pass
1067910 | 197868 | ACC | CMPLT | CMPLT | 2023-09-18 09:32 | 2023-09-29 13:26 | Incoming inspection | (TC 09/29/2023 01:24:25 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067911 | 197869 | CM | CMPLT | CMPLT | 2023-09-18 09:34 | 2023-09-29 14:57 | Incoming inspection | (TC 09/29/2023 02:57:03 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067912 | 197870 | ACC | CMPLT | CMPLT | 2023-09-18 09:35 | 2023-09-29 13:49 | Incoming inspection | (TC 09/29/2023 01:48:25 PM, Canjimavadivel, Kamallesh ) Functional test pass
| 197871 | RDS | CMPLT | CMPLT | 2023-09-18 09:39 | 2023-09-20 12:57 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
| 197872 | CM | CMPLT | CMPLT | 2023-09-18 09:44 | 2023-09-18 09:44 | Client down Safetynetview 16N1 | reset both units at 16N. They are working fine. Informed Biomed dept.
| 197873 | RDS | CMPLT | CMPLT | 2023-09-18 09:52 | 2023-09-18 09:54 | Weekly Check - NICU | (TC 09/18/2023 09:54:01 AM, Ferreira, Qwynn) weekly check
| 197874 | RDS | CMPLT | CMPLT | 2023-09-18 09:55 | 2023-09-18 09:57 | Weekly Check - 12S | (TC 09/18/2023 09:55:07 AM, Ferreira, Qwynn) weekly check
| 197875 | RDS | CMPLT | CMPLT | 2023-09-18 09:58 | 2023-09-18 09:59 | Weekly Check - Oncology | (TC 09/18/2023 09:58:50 AM, Ferreira, Qwynn) weekly check
68641 | 197876 | CM | CMPLT | CMPLT | 2023-09-18 10:01 | 2023-09-18 10:03 | Created by Failed PM WO # 197188 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 10:02:13 AM, Campbell, Connor) replaced bottle pressure sensor Performed PM, refer to ASM Pm complete/passed | (TC 09/18/2023 09:58:17 AM, Campbell, Connor) suspected the pressure sensor was fauly, confirmed bottle pressure sensor fatal error in the logs
| 197877 | RDS | CMPLT | CMPLT | 2023-09-18 11:02 | 2023-09-18 11:05 | Weekly Checks For PACU | (TC 09/18/2023 11:03:00 AM, Reid, Emmerson ) Completed checks, PHI was vidible on central monitors. The Central in POD E was was powered down, so I restarted it.
68657 | 197878 | CM | CMPLT | CMPLT | 2023-09-18 11:38 | 2023-09-18 11:41 | Created by Failed PM WO # 197192 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 11:39:17 AM, Campbell, Connor) replaced both pressure sensors PM passed/complete | (TC 09/18/2023 11:32:44 AM, Campbell, Connor) Pm failed due to "check IV set" error and pump sensor fatal errors on both pressure sensors
69456 | 197879 | CM | CMPLT | CMPLT | 2023-09-18 11:55 | 2023-09-18 11:59 | Created by Failed PM WO # 197269 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 11:55:59 AM, Campbell, Connor) calibrated rate PM passed/complete | (TC 09/18/2023 11:54:58 AM, Campbell, Connor) rate test failed
69294 | 197880 | CM | CMPLT | CMPLT | 2023-09-18 12:05 | 2023-09-18 12:11 | Created by Failed PM WO # 191684 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 12:09:26 PM, Campbell, Connor) replaced the bottle pressure sensor membrane housing was cracked, replaced membrane with an already charged, said no charge on the bag it was in PM passed/complete | (TC 09/18/2023 12:03:36 PM, Campbell, Connor) bottle pressure sensor error rate needs calibration after new bottle pressure sensor installed
65387 | 197881 | CM | CMPLT | CMPLT | 2023-09-18 12:39 | 2023-09-18 12:39 | broken door | replaced broken left housing door.
68660 | 197882 | CM | CMPLT | CMPLT | 2023-09-18 12:40 | 2023-09-18 12:50 | Created by Failed PM WO # 193918 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/18/2023 12:48:01 PM, Campbell, Connor) replaced bottle pressure sensor PM still failed due to failing pressure test, even after rate calibration replaced the patient pressure sensor now the PM passed/complete | (TC 09/18/2023 12:39:18 PM, Campbell, Connor) bottle pressure sensor fatal error
| 197883 | RDS | CMPLT | CMPLT | 2023-09-18 12:44 | 2023-09-18 12:45 | Weekly checks in CCU | (TC 09/18/2023 12:44:33 PM, Scero, Tomlyn) Weekly checks complete
1048158 | 197884 | CM | CMPLT | CMPLT | 2023-09-18 12:56 | 2023-09-18 12:57 | Battery replaced still not working | (TC 09/18/2023 12:56:13 PM, Scero, Tomlyn) Blue piece that holds the probe in place and covers the battery wasn't quite on the tracks properly. Causing it to work intermittently. Reseated and tested unit, working fine.
| 197885 | CM | CMPLT | CMPLT | 2023-09-18 14:06 | 2023-09-18 14:06 | Biomed meeting | Biomed meeting
67433 | 197886 | CM | CMPLT | CMPLT | 2023-09-18 14:10 | 2023-09-18 14:11 | Scale not powering on | (TC 09/18/2023 02:09:50 PM, Ferreira, Qwynn) Tested the scale and it was not powering on Checked the connections everything seemed to be good Tested the batteries- low voltage Replaced batteries and working fine. Tested with 5kg = 5kg
1048435 | 197887 | CM | CMPLT | CMPLT | 2023-09-18 15:16 | 2024-04-26 13:45 | MSH NICU MINDRAY ZONARE Z63 SOFTWARE UPGRADE | (TC 09/18/2023 03:16:13 PM, Whittington, Scott) MSH NICU MINDRAY ZONARE Z63 SOFTWARE UPGRADE SOFTWARE ORDERED AND DELIVERED WORK WITH JENNA IN NICU AND BOMIMEND TO ENSURE ALL CLINICAL RESEARCH DATA BEPRESERVED SETTINGS BE PRESERVED (TC 10/25/2023 02:06:31 PM, Whittington, Scott) MSH NICU MINDRAY ZONARE Z63 SOFTWARE UPGRADE JENNA CALLED ME OVER TO START UPGRADE, BUT SYSTEM NEEDED FOR CLIICAL USE (TC 12/04/2023 11:47:19 AM, Whittington, Scott) MSH NICU MINDRAY ZONARE Z63 SOFTWARE UPGRADE SYSTEM BROUGHT TO SHOP, ATTEMPTED TO CONTACT BOMIMED SERVICE, BUT NO RESPONSE (TC 01/31/2024 04:10:56 PM, Whittington, Scott) MSH NICU MINDRAY ZONARE Z63 SOFTWARE UPGRADE AHMAD YOUSIF ON SITE WITH MICHELLE LEE SW ON USB FAULTY, MUST RETURN C 289-541-8087 T 800-667-6276 AYOUSIF@BOMIMED.COM (TC 04/26/2024 01:39:40 PM, Whittington, Scott) SYSTEM HAS TECHNICAL ISSUE, BOMIMED QUITED REPAIR TO JENNA SYSTEM CANNOT RECEIVE SW UPGRADE IN CURRENT FORM | MSH NICU MINDRAY ZONARE Z63 SOFTWARE UPGRADE
1066629 | 197888 | CM | CMPLT | CMPLT | 2023-09-18 15:24 | 2024-09-25 10:41 | MSH NICU MINDRAY ZONARE ZS3 (NEW) NETWORK & PACS CONNECTIONS | (TC 09/20/2023 02:40:52 PM, Whittington, Scott) MSH NICU MINDRAY ZONARE ZS3 (NEW) NETWORK & PACS CONNECTIONS (TC 09/21/2023 01:59:08 PM, Whittington, Scott) MSH NICU MINDRAY ZONARE ZS3 (NEW) NETWORK & PACS CONNECTIONS WIRELESS BRIDGE CEASED FUNCTIONING JUSTIN FIEGE CALLED FROM MSH IT, WANTS FULL NAC & MDSS ON NETWORK CONNECTIONS GOING FWD (TC 01/18/2024 11:41:47 AM, Whittington, Scott) MSH NICU MINDRAY ZONARE ZS3 (NEW) NETWORK & PACS CONNECTIONS MSH NICU HAS PURCHASED AN ADDITIONAL MINDRAY/ZONARE ZS3 ULTRASOUND FOR NEONATAL LUNG SCANNING AND REQUIRES STORAGE OF STUDIES ON JDMIPACS CR 1732 MSH_NICU_Z2 PLEASE REFER TO CR 1122 FOR MSH_NICU_Z1 JAN 18 2024 DISCUSSED WITH DANIEL REYES IT IS P[OSSIBLE THAT GRANTING CR 1122 FOR MSH_NICU_Z1 WAS NOT IN LINE WITH JDMIPACS POLICY, RE THE REPORTING OF CASES FROM POCUS THE VISUALSONICS CR AND THE MSH_NICU_Z2 CR WERE BOTH BASED ON THE PREMISE THAT CR 1122 WAS APPROPRIATE (TC 01/31/2024 04:03:05 PM, Whittington, Scott) MSH NICU MINDRAY ZONARE ZS3 (NEW) NETWORK & PACS CONNECTIONS MSH NICU HAS PURCHASED AN ADDITIONAL MINDRAY/ZONARE ZS3 ULTRASOUND FOR NEONATAL LUNG SCANNING AND REQUIRES STORAGE OF STUDIES ON JDMIPACS CR 1732 MSH_NICU_Z2 PLEASE REFER TO CR 1122 FOR MSH_NICU_Z1 CONNECTED TO WIRELESS NETOWRK VERFIED ALL DICOM DESTINATIONS | MSH NICU MINDRAY ZONARE ZS3 (NEW) NETWORK & PACS CONNECTIONS MSH NICU HAS PURCHASED AN ADDITIONAL MINDRAY/ZONARE ZS3 ULTRASOUND FOR NEONATAL LUNG SCANNING AND REQUIRES STORAGE OF STUDIES ON JDMIPACS CR 1732 MSH_NICU_Z2 PLEASE REFER TO CR 1122 FOR MSH_NICU_Z1
32510 | 197889 | CM | CMPLT | CMPLT | 2023-09-18 16:15 | 2023-09-18 16:15 | MSH DR1 OPG & Guard error. | (TC 09/20/2023 07:31:20 PM, Whittington, Scott) MSH DR1 OPG & Guard error. | MSH DR1 OPG & Guard error. Centred table and did a generator reset.
| 197890 | CM | CMPLT | CMPLT | 2023-09-18 16:17 | 2023-09-18 16:17 | Site Coverage
1050056 | 197891 | CM | CMPLT | CMPLT | 2023-09-19 08:28 | 2023-12-12 08:59 | pen has to be held up on connector | (TC 09/18/2023 08:46:42 AM, Rendon Lopez , Elizabeth) Removed device from OR 1 and placed spare, testing device with student, all values in range, but output connector feels loose. Sending out for repair under consignment. (TC 09/21/2023 03:26:50 PM, Rendon Lopez , Elizabeth) Sent out for repair (TC 11/10/2023 09:35:19 AM, Rendon Lopez , Elizabeth) Device is no longer under consignment. Received quote, approved repair. (TC 12/07/2023 08:58:18 AM, Rendon Lopez , Elizabeth) Received device. Incoming inspection-OK. See attached report Returned to L&D OR5
1067915 | 197892 | CM | CMPLT | CMPLT | 2023-09-19 09:21 | 2023-09-19 09:24 | Alarms when using foot pedal | (TC 09/19/2023 09:21:14 AM, Reid, Emmerson ) Was emailed that this device was alarming when they tried to move it up/down utill it was unplugged. I confirmed this was happening and spoke with tech support. They told me that this was due to a plate getting caught pressed in from under the drawer. This resolved the issue. (TC 09/19/2023 09:21:14 AM, Reid, Emmerson ) Was emailed that this device was alarming when they tried to move it up/down utill it was unplugged. I confirmed this was happening and spoke with tech support. They told me that this was due to a plate getting caught pressed in from under the drawer. This resolved the issue. This device was not asseted either so I created the asset for it
| 197893 | TEDU | CMPLT | CMPLT | 2023-09-19 09:26 | 2023-09-19 09:26 | Patient monitoring scenario | (TC 09/19/2023 09:25:56 AM, Scero, Tomlyn) Patient monitoring scenario complete and passed.
1061214 | 197894 | CM | CMPLT | CMPLT | 2023-09-19 09:41 | 2023-09-19 09:41 | iPad exam light is not working | (TC 09/19/2023 09:40:41 AM, Samanian, Kamran) verified the problem, black screen on the iPad. pressed the home button a couple of times and iPad showed up the program for iLed control, Battery was OK too,
69423 | 197895 | CM | CMPLT | CMPLT | 2023-09-19 09:58 | 2023-09-19 10:06 | Broken door harness assembly, broken door and bezel | (TC 09/19/2023 09:57:38 AM, Campbell, Connor) replaced the bezel because the door harness assembly is part of the bezel replaced the door because the door hinge part was also broken
69173 | 197896 | CM | CMPLT | CMPLT | 2023-09-19 10:44 | 2023-09-19 10:50 | Created by Failed PM WO # 194436 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 10:48:05 AM, Campbell, Connor) replaced bottle pressure sensor replaced male IUI connector PM now complete/passed, refer to ASM | (TC 09/19/2023 10:43:39 AM, Campbell, Connor) IUI male connector severly corroded bottle pressure sensor broken
1064421 | 197897 | CM | CMPLT | CMPLT | 2023-09-19 10:54 | 2023-09-19 11:08 | Patient data cable ripped | (TC 09/19/2023 10:59:01 AM, Canjimavadivel, Kamallesh ) Replaced patient data cable. Tested with simulator, tested ok
69412 | 197898 | PM | CMPLT | CMPLT | 2023-09-19 12:02 | 2023-09-19 12:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 12:01:53 PM, Campbell, Connor) Pump came down with no tag, not sure what the problem was performed a PM to check function and everything was working as expected. PM complete/passed, refer to ASM
69520 | 197899 | PM | CMPLT | CMPLT | 2023-09-19 12:35 | 2023-09-19 12:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/19/2023 12:35:48 PM, Campbell, Connor) came in with no tag performed a PM to check functionality rate needed calibration which was done IUI female connector was not working, replaced it PM now passed/complete, refer to ASM
1055108 | 197900 | CM | CMPLT | CMPLT | 2023-09-19 13:15 | 2023-09-19 13:20 | Not turning on | (TC 09/19/2023 01:18:34 PM, Campbell, Connor) replaced the front panel as per the recall performed a check in, refer to ASM refer to procedure for ESA check in complete/passed | problem is likely due to the front panel recall
1046439 | 197901 | CM | CMPLT | CMPLT | 2023-09-19 14:37 | 2023-09-19 14:39 | Message "Acc F5 0" not working | (TC 09/19/2023 02:37:02 PM, Reid, Emmerson ) ESU was not working and had a message on the screen "Acc F5 0" This was an error relating to the foot pedals, I found after clearing the fuses and MOSFETs. I cleared the errors and tested the unit, also attached now foot pedals. Working now. When putting back, I noticed that the original foot pedal was stuck pressed in by the way it was hanging. This was what caused the issue.
66719 | 197902 | CM | CMPLT | CMPLT | 2023-09-19 15:51 | 2023-09-19 15:51 | Monitor not connecting to fetalink | (TC 09/18/2023 03:50:47 PM, Rendon Lopez , Elizabeth) Troubleshooted, resterted CCE, changed USB plug, tested, association working now. Moniotr-OK
| 197903 | CM | CMPLT | CMPLT | 2023-09-19 16:47 | 2023-09-19 16:47 | Site Coverage.
1066454 | 197904 | FC | CMPLT | CMPLT | 2023-09-20 07:42 | 2023-09-20 07:42 | passed QC check | passed QC check
1066462 | 197905 | FC | CMPLT | CMPLT | 2023-09-20 07:43 | 2023-09-20 07:43 | passed QC check | passed QC check
65903 | 197906 | CM | CMPLT | CMPLT | 2023-09-20 08:44 | 2023-09-20 08:47 | Created by Failed PM WO # 197110 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/15/2023 08:44:24 AM, Rendon Lopez , Elizabeth) Replaced front case (TC 09/18/2023 08:44:24 AM, Rendon Lopez , Elizabeth) During PM fails barrel calmp position test, Replaced sizer sensor under recall (TC 09/19/2023 08:44:24 AM, Rendon Lopez , Elizabeth) Fails plunger sensor, checked, sensor stocked, cleaned, moving freely now. PM completed, all tests passed. | (TC 09/14/2023 10:45:42 AM, Rendon Lopez , Elizabeth) Front case cracked.
69536 | 197907 | CM | CMPLT | CMPLT | 2023-09-20 09:01 | 2023-09-25 11:41 | door latch spring un-seated | (TC 09/20/2023 09:00:37 AM, Campbell, Connor) replaced the door latch spring and screw with new ones door latch now functions as expected
1048500 | 197908 | CM | CMPLT | CMPLT | 2023-09-20 10:08 | 2023-10-02 12:39 | MSH 5 US EPIQ 7 DONNA WATT REPORTED SEVERAL FAULTS | (TC 09/20/2023 10:08:36 AM, Whittington, Scott) MSH 5 US EPIQ 7 DONNA WATT REPORTED SEVERAL FAULTS SITE 75025255 REF# -----Original Message----- From: Watt, Donna <donna.watt@philips.com> Sent: September 18, 2023 12:15 PM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Laing, Cody <cody.laing@philips.com>; Adamjee, Ammar <Ammar.Adamjee@philips.com> Subject: [External] EPIQ Mount Sinai 5th floor main dept Hi Scott/Cody, I am on site at Sinai today as the department trials the mC12-3 transducer for peds. I noticed a few things that need to be adjusted: -The control panel drops on its own when the centre button on the front handle is activated. -The articulating arm does not hold the monitor in place. It needs tightening or replacing. - The steering mechanism does not work properly with the brake pedal is in the lifted position. -I have instructed some of the sonographers on how to clean the trackball and the “cup” that it sits in. - The monitor has a green hue I’m on site until the end of the day. On portables this afternoon. Regards, Donna Donna Watt RDMS, RDCS, RVT, GI Clinical Specialist Philips Healthcare Canada donna.watt@philips.com 416-587-6798 (TC 09/26/2023 08:00:00 AM, Whittington, Scott) MONITOR GREEN TINGE REF# AARON 0122319245 (TC 09/26/2023 01:00:59 PM, Whittington, Scott) MONITOR GREEN TINGE REF# AARON 0122319245 (TC 09/26/2023 08:00:00 PM, Whittington, Scott) MONITOR GREEN TINGE REF# AARON 0122319245 CASTER REPLACEMENT MONITOR/CONTROL PANEL ARM MOVEMEMT REPAIRS | MSH 5 US EPIQ 7 DONNA WATT REPORTED SEVERAL FAULTS SITE 75025255 REF# BRUCE 0122309728 MONITOR GREEN TINGE REF# AARON 0122319245 WHEELS REF# ERAN AARON 0122319246 -----Original Message----- From: Watt, Donna <donna.watt@philips.com> Sent: September 18, 2023 12:15 PM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Laing, Cody <cody.laing@philips.com>; Adamjee, Ammar <Ammar.Adamjee@philips.com> Subject: [External] EPIQ Mount Sinai 5th floor main dept Hi Scott/Cody, I am on site at Sinai today as the department trials the mC12-3 transducer for peds. I noticed a few things that need to be adjusted: -The control panel drops on its own when the centre button on the front handle is activated. -The articulating arm does not hold the monitor in place. It needs tightening or replacing. - The steering mechanism does not work properly with the brake pedal is in the lifted position. -I have instructed some of the sonographers on how to clean the trackball and the “cup” that it sits in. - The monitor has a green hue I’m on site until the end of the day. On portables this afternoon. Regards, Donna Donna Watt RDMS, RDCS, RVT, GI Clinical Specialist Philips Healthcare Canada donna.watt@philips.com 416-587-6798
| 197909 | CM | CMPLT | CMPLT | 2023-09-20 11:06 | 2023-09-20 11:06 | P3A meeting | (TC 09/20/2023 11:05:53 AM, Samanian, Kamran) done!
65921 | 197910 | PM | CMPLT | CMPLT | 2023-09-20 11:20 | 2023-09-20 11:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/20/2023 11:20:40 AM, Campbell, Connor) PM complete/passed, refer to ASM
| 197911 | CM | CMPLT | CMPLT | 2023-09-20 12:51 | 2023-09-20 12:51 | Biomed meeting | Biomed meeting
| 197912 | CM | CMPLT | CMPLT | 2023-09-20 12:53 | 2023-09-20 12:53 | BBQ | BBQ
| 197913 | CM | CMPLT | CMPLT | 2023-09-20 14:24 | 2023-09-21 14:49 | Order items from U-line to Courtney | (TC 09/21/2023 02:48:38 PM, Chukri, Samir) received all items
| 197914 | CM | CMPLT | CMPLT | 2023-09-20 14:42 | 2023-10-13 11:31 | IBE down, Case# 0122310553 | (TC 09/19/2023 02:42:00 PM, Samanian, Kamran) identified no data going to powerchart. IBE was down. trobleshooted and noticed drive D is full. some data deleted from data-messagestore folder to make space in the D drive. It seemed from AUG 23 files were getting created. (TC 09/20/2023 02:43:20 PM, Samanian, Kamran) contacted philips tech support and opened the case # 0122310553 to aks for looking into the IBE server and drive D. (TC 10/13/2023 11:29:44 AM, Samanian, Kamran) the root cause was to have pile of data sitting on old TEST port to old iBUS TEST which was turned off. deleted data. Auto archive changed to two days to keep the log/data in low volume. still manual delete would be needed
501888 | 197915 | CM | CMPLT | CMPLT | 2023-09-20 16:20 | 2023-09-26 19:46 | MSH 5 US RM 2 APLIO 500 REPORTED THAT THE SYSTEM LOADED A WORKLIST DOWNSTAIRS, AND WHEN ROAMED, NO ONGER HAD THE WORKLIST | (TC 09/21/2023 07:45:25 PM, Whittington, Scott) MSH 5 US RM 2 APLIO 500 REPORTED THAT THE SYSTEM LOADED A WORKLIST DOWNSTAIRS, AND WHEN ROAMED, NO ONGER HAD THE WORKLIST | MSH 5 US RM 2 APLIO 500 REPORTED THAT THE SYSTEM LOADED A WORKLIST DOWNSTAIRS, AND WHEN ROAMED, NO ONGER HAD THE WORKLIST
1049901 | 197916 | CM | CMPLT | CMPLT | 2023-09-20 16:52 | 2023-12-28 12:40 | MSH R61 TUBE HEAD COVER MELTED/BURNED UNDER NICU INFANT WARMING SYSTEM LAMP | (TC 09/20/2023 04:53:02 PM, Whittington, Scott) MSH R61 TUBE HEAD COVER MELTED/BURNED UNDER NICU INFANT WARMING SYSTEM LAMP JOHN ORDERED PARTS, DELIVERED, WENT TO NICU SYSTEM IN USE | MSH R61 TUBE HEAD COVER MELTED/BURNED UNDER NICU INFANT WARMING SYSTEM LAMP Part Number Quantity Brief Part Description SPAA9679 1 Cover Assembly
68850 | 197917 | CM | CMPLT | CMPLT | 2023-09-21 09:07 | 2023-09-27 14:29 | missing front panel, display and power cord | (TC 09/21/2023 09:07:27 AM, Campbell, Connor) We have the power cord, front panel and the LCD but some parts are still missing. Missing the molded LCD gasket, Display retainer, and the door harness (ribbon cable). we do not have these parts available to replace them. Emailed BD to enquire about part prices. (TC 09/21/2023 03:04:00 PM, Campbell, Connor) TC10006117 - Display Flex Cable - $51.87 TC10012692 - Molded Gasket - $15.00 TC10007990 - Display retainer - $27.30 (TC 09/27/2023 02:26:49 PM, Campbell, Connor) replaced: power cord LCD LCD gasket Display ribbon cable Display retainer Front case Will perform PM which is overdue to check functionality
1067935 | 197918 | ACC | CMPLT | CMPLT | 2023-09-21 09:18 | 2023-09-29 13:51 | Incoming inspection | (TC 09/29/2023 01:50:54 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067936 | 197919 | ACC | CMPLT | CMPLT | 2023-09-21 09:27 | 2023-09-29 13:52 | Incoming inspection | (TC 09/29/2023 01:51:30 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067937 | 197920 | ACC | CMPLT | CMPLT | 2023-09-21 09:30 | 2023-09-29 13:52 | Incoming inspection | (TC 09/29/2023 01:52:09 PM, Canjimavadivel, Kamallesh ) Functional test pass
| 197921 | CM | CMPLT | CMPLT | 2023-09-21 09:38 | 2023-09-21 09:38 | Ran the fiber cable E in DataCenter to be used for ER East soon | (TC 09/21/2023 09:38:28 AM, Samanian, Kamran) done. removed the older defective one
1053349 | 197922 | CM | CMPLT | NFF | 2023-09-21 10:02 | 2023-11-21 12:00 | MSH BI 1274 H2 CAR SUBMISSION | (TC 09/21/2023 10:03:36 AM, Whittington, Scott) MSH BI 1274 H2 CAR SUBMISSION | MSH BI 1274 H2 CAR SUBMISSION
1057286 | 197923 | CM | CMPLT | CMPLT | 2023-09-21 11:02 | 2023-09-21 11:02 | Replaced battery | (TC 09/21/2023 11:01:52 AM, Chukri, Samir) Replaced battery
| 197924 | CM | CMPLT | CMPLT | 2023-09-21 11:08 | 2023-09-21 11:11 | blood pressure cuff leaking | (TC 09/21/2023 11:08:36 AM, Campbell, Connor) tested the cuff with the pressure guage to determine leakage rate, it passed with less than 6mmHg in a minute cable was damaged, it was disposed of cuff is good
1061333 | 197925 | FC | CMPLT | CMPLT | 2023-09-21 11:18 | 2023-09-21 11:18 | QC check | passed QC check
64293 | 197926 | CM | CMPLT | CMPLT | 2023-09-21 11:23 | 2023-09-21 11:28 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/21/2023 11:27:15 AM, Samanian, Kamran) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609357 Put monitor back in room 84. Tested connection and passed
64262 | 197927 | CM | CMPLT | CMPLT | 2023-09-21 11:31 | 2023-09-21 11:31 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/21/2023 11:30:55 AM, Samanian, Kamran) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609375 Put monitor back in room 82. Tested connection and passed
| 197928 | CM | CMPLT | CMPLT | 2023-09-21 11:31 | 2023-09-21 14:45 | 6th floor urology clinic exam rooms need BP cuffs checked | (TC 09/20/2023 11:33:20 AM, Rendon Lopez , Elizabeth) Dr. Adam Miller reached out, 60 Murray- 6th floor clinic needed to check BP cuffs. (TC 09/21/2023 02:44:24 PM, Rendon Lopez , Elizabeth) Checking and repairing BP devices in clinic
67215 | 197929 | CM | CMPLT | CMPLT | 2023-09-21 11:32 | 2023-09-21 11:32 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/21/2023 11:32:13 AM, Samanian, Kamran) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609383 Put monitor back in room 68. Tested connection and passed
| 197930 | CM | CMPLT | CMPLT | 2023-09-21 12:09 | 2023-09-21 12:09 | patient monitoring scenario/training | (TC 09/21/2023 12:09:18 PM, Campbell, Connor) completed the scenario as supervised by Kam
| 197931 | CM | CMPLT | CMPLT | 2023-09-21 14:42 | 2023-09-21 14:43 | work on stryker lifenet system and correct error in 10 level. | (TC 09/21/2023 02:42:34 PM, Chukri, Samir) working on LifeNet system and working with John to rest errors in 10 levels beds
1050088 | 197932 | CM | CMPLT | CMPLT | 2023-09-21 14:46 | 2023-09-21 14:52 | Breathing bag compressed not inflating | (TC 09/21/2023 02:49:44 PM, Rendon Lopez , Elizabeth) Device is being used for surgery working okay at the time, patient circuit was replaced. (TC 09/21/2023 02:49:44 PM, Rendon Lopez , Elizabeth) Came back to check functions, run leakage test-OK. Self tests passed. No issues found
| 197933 | CM | CANCL | CANCL | 2023-09-21 15:10 |  | Monthly Department meeting
| 197934 | CM | ACTIV | ISSUE | 2023-09-21 15:15 |  | Monthly Meeting | (TC 09/20/2023 03:14:58 PM, Chukri, Samir) Meeting (TC 09/27/2023 12:39:08 PM, Chukri, Samir) Biomed General Meeting (TC 11/20/2023 03:00:11 PM, Chukri, Samir) Cerner RHO meeting (TC 12/20/2023 12:13:22 PM, Chukri, Samir) monthly meeting (TC 02/28/2024 02:46:38 PM, Chukri, Samir) meeting (TC 03/27/2024 01:07:19 PM, Chukri, Samir) PM meeting . (TC 04/24/2024 01:57:03 PM, Chukri, Samir) Monthly PM meeting
| 197935 | CM | CMPLT | CMPLT | 2023-09-21 15:34 | 2023-09-27 08:44 | Spare pump parts | (TC 09/21/2023 03:39:13 PM, Rendon Lopez , Elizabeth) Spare pump parts, will be charged when used Received
1061985 | 197936 | CM | CMPLT | CMPLT | 2023-09-21 15:44 | 2023-09-21 15:46 | Active error code | (TC 09/21/2023 03:44:19 PM, Cortez, John) Bed functional, but has active error code. Power cycled bed and cleared error log
1062023 | 197937 | CM | CMPLT | CMPLT | 2023-09-21 15:50 | 2023-09-21 15:51 | Active error code | (TC 09/21/2023 03:49:40 PM, Cortez, John) Bed functional, but has active error code according to Stryker dashboard. Power cycled bed and cleared error log
1067962 | 197938 | CM | CMPLT | CMPLT | 2023-09-22 08:42 | 2023-10-19 13:28 | Incoming inspection | (TC 09/22/2023 09:40:23 AM, Canjimavadivel, Kamallesh ) Added to TMS and looked for warranty/support information. Need to contact vendor
| 197939 | CM | CMPLT | CMPLT | 2023-09-22 09:49 | 2023-09-22 09:49 | labeling Access Points in ER EAST | (TC 09/22/2023 09:49:05 AM, Samanian, Kamran) done!
1050617 | 197940 | CM | CMPLT | CMPLT | 2023-09-22 10:43 | 2023-09-22 10:45 | does not recognize syringe sizes | (TC 09/22/2023 10:43:38 AM, Campbell, Connor) performed a functional check-in, calibrated the barrel clamp performed a test infusion, no errors, everything functions as expected noticed the housing was cracked on the bottom, replaced with a new housing as per recall
1057511 | 197941 | CM | CMPLT | CMPLT | 2023-09-22 11:00 | 2023-11-06 15:29 | Bipolar Footswitch
65905 | 197942 | CM | CMPLT | CMPLT | 2023-09-22 11:59 | 2023-09-22 12:04 | Channel error message | (TC 09/22/2023 12:00:46 PM, Campbell, Connor) channel error message confirmed error logs show a pressure sensor fatal error sensor had already been reseated 2x in last 2 years due to this issue replaced the pressure sensor ran a test infusion, no more error message functions as expected now
1061210 | 197943 | CM | CMPLT | CMPLT | 2023-09-22 12:20 | 2023-09-22 12:20 | Foot Boom is drifted when moves | (TC 09/22/2023 12:20:04 PM, Samanian, Kamran) verified the problem, performed adjustment
1055803 | 197944 | CM | CMPLT | CMPLT | 2023-09-22 12:39 | 2023-09-22 12:40 | Software upgrade to 02.36.21 | (TC 09/22/2023 12:39:23 PM, Canjimavadivel, Kamallesh ) Upgrade software with the help of Zoll
1055807 | 197945 | CM | CMPLT | CMPLT | 2023-09-22 12:41 | 2023-09-22 12:52 | Software upgrade to 02.36.21 | (TC 09/22/2023 12:41:10 PM, Canjimavadivel, Kamallesh ) Upgraded software with help of Zoll
1061207 | 197946 | CM | CMPLT | CMPLT | 2023-09-22 14:02 | 2023-09-22 14:05 | iPad went to sleep | (TC 09/22/2023 02:02:44 PM, Reid, Emmerson ) iPad was set to sleep after 2 minutes. I changed it to never sleep.
66574 | 197947 | CM | CMPLT | CMPLT | 2023-09-22 14:17 | 2023-09-22 14:20 | Unable to see X2 information | (TC 09/22/2023 02:18:29 PM, Rendon Lopez , Elizabeth) Checked monitor in Procedure room. Found loose connection in FMS cable, reseated. Tested-OK
| 197948 | PA | CMPLT | CMPLT | 2023-09-22 15:22 | 2023-09-28 12:01 | Sterile Water
| 197949 | RDS | CMPLT | CMPLT | 2023-09-25 07:51 | 2023-09-28 10:51 | Weekly check for ED
1054522 | 197950 | CM | CMPLT | CMPLT | 2023-09-25 08:03 | 2023-09-28 07:51 | Screen not working | (TC 09/25/2023 03:00:18 PM, Canjimavadivel, Kamallesh ) Tried to turn on, didn't work. Plugged in and turned on. Battery was dead. (TC 09/26/2023 01:23:10 PM, Canjimavadivel, Kamallesh ) Replaced ripped keyboard cover. Still experiencing charging issues as batteries did not charge. Battery status all good. (TC 09/27/2023 01:59:04 PM, Canjimavadivel, Kamallesh ) Intermittent issue with the buttons on keyboard. Requested quote for replacement. (TC 09/28/2023 07:50:21 AM, Canjimavadivel, Kamallesh ) Reseated keyboard connector. Buttons working okay, to be monitored long term. Returned to unit.
| 197951 | RDS | CMPLT | CMPLT | 2023-09-25 08:11 | 2023-09-25 08:14 | Weekly Checks For PACU | (TC 09/25/2023 08:13:01 AM, Reid, Emmerson ) Had to Remove PHI from POD B and C, as well as turn on two monitors from Pod E. Everything else looks good.
| 197952 | RDS | ACTIV | ISSUE | 2023-09-25 08:15 |  | Weekly Checks For OR
| 197953 | RDS | CMPLT | CMPLT | 2023-09-25 08:16 | 2023-09-25 08:18 | Weekly Checks For OR Pod 2 | (TC 09/25/2023 08:17:39 AM, Reid, Emmerson ) Tightned one monitor in OR8, everything else was good
| 197954 | CM | CMPLT | CMPLT | 2023-09-25 09:06 | 2023-10-10 13:09 | Weekly checks in CCU | (TC 09/25/2023 01:08:35 PM, Scero, Tomlyn) weekly checks complete
69580 | 197955 | CM | CMPLT | CMPLT | 2023-09-25 09:09 | 2023-09-25 09:25 | Created by Failed PM WO # 197283 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/25/2023 09:24:11 AM, Campbell, Connor) Replaced display board Replaced front case PM complete/passed, refer to ASM | (TC 09/25/2023 09:07:46 AM, Campbell, Connor) Display board top right segment is not working, need display replaced Front case has cracked plastic from being dropped, needs replacement to prevent liquid ingress
1062005 | 197956 | CM | CMPLT | CMPLT | 2023-09-25 10:09 | 2024-01-12 10:00 | error 3009-5-7-1203 | (TC 09/25/2023 10:08:57 AM, Cortez, John) "Electric brakes-motor drive Limited brake functionality" Found error in dashboard, Power cycled bed, Will monitor to see if error comes back
| 197957 | CM | CMPLT | CMPLT | 2023-09-25 10:19 | 2024-05-28 13:35 | EMG troubleshooting | (TC 09/25/2023 10:19:07 AM, Baghirzade, Nara) Ran the HW programmer on EMG MAIN cart; ran HW diagnostics; all tests passed. Saved log to local database. Checked the split drives; reports are being stored locally and on the server but the .doc format is only stored locally. Tech and Dr. both need access to the reports folder. Data is pushed to the server in SQL and info is altered on SQL level but the reports are not getting updated on the server. Checked the report export path and it is being saved the EmgNetData folder on the server. Created a test patient and report; confirmed that the report is saved locally only on the cart that the test was done on. Could not see the report or patient folder on the server however i could access the test patient from both carts. Will troubleshoot this with Kamran first and then possibly get IT and Natus involved as well. (TC 09/25/2023 11:03:57 AM, Baghirzade, Nara) Used the same test patient to put in a comment in the conclusion; it saved and i was able to access the report with the new comment on the second EMG cart. After editing the report again, it prompts to save it as a second file in any location, does not automatically override and update in the reports folder. Tried this on both carts and had the same outcome. Checked the report export path mapping on both carts and they are the same: \\mshemg01\emgnetdata\reports however any reports AFTER june 2023 (new install) are not there and not in patients folder either. (TC 09/25/2023 02:00:00 PM, Baghirzade, Nara) I was able to edit reports on individual carts - confirmed that the changes are being autosaved in to the report and accessible from both carts. Set the path to C:\ drive and created a new test however the report does not get exported to the set path. Auto-export did not work. Checked all the options on Word document; ensured that all macros are enabled and checked Word settings on both carts to ensure all the settings are the same. Will contact Natus for further support on auto-export. (TC 02/13/2024 02:22:41 PM, Baghirzade, Nara) Vinod called to report that on 3 separate times the amplifier was stalling and required multiple reboots to work with the needle already inserted in the patient. Followed up with Andrew to let him know of the issues as well as to find out about the progression of the request to purchase on capital. EMG is on one of the highest priorities, most likely in the summer time will purchase new amplifiers. Called vinod back to update. (TC 05/02/2024 11:58:52 AM, Baghirzade, Nara) Meeting with Vinod and Andrew regarding amplifier issues on both EMGs (TC 05/03/2024 01:06:05 PM, Baghirzade, Nara) Met with Vinod to address his concerns with amplifier in room 1643; took the amplifier down to biomed, opened it, and re-soldered the on/off connector site on the board. Vinod to test and let me know if its working. (TC 05/28/2024 01:34:49 PM, Baghirzade, Nara) Capital was approved for EMG hardware replacement.
| 197958 | RDS | CMPLT | CMPLT | 2023-09-25 10:54 | 2023-09-25 10:56 | Weekly Check - NICU | (TC 09/25/2023 10:55:28 AM, Ferreira, Qwynn) Weekly Check for the NICU and Resus
1054039 | 197959 | CM | CMPLT | CMPLT | 2023-09-25 10:55 | 2023-10-19 15:03 | needs calibration - RMA 316678 | (TC 09/25/2023 10:55:36 AM, Campbell, Connor) talked and emailed with smiths medical cutomer service got an RMA form sent to me, will forward to samir once the helpdesk releases the email attachment to me (TC 09/28/2023 03:32:18 PM, Chukri, Samir) Called and received email from Eric, RMA 316678 waiting for the PO # (TC 10/18/2023 02:03:57 PM, Campbell, Connor) Recieved device back from ICU medical, it was calibrated refer to the documents attached to this WO for the documentation
| 197960 | RDS | CMPLT | CMPLT | 2023-09-25 10:57 | 2023-09-25 10:58 | Weekly Check - Oncology | (TC 09/25/2023 10:57:54 AM, Ferreira, Qwynn) Weekly Check Oncology
| 197961 | RDS | CMPLT | CMPLT | 2023-09-25 10:59 | 2023-09-25 11:01 | Weekly Check - 12South | (TC 09/25/2023 11:00:16 AM, Ferreira, Qwynn) Weekly Check - 12south
| 197962 | CM | CMPLT | CMPLT | 2023-09-25 11:04 | 2023-09-25 11:04 | biomed meeting | meeting
| 197963 | RDS | CMPLT | CMPLT | 2023-09-25 11:07 | 2023-09-25 11:07 | weekly check for 14S and 16N | weekly check for 14S and 16N
| 197964 | CM | CMPLT | CMPLT | 2023-09-25 11:13 | 2023-09-25 11:13 | 14S suction issue | 1410B intermittent suction is not working. did troubleshooting with nurses. at the end I gave them a brand new suction regulator and explained them if it is still not working there is a leak on patient set.
1059828 | 197965 | CM | CMPLT | CMPLT | 2023-09-25 11:26 | 2023-09-25 11:30 | Cleaning | (TC 09/25/2023 11:29:07 AM, Campbell, Connor) Disinfected, cleaned, rinsed and refilled
501585 | 197966 | CM | CMPLT | CMPLT | 2023-09-25 12:03 | 2024-02-26 13:34 | MSH OPG SPP ROOM C2 IU22 TCS GOES OUT OF CALIBRATION THROUGH THE DAY | (TC 09/25/2023 12:04:02 PM, Whittington, Scott) MSH OPG SPP ROOM C2 IU22 TCS GOES OUT OF CALIBRATION THROUGH THE DAY SITE# 52668736 (TC 09/25/2023 12:04:02 PM, Whittington, Scott) MSH OPG SPP ROOM C2 IU22 TCS GOES OUT OF CALIBRATION THROUGH THE DAY SITE# 52668736 DELETE PT DB DEFRAG HDD UNABLE TO ACCESS SOFT KEYS AT BOTTOM OF TCS, LATER IN DAY | MSH OPG SPP ROOM C2 IU22 TCS GOES OUT OF CALIBRATION THROUGH THE DAY SITE# 52668736 REF# 0122333572
68034 | 197967 | CM | CMPLT | CMPLT | 2023-09-25 12:57 | 2023-09-26 08:18 | Latch is broken and missing door | (TC 09/25/2023 08:16:38 AM, Ferreira, Qwynn) Replaced latch and door Tilt was tight so lubricated the tilt. Functional Check - OK Temp 35 - 35.2 RH 75 - 75.8%
67813 | 197968 | PM | CMPLT | CMPLT | 2023-09-25 13:21 | 2023-09-25 13:23 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/25/2023 01:21:20 PM, Scero, Tomlyn) PM tests complete. Leakage tests passed. Battery replaced.
66791 | 197969 | PM | CMPLT | CMPLT | 2023-09-25 13:24 | 2023-09-25 13:26 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/25/2023 01:24:39 PM, Scero, Tomlyn) PM tests done. Leakage tests passed. battery replaced.
68975 | 197970 | CM | CMPLT | CMPLT | 2023-09-25 14:48 | 2023-09-25 14:50 | IUI connector release button was sticking | (TC 09/25/2023 02:48:21 PM, Campbell, Connor) Took the connector off replaced the metal bracket with a spare part, no charge reseated the connector release button and tested it, works much better
| 197971 | CM | CMPLT | CMPLT | 2023-09-25 14:55 | 2023-09-25 14:55 | working with Nara on EMG and Report | (TC 09/25/2023 02:55:33 PM, Samanian, Kamran) done!
64088 | 197972 | CM | CMPLT | CMPLT | 2023-09-25 14:56 | 2023-09-25 15:02 | BAT replacement | (TC 09/25/2023 03:02:20 PM, Samanian, Kamran) done!
| 197973 | CM | CMPLT | CMPLT | 2023-09-25 15:03 | 2023-10-13 11:32 | working with Brandon on Philips ER EAST | (TC 09/25/2023 03:03:19 PM, Samanian, Kamran) - Fiber connected to Core Routers in MDC - switch, UPS and Sync were installed and pacthed - Switched was configured - APs are good. - 3 x PICiX were installed and configured - Monitors were configured to N.04.104 (TC 10/02/2023 11:32:08 AM, Samanian, Kamran) labeling done
65985 | 197974 | CM | CMPLT | CMPLT | 2023-09-25 15:54 | 2023-09-26 13:38 | Pressure sensor fatal error | (TC 09/25/2023 03:54:25 PM, Campbell, Connor) determined the pressure sensor was faulty from the error log, it couldnt be calibrated at all replaced the pressure sensor performed a check in to ensure function need to run functional test in infusion mode to double check no errors, will update once complete (TC 09/26/2023 08:35:13 AM, Campbell, Connor) calibrated pressure sensor and ran check in again ran test infusion, no issues found now
1067975 | 197975 | ACC | CMPLT | CMPLT | 2023-09-25 16:05 | 2023-12-28 12:39 | MSH ER US CANON PRISM i800 ACCEPTANCE | (TC 09/25/2023 04:05:39 PM, Whittington, Scott) MSH ER US CANON PRISM i800 ACCEPTANCE (TC 09/25/2023 04:05:39 PM, Whittington, Scott) MSH ER US CANON PRISM i800 ACCEPTANCE PO# 7600042543 (TC 09/26/2023 12:52:05 PM, Whittington, Scott) MSH ER US CANON PRISM i800 ACCEPTANCE (TC 10/02/2023 05:14:34 PM, Heron, Chris) MSH ER US CANON PRISM i800 ACCEPTANCE WORK WITH ASHLEY SISETO & RYAN FRANCIS FOR PREP FOR LUNCH & LEARN (TC 10/17/2023 06:49:13 PM, Whittington, Scott) MSH ER US CANON PRISM i800 ACCEPTANCE CHECK ROOM G7 IN NEW MSH ER FOR US USE (TC 10/24/2023 07:36:38 AM, Whittington, Scott) MSH ER US CANON PRISM i800 ACCEPTANCE SYSTEM MOVING TO ER G7 | MSH ER US CANON PRISM i800 ACCEPTANCE 57 mOHM 270uA
1059711 | 197976 | CM | CMPLT | CMPLT | 2023-09-26 07:56 | 2023-09-26 08:00 | Arm loose | (TC 09/26/2023 07:56:37 AM, Reid, Emmerson ) The arm was loose as one of the screws came out. I tightened and added locktight but the brushing inside has come loose. I added two more screws with locktight to compensate. It is secured now.
1060032 | 197977 | CM | CMPLT | CMPLT | 2023-09-26 09:15 | 2024-05-02 10:25 | DLCO is reading high | (TC 09/25/2023 02:30:00 PM, Baghirzade, Nara) When doing diffusion, the DLCO values come back high. DLCO should not be higher than 0.5 mL/min/mmHg however after checking the results dating back to April, it was noted that DLCO varies from 0.7 to 1.5 mL/min/mmHg. The reading is too high. There could be a potential leak in the umbilical or the housing where the pneuomatch is housed in. Gaby will coordinate with me when there is availability to further troubleshoot the system. (TC 10/05/2023 11:46:07 AM, Baghirzade, Nara) Spoke to Antony from Novus and had sent him an email regarding all the issues that are occurring on PFT devices. He will forward my email to Medisoft. Looked for fluid in tubing as Antony suggested but all tubing looks clear. Also filter is used on all patient tests so there should be nothing passing to the tubings. Waiting for further instructions from Antony and Medisoft. (TC 11/08/2023 08:51:31 AM, Baghirzade, Nara) Smain remoted to the micro6000 station as its networked with all PFT stations; found the inconsistencies through the calibrations results. Possible that the analyzer does not work. Gaby was also present for these observations, she showed him the exact days when they do have issues with the DLCO. Monday november 13, will be remoting to Hypair station specifically to test the analyzer. (TC 11/13/2023 09:17:38 AM, Baghirzade, Nara) Smain remoted in to the Hypair station. Gaby performed the daily and weekly cal; the daily cal passes however the weekly cal continues to fail. Looking at the inspired/expired He and CO, Smain deduced that there could be a leak either internally or externally (patient box). Antony to send over instructions from the manual on leakage tests and Samir and I to perform the leakage tests. Will coordinate the work to be done after hours. (TC 01/11/2024 10:13:19 AM, Baghirzade, Nara) Troubleshooting with Smain, Samir, and Antony. Reviewed past calibrations and bodybox calibrations as well. Asked Gaby to perform SVC and DLCO cals; at first it was reading high, then the CO was down to 0.7 however the analyser stops analysing gas at the beginning of the test. Also the diffusion is higher on Hypair than on DLCO. Also let Smain know that we had positive pressure when doing a leakage test which he suspects is from the Pistons. Requested that we repeated the leakage test. Also will order a new patient valve. (TC 01/11/2024 10:25:09 AM, Chukri, Samir) Troubleshooting with Medisoft and Novus. Reviewed past calibrations and bodybox calibrations as well. Asked Gaby to perform SVC and DLCO cals; at first it was reading high, then the CO was down to 0.7 however the analyser stops analysing gas at the beginning of the test. Also the diffusion is higher on Hypair than on DLCO. Also let Smain know that we had positive pressure when doing a leakage test which he suspects is from the Pistons. Requested that we repeated the leakage test. Also will order a new patient valve. (TC 01/11/2024 01:13:00 PM, Baghirzade, Nara) Received quote from Chantal. Had to send picture of the patient valve. (TC 01/25/2024 09:26:40 AM, Baghirzade, Nara) Replaced patient valve. Jen did the calibration and all values are within specifications as well as repeatable. None of the values for DLCO were above 1. He&CO inspired and expired were correct. Gaby then did 2 patient tests on herself, both were within 1L for DLCO (basically repeatable values) and she mentioned it was working as it should be. The analyzer circuit was working in all cals and tests; no drop offs for the analyzer circuit. John and I noted that there is a slight gap between the heater and patient valve which could be causing the slight increase in DLCO values (0.8 was the resullt however 0.5 or lower is the optimal). The next day John and I will replace the heater and Gaby will do another full cal along with a DLCO patient test to observe if the values are staying the same or lower. (TC 01/26/2024 11:19:26 AM, Baghirzade, Nara) Replaced the heater on the pneumotachograph. Updated the tech at PFT lab as well as Gaby that the heater requires 10-20 minutes to heat up and then a full cal (SVC & DLCO) needs to be performed on the system again. (TC 05/02/2024 10:23:47 AM, Baghirzade, Nara) Followed up with Gaby and other PFT techs who work on the Hypair system. Since the replacement of the patient valve, the discrepancy in DLCO values has been minimal. Ranges stay below 1 and does not exceed 1. No issues with the system.
1060030 | 197978 | CM | CMPLT | CMPLT | 2023-09-26 09:24 | 2024-02-02 07:22 | Possible Helium leak and loose arm | (TC 09/25/2023 03:30:00 PM, Baghirzade, Nara) When doing daily calibrations, the He sensor keeps giving a red light and failing the calibrations. As a result, Gaby has to wiggle the gas tank connections as well as put the arm in a specific position for the He to pass the daily cal. There could be a possible leak in the tank or connectors. The arm also fell off during a patient test however Gaby was able to place it back on but it is still too loose/wobbly. Gaby to coordinate with me when the system becomes available for troubleshooting/repairs. (TC 10/04/2023 07:55:24 AM, Baghirzade, Nara) Replaced the fuse on the heater board. Original fuse was blown. Powered the entire system down. The fuse I placed on the board is 800mA / 250V Booted the system back up. Will check the fuse at a later date to ensure that it is still functional. (TC 10/05/2023 12:39:26 PM, Baghirzade, Nara) When doing DLCO, during breath hold and exhale the He and CO do not diffuse in a linear fashion. Another issue could be that there is not enough pressure coming from device to fully close the pistons. Jen did a bio-test and said there was barely any resistance when trying to exhale when the pistons are closed. Removed the patient valve, samir cleaned it with alcohol, we cut the sample line tube as it was getting kinked. After placing the valve back, did another bio-test and the He and CO were diffusing properly but there is still a very slight leak. Will stay late one night to do a leak test and pressure test on the board side. (Samir and John will both work with me on this) (TC 10/05/2023 04:14:39 PM, Cortez, John) Went with Nara and Samir to assess issue (TC 02/02/2024 07:20:06 AM, Baghirzade, Nara) Met with Gaby to check on the body box system. Stayed during her cal and all the cal numbers were within nominal range. Looked at some previous tests with her and did not find discrepancies in their test numbers. The full monday cal is also within range; no complaints about the calibration numbers or patient test numbers.
| 197979 | CM | CMPLT | CMPLT | 2023-09-26 09:27 | 2023-09-26 09:29 | Olympus camerahead missing screws | (TC 09/25/2023 12:00:00 PM, Baghirzade, Nara) The camera head was missing 2/3 screws with the springs still in tact. Tried to look for the screws on the ground but was unable to locate them. Checked in biomed but we do not have screws that are small enough. Also noted that the cord is damaged at the side closest to the camera head. Returned the camera head back to Ana and had informed them that they need to send it out for repair or purchase a new one.
1050479 | 197980 | PM | CMPLT | CMPLT | 2023-09-26 11:23 | 2023-09-26 11:23 | PM-MAJOR- Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | performed annual PM. Tested. Functional.
66466 | 197981 | CM | CMPLT | CMPLT | 2023-09-26 11:25 | 2023-10-31 11:56 | unit is alarming | C500 unit here it is not booting up at all. Infinity C500 MK31500-07, sn.350145068 from VN500, sn.ASFH-0024, asset#66466 I replaced the cable first(from C500 to VN500) then installed a new SSD and replaced BIOS coin battery but it is still dead. I contacted Draeger to ship C500 for repair (TC 09/29/2023 12:07:46 PM, Rendon Lopez , Elizabeth) Shipped device. Return Reference number is 308620811 Received unit. PO#7500195322. installed and tested 24 hours. no issue. returned back to RT. informed Rex. Draeger replaced Basis PCB and SSD. Attached service report.
| 197982 | CM | CMPLT | CMPLT | 2023-09-26 11:28 | 2023-09-25 11:28 | O2 regulator not working | Checked the unit with the nurse. they removed and installed another one
68598 | 197983 | CM | CMPLT | CMPLT | 2023-09-26 11:37 | 2023-09-26 11:40 | Plunger is sticky and the bottom of the case is cracked | (TC 09/26/2023 11:37:01 AM, Campbell, Connor) replaced the back case as per the recall dissasembled the entire pump, cleaned the guide rail and syringe pump mechanism, lubricated it very slightly with WD 40 no change was noticed in the smoothness of the plunger mechanism even after cleaning every part after spending so much time on it, elizabeth and I decided to put it back together and go ahead with the PM to see if it passes, as nothing more can be done about the smoothness of the plunger mechanism
1067979 | 197984 | ACC | CMPLT | CMPLT | 2023-09-26 12:27 | 2024-02-26 13:34 | MSH 20 DOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE | (TC 09/26/2023 12:48:03 PM, Whittington, Scott) MSH 20 DIOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE ASSET MIF CR (TC 09/26/2023 12:48:03 PM, Whittington, Scott) MSH 20 DIOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE ASSET MIF CR (TC 12/20/2023 10:20:54 AM, Whittington, Scott) MSH 20 DIOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE SHOW ASHLEY SISETO LOCATION MEET WITH GABBY GIANNETTI, MGR (TC 12/28/2023 08:50:29 AM, Whittington, Scott) MSH 20 DIOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE (TC 01/08/2024 04:08:00 PM, Whittington, Scott) MSH 20 DOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE NETWORK JACKS ARE DISCONNECTED PERFORMED NETWORK , PACS TEST IN SHOP OK (TC 01/08/2024 05:10:00 PM, Whittington, Scott) MSH 20 DOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE NETWORK JACKS ARE DISCONNECTED PERFORMED NETWORK , PACS TEST IN SHOP OK (TC 01/09/2024 10:18:27 AM, Whittington, Scott) MSH 20 DOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE ASHLEY SISETO ON SITE FOR APPS | MSH 20 DOVIGI SPORTS CLINIC CANON APLIO i800 PRISM ULTRASOUND ACCEPTANCE From: Whittington, Scott Sent: January 8, 2024 4:57 PM To: HELP Desk (MSH) <helpdesk.MSH@sinaihealth.ca> Cc: Riaz, Mirza - Sinai Health System <Mirza.Riaz@sinaihealth.ca> Subject: MSH 20 Sports clinic. Connect Network Jacks and activate on MEDICAL IMAGING VLAN MSH 20 Sports clinic. ReConnect previously connected Network Jacks and activate on MEDICAL IMAGING VLAN 20SW-043-D 20SW-044-D Thanks Scott 416-402-2462
68598 | 197985 | CM | CMPLT | CMPLT | 2023-09-26 12:48 | 2023-09-29 11:05 | Created by Failed PM WO # 196404 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/29/2023 10:21:45 AM, Campbell, Connor) replaced the syringe plunger assembly plunger now moves much more freely and is useable again calibrated the unit check in the unit to clear the error PM complete/passed, refer to ASM | (TC 09/26/2023 12:47:00 PM, Campbell, Connor) Plunger position failed in PM Plunger position cannot be calibrated The plunger is much too sticky and makes it impossible to use Ordering a new plunger, will perform the PM and functional check when it arrives
66056 | 197986 | PM | CMPLT | CMPLT | 2023-09-26 13:22 | 2023-09-26 13:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 09/26/2023 01:30:04 PM, Campbell, Connor) PM complete/passed, refer to ASM For leakage current and resistance refer to procedure Conditioned the battery
64203 | 197987 | CM | CMPLT | CMPLT | 2023-09-26 13:25 | 2023-09-26 13:26 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:25:25 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609350 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64274 | 197988 | CM | CMPLT | CMPLT | 2023-09-26 13:28 | 2023-09-26 13:28 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:27:40 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609326 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64302 | 197989 | CM | CMPLT | CMPLT | 2023-09-26 13:28 | 2023-09-26 13:29 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:28:34 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609382 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64206 | 197990 | CM | CMPLT | CMPLT | 2023-09-26 13:32 | 2023-09-26 13:32 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:31:57 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609343 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64257 | 197991 | CM | CMPLT | CMPLT | 2023-09-26 13:33 | 2023-09-26 13:33 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:32:41 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609374 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64285 | 197992 | CM | CMPLT | CMPLT | 2023-09-26 13:33 | 2023-09-26 13:34 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:33:25 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609380 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
1054741 | 197993 | CM | CMPLT | CMPLT | 2023-09-26 13:35 | 2023-09-26 13:36 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:35:02 PM, Baghirzade, Nara) Monitor already had existing MIB RS232 board. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232.
64298 | 197994 | CM | CMPLT | CMPLT | 2023-09-26 13:36 | 2023-09-26 13:37 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:36:24 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609381 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64207 | 197995 | CM | CMPLT | CMPLT | 2023-09-26 13:37 | 2023-09-26 13:37 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:37:08 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609318 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64265 | 197996 | CM | CMPLT | CMPLT | 2023-09-26 13:38 | 2023-09-26 13:38 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:38:04 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609325 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64275 | 197997 | CM | CMPLT | CMPLT | 2023-09-26 13:39 | 2023-09-26 13:39 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:38:58 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609328 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64294 | 197998 | CM | CMPLT | CMPLT | 2023-09-26 13:40 | 2023-09-26 13:42 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:40:18 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609379 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64276 | 197999 | CM | CMPLT | CMPLT | 2023-09-26 13:45 | 2023-09-26 13:46 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:45:59 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609355 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64296 | 198000 | CM | CMPLT | CMPLT | 2023-09-26 13:49 | 2023-09-26 13:50 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:49:38 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609313 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64295 | 198001 | CM | CMPLT | CMPLT | 2023-09-26 13:50 | 2023-09-26 14:02 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 01:54:28 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609384 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
1046274 | 198002 | CM | CMPLT | CMPLT | 2023-09-26 14:01 | 2023-09-26 14:02 | MSH IVF - Aplio 300 - Set Button Malfunction | (TC 09/26/2023 02:01:02 PM, Calauna, Carlos) replaced operator membrane and tested okay
64279 | 198003 | CM | CMPLT | CMPLT | 2023-09-26 14:03 | 2023-09-26 14:03 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:02:37 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609356 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64283 | 198004 | CM | CMPLT | CMPLT | 2023-09-26 14:03 | 2023-09-26 14:04 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:03:41 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609373 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64202 | 198005 | CM | CMPLT | CMPLT | 2023-09-26 14:05 | 2023-09-26 14:05 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:05:00 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609334 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64300 | 198006 | CM | CMPLT | CMPLT | 2023-09-26 14:06 | 2023-09-26 14:06 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:05:47 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609315 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64286 | 198007 | CM | CMPLT | CMPLT | 2023-09-26 14:07 | 2023-09-26 14:07 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:06:43 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609349 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64282 | 198008 | CM | CMPLT | CMPLT | 2023-09-26 14:07 | 2023-09-26 14:09 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:07:33 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609336 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64247 | 198009 | CM | CMPLT | CMPLT | 2023-09-26 14:10 | 2023-09-26 14:11 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:10:27 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609320 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
65974 | 198010 | CM | CMPLT | CMPLT | 2023-09-26 14:17 | 2023-09-27 09:36 | Created by Failed PM WO # 194348 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/26/2023 03:36:50 PM, Campbell, Connor) replaced front case replaced pressure sensor calibrated the unit (TC 09/27/2023 09:34:51 AM, Campbell, Connor) had to re-seat the gasket over the pressure sensor re-calibrated the device did a check in to remove the software error after calibration PM complete/passed, refer to ASM | (TC 09/26/2023 02:16:08 PM, Campbell, Connor) Pressure sensor not working at all front case standoff is cracked on the inside near the pressure sensor
64299 | 198011 | CM | CMPLT | CMPLT | 2023-09-26 14:20 | 2023-09-26 14:20 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:19:45 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609335 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64255 | 198012 | CM | CMPLT | CMPLT | 2023-09-26 14:21 | 2023-09-26 14:21 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:20:36 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609365 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64277 | 198013 | CM | CMPLT | CMPLT | 2023-09-26 14:22 | 2023-09-26 14:22 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:21:41 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609331 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
67214 | 198014 | CM | CMPLT | CMPLT | 2023-09-26 14:22 | 2023-09-26 14:23 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:22:30 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609385 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64271 | 198015 | CM | CMPLT | CMPLT | 2023-09-26 14:23 | 2023-09-26 14:24 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:23:14 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609363 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64281 | 198016 | CM | CMPLT | CMPLT | 2023-09-26 14:24 | 2023-09-26 14:24 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:23:54 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609344 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64287 | 198017 | CM | CMPLT | CMPLT | 2023-09-26 14:24 | 2023-09-26 14:25 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/26/2023 02:24:27 PM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609369 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
1061211 | 198018 | CM | ACTIV | ISSUE | 2023-09-26 15:12 |  | Recorder Memory Full | (TC 09/26/2023 03:12:14 PM, Reid, Emmerson ) nCare was not able to start a recording because there was no available space. It looked like there was 46 consecutive blank-12 hour recordings that filled up the local memory. I worked with tech support after the case to locally delete files from 2021 to clear up some space for now.
1062999 | 198019 | PM | CMPLT | CMPLT | 2023-09-26 15:52 | 2023-09-26 15:53 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 09/26/2023 03:52:34 PM, Cortez, John) Completed cleaning cycle
1062028 | 198020 | CM | CMPLT | CMPLT | 2023-09-26 16:11 | 2023-09-26 16:12 | error 3009-3-5-1370 | (TC 09/26/2023 04:10:50 PM, Cortez, John) Bed functional, but has active error code according to Stryker dashboard. Power cycled bed and cleared error log Will monitor dashboard to see if error returns
1055644 | 198021 | CM | CNP | EIU | 2023-09-26 16:15 |  | Active error, wrench (service) LED lit up | (TC 09/25/2023 04:15:50 PM, Cortez, John) error 1041 found in logs. Refers to low battery voltage Bed was moved before I could replace battery
1067052 | 198022 | CM | CMPLT | CMPLT | 2023-09-26 16:18 | 2023-10-18 09:13 | Active error, wrench (service) LED lit up | (TC 09/25/2023 04:18:26 PM, Cortez, John) error 1041 found in logs. Refers to low battery voltage Power cycling bed doesn't resolve issue (TC 10/17/2023 09:12:29 AM, Cortez, John) Resolved during PM, WO 198534
1067048 | 198023 | CM | CMPLT | CMPLT | 2023-09-26 16:20 | 2023-10-05 16:10 | Active error, wrench (service) LED lit up | (TC 09/25/2023 04:20:25 PM, Cortez, John) error 1041 found in logs. Refers to low battery voltage Power cycling bed doesn't resolve issue Patient on bed (TC 10/05/2023 04:10:00 PM, Cortez, John) see wo #198372
1055647 | 198024 | CM | CMPLT | CMPLT | 2023-09-26 16:21 | 2023-10-13 15:27 | Active error, wrench (service) LED lit up | (TC 09/26/2023 04:21:36 PM, Cortez, John) error 1041 found in logs. Refers to low battery voltage Cleared error logs. Will see if error returns tomorrow (TC 10/13/2023 03:24:43 PM, Cortez, John) Error returned, but replaced battery as part of PM Tested ok
66680 | 198025 | CM | CMPLT | CMPLT | 2023-09-27 07:36 | 2023-09-27 09:17 | "air probe failure" | (TC 09/27/2023 09:15:37 AM, Rhuland, Courtney) Air probe failure alarm ocurred during functional check Replaced air probe Replaced 4 porthole door springs RH calibrated Functional check - OK
1063746 | 198026 | CM | CMPLT | CMPLT | 2023-09-27 09:22 | 2023-09-27 09:23 | Cleaning | (TC 09/27/2023 09:22:19 AM, Campbell, Connor) Drained, disinfected, rinsed and refilled with sterile water
1064418 | 198027 | PA | CMPLT | CMPLT | 2023-09-27 09:24 | 2024-04-03 10:51 | MRI Doors
1065639 | 198028 | ACC | CMPLT | CMPLT | 2023-09-27 10:01 | 2023-09-27 10:04 | Incoming Inspection | (TC 09/27/2023 10:02:48 AM, Reid, Emmerson ) Incoming complete, spoke with Maureen who said she wanted to put the ESU on a cart instead of the tower or boom. I found a cart in OR12 and removed the ESU from the tower. Did electrical safety, created PM schedule and added assignments. Incoming complete.
64258 | 198029 | CM | CMPLT | CMPLT | 2023-09-27 10:09 | 2023-09-27 10:13 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:09:39 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609316 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64256 | 198030 | CM | CMPLT | CMPLT | 2023-09-27 10:14 | 2023-09-27 10:27 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:26:30 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609360 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
67212 | 198031 | CM | CMPLT | CMPLT | 2023-09-27 10:27 | 2023-09-27 10:28 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:27:14 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609323 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
67213 | 198032 | CM | CMPLT | CMPLT | 2023-09-27 10:28 | 2023-09-27 10:29 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:28:31 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609376 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
67210 | 198033 | CM | CMPLT | CMPLT | 2023-09-27 10:29 | 2023-09-27 10:30 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:29:28 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609364 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64253 | 198034 | CM | CMPLT | CMPLT | 2023-09-27 10:30 | 2023-09-27 10:30 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:30:09 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609348 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64284 | 198035 | CM | CMPLT | CMPLT | 2023-09-27 10:31 | 2023-09-27 10:31 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:31:13 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609327 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64292 | 198036 | CM | CMPLT | CMPLT | 2023-09-27 10:32 | 2023-09-27 10:32 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:32:01 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609341 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64251 | 198037 | CM | CMPLT | CMPLT | 2023-09-27 10:35 | 2023-09-27 10:35 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:34:53 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609346 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64301 | 198038 | CM | CMPLT | CMPLT | 2023-09-27 10:36 | 2023-09-27 10:36 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:35:41 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609322 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64205 | 198039 | CM | CMPLT | CMPLT | 2023-09-27 10:37 | 2023-09-27 10:37 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:36:38 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609312 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64280 | 198040 | CM | CMPLT | CMPLT | 2023-09-27 10:37 | 2023-09-27 10:37 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:37:16 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609321 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64291 | 198041 | CM | CMPLT | CMPLT | 2023-09-27 10:40 | 2023-09-27 10:40 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:39:53 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609339 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64264 | 198042 | CM | CMPLT | CMPLT | 2023-09-27 10:40 | 2023-09-27 10:41 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:40:33 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609353 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64273 | 198043 | CM | CMPLT | CMPLT | 2023-09-27 10:41 | 2023-09-27 10:41 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:41:10 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609337 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64249 | 198044 | CM | CMPLT | CMPLT | 2023-09-27 10:42 | 2023-09-27 10:42 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:42:03 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609317 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64250 | 198045 | CM | CMPLT | CMPLT | 2023-09-27 10:43 | 2023-09-27 10:43 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:42:40 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609345 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64263 | 198046 | CM | CMPLT | CMPLT | 2023-09-27 10:48 | 2023-09-27 10:49 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 10:48:37 AM, Baghirzade, Nara) Brandon disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609342 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
1050106 | 198047 | CM | CMPLT | CMPLT | 2023-09-27 12:34 | 2023-09-27 12:35 | Frame was broken off the hinge | (TC 09/27/2023 12:34:02 PM, Baghirzade, Nara) Reseated the frame back in the hinge. Tightened the nut enough that they can still move the frame as needed.
1066460 | 198048 | FC | CMPLT | CMPLT | 2023-09-27 12:36 | 2023-09-27 12:37 | QC checklist | (TC 09/27/2023 12:36:39 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
| 198049 | CM | CMPLT | CMPLT | 2023-09-27 12:39 | 2023-09-27 12:39 | Biomed staff meeting | (TC 09/27/2023 12:38:52 PM, Baghirzade, Nara) Biomed staff meeting
| 198050 | CM | CMPLT | CMPLT | 2023-09-27 13:11 | 2023-10-03 09:31 | Nitronox parts | (TC 09/27/2023 01:16:17 PM, Rendon Lopez , Elizabeth) Received quote from mcArthur Medical, ordered parts (TC 10/02/2023 09:30:21 AM, Rendon Lopez , Elizabeth) Received
| 198051 | TEDU | CMPLT | CMPLT | 2023-09-27 13:48 | 2023-09-27 13:48 | Biomedical Engineering meeting | (TC 09/27/2023 01:47:42 PM, Campbell, Connor) attended meeting
1055800 | 198052 | CM | CMPLT | CMPLT | 2023-09-27 13:57 | 2023-09-27 13:58 | Software upgrade to 02.36.21 | (TC 09/27/2023 01:56:56 PM, Scero, Tomlyn) Software upgraded from 2.32.06.00 to the current version 2.36.21.00. Also uploaded configuration. (TC 09/27/2023 02:08:45 PM, Canjimavadivel, Kamallesh ) Assisted with upgrade
67654 | 198053 | CM | CMPLT | CMPLT | 2023-09-27 14:30 | 2023-09-27 14:31 | Battery replacement | (TC 09/27/2023 02:30:10 PM, Rendon Lopez , Elizabeth) Battery replacement
67653 | 198054 | CM | CMPLT | CMPLT | 2023-09-27 14:31 | 2023-09-27 14:31 | Battery replacement | (TC 09/27/2023 02:30:50 PM, Rendon Lopez , Elizabeth) Battery replacement
67638 | 198055 | CM | CMPLT | CMPLT | 2023-09-27 14:31 | 2023-09-27 14:32 | Battery replacement | (TC 09/27/2023 02:31:37 PM, Rendon Lopez , Elizabeth) Battery replacement
| 198056 | CM | CMPLT | CMPLT | 2023-09-27 14:32 | 2023-09-27 14:32 | Biomed meeting | Biomed meeting
67640 | 198057 | CM | CMPLT | CMPLT | 2023-09-27 14:32 | 2023-09-27 14:33 | Battery replacement | (TC 09/27/2023 02:32:24 PM, Rendon Lopez , Elizabeth) Battery replacement
67650 | 198058 | CM | CMPLT | CMPLT | 2023-09-27 14:33 | 2023-09-27 14:34 | Battery replacement | (TC 09/27/2023 02:33:38 PM, Rendon Lopez , Elizabeth) Battery replacement
1067981 | 198059 | CM | CMPLT | CMPLT | 2023-09-27 14:34 | 2023-09-27 14:35 | Battery replacement | (TC 09/27/2023 02:34:35 PM, Rendon Lopez , Elizabeth) Added to TMS. Battery replacement
| 198060 | CM | CMPLT | CMPLT | 2023-09-27 14:35 | 2023-09-27 14:36 | Department meeting | (TC 09/27/2023 02:35:09 PM, Rhuland, Courtney) general staff meeting
1067982 | 198061 | CM | CMPLT | CMPLT | 2023-09-27 14:35 | 2023-09-27 14:36 | Battery replacement | (TC 09/27/2023 02:35:29 PM, Rendon Lopez , Elizabeth) Added to TMS. Battery replacement
1067983 | 198062 | CM | CMPLT | CMPLT | 2023-09-27 14:38 | 2023-09-27 14:41 | Battery replacement | (TC 09/27/2023 02:38:09 PM, Rendon Lopez , Elizabeth) Added to TMS. Battery replacement
1067984 | 198063 | CM | CMPLT | CMPLT | 2023-09-27 14:42 | 2023-09-27 14:43 | Battery replacement | (TC 09/27/2023 02:41:59 PM, Rendon Lopez , Elizabeth) Added to TMS. Battery replacement
64204 | 198064 | CM | CMPLT | CMPLT | 2023-09-27 15:28 | 2023-09-27 15:28 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:27:51 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609372 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64260 | 198065 | CM | CMPLT | CMPLT | 2023-09-27 15:29 | 2023-09-27 15:29 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:28:46 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609330 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
1049532 | 198066 | CM | CMPLT | CMPLT | 2023-09-27 15:29 | 2023-09-27 15:30 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:29:33 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609329 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64272 | 198067 | CM | CMPLT | CMPLT | 2023-09-27 15:30 | 2023-09-27 15:31 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:30:20 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609373 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64259 | 198068 | CM | CMPLT | CMPLT | 2023-09-27 15:31 | 2023-09-27 15:31 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:31:11 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609332 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
67209 | 198069 | CM | CMPLT | CMPLT | 2023-09-27 15:32 | 2023-09-27 15:32 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:31:55 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609371 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64297 | 198070 | CM | CMPLT | CMPLT | 2023-09-27 15:33 | 2023-09-27 15:33 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:32:43 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609354 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64288 | 198071 | CM | CMPLT | CMPLT | 2023-09-27 15:33 | 2023-09-27 15:34 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:33:24 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609352 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64289 | 198072 | CM | CMPLT | CMPLT | 2023-09-27 15:34 | 2023-09-27 15:34 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/27/2023 03:34:10 PM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609351 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
1056209 | 198073 | CM | CMPLT | CMPLT | 2023-09-27 16:02 | 2023-09-29 15:59 | BP not working | (TC 09/27/2023 04:02:25 PM, Cortez, John) Confirmed issue, O-ring on BP hose connnector cracked Replaced hose with spare
1054221 | 198074 | CM | CMPLT | CMPLT | 2023-09-28 08:43 | 2023-09-28 08:49 | Not working | (TC 09/28/2023 08:46:30 AM, Canjimavadivel, Kamallesh ) Docked thermometer into base to reset security counter. | "SEC" displaying on screen when attempting to use
64953 | 198075 | CM | CMPLT | CMPLT | 2023-09-28 08:44 | 2023-09-28 08:46 | Errors E6 & E8 | (TC 09/28/2023 08:44:53 AM, Rhuland, Courtney) Errors E6 and E8 showing up when trying to take BP reading Leak found in tubing Tubing replaced, no errors after Ran multiple test readings - all passed
65888 | 198076 | CM | CMPLT | CMPLT | 2023-09-28 09:15 | 2023-10-30 10:36 | Created by Failed PM WO # 195663 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 10/30/2023 10:31:05 AM, Campbell, Connor) PM failed again at the plunger accuracy calibrated the unit to see if that fixed it did a check in to clear the error and test the plunger calibration Unit passed check in and PM PM complete/passed, refer to ASM | (TC 09/28/2023 09:13:39 AM, Campbell, Connor) Failed plunger position accuracy calibration of plunger position failed fatal error in syringe motor, needs new syringe assembly refer to WO:198087
64252 | 198077 | CM | CMPLT | CMPLT | 2023-09-28 09:21 | 2023-09-28 09:21 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/28/2023 09:20:44 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609361 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64254 | 198078 | CM | CMPLT | CMPLT | 2023-09-28 09:21 | 2023-09-28 09:22 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/28/2023 09:21:28 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609324 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64248 | 198079 | CM | CMPLT | CMPLT | 2023-09-28 09:22 | 2023-09-28 09:23 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/28/2023 09:22:15 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609366 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
64270 | 198080 | CM | CMPLT | CMPLT | 2023-09-28 09:23 | 2023-09-28 09:23 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 09/28/2023 09:22:58 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609314 Put monitor back in room. Confirmed with the nurse in charge that all patient data is being displayed on the monitor.
68583 | 198081 | CM | CMPLT | CMPLT | 2023-09-28 10:11 | 2023-09-28 12:10 | Created by Failed PM WO # 193913 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 09/28/2023 10:58:27 AM, Campbell, Connor) Replaced rear cover (recall) Replaced front cover (recall) replaced mylar gasket replaced barrel clamp sensor (recall) replaced pressure sensor front cover (TC 09/28/2023 12:02:50 PM, Campbell, Connor) knob wasnt fully closing, opened the plunger head and re-seated the spring, knob now works properly and closes snappily calibrated the unit ran a check in to get rid of the error from calibration PM complete/passed, refer to ASM | (TC 09/28/2023 10:09:49 AM, Campbell, Connor) Cracked rear case near pressure sensor cracked front case in the same spot barrel clamp not working
| 198082 | CM | CMPLT | CMPLT | 2023-09-28 10:39 | 2023-11-21 06:48 | Spare pump parts | Received all parts Nov21
1050580 | 198083 | CM | CMPLT | CMPLT | 2023-09-28 10:41 | 2023-11-16 13:37 | warranty replacement
1055799 | 198084 | CM | CMPLT | CMPLT | 2023-09-28 10:47 | 2023-09-28 10:48 | Software upgrade to 02.36.21 | (TC 09/28/2023 10:47:22 AM, Scero, Tomlyn) software upgraded from 02.32.06.00 to 02.36.21.00 New configuration also uploaded.
| 198085 | CM | CMPLT | CMPLT | 2023-09-28 11:50 | 2023-11-22 15:58 | Order for Alaris Syringe plastic parts, biomed stock | (TC 09/28/2023 11:58:52 AM, Cortez, John) Placed order for 10x Front Case 49000218 10x Rear Case 10013671 10x Handle, Left TC10007985 10x Handle, Right TC10007986 (TC 11/03/2023 04:04:00 PM, Cortez, John) Received left handles, and rear cases (TC 11/16/2023 03:57:28 PM, Cortez, John) Rest of order came in
| 198086 | CM | CMPLT | CMPLT | 2023-09-28 13:16 | 2023-09-28 13:19 | Meeting with Dieter from Rheumatology re: VSMs in clinic and research | (TC 09/28/2023 01:16:13 PM, Baghirzade, Nara) Met with Dieter to discuss the out of support VSMs in the clinic and research area as well as to locate the VSMs. Becky requested that we find out what the needs are for the clinic. As per Dieter, it would be ideal to have 2 VSMs, 1 for each wing and 1 device for the sclero-derma-vasculitis research/infusion research (shared). Out of the 7 VSMs in the clinic, only was able to find 5. Dieter to look for the other two and give me an update.
65888 | 198087 | CM | CMPLT | CMPLT | 2023-09-28 13:32 | 2023-10-30 09:34 | plunger assembly housing kit | (TC 10/30/2023 09:33:13 AM, Campbell, Connor) Assembly intstalled, will proceed to PM, WO 198076
1067877 | 198088 | ACC | CMPLT | CMPLT | 2023-09-28 14:28 | 2023-09-28 14:28 | Incoming inspection | (TC 09/28/2023 02:27:49 PM, Canjimavadivel, Kamallesh ) Functional test pass
1067990 | 198089 | ACC | CMPLT | CMPLT | 2023-09-28 14:48 | 2023-12-28 12:41 | MSH ER GE VIVID iQ ACCEPTANCE | (TC 09/28/2023 03:05:13 PM, Whittington, Scott) MSH ER GE VIVID iQ ACCEPTANCE (TC 09/29/2023 04:23:33 PM, Whittington, Scott) MSH ER GE VIVID iQ ACCEPTANCE SYSTEM RECEIVED ON SITE (TC 10/24/2023 10:25:03 AM, Whittington, Scott) MSH ER GE VIVID iQ ACCEPTANCE OCT 23 TONY & DEO SETUP SYSTEM (TC 11/02/2023 02:07:04 PM, Whittington, Scott) MSH ER GE VIVID IQ ACCEPTANCE MET WITH MIKE MERCHER AND MEAGAN FROM APPS SYSTEM TAKEN TO ER MET WIH ER MD CESAR INTRODUCED GE TO JOHN AYRES FROM MDRD, FOR THEM TO DISCUSS REPROCESSING OF TEE& ENDO CAVITY TXDCR (TC 11/28/2023 04:14:41 PM, Whittington, Scott) MSH ER GE VIVID iQ ACCEPTANCE msh er vivid iq b/u materials to er admin &ge inservice | MSH ER GE VIVID iQ ACCEPTANCE TUV 46 mOHM 175uA H40402LY (12L-RS Probe) 1 1272907WX5 Ì1272907WX5sÎ H40462LA (C1-5-RS Probe) 1 1275523WX7 Ì1275523WX7oÎ H44901AG (M5Sc-RS PROBE) 1 470245YP2 Ì470245YP23Î H45551ZE (6Tc-RS Probe) 2 NT105505 ÌNT105505XÎ NT105527 ÌNT105527vÎ H48062AF (E8Cs-RS Probe) 2 1268200WX1 Ì1268200WX1sÎ 1268223WX3 Ì1268223WX3AÎ H487
| 198090 | CM | CMPLT | CMPLT | 2023-09-28 14:56 | 2023-09-28 14:56 | cleaning and organizing battery shelf
1061207 | 198091 | CM | CMPLT | CMPLT | 2023-09-28 15:12 | 2023-09-28 15:13 | PACS PC not connecting to monitors | (TC 09/28/2023 03:12:32 PM, Reid, Emmerson ) Went up and reseated the display cable, working now. I also reminded them that this is a help desk issue not biomed.
1062546 | 198092 | CM | CMPLT | CMPLT | 2023-09-28 15:26 | 2023-09-28 15:27 | Cannot use functions on bed; locked out due to headboard size not matching up | (TC 09/28/2023 03:25:39 PM, Baghirzade, Nara) Cannot use functions on bed; locked out due to headboard size not matching up with the rest of the bed. The sizer on the headboard was inbetween sizes. Showed Iliana how to ensure that she knows how to adjust it. Worked with Courtney.
1061435 | 198093 | CM | CMPLT | CMPLT | 2023-09-28 15:29 | 2023-09-28 15:30 | Bed motion controls locked out | (TC 09/28/2023 03:28:39 PM, Baghirzade, Nara) Bed motion functions were locked out and the user did not know how to unlock it. Unlocked the motion controls and showed Iliana how to unlock it and to pass this on to her staff as well. Worked with Courtney.
1067992 | 198094 | ACC | CMPLT | CMPLT | 2023-09-28 16:19 | 2023-12-28 12:41 | MSH 5TH FloorUS CANON APLIO i800 PRISM ACCEPTANCE sn# 5HI2332063 | (TC 10/04/2023 06:41:46 PM, Whittington, Scott) MSH 5TH Floor US CANON I800 PRISM ACCEPTANCE (TC 10/04/2023 06:41:46 PM, Whittington, Scott) MSH 5TH Floor US CANON I800 PRISM ACCEPTANCE (TC 11/29/2023 12:20:48 PM, Chau, Ken) Setup wireless, setup tablet, added certificates (TC 12/12/2023 12:42:41 PM, Whittington, Scott) MSH 5TH FloorUS CANON APLIO i800 PRISM ACCEPTANCE sn# 5hi233263 SYSTEM TO BE USED HAD TO MATCH ASSIGNED LICENCES PRISM sn# 5hi233261 REPLACED BY sn# 5hi233263 FOR MSH5_US_P1 ASSIGNMENT TAKE BACKUP FROM SN# sn# 5HI2332061 AND SET INTO sn# 5HI2332063 (TC 12/15/2023 11:57:49 AM, Whittington, Scott) SYSTEM TO BE USED HAD TO MATCH ASSIGNED LICENCES PRISM sn# 5hi233261 REPLACED BY sn# 5hi233263 FOR MSH5_US_P1 ASSIGNMENT TAKE BACKUP FROM SN# sn# 5HI2332061 AND SET INTO sn# 5HI2332063 (TC 12/18/2023 04:27:57 PM, Fan, Tak) MSH 5TH FLR CANON APLIO i800 PRISM ACCEPTANCE SETUP | MSH 5TH Floor US APLIO i800 PRISM ACCEPTANCE
1067993 | 198095 | ACC | CMPLT | CMPLT | 2023-09-29 08:24 | 2023-09-29 08:46 | Incoming inspection | (TC 09/29/2023 08:45:23 AM, Canjimavadivel, Kamallesh ) Electrical safety pass Functional test pass
1067994 | 198096 | CM | CMPLT | CMPLT | 2023-09-29 08:30 | 2023-09-29 08:45 | Incoming inspection | (TC 09/29/2023 08:44:51 AM, Canjimavadivel, Kamallesh ) Electrical safety pass Functional test pass
1067995 | 198097 | ACC | CMPLT | CMPLT | 2023-09-29 08:31 | 2023-09-29 08:44 | Incoming inspection | (TC 09/29/2023 08:44:04 AM, Canjimavadivel, Kamallesh ) Electrical safety pass Functional test pass
1067996 | 198098 | ACC | CMPLT | CMPLT | 2023-09-29 08:34 | 2023-09-29 08:43 | Incoming inspection | (TC 09/29/2023 08:43:14 AM, Canjimavadivel, Kamallesh ) Electrical safety pass Functional test pass
1067997 | 198099 | ACC | CMPLT | CMPLT | 2023-09-29 08:36 | 2023-09-29 08:43 | Incoming inspection | (TC 09/29/2023 08:42:07 AM, Canjimavadivel, Kamallesh ) Electrical safety pass Functional test pass
1067998 | 198100 | ACC | CMPLT | CMPLT | 2023-09-29 08:38 | 2023-09-29 08:39 | Incoming inspection | (TC 09/29/2023 08:38:43 AM, Canjimavadivel, Kamallesh ) Electrical safety pass Functional tests pass
65641 | 198101 | CM | CMPLT | CMPLT | 2023-09-29 09:16 | 2023-09-29 09:18 | Blood pressure cuff check | (TC 09/29/2023 09:16:21 AM, Rhuland, Courtney) Checked all sphygmometers in urology clinic for leaks/verified pressure readings. Repaired hoses with leaks and replaced one cuff with clinic spare.
67971 | 198102 | CM | CMPLT | CMPLT | 2023-09-29 09:24 | 2023-09-29 14:20 | Humidifer failure | (TC 09/29/2023 01:41:55 PM, Rhuland, Courtney) Humidifier failure present when switched on; X ovet humidity icon Tested resistance on relay board and result showed open line Open line resistance persisted through one of the connectors to the humidifier Humidifier replaced, error cleared RH calibrated Functional check - OK
1059778 | 198103 | FC | CMPLT | CMPLT | 2023-09-29 11:50 | 2023-09-29 11:50 | QC checklist | (TC 09/29/2023 11:50:00 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1063140 | 198104 | CM | CMPLT | CMPLT | 2023-09-29 12:12 | 2023-09-29 12:13 | Error code | (TC 09/29/2023 12:11:44 PM, Rendon Lopez , Elizabeth) Confirmed, error code E 5.0, faulty temp probe. Replaced temperature probe, tested-OK
| 198105 | CM | CMPLT | CMPLT | 2023-09-29 12:15 | 2023-09-29 12:15 | PM update meeting | (TC 09/29/2023 12:14:56 PM, Baghirzade, Nara) PM update meeting
| 198106 | CM | CMPLT | CMPLT | 2023-09-29 12:16 | 2023-09-29 12:16 | Patient monitoring scenario | (TC 09/29/2023 12:15:35 PM, Baghirzade, Nara) Patient monitoring scenario
65970 | 198107 | CM | CMPLT | CMPLT | 2023-09-29 12:16 | 2023-09-29 12:18 | Channel error 6990 | (TC 09/29/2023 12:16:15 PM, Baghirzade, Nara) Downloaded error log. Noted errors pertaining to pressure sensor disc failures. Looked over the old WOs and the pressure sensor disc has been re-seated multiple times in the past. A couple of screws were missing for the board/bezel; found Alaris screws from spare parts box. Replaced the pressure sensor disc. Calibrated the pressure sensor disc. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1049885 | 198108 | CM | CMPLT | CMPLT | 2023-09-29 12:20 | 2023-10-04 14:14 | Date and time not updating | (TC 09/29/2023 12:20:04 PM, Baghirzade, Nara) Sue requested to have this monitor looked at; the date and time is not updating/retaining even after changing the date and time. Brought the device down to biomed, opened it up to check for the CMOS battery (requires 1220 battery). Biomed does not have that battery in stock. Will order the battery and replace it. Sue requested to put the monitor back in the procedure room; placed all connectors back on the monitor (ECG, SpO2, NIBP). (TC 10/04/2023 02:11:52 PM, Baghirzade, Nara) Opened the device to board level; replaced the CMOS battery (1220); checked that date and time is updating and being retained on the monitor. Did a functional check of ECG, NIBP, and CO2. All functional tests passed. Passed the system back to Sue. (TC 10/04/2023 02:13:06 PM, Rhuland, Courtney) Opened the device to board level; replaced the CMOS battery (1220); checked that date and time is updating and being retained on the monitor. Did a functional check of ECG, NIBP, and CO2. All functional tests passed.
65767 | 198109 | CM | CMPLT | CMPLT | 2023-09-29 12:23 | 2023-12-21 09:14 | Dr. Witterick requested to have this unit functional | (TC 09/29/2023 12:22:46 PM, Baghirzade, Nara) Device was disposed of by the OR team however Dr. Witterick wants to make it functional for the ENT clinic for students. Researched a bit about this device and Storz and Olympus are both compatible with this device. Dr. requested that it operates as a headlight. Will email Olympus to see if they have a compatible headlight for this device. Will forward the quote to Dr. Witterick once there is more information from the vendor. (TC 10/04/2023 02:14:12 PM, Baghirzade, Nara) Tested a welch-allyn headlight but the connector type doesn't fit the light source. From the OR, I tested the luxtec integra headlight on the light source and it fit on the ACMI connector however the luxtec headlight is too expensive for what Dr. Witterick will be using it for. Will inquire about a different model headlight. (TC 10/25/2023 01:52:39 PM, Baghirzade, Nara) After doing some research on headlights that would fit the light source model at the ENT clinic, found 2 from Valleymed. Contacted valleymed for a quote and sent an email to follow up on the quote. Waiting for a response.
| 198110 | CM | CMPLT | CMPLT | 2023-09-29 12:25 | 2023-09-29 12:26 | Meeting with Nima re: VSMs in Rheumatology | (TC 09/29/2023 12:25:03 PM, Baghirzade, Nara) Met with Nima to discuss the VSM needs for the Rheumatology clinic and research.
| 198111 | RDS | CMPLT | CMPLT | 2023-09-29 12:26 | 2023-09-29 12:28 | Weekly Rounds: 11 South | (TC 09/29/2023 12:26:24 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 198112 | RDS | CMPLT | CMPLT | 2023-09-29 12:29 | 2023-09-29 12:33 | Weekly Rounds: 18 South and PACU POD A | (TC 09/29/2023 12:28:45 PM, Baghirzade, Nara) Weekly rounds on 18 South and PACU POD A completed. See procedure for results.
| 198113 | RDS | CMPLT | CMPLT | 2023-09-29 12:34 | 2023-09-29 12:36 | Weekly Rounds: NICU | (TC 09/29/2023 12:33:58 PM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure.
| 198114 | RDS | CMPLT | CMPLT | 2023-09-29 12:36 | 2023-09-29 12:37 | Weekly Rounds: RESUS | (TC 09/29/2023 12:36:07 PM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure.
66683 | 198115 | CM | CMPLT | CMPLT | 2023-09-29 13:38 | 2023-09-29 13:40 | Reported that incubator beeps when switching on | (TC 09/29/2023 01:37:56 PM, Baghirzade, Nara) Back switch was in off position. Switch is now in on position. During functional test, the RH was reading higher than expected (set 75%, it was reading 80%). Calibrated the RH sensor. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and humidity 75%. Incubator is fully functional.
1059981 | 198116 | PM | CMPLT | PMFAIL | 2023-09-29 15:50 | 2023-09-29 15:57 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 09/29/2023 03:50:33 PM, Cortez, John) 2 wires with connectors loose on left foot section side, pendant clip/headboard missing, 70mohms, 21uA 2-69 deg head section chair egress failed "doesnt sense footboard is removed might be related to loose wire pin missing on bottom right rail dampener, 90.7 kg scale verification reading Pm failed, will request service when bed is free (TC 09/28/2023 03:50:33 PM, Cortez, John) 2 wires with connectors loose on left foot section side, pendant clip/headboard missing, 70mohms, 21uA 2-69 deg head section chair egress failed "doesnt sense footboard is removed might be related to loose wire pin missing on bottom right rail dampener, 90.7 kg scale verification reading Pm failed, will request service when bed is free
1059981 | 198117 | CM | CMPLT | CMPLT | 2023-09-29 15:57 | 2024-01-10 10:38 | Created by Failed PM WO # 198116 : PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/13/2023 03:10:49 PM, Cortez, John) Swapped bed and requested service from hillrom (TC 12/27/2023 02:59:29 PM, Canjimavadivel, Kamallesh ) Found bed in room 34. Swapped out with repaired bed from 18 (TC 01/09/2024 10:36:27 AM, Cortez, John) Bed+Intellidrive batteries were replaced last PM attempt 80 mOhms, 24 uAmps PM passed | (TC 09/29/2023 03:50:33 PM, Cortez, John) 2 wires with connectors loose on left foot section side, pendant clip/headboard missing, 70mohms, 21uA 2-69 deg head section chair egress failed "doesnt sense footboard is removed might be related to loose wire pin missing on bottom right rail dampener, 90.7 kg scale verification reading Pm failed, will request service when bed is free
1055911 | 198118 | CM | CMPLT | CMPLT | 2023-09-29 16:00 | 2023-09-29 16:01 | "bp not working" | (TC 09/29/2023 04:00:05 PM, Cortez, John) BP hose connector has leak, replaced with spare
68369 | 198119 | CM | CMPLT | CMPLT | 2023-09-29 16:37 | 2023-09-29 16:41 | Oxygen reading issues | (TC 09/28/2023 04:37:14 PM, Rendon Lopez , Elizabeth) Complaint, device was misreading oxygen concentration. Downloaded error logs, no error found, only regular clinical alarms. Performed self test and leakage, all passed. (TC 09/29/2023 04:37:14 PM, Rendon Lopez , Elizabeth) Scheduled self test passed in the morning. Tested in pressure control mode at different O2 levels, all readings in range.
1050727 | 198120 | CM | CMPLT | CMPLT | 2023-09-29 16:43 | 2023-12-12 14:36 | Not working | (TC 09/29/2023 04:43:58 PM, Rendon Lopez , Elizabeth) Confirmed issue, device not activating, front commands not working. To be sent out for repair under consignment (TC 12/12/2023 02:35:19 PM, Rendon Lopez , Elizabeth) Incoming inspection, Tested-OK. Returned to OR 1 | Coordinator- The fan turns on initially but after that even when pen is in use it does not turn on again unless you unplug it
52109 | 198121 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
52428 | 198122 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Lasers, Nd:YAG, Ophthalmic
62485 | 198123 | PM | ACTIV | ISSUE | 2023-10-01 03:16 |  | ANNUAL PM FOR: Electrosurgical Unit
62976 | 198124 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-06-26 12:59 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/26/2024 12:58:46 PM, Rhuland, Courtney) PM completed All tests passed Irradiance check - OK Electrical safety - OK
63019 | 198125 | PM | CANCL | CANCL | 2023-10-01 03:16 | 2024-07-08 10:05 | PM-MAJOR-Annual PM for:MR-Diagnostics Incubator | Device is sitting in NICU Storage for parts/removal
63025 | 198126 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-05-01 13:15 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/01/2024 01:10:20 PM, Rhuland, Courtney) PM completed Irradiance - 50.1 uW/cm^2nm Hours - 7366.7 Current leakage - 9.8 uA Ground resistance - 145 mOhms
63787 | 198127 | PM | CANCL | CANCL | 2023-10-01 03:16 | 2024-06-27 10:04 | PM-MAJOR - Annual PM For: Testers, Ventilator | Its is old tester - retired.
64060 | 198128 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-06-24 14:59 | PM MAJOR ANnual: Power Systems, Uninterruptible | (TC 06/24/2024 02:58:36 PM, Cortez, John) Replaced battery
64062 | 198129 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-07-02 09:56 | PM MAJOR Annual: Power Systems, Uninterruptible | (TC 07/02/2024 09:56:25 AM, Canjimavadivel, Kamallesh ) Battery replaced
64065 | 198130 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-06-24 15:04 | PM Major Annual - Power Systems, Uninterruptible | (TC 06/24/2024 03:04:17 PM, Cortez, John) Replaced battery
64992 | 198131 | PM | CANCL | CANCL | 2023-10-01 03:16 | 2024-06-18 13:53 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
64995 | 198132 | PM | CMPLT | PMFAIL | 2023-10-01 03:16 | 2024-03-01 13:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 02/29/2024 01:54:35 PM, Rendon Lopez , Elizabeth) performing PM, pumps is overdelivering. Tried different cassettes, same issue
65131 | 198133 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-11-16 11:14 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 11/16/2023 11:13:14 AM, Reid, Emmerson ) Downloaded firmware V040201 and Pushed APS IV PCA MAR 2022. PM Complete, see procedure for details.
65163 | 198134 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-27 14:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 10/27/2023 02:56:06 PM, Reid, Emmerson ) PM Complete, see attached proceedure for details.
65164 | 198135 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-30 15:53 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 10/30/2023 03:50:18 PM, Cortez, John) PM passed (see procedure for results)
65264 | 198136 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-11-10 14:23 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 83uA and 91 ohm.
65292 | 198137 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-19 11:12 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/19/2023 11:10:55 AM, Campbell, Connor) Pm passed/complete, refer to procedure Device disinfected Fluid replaced with fresh solution
65572 | 198138 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-05-03 08:55 | PM-MAJOR - Annual PM For: YAG laser, gyne | (TC 05/03/2024 08:53:20 AM, Baghirzade, Nara) Received the correct PM service report from Steve (Dornier service tech) as he initially sent the one for L&D. Attached the service report.
65876 | 198139 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65895 | 198140 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65900 | 198141 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-11-30 09:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 11/30/2023 09:47:11 AM, Campbell, Connor) PM passed/complete, refer to ASM
65920 | 198142 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-03-28 12:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 12:54:36 PM, Rhuland, Courtney) PM completed by BD
65923 | 198143 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65960 | 198144 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65962 | 198145 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-25 13:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 10/25/2023 01:25:44 PM, Rhuland, Courtney) Performed calibration after pressure verification failed PM complete All tests passed post-calibration Refer to ASM for details
65964 | 198146 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65966 | 198147 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-03-19 09:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/19/2024 09:04:41 AM, Rhuland, Courtney) Replaced rear case Replaced iui Calibration and PM completed All tests passed Refer to ASM for details
66003 | 198148 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-03-25 09:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66018 | 198149 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66029 | 198150 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66463 | 198151 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-11 11:50 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Tested. Functional. SW.2.51 Leakage: 70uA and 85 ohm.
66493 | 198152 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-11-17 15:28 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 11/16/2023 03:27:25 PM, Rendon Lopez , Elizabeth) Functional check- all passed
66680 | 198153 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-23 10:58 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 10/23/2023 10:54:36 AM, Baghirzade, Nara) PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80%. Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 29.5 uA Power cord resistance = 242 mohms
66770 | 198154 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66817 | 198155 | PM | CANCL | CANCL | 2023-10-01 03:16 | 2024-07-25 12:38 | BP Check
66825 | 198156 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | BP Check
66829 | 198157 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | BP Check
66855 | 198158 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | BP Check
67315 | 198159 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-01-12 09:56 | PM-MAJOR - Annual PM For: High Frequency Desiccator with Cut | (TC 01/12/2024 09:37:19 AM, Rhuland, Courtney) PM completed - see attached document for details Noted fulguration could not get to wattage past 33W - Unit does not use at this power setting Electrical safety: leakage 9.4uA Resistance 181mOhms (TC 01/19/2024 02:25:42 PM, Rhuland, Courtney) Retested fulguration at 35W and issue is resolved. Updated checklist attatched
67764 | 198160 | PM | ACTIV | ISSUE | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67770 | 198161 | PM | ACTIV | ISSUE | 2023-10-01 03:16 |  | ANNUAL PM - Catarac
67918 | 198162 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
68098 | 198163 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Laryngoscopes, Video
68367 | 198164 | PM | ACTIV | ISSUE | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68545 | 198165 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Anesthesia Units | PM done March 2023, Rescheduled for March 2024
68553 | 198166 | PM | ACTIV | ISSUE | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68608 | 198167 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-20 11:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 10/20/2023 11:22:26 AM, Campbell, Connor) PM passed/complete, refer to ASM
68618 | 198168 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68619 | 198169 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-04-03 13:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:22:10 PM, Rhuland, Courtney) PM completed by BD
68679 | 198170 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68689 | 198171 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68690 | 198172 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68700 | 198173 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-04-03 13:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:14:34 PM, Rhuland, Courtney) PM completed by BD
68723 | 198174 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68769 | 198175 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68803 | 198176 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68829 | 198177 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68861 | 198178 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-03-12 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 01:11:44 PM, Reid, Emmerson ) PM Complete.
68865 | 198179 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68873 | 198180 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-04-24 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68880 | 198181 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68885 | 198182 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-03-25 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68893 | 198183 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68907 | 198184 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-04-03 14:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:01:02 PM, Rhuland, Courtney) PM completed by BD
68915 | 198185 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68921 | 198186 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68922 | 198187 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-27 13:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/27/2023 01:38:13 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 35 mohms Leakage current: 74.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68932 | 198188 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68938 | 198189 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68971 | 198190 | PM | CNP | ENL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69000 | 198191 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-04-03 14:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:05:36 PM, Rhuland, Courtney) PM completed by BD
69011 | 198192 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-04-05 14:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:39:24 PM, Rhuland, Courtney) PM completed by BD
69028 | 198193 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69119 | 198194 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-04-09 10:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/09/2024 10:45:59 AM, Reid, Emmerson ) Downloaded logs. Ran ASM Calibration and PM, all tests passed. PM Complete.
69138 | 198195 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2023-10-17 11:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/17/2023 11:43:00 AM, Campbell, Connor) pm complete/passed, refer to ASM
69215 | 198196 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69236 | 198197 | PM | CMPLT | CMPLT | 2023-10-01 03:16 | 2024-03-28 11:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/21/2024 03:15:18 PM, Reid, Emmerson ) Replaced the front casing
69265 | 198198 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69298 | 198199 | PM | CMPLT | PMFAIL | 2023-10-01 03:16 | 2024-03-13 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:11:39 PM, Baghirzade, Nara) Failed rate calibration.
69308 | 198200 | PM | CANCL | CANCL | 2023-10-01 03:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69350 | 198201 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69364 | 198202 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-12-13 08:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 12/13/2023 08:32:06 AM, Campbell, Connor) Pm complete/passed, refer to ASM unit was very dirty, had to clean IUI and the locking tab really well as they were covered in sticky stuff
69430 | 198203 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69434 | 198204 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69446 | 198205 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69480 | 198206 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69489 | 198207 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-03-25 11:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69533 | 198208 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-03-19 21:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69539 | 198209 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69590 | 198210 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69610 | 198211 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-03-25 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69635 | 198212 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69637 | 198213 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
32540 | 198214 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-06 14:11 | PM - Monthly | PM completed Mechanical PM done sc-3518 TOS phantom a(-996.0,10.0) b(331.1,13.1) c(132.2,13.5) d(97.5,13.6) e(-105.1,12.8) f(0.3,12.7) -passed all phantom tests
32560 | 198215 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM - Monthly | PM Canceled due to service work at WCH (system down)
32580 | 198216 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-01 14:29 | PM - Monthly | PM completed Mechanical PM done sc-133083 TOS phantom a(-991.5,10.6) b(335.5,15.1) c(134.6,13.5) d(94.7,12.9) e(-106.7,12.0) f(0.4,14.7) -passed all phantom tests
501648 | 198217 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-01-10 15:17 | PM - Annual | (TC 10/31/2023 03:07:49 PM, Green, Paul) Performed PM. Completed most of PM. When calibration of position was done, failed on B side stating out of tolerance. Kept spare installed at MSH 1.5T and informed the customer. Getting a quote on potentiometers. (TC 11/01/2023 03:06:13 PM, Green, Paul) Looking into PCU contract and ordering parts to correct postion calibration problem. (TC 01/10/2024 03:16:31 PM, Green, Paul) Mechanical Drive assembly defective. Replaced and full calibrations performed. Reinstalled back in room.
502088 | 198218 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-14 12:59 | PM - Semi-Annual | (TC 11/14/2023 12:57:11 PM, Paz, Ramon (Tino)) Preventive maintenance PM1 performed. No defficiencies. All system passed. Performed PM as per manufacturer's intructions.
502089 | 198219 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-09 14:38 | PM - Semi-Annual | (TC 11/09/2023 02:37:33 PM, Green, Paul) PM1 completed as per procedure.
502231 | 198220 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-07 16:40 | PM - Semi-Annual | A complete PM was done.
502231 | 198221 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-17 20:32 | HARP - Semi Annual | HARP
1048154 | 198222 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | ANNUAL PM - Microscope
1048660 | 198223 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049075 | 198224 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | Planned Maintenance
1049878 | 198225 | PM | CMPLT | PMFAIL | 2023-10-01 03:17 | 2023-11-28 16:41 | PM-MAJOR - Annual PM For: YAG laser, gyne | (TC 11/21/2023 04:08:35 PM, Rendon Lopez , Elizabeth) laser PM scheduled with Dormier Technician. Reported failure during PM, laser needs new part, techcnician will order part and come repair next week NOV 27-28. laser kept in 18fl office (TC 11/28/2023 04:40:33 PM, Rendon Lopez , Elizabeth) Device needs to be sent out for repair
1050273 | 198226 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050615 | 198227 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-04-05 12:17 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/05/2024 12:16:36 PM, Rhuland, Courtney) Calibration and PM completed All tests passed Refer to ASM for details
1051369 | 198228 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-04-30 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1053347 | 198229 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-08 17:18 | MSH BI 1270 H1 HARP NOV 8 TRGT SEMI-ANNUAL | (TC 11/08/2023 05:15:16 PM, Dolinar, John) MSH BI 1270 H1 HARP NOV 8 TRGT SEMI-ANNUAL (TC 11/08/2023 05:15:36 PM, Whittington, Scott) MSH BI 1270 H1 HARP NOV 8 TRGT SEMI-ANNUAL (TC 11/08/2023 05:15:56 PM, Whittington, Scott) MSH BI 1270 H1 HARP NOV 8 TRGT SEMI-ANNUAL | MSH BI 1270 H1 HARP NOV 8 TRGT SEMI-ANNUAL
1053347 | 198230 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-10 17:22 | MSH BI 1270 -PM SEMI-ANNUAL |  | PM minor
1053379 | 198231 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-12-05 11:17 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | completed annual PM. Replaced valve on cassette too. tested. funtional. | 2 years warranty covers PMs and repairs
1054460 | 198232 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-10-17 08:22 | Lubrication cycle | (TC 10/17/2023 08:21:39 AM, Canjimavadivel, Kamallesh ) Lubrication complete
1054462 | 198233 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-10-13 09:44 | Lubrication cycle | (TC 10/13/2023 09:42:50 AM, Cortez, John) Lubricated o-rings
1054468 | 198234 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-10-26 10:55 | Lubrication cycle | (TC 10/26/2023 10:54:36 AM, Canjimavadivel, Kamallesh ) Lubrication of O-rings complete
1055104 | 198235 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-03-14 10:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 10:24:10 AM, Huang, Jessica) replaced the power cord wrap and completed PM
1055107 | 198236 | PM | CNP | ENL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055109 | 198237 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-10-26 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/25/2023 01:44:28 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 196 mohms Leakage current: 59.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 10/26/2023 09:48:41 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 4103 mamphr PM completed.
1055118 | 198238 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-03-25 10:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055788 | 198239 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-22 10:14 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 11/22/2023 10:11:47 AM, Campbell, Connor) PM complete/passed, see attached PM sheet for more info
1055789 | 198240 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-08 10:22 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 11/08/2023 10:21:21 AM, Reid, Emmerson ) PM Complete, showed Jessica and Connor how to do the PM. See attached PM Report.
1055790 | 198241 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-08 14:27 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 11/08/2023 02:26:23 PM, Reid, Emmerson ) OM Complete, see attached PM Report.
1055792 | 198242 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1056055 | 198243 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-24 20:34 | HARP - Semi Annual | HARP
1056055 | 198244 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-16 20:28 | PM - Semi Annual | PM was done.
1057388 | 198245 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-06 09:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/01/2023 09:08:21 AM, Rendon Lopez , Elizabeth) PM complete, all tests passed
1057513 | 198246 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-03 10:31 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 11/03/2023 10:30:56 AM, Reid, Emmerson ) PM Complete, see attached report.
1057537 | 198247 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057538 | 198248 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057539 | 198249 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | Device in use
1057540 | 198250 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057541 | 198251 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | batteries weren't available
1057542 | 198252 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057543 | 198253 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-01-31 10:50 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 01/31/2024 10:49:39 AM, Scero, Tomlyn) replaced battery (every 3 years) replaced missing bezel, already charged to OR(RT)
1057544 | 198254 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-11-07 09:31 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/07/2024 09:31:00 AM, Scero, Tomlyn) All PM tests complete. Leak test passed. Battery replaced.
1057545 | 198255 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057546 | 198256 | PM | CANCL | CANCL | 2023-10-01 03:17 | 2024-05-30 11:24 | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057547 | 198257 | PM | CANCL | CANCL | 2023-10-01 03:17 | 2024-10-22 11:44 | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057548 | 198258 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057549 | 198259 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-06 09:03 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 10/31/2023 09:02:12 AM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed
1057550 | 198260 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057551 | 198261 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-06 08:55 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 10/31/2023 08:52:55 AM, Rendon Lopez , Elizabeth) Replaced battery. PM completed, all tests passed
1057552 | 198262 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057553 | 198263 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057554 | 198264 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-23 16:14 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/20/2023 04:13:20 PM, Rendon Lopez , Elizabeth) Battery replaced. PM completed, all tests passed. Returned to L&D OR
1057555 | 198265 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057556 | 198266 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057557 | 198267 | PM | CANCL | CANCL | 2023-10-01 03:17 | 2024-06-03 11:13 | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057558 | 198268 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057559 | 198269 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-05-15 11:46 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/15/2024 11:44:19 AM, Scero, Tomlyn) PM tests complete and passed. no leaks and battery replaced.
1057560 | 198270 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-06 09:14 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/02/2023 09:13:55 AM, Rendon Lopez , Elizabeth) Replaced battery and side panel. PM completed, all tests passed
1057571 | 198271 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-02-09 09:12 | PM-MAJOR - Annual PM For: Electrosurgical Units
1057572 | 198272 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-03 10:24 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 11/03/2023 10:23:18 AM, Reid, Emmerson ) PM Complete, see attached report.
1057599 | 198273 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-12-28 10:53 | PM-MAJOR ANNUAL PM for: Ventilators, Intensive Care, Adult | (TC 12/20/2023 11:26:36 AM, Cortez, John) Replaced filters and nozzles Pre check pass Total op time: 11418 hrs Total vent time: 8252 hrs PM complete (TC 12/27/2023 03:08:55 PM, Canjimavadivel, Kamallesh ) Replaced cable guide
1057600 | 198274 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-12-20 11:18 | PM-MAJOR ANNUAL PM for: Ventilators, Intensive Care, Adult | (TC 12/20/2023 11:16:55 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 13654 hrs Total vent time: 9073 hrs PM complete (TC 12/20/2023 11:17:50 AM, Campbell, Connor) Worked with Kam
1057602 | 198275 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-01-09 10:46 | PM-MAJOR ANNUAL PM for: Ventilators, Intensive Care, Adult | (TC 01/09/2024 10:45:10 AM, Canjimavadivel, Kamallesh ) Replaced damaged cable guide Replaced filters and nozzles Pre check pass Total op time: 8465 hrs Total vent time: 3016 hrs
1058253 | 198276 | PM | CANCL | CANCL | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058266 | 198277 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2024-03-12 09:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:33:48 AM, Rendon Lopez , Elizabeth) PM completed by BD
1060518 | 198278 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-24 20:47 | HARP - Semi Annual | HARP
1060518 | 198279 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-23 22:48 | PM - Semi Annual | PM
1062586 | 198280 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | Planned Maintenance
1062603 | 198281 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM MAJROR - ANNUAL PM For: Warming Unit
1062660 | 198282 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - ANNUAL PM for: Insufflators | (TC 07/03/2024 01:37:17 PM, Patel, Rutvik) Contacted vendor for CO2 Gas connector purchase info. Received. Created SC. Requested approval. Received. Placed order via MMC.
1062661 | 198283 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-10-27 09:51 | PM-MAJOR - ANNUAL PM for: Insufflators | (TC 10/27/2023 09:50:22 AM, Reid, Emmerson ) Moved from the cart to the boom in OR14. Changed the hose adaptor and did self test and physical inspection. PM Complete.
1064495 | 198284 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-08 09:08 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 11/08/2023 09:07:23 AM, Reid, Emmerson ) PM Complete, see attached PM Report.
1064583 | 198285 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-11-07 13:25 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 11/07/2023 01:24:26 PM, Reid, Emmerson ) PM Complete, see attached PM Report.
1065114 | 198286 | PM | CANCL | CANCL | 2023-10-01 03:17 | 2024-10-01 07:29 | Annual PM
1065115 | 198287 | PM | CANCL | CANCL | 2023-10-01 03:17 | 2024-10-01 07:29 | Annual PM
1065167 | 198288 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065168 | 198289 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065169 | 198290 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065170 | 198291 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065171 | 198292 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065173 | 198293 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065174 | 198294 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065176 | 198295 | PM | ACTIV | ISSUE | 2023-10-01 03:17 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065181 | 198296 | PM | CANCL | CANCL | 2023-10-01 03:17 | 2024-07-18 08:56 | PM-MAJOR- Annual PM for :Chairs, Examination/Treatment,Bariatric
1066756 | 198297 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-10-26 10:56 | Lubrication cycle | (TC 10/26/2023 10:55:27 AM, Canjimavadivel, Kamallesh ) Lubrication of O-rings complete
1066757 | 198298 | PM | CMPLT | CMPLT | 2023-10-01 03:17 | 2023-10-26 10:56 | Lubrication cycle | (TC 10/26/2023 10:55:53 AM, Canjimavadivel, Kamallesh ) Lubrication of O-rings complete
63222 | 198299 | CM | CMPLT | CMPLT | 2023-10-02 07:32 | 2023-10-02 07:33 | missing power supply plug | (TC 10/02/2023 07:31:54 AM, Campbell, Connor) replaced power supply plug (no charge) performed a functional rate check, no issues found
66817 | 198300 | CM | CMPLT | CMPLT | 2023-10-02 07:41 | 2023-10-02 07:42 | missing battery | (TC 10/02/2023 07:41:24 AM, Campbell, Connor) charged and replaced battery
63209 | 198301 | CM | CMPLT | CMPLT | 2023-10-02 08:16 | 2023-10-02 08:17 | "not working" | (TC 10/02/2023 08:16:02 AM, Campbell, Connor) performed a visual inspection and a functional test test passed, no issues found
1054697 | 198302 | CM | CMPLT | CMPLT | 2023-10-02 08:54 | 2024-05-16 13:30 | tech support required error 23 | (TC 10/02/2023 08:54:59 AM, Campbell, Connor) tried to perform functional check when priming, error 23 confirmed error 23 - Motor Failsafe Time Out Error. Possible damaged gearbox encoder or flex cable. gave to john to send to company for repairs (TC 05/16/2024 01:23:30 PM, Cortez, John) Replacement arrived, Unboxed and completed functional check
67962 | 198303 | CM | CMPLT | CMPLT | 2023-10-02 09:23 | 2023-10-02 09:24 | Reported "system failure" | (TC 10/02/2023 09:23:25 AM, Baghirzade, Nara) System failure upon boot up; fan was improperly seated. Re-seated the fan. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and approx. 74.8% RH. Incubator is fully functional.
67916 | 198304 | CM | CMPLT | CMPLT | 2023-10-02 09:25 | 2023-10-02 09:25 | Latch recall replacement | (TC 10/02/2023 09:24:44 AM, Baghirzade, Nara) Replaced porthole latches and NW/NE latches for recall.
67913 | 198305 | PA | CMPLT | CMPLT | 2023-10-02 09:26 | 2023-10-05 07:43 | Ordering incubator hood | (TC 10/02/2023 09:28:35 AM, Baghirzade, Nara) SC 1003065271 (TC 10/05/2023 07:42:15 AM, Baghirzade, Nara) PO 7202294077
| 198306 | CM | CMPLT | CMPLT | 2023-10-02 10:06 | 2023-12-12 08:54 | Spare pump parts | (TC 10/02/2023 10:08:13 AM, Rendon Lopez , Elizabeth) Ordering spare pump parts, will be charged when used Received
| 198307 | CM | CMPLT | CMPLT | 2023-10-02 10:15 | 2023-10-02 10:15 | ER-EAST philips documentation- going through all rooms and stations with Kam | (TC 10/02/2023 10:14:47 AM, Samanian, Kamran) done!
| 198308 | RDS | CMPLT | CMPLT | 2023-10-02 10:41 | 2023-10-02 14:33 | Weekly check for ED
| 198309 | MEP | CMPLT | CMPLT | 2023-10-02 10:55 | 2023-10-02 10:56 | Bookings for training and expenses | (TC 10/02/2023 10:55:18 AM, Baghirzade, Nara) Bookings for the EEG training (OCT 2023) with Adeel. Added in all expenses to Concur for the EMG training (SEPT 2023).
1056125 | 198310 | PM | CMPLT | CMPLT | 2023-10-02 11:05 | 2023-10-02 11:21 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/02/2023 11:19:40 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
63219 | 198311 | CM | CMPLT | NFF | 2023-10-02 11:20 | 2023-10-02 11:21 | "false alarm" | (TC 10/02/2023 11:20:39 AM, Campbell, Connor) performed a visual inspection performed a functional test no fault found at any point
| 198312 | CM | CMPLT | CMPLT | 2023-10-02 14:17 | 2023-10-02 14:17 | Monday huddle | (TC 10/02/2023 02:16:46 PM, Baghirzade, Nara) Biomed huddle
| 198313 | CM | ACTIV | READY | 2023-10-02 14:57 |  | Training/Education/Meetings Etc. | (TC 10/02/2023 02:57:09 PM, Campbell, Connor) Huddle (TC 10/20/2023 11:00:18 AM, Campbell, Connor) Full team meeting (TC 10/24/2023 03:15:41 PM, Campbell, Connor) huddle (TC 10/24/2023 03:15:41 PM, Campbell, Connor) huddle (TC 10/27/2023 10:41:52 AM, Campbell, Connor) Learning and practicing 3D modeling (TC 10/27/2023 10:41:52 AM, Campbell, Connor) Learning and practicing 3D modeling (TC 10/27/2023 10:41:52 AM, Campbell, Connor) Learning and practicing 3D modeling (TC 11/01/2023 02:40:46 PM, Campbell, Connor) huddle (TC 11/01/2023 02:40:46 PM, Campbell, Connor) huddle (TC 11/13/2023 02:02:05 PM, Campbell, Connor) Huddle (TC 11/13/2023 02:02:05 PM, Campbell, Connor) Huddle (TC 11/15/2023 09:54:46 AM, Campbell, Connor) CALM meeting (TC 11/16/2023 10:02:09 AM, Campbell, Connor) 3D modelling a tray mold for Qwynn (TC 11/17/2023 08:36:50 AM, Campbell, Connor) 3D modelling (TC 11/17/2023 08:36:50 AM, Campbell, Connor) 3D modelling (TC 11/17/2023 10:37:35 AM, Campbell, Connor) Full team meeting (TC 11/20/2023 02:19:42 PM, Campbell, Connor) Huddle (TC 11/22/2023 08:56:38 AM, Campbell, Connor) Learning about philips patient monitoring to slightly prepare for the study at UHN (TC 11/23/2023 12:29:17 PM, Campbell, Connor) Philips Patient Monitor study @UHN (TC 11/24/2023 11:48:26 AM, Campbell, Connor) Helping Qwynn to make silicon molds (TC 11/28/2023 12:13:41 PM, Campbell, Connor) Silicon suture pad instruction write-up for Qwynn (TC 11/28/2023 02:03:06 PM, Campbell, Connor) working on co-op reflection documents (TC 11/28/2023 02:37:25 PM, Campbell, Connor) pouring another version of the silicon suture pad (TC 11/29/2023 09:56:03 AM, Campbell, Connor) silicon suture pad testing (TC 11/29/2023 12:26:21 PM, Campbell, Connor) Huddle (TC 11/29/2023 01:40:23 PM, Campbell, Connor) Silicon suture pad v4 (TC 11/30/2023 12:07:31 PM, Campbell, Connor) silicon suture pad v5 (TC 11/30/2023 12:07:54 PM, Campbell, Connor) reading about rhinoplasty silicon models (TC 11/30/2023 12:07:54 PM, Campbell, Connor) reading about rhinoplasty silicon models (TC 12/01/2023 09:37:43 AM, Campbell, Connor) helping jessica take inventory in OR (TC 12/01/2023 01:55:52 PM, Campbell, Connor) reading about silicon modelling for rhinoplasty (TC 12/04/2023 12:21:26 PM, Campbell, Connor) 3D modelling reading (TC 12/05/2023 10:47:07 AM, Campbell, Connor) 3d printed mandible model processing (TC 12/05/2023 07:51:18 AM, Campbell, Connor) helping qwynn move transport incubator and set it up in NUCU (TC 12/07/2023 12:44:47 PM, Campbell, Connor) Huddle (TC 12/08/2023 09:20:20 AM, Campbell, Connor) testing SpO2 Cables (TC 12/08/2023 01:08:56 PM, Campbell, Connor) Helping Kamran put back together and pack up 2 light sources to send back to olympus (TC 12/12/2023 09:21:46 AM, Campbell, Connor) testing SpO2 cables (TC 12/12/2023 10:00:01 AM, Campbell, Connor) Learning about anaethesia machine with Tomi (TC 12/13/2023 09:09:49 AM, Campbell, Connor) testing SpO2 cables (TC 12/13/2023 10:05:56 AM, Campbell, Connor) helping maisha unpack and clean up from uline boxes shipment (TC 12/13/2023 10:56:06 AM, Campbell, Connor) Helping Nara install bed locators on 11 (TC 12/13/2023 12:05:27 PM, Campbell, Connor) performance meeting with Adeel and Samir (TC 12/13/2023 02:43:28 PM, Campbell, Connor) Huddle (TC 12/13/2023 12:05:27 PM, Campbell, Connor) performance meeting with Adeel and Samir (TC 12/13/2023 02:43:58 PM, Campbell, Connor) trying to set up OR 12 tv, failed again due to wrong mounting bracket (TC 12/13/2023 02:44:27 PM, Campbell, Connor) setting up features on oculus (TC 12/14/2023 09:06:09 AM, Campbell, Connor) cleaning random e waste off the biomed shelf (TC 12/15/2023 11:03:53 AM, Campbell, Connor) Full team meeting (TC 12/18/2023 10:45:54 AM, Campbell, Connor) checking for beds to pm in the ICU (TC 12/19/2023 09:54:04 AM, Campbell, Connor) helping john move a bed mattress from the shop to 10 (TC 12/19/2023 09:54:31 AM, Campbell, Connor) looking for beds that need PM in ICU with john (TC 12/19/2023 02:41:44 PM, Campbell, Connor) Testing SpO2 Cables (TC 12/20/2023 08:14:08 AM, Campbell, Connor) Testing SpO2 cables (TC 12/20/2023 09:49:38 AM, Campbell, Connor) CALM meeting (TC 12/22/2023 09:50:09 AM, Campbell, Connor) troubleshooting some genius thermometers, neither of them worked, will try again next week
32560 | 198314 | CM | CMPLT | CMPLT | 2023-10-02 16:01 | 2023-10-02 16:01 | Patient strap stuck under table top | -Patient strap stuck under table top -Table top was taken apart to remove straps -Table was moving smooth and at normal speed after.
32510 | 198315 | CM | CMPLT | CMPLT | 2023-10-02 17:16 | 2023-11-01 14:29 | MSH DR1 PARTS | (TC 10/02/2023 05:15:54 PM, Heron, Chris) MSH DR1 PARTS (TC 09/30/2023 05:17:38 PM, Whittington, Scott) MSH DR1 PARTS | MSH DR1 PARTS
32520 | 198316 | CM | ACTIV | ISSUE | 2023-10-02 17:17 |  | MSH DR2 PARTS |  | MSH DR2 PARTS
32520 | 198317 | CM | ACTIV | ISSUE | 2023-10-02 17:23 |  | MSH DR2 COLLIMATOR LAMP BUTTON INTERMITTENT |  | MSH DR2 COLLIMATOR LAMP BUTTON INTERMITTENT
1061333 | 198318 | FC | CMPLT | CMPLT | 2023-10-03 07:24 | 2023-10-03 07:24 | QC checklist | (TC 10/03/2023 07:23:38 AM, Baghirzade, Nara) QC checklist completed. Returned to 18 South
| 198319 | PA | CMPLT | CMPLT | 2023-10-03 07:26 | 2024-05-02 08:11 | Ordering pressure sensors for stock | (TC 10/03/2023 07:28:19 AM, Baghirzade, Nara) SC 1003066501 (TC 02/02/2024 12:02:00 PM, Baghirzade, Nara) E-mailed Fan and Claire from mohawkmedbuy to cancel this order/shopping cart as the order was never received and one of the pump team members had already placed and received an order for syringe pressure sensors. Awaiting confirmation from Fan and Claire to close this WO. (TC 02/15/2024 08:22:55 AM, Baghirzade, Nara) Received a pack list and fedex delivery confirmation from BD for 1 pressure sensor. Printed pack list and gave it to Robin from receiving. Requested that he locate the signed packing list and then will further look into where the part is. (TC 02/15/2024 11:12:53 AM, Baghirzade, Nara) Robin from receiving provided a signed packing list; signed by Samir. Received 1 pressure sensor. PO 7202294079 (TC 05/02/2024 08:10:05 AM, Baghirzade, Nara) Sent email to Fan and Claire, confirming only 1 was received. Rest of the order was cancelled.
| 198320 | RDS | CMPLT | CMPLT | 2023-10-03 07:32 | 2023-10-02 07:31 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
| 198321 | CM | CMPLT | CMPLT | 2023-10-03 07:35 | 2023-10-02 07:35 | Biomed meeting | Biomed meeting
| 198322 | CM | CMPLT | CMPLT | 2023-10-03 07:37 | 2023-10-02 07:36 | Power adapter issue at 16N | Helped 16N to find a power adater for their patient chair
| 198323 | CM | CMPLT | CMPLT | 2023-10-03 07:52 | 2023-10-03 07:52 | worked on TMS | worked on TMS
1048249 | 198324 | CM | CMPLT | CMPLT | 2023-10-03 08:03 | 2023-10-03 08:04 | Intermittently switches on/off, error message "4-DP ON" | (TC 10/03/2023 08:03:28 AM, Baghirzade, Nara) Unable to bypass the error message. Re-seated the probe, changed batteries, cleaned the battery contacts, and opened up the device momentarily to check for loose connections. Could not mitigate the error message. Device will be retired and disposed of; repurposed a spare biomed thermometer for NICU.
| 198325 | CM | CMPLT | CMPLT | 2023-10-03 08:37 | 2023-10-03 08:37 | suction regulator issue | worked on suction regulator issues
1051699 | 198326 | CM | CMPLT | CMPLT | 2023-10-03 08:56 | 2023-10-03 08:57 | Repurposed for NICU | (TC 10/03/2023 08:55:54 AM, Campbell, Connor) Changed account and location; set status to active. Reprinted new asset tag. Repurposed spare thermometer from biomed. All functional tests passed. (TC 10/03/2023 08:57:01 AM, Baghirzade, Nara) Changed account and location; set status to active. Reprinted new asset tag. Repurposed spare thermometer from biomed. All functional tests passed.
| 198327 | CM | CMPLT | CMPLT | 2023-10-03 08:58 | 2023-10-03 09:00 | Slit lamp not working; no light | (TC 10/03/2023 08:58:26 AM, Baghirzade, Nara) Slit lamp in exam room 8 not working. Powered on the entire stand, all functions such as chair and external light are functional. There is power going to the stand and out of the stand as well. Re-seated the power cord on both ends of the slit lamp main switch box; AC light not coming on. Used a different power cord, same outcome. Informed Henna that she has to call the company to come in as this is the same slit lamp that has repeated problems.
65990 | 198328 | CM | CMPLT | CMPLT | 2023-10-03 09:04 | 2023-10-03 09:06 | Repeated 6990 error when connecting to PCU | (TC 10/03/2023 09:04:07 AM, Baghirzade, Nara) Pressure sensor was previously re-seated. Replaced the pressure sensor. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
| 198329 | CM | CMPLT | CMPLT | 2023-10-03 09:34 | 2023-10-03 09:34 | researching headlights for nara | (TC 10/03/2023 09:33:48 AM, Campbell, Connor) compiled a list of possible products for nara to browse
1057283 | 198330 | CM | CMPLT | NFF | 2023-10-03 09:54 | 2023-10-03 09:55 | Reported that thermometer is not measuring | (TC 10/03/2023 09:53:45 AM, Baghirzade, Nara) Checked the probe; probe in good condition. Took multiple temperature measurements successfully; no issues/errors. Checked with a different temperature probe; temperature module is fully functional. Functional tests for NIBP and SpO2 passed. No fault found.
| 198331 | CM | CMPLT | CMPLT | 2023-10-03 11:02 | 2023-10-03 11:03 | Masimo to Philips connectivity | (TC 10/03/2023 11:02:24 AM, Baghirzade, Nara) Worked with Kamran to connect the Masimo Root to Philips MX800. Successfully connected by using port#2 connecting to EC10, adjust the baudrate for the port, and used connectivity Intellivue Bridge > Radical
1059208 | 198332 | CM | CMPLT | NFF | 2023-10-03 13:00 | 2023-10-03 13:53 | SPO2 port not working | (TC 10/03/2023 01:52:03 PM, Scero, Tomlyn) Tested spo2 with a new sensor and using a simulator. No issues were found getting the proper reading.
1055949 | 198333 | CM | CMPLT | NFF | 2023-10-03 13:42 | 2023-10-03 13:44 | BP Cuff not working | (TC 10/03/2023 01:42:26 PM, Campbell, Connor) suspected the BP hose connector was improperly seated and slightly leaking reseated connector BP, SpO2 and temperature all functioning normally No fault found, returned to service
| 198334 | RDS | CMPLT | CMPLT | 2023-10-03 13:54 | 2023-10-03 14:29 | Weekly checks in the CCU | (TC 10/03/2023 02:02:37 PM, Scero, Tomlyn) Weekly checks in CCU complete
| 198335 | CM | CMPLT | CMPLT | 2023-10-03 13:57 | 2023-10-03 13:58 | Huddle | (TC 10/03/2023 01:57:13 PM, Baghirzade, Nara) NICU Huddle
| 198336 | CM | ACTIV | ISSUE | 2023-10-03 13:58 |  | Dr. Blazer requested for GLI 2021 references | (TC 10/03/2023 01:58:12 PM, Baghirzade, Nara) GLI 2012 references were requested for the pulmonary function devices; went into expair in admin mode and found the predicted tab to be greyed out. Took the system off the network but still was unable to access the predicted tab. Went during the lunch hour when there are no tests/patients on the devices but predicted tab was still not accessible. Called Antony from Novus and he suggested to check the software versions; the CPET device is 1.35 SW and the rest of the devices are 1.34. The issue might be in the software differences. Will send an email to Antony outlining the findings and possibly to upgrade the software version. (TC 11/06/2023 02:43:01 PM, Baghirzade, Nara) Confirmed that there is Support Assist on all PFT stations for Medisoft to remote in. CPET has teamviewer. Organized with Antony for Wednesday morning remote support, earliest time possible. Waiting on his confirmation. (TC 11/08/2023 08:55:55 AM, Baghirzade, Nara) Troubleshooting with Smain remotely: demonstrated how to change the predicteds and make it available for selection on the software. Will show the equations and predicteds to Dr. Blazer and Gaby, and with their decision, will change the predicteds across all stations including CPET. (TC 11/15/2023 09:03:28 AM, Baghirzade, Nara) Renamed the normes folder on the server to access the predicteds on local. Renamed it to "normes111" Went through all GLI references/predicteds and equations; took screenshots of every references and correlating equations for Dr. Blazer. Additionally, took screenshots of all existing equations that are currently used across all PFT stations to confirm if Dr. Blazer wants any of them removed. Went back on the server, renamed the folder back to "normes". Launched expair to ensure that the predicteds are the current ones that are used. Emailed Dr. Blazer, Gaby, and Steve all the references and equations. Waiting on confirmation from Dr. Blazer to apply the GLI references across all stations. (TC 11/16/2023 02:11:57 PM, Baghirzade, Nara) Dr. Blazer does not want GLI equations with race and ethnic corrections; as per Dr. Blazer's request, went to check all available equations on the software. Took screenshots of all available equations and found that all GLI equations that are available have race and ethnic corrections. Dr. Blazer also requested to find an equation marked as "GLI-other" however I was unable to find it on PFT stations. Checked the CPET statio and it also did not have it. Contacted Novus for support. Antony to find out if he can send me the equations directly. Met with Gaby and showed her the equations as well; she works at another pulmonary function clinic that is utilizing the GLI equations and will check about getting them from there as well. (TC 02/13/2024 11:35:09 AM, Baghirzade, Nara) Sent the 2nd follow up email to Dan Pinard regarding predicteds. (TC 05/08/2024 12:40:40 PM, Baghirzade, Nara) Received "normes" folder from Novus; dropped it in the C: drive Expair folder. Was able to access the predicteds from expair 3.14. Met with Gaby to test if it works on a PFT station; was able to access the predicteds as well however there is still ethnic corrections in the equations. Talked it over with Gaby and was able to get the normes folder from TGH. They have the "neutral" GLI that Dr. Blazer requires. Will follow up with Chris LaPlante from Novus regarding this and if we could use the same normes. (TC 05/13/2024 12:02:33 PM, Baghirzade, Nara) Received "normes" folder from TGH from Gaby. Compared this folder to the one that Novus sent; took snippings of the equations from the TGH normes and compared it to the Novus ones. Most of the equations are the same but require confirmation from Novus about this. Sent them the equations for both. All the ethnic corrections on the Novus normes are "caucasian"; requiring confirmation from Novus that this is just the naming and does not actually take in the caucasian ethnic correction. Also inquired if the predicteds can be used on version 1.34 (TC 09/05/2024 07:19:36 AM, Baghirzade, Nara) Remote session with Smain to go over the GLI predicteds that Novus and TGH sent over. Smain concluded that Novus and TGH predicteds were the old version and still accounted for race/ethnic correction. Smain then provided us with the most recent (updated in August 2024) GLI references. Confirmed that there are no race corrections. Smain also went over about how to apply the predicteds and any changes that are required to be made. Will follow up with Dr. Blazer and with her approval, will apply the new predicteds on all PFT/CPEt stations, physician office computers, and south ambulatory pod computers.
| 198337 | CM | CMPLT | CMPLT | 2023-10-03 14:17 | 2023-10-03 14:19 | Meeting with Adeel re: PFT values | (TC 10/03/2023 02:18:20 PM, Baghirzade, Nara) Meeting re: GLI references for PFT
1062467 | 198338 | CM | CMPLT | CMPLT | 2023-10-03 16:08 | 2023-10-03 16:09 | "scale not working" | (TC 10/03/2023 04:07:59 PM, Cortez, John) Confirmed issue Power cycled bed, issue resolved
1054469 | 198339 | CM | CMPLT | CMPLT | 2023-10-03 16:10 | 2023-11-09 13:28 | Right side head rail bent | (TC 10/03/2023 04:10:13 PM, Cortez, John) Swapped bed with spare, Requested service from Hillrom (TC 11/08/2023 01:26:59 PM, Cortez, John) Bed repaired by Vince
1054960 | 198340 | CM | CMPLT | CMPLT | 2023-10-03 16:13 | 2023-10-03 16:14 | Radical battery swelling | (TC 10/03/2023 04:13:14 PM, Cortez, John) Removed Radical battery
| 198341 | CM | CMPLT | CMPLT | 2023-10-03 16:14 | 2023-11-01 16:05 | misc biomed - oct 2023 | (TC 10/03/2023 04:14:28 PM, Cortez, John) Set up demo mode on test Philips monitor for Kamran (TC 10/03/2023 04:15:26 PM, Cortez, John) Gave spare suction regulator to Ivy (Manager, IR) for room 508, and tested Suggested to her to use yellow regulator instead, which has a connection similar to the regulators they're currently using (TC 10/02/2023 04:17:36 PM, Cortez, John) Installed tightvnc on Masimo viewstations on 15M and 16M (TC 10/03/2023 04:22:16 PM, Cortez, John) Helped gather pumps for 17M surg (TC 10/02/2023 04:23:05 PM, Cortez, John) huddle (TC 10/03/2023 04:29:06 PM, Cortez, John) Worked on template for device issue log document (TC 10/02/2023 03:51:52 PM, Cortez, John) Showed Barb (10S) how to correctly power cycle Procuity beds. She videotaped process for night staff (TC 10/13/2023 04:00:32 PM, Cortez, John) Addressed call from PACU, they were having problems admitting patient, Eventually was able to help them admit patient. Not sure what was causing the patient name search via MRN, to not work ealier. May be IT issue. (TC 10/18/2023 04:02:40 PM, Cortez, John) Nigel from RT brought a faulty flowmeter extension to biomed. Does not allow flow. Rod is bent which causing the flowmeter to remain in closed state. Unbent and it tested ok (TC 10/23/2023 01:44:27 PM, Cortez, John) Dianne Locke (CCU) asked about telepack coverage in Pod A PACU, Verified that it works and informed her (TC 10/24/2023 03:43:11 PM, Cortez, John) huddle (TC 10/25/2023 04:18:19 PM, Cortez, John) ICU code blue defib not plugged in, plugged in and ran 30J test (TC 10/31/2023 04:15:49 PM, Cortez, John) Setup Service Connect account for in preparation for Draeger training
1062263 | 198342 | CM | CMPLT | CMPLT | 2023-10-03 16:20 | 2023-10-03 16:21 | "scale not working" | (TC 10/02/2023 04:20:23 PM, Cortez, John) Confirmed issue, Power cycled bed, issue resolved
1062143 | 198343 | CM | CMPLT | CMPLT | 2023-10-03 16:21 | 2023-10-03 16:22 | "scale not working" | (TC 10/02/2023 04:21:31 PM, Cortez, John) Confirmed issue, Power cycled bed, issue resolved
32510 | 198344 | CM | ACTIV | ISSUE | 2023-10-03 21:35 |  | MSH DR2 TUBE LONGITUDIANL TRAVEL |  | MSH DR2 TUBE LONGITUDINAL TRAVEL
1049902 | 198345 | CM | CMPLT | CMPLT | 2023-10-03 21:50 | 2023-10-04 14:04 | MSH R62 MSH 7 OR top charger slot not working | (TC 10/04/2023 11:02:00 AM, Fan, Tak) Tested batteries on both slots. Charge slots are working but one of the batteries is not. Cannot get battery from warranty. Replaced battery with a spare. No detector issues but one of the pins was slightly bent on the detector. Straightened pin and tested, could not replicate any disconnections. | MSH R62 MSH 7 OR top charger slot not working
1066115 | 198346 | FC | CMPLT | CMPLT | 2023-10-04 07:50 | 2023-10-04 08:03 | QC check | (TC 10/04/2023 08:02:12 AM, Canjimavadivel, Kamallesh ) QC check complete
68899 | 198347 | CM | CMPLT | NFF | 2023-10-04 09:08 | 2023-10-04 09:11 | Battery not charging | (TC 10/04/2023 09:07:59 AM, Campbell, Connor) checked battery runtime, battery was low charged battery for an hour performed test infusion to fully drain battery and test capability no fault found charged battery fully
68913 | 198348 | CM | CMPLT | NFF | 2023-10-04 12:25 | 2023-10-04 12:26 | channel disconnected error | (TC 10/04/2023 12:25:04 PM, Campbell, Connor) powered system off no error seen after power off performed a check in performed a test infusion no faults found returned to service
1054537 | 198349 | PM | CMPLT | CMPLT | 2023-10-04 13:38 | 2023-10-04 13:38 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/04/2023 01:37:57 PM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
1067049 | 198350 | PM | CMPLT | CMPLT | 2023-10-04 13:39 | 2023-10-04 14:53 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/04/2023 01:45:37 PM, Cortez, John) Bed+Intellidrive batteries were replaced 70 mOhms, 24 uAmps Pendant missing clip 0-67 deg head section. At first, reading seemed to be delayed/inaccurate. Restarting bed seems to resolve issue. Scale verification failed, readings 5-10kg below acceptable values. Scale calibration not completing successfully. With Vince's (Hill-rom) assistance, found out that spacers were still installed on frame near head section. Vince mentioned that these should have been removed during installation. Scale successfully calibrated after removal. PM passed
1049903 | 198351 | CM | CMPLT | CMPLT | 2023-10-04 14:05 | 2023-10-04 14:06 | Bottom charger slot not charging. | (TC 10/04/2023 02:05:13 PM, Fan, Tak) Checked slots. Confirmed issue. Pins bent on charger. Replaced pins. Returned system.
1046206 | 198352 | CM | CMPLT | CMPLT | 2023-10-04 14:12 | 2023-10-10 14:20 | CW noisy. | (TC 10/04/2023 09:12:00 AM, Fan, Tak) Rashmi had an issue with CW, tested, found problem is on the probe. Ordered new probe. (TC 10/10/2023 02:18:47 PM, Fan, Tak) 232420YP0 replaced by 461106YP7. Tested, seems OK.
| 198353 | CM | CMPLT | CMPLT | 2023-10-04 14:20 | 2023-10-04 14:43 | NICU room 42, the nurse call was stuck and unable to silence it | (TC 10/04/2023 02:39:30 PM, Baghirzade, Nara) When nurse call was initiated, they were not able to silence or accept it on their phone. The responder was power cycled; tested that nurse call is working and they are able to silence/accept it. Nurse call is fully functional. (TC 10/04/2023 02:42:17 PM, Rhuland, Courtney) When nurse call was initiated, they were not able to silence or accept it on their phone. The responder was power cycled; tested that nurse call is working and they are able to silence/accept it. Nurse call is fully functional.
1047849 | 198354 | CM | CMPLT | CMPLT | 2023-10-04 14:31 | 2023-10-04 14:35 | Loud motor drive train | (TC 10/04/2023 02:31:28 PM, Rhuland, Courtney) Maintenance due message when powered on Changed maintenance date to clear message Loud motor drive train when test infusion Clutch assembly worn out Replaced clutch and greased lead screw Check in compelete Functional check - OK
1061960 | 198355 | CM | CMPLT | CMPLT | 2023-10-04 14:43 | 2023-10-04 14:45 | Limited motion functionality error | (TC 10/04/2023 02:43:10 PM, Baghirzade, Nara) Sri reported that they cannot use any of the motion functions on the bed and there is a patient on the bed. They are unable to move the patient off the bed so I had to work on the bed with the patient on it. Disconnected the AC power plug ONLY and turned off the battery. Waited for the bed to power cycle, and then booted it back up. Error went away and I had asked the nurse to test the motion controls; all motion functions are working. Also had the nurse confirm that all devices connected to the auxiliary port on the bed did not shut off during the power cycle. Bed is fully functional.
| 198356 | CM | CMPLT | CMPLT | 2023-10-04 14:47 | 2023-10-04 14:48 | Meeting with Kamran, Jack, and Milad | (TC 10/04/2023 02:46:53 PM, Baghirzade, Nara) Meeting with Kamran, Milad, and Jack from UHN regarding gathering high-resolution data from the Philips monitor.
66711 | 198357 | PM | CMPLT | CMPLT | 2023-10-05 07:28 | 2023-10-05 12:46 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 10/05/2023 12:43:47 PM, Rhuland, Courtney) Recall latches replaced 2 porthole doors and hinges replaced Power cord replaced RH calibrated with calibration salt Scale calibrated LV and main voltage calibrated Functional check - OK Electrical safety - OK
67631 | 198358 | CM | CMPLT | CMPLT | 2023-10-05 08:55 | 2023-10-05 08:56 | Recorder not printing clearly | (TC 10/04/2023 08:55:16 AM, Rendon Lopez , Elizabeth) Cleaned printerhead, ran selftest-OK
66516 | 198359 | CM | CMPLT | CMPLT | 2023-10-05 08:56 | 2023-10-05 08:56 | Recorder not printing clearly | (TC 10/04/2023 08:56:02 AM, Rendon Lopez , Elizabeth) Cleaned printerhead, ran selftest-OK
66502 | 198360 | CM | CMPLT | CMPLT | 2023-10-05 08:57 | 2023-10-05 08:57 | Recorder not printing clearly | (TC 10/04/2023 08:56:48 AM, Rendon Lopez , Elizabeth) Cleaned printerhead, ran selftest-OK
66501 | 198361 | CM | CMPLT | CMPLT | 2023-10-05 08:57 | 2023-10-05 08:58 | Recorder not printing clearly | (TC 10/04/2023 08:57:33 AM, Rendon Lopez , Elizabeth) Cleaned printerhead, ran selftest-OK
530510 | 198362 | CM | CMPLT | CMPLT | 2023-10-05 08:58 | 2023-10-05 08:59 | Missing valve | (TC 10/03/2023 08:58:35 AM, Rendon Lopez , Elizabeth) Replaced with spare, tested with tanks-OK
67413 | 198363 | CM | CMPLT | CMPLT | 2023-10-05 08:59 | 2023-10-05 09:01 | Missing valve, broken piece in tubing | (TC 10/03/2023 08:59:53 AM, Rendon Lopez , Elizabeth) Replaced valve, ring, and magenta scavenging tube. Checks - OK
1051378 | 198364 | CM | CMPLT | CMPLT | 2023-10-05 09:55 | 2023-10-05 10:08 | Created by Failed PM WO # 192900 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/05/2023 10:05:03 AM, Campbell, Connor) network card is actually good, the computer was causing the issue. PM shouldnt have failed the first time refer to procedure for ESA PM complete/passed, refer to ASM | (TC 10/05/2023 09:54:23 AM, Campbell, Connor) replaced battery attempted to connect to ASM, not working faulty network board waiting for order of replacements
62623 | 198365 | PA | CMPLT | CMPLT | 2023-10-05 11:16 | 2023-10-23 08:18 | Ordering parts for repairs/stock | (TC 10/05/2023 11:21:43 AM, Baghirzade, Nara) SC 1003069365 (TC 10/11/2023 07:56:44 AM, Baghirzade, Nara) Received batteries only. PO 7202295876 (TC 10/23/2023 08:14:24 AM, Baghirzade, Nara) Received 5 humidity sensor assembly and 2 humidifer kits: PO 7202295877
1047850 | 198366 | CM | CMPLT | CMPLT | 2023-10-05 11:26 | 2024-03-04 11:28 | Pump motor drive phase B post error | (TC 10/05/2023 11:26:18 AM, Baghirzade, Nara) As per service manual troubleshooting steps; allowed the device to charge for a few hours, still comes back with the same error. Re-seated the battery and still same issue persists. When plugging to AC, battery light does not come on. Replaced the battery. Allowed it charge for a few hours. Error still persists. Opened up the device, checked that all capacitors are seated properly and no physical damage observed on the boards. Next step is to replace the motor. (TC 10/11/2023 01:41:50 PM, Baghirzade, Nara) Replaced the motor. At first boot up, it stated the phase B post error and asked to input biomed password. After inputting the password, the error disappeared from the displayed. Switched off/on the pump multiple times and error does not show on display. Will leave pump to charge before attempting to test it. (TC 12/20/2023 02:29:11 PM, Rhuland, Courtney) When reassembling device, noticed the pressure sensor cable connector was damaged With help from Samir, the pin was resoldered and the connector was replaced Plunger cable was stuck in assembly, and when released, there was damage to the outside Repaired using heatshrink Reassembled device and no errors when turned on (TC 12/21/2023 08:50:27 AM, Baghirzade, Nara) Connected pump to pharmguard toolbox 2 software; updated the date/time. Uploaded the configuration. When trying to run an infusion, "syringe plunger not in place" error is recurring. (TC 02/16/2024 01:23:21 PM, Rhuland, Courtney) Only one plunger flipper was registering when opened Replaced plunger cable with one from another device Issue resolved When running test, new error came up: Check Clutch / Plunger Lever Replaced and greased clutch as plunger lever was working as expected Began rebuilding device (TC 02/20/2024 03:25:38 PM, Rhuland, Courtney) Finished rebuild of device New error popped up saying motor not running Took device apart again to check the connection to motor and gear No issue found Rebuilt device with same issue Potenially software issue associated with recall (TC 03/04/2024 11:26:55 AM, Rhuland, Courtney) Leadscrew nut was too tight Loosened screw and motor error cleared PM completed - all tests passed Refer to procedure for details
1063000 | 198367 | PM | CMPLT | CMPLT | 2023-10-05 11:38 | 2023-10-05 11:39 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/05/2023 11:37:54 AM, Cortez, John) Completed cleaning cycle
| 198368 | PA | CMPLT | CMPLT | 2023-10-05 11:41 | 2023-10-23 08:19 | Ordering sterile water | (TC 10/05/2023 11:43:25 AM, Baghirzade, Nara) SC 1003069421 (TC 10/23/2023 08:19:03 AM, Baghirzade, Nara) 4150264099
| 198369 | CM | CMPLT | CMPLT | 2023-10-05 14:46 | 2023-10-05 14:46 | 16 North work | worked with O'connor
395000 | 198370 | CM | CMPLT | CMPLT | 2023-10-05 15:04 | 2023-10-05 15:04 | Injector Piston end broken | -Injector Piston end broken -Installed spare injector while main injector was being repaired -After repair main injector was returned for use and spare injector moved back CR room
67913 | 198371 | CM | CMPLT | CMPLT | 2023-10-05 15:28 | 2023-10-06 10:24 | replacing incubator hood and door latch recall | (TC 10/05/2023 03:28:25 PM, Campbell, Connor) replaced incubator hood replaced door latches as per recall refer to next WO (198379) for PM and functional check
1067048 | 198372 | PM | CMPLT | CMPLT | 2023-10-05 15:54 | 2023-10-05 16:09 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/05/2023 03:57:08 PM, Cortez, John) 1041 error active, related to battery Bed scale also needs to be checked because spacer found still installed on another bed delivered at same time Bed+Intellidrive batteries were replaced 60 mOhms, 21 uAmps Bottom right rail, bed down button label ripped Ecountered 1155 error "flex a position error" while testing foot raising/lowering test. Error cleared after I retracted foot section a bit. Will monitor. PM passed
1066114 | 198373 | CM | CMPLT | CMPLT | 2023-10-05 16:59 | 2023-10-06 09:48 | "keeps indicating leak detected" | (TC 10/05/2023 04:59:16 PM, Cortez, John) QC check passed, Couldn't find any sign of a leak (TC 10/06/2023 09:47:17 AM, Cortez, John) Also ran seal check, tested ok
65119 | 198374 | CM | CMPLT | CMPLT | 2023-10-05 17:01 | 2023-10-25 09:06 | "button doesn't work" | (TC 10/05/2023 05:00:52 PM, Cortez, John) Confirmed issue on remote dose button, Pins are bent on pump, and dose cord connector side. Attempted to bend pins back, but getting a "bad button" state in status screen (TC 10/25/2023 09:06:00 AM, Reid, Emmerson ) Removed the dose cord and sending back
| 198375 | OC | CMPLT | CMPLT | 2023-10-06 08:57 | 2023-10-19 16:00 | On-call week, Oct 6-12 | (TC 10/06/2023 08:56:50 AM, Cortez, John) Daily check (TC 10/10/2023 03:15:12 PM, Cortez, John) Daily check (TC 10/11/2023 03:15:36 PM, Cortez, John) Daily check (TC 10/12/2023 09:59:12 AM, Cortez, John) Daily check
67931 | 198376 | CM | CMPLT | CMPLT | 2023-10-06 09:42 | 2023-10-06 09:43 | broken door hinge on porthole and SW door | (TC 10/06/2023 09:41:48 AM, Campbell, Connor) replaced the door hinges with new ones performed a functional check functional check passed, returned to service
65945 | 198377 | CM | CMPLT | CMPLT | 2023-10-06 09:47 | 2023-10-06 09:49 | channel error | (TC 10/06/2023 09:47:18 AM, Campbell, Connor) performed a functional check, no errors performed a check in, no errors noticed a fatal pressure sensor error in logs reseated pressure sensor cable performed another check in, no errors performed another functional test, no errors no fault found, returned to service
67097 | 198378 | CM | CMPLT | CMPLT | 2023-10-06 09:55 | 2023-10-06 09:57 | "BP not working" | (TC 10/02/2023 09:55:20 AM, Cortez, John) BP seems to work fine, but tab on connector that attaches to cuff is broken off. Replaced connector Hose connector also a bit loose on end that connects to machine
67913 | 198379 | PM | CMPLT | CMPLT | 2023-10-06 10:24 | 2023-10-06 10:57 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 10/06/2023 10:24:38 AM, Campbell, Connor) PM Passed Calibrated Scale Calibrated System votlage Calibrated RH Calibrated Line Voltage ESA failed due to broken power cord Took cord from another broken system (Asset 67956) becuase we were out of cords ESA now passed Refer to procedure for Electrical Safety Functional Test - OK Returned to service
1055901 | 198380 | CM | CMPLT | CMPLT | 2023-10-06 10:30 | 2023-10-06 10:33 | Screen damaged | (TC 10/06/2023 10:30:11 AM, Cortez, John) Visible artifacts on display Called Masimo for RMA, RMA# 1041064
1055901 | 198381 | CM | CMPLT | CMPLT | 2023-10-06 10:37 | 2023-10-24 13:34 | Broken screen | (TC 10/06/2023 10:36:57 AM, Cortez, John) Accidently closed previous WO (TC 10/06/2023 10:36:57 AM, Cortez, John) RMA 1041064 (TC 10/11/2023 03:32:16 PM, Cortez, John) shipped (TC 10/23/2023 09:31:04 AM, Cortez, John) Replacement was received, has software version 2.1.7.0 US Writing config the usual way results in "invalid profile setting" message, Called stryker and they provided newer version of MICT software, Exported config from working Root and wrote to replacement Root. Config was written successfully Supplied BP hose, All parameters tested ok
| 198382 | CM | CMPLT | CMPLT | 2023-10-06 11:08 | 2023-12-05 12:56 | power cord for Giraffe incubator. ordered from ACworks company | (TC 10/06/2023 11:10:38 AM, Chukri, Samir) Ordered parts for Giragge incubator. (TC 11/29/2023 02:28:47 PM, Chukri, Samir) Canceled the order and replaced order from acworks.ca web side use pro-card. (TC 12/05/2023 12:54:25 PM, Chukri, Samir) Received power cords and attached invoice to Samir Pro-card. keep them on the shelf behind Courtney
1054960 | 198383 | CM | CMPLT | CMPLT | 2023-10-06 13:37 | 2023-10-06 13:39 | "Temp probe not working" | (TC 10/06/2023 01:37:32 PM, Cortez, John) Temp probe well missing, Updated software to 2149 Added Radical battery as per recall All parameters tested ok
1059778 | 198384 | CM | CMPLT | CMPLT | 2023-10-06 15:38 | 2023-10-06 15:42 | QC check | (TC 10/06/2023 03:37:53 PM, Cortez, John) QC check passed Note mentioned that it was "not working, beeping" Pump was sent with power cord detached, and low battery alarms were found in alarm history
| 198385 | RDS | CMPLT | CMPLT | 2023-10-06 15:43 | 2023-10-06 15:45 | Weekly rounds- Oct 6, 2023 | (TC 10/06/2023 03:43:07 PM, Cortez, John) Showed Connor around ICU space as well
| 198386 | RDS | CMPLT | CMPLT | 2023-10-06 15:47 | 2023-10-06 15:49 | Weekly rounds 10S- Oct 6 2023 | (TC 10/02/2023 03:47:14 PM, Cortez, John) No issues
| 198387 | RDS | CMPLT | CMPLT | 2023-10-06 15:50 | 2023-10-06 15:51 | Weekly rounds 10N- Oct 6, 2023 | (TC 10/02/2023 03:49:36 PM, Cortez, John) No issues
63874 | 198388 | CM | CMPLT | CMPLT | 2023-10-10 07:38 | 2023-10-10 07:39 | doors wont latch | (TC 10/10/2023 07:38:30 AM, Campbell, Connor) doors were on the wrong sides of the incubator switched them performed functional test test passed, returned to service
| 198389 | RDS | CMPLT | CMPLT | 2023-10-10 08:25 | 2023-10-10 08:25 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
65986 | 198390 | CM | CMPLT | CMPLT | 2023-10-10 09:53 | 2023-11-16 10:02 | error, does not recognize the correct medication/ not programmable | (TC 10/10/2023 09:53:39 AM, Campbell, Connor) error produced when performing check in during the plunger force accuracy test, does not pass error 354.6690, suspect need pressure sensor reseat or new sensor, or could be logic board went to open and noticed the front cases is cracked, cant fix until we recieve the new front covers (TC 10/11/2023 09:01:19 AM, Campbell, Connor) found a front case in a spare parts bin reseated the pressure sensor and changed the front case to the spare part performed a check in check in failed at the button test, the front case spare is broken wasted time, still need to wait for front cases to arrive before the unit can be functioning (TC 11/16/2023 09:58:46 AM, Campbell, Connor) replaced the front case replaced the pressure sensor and mylar gasket attempted check in, it failed calibrated the pump re-attempted check in, check in passed pump is functional now returned to service
1053876 | 198391 | CM | CMPLT | NFF | 2023-10-10 10:16 | 2023-10-10 10:18 | Cut not working | (TC 10/10/2023 10:15:55 AM, Reid, Emmerson ) Anita called to say that the ESU on the emergency cart was not working for CUT. I gathered up tools and went to check it out. I turned the ESU on and connected with the tester. I saw "--" for power under cut so I changed to basic mode and that allowed me to change the power level. With the tester it read that power was being delivered. NFF
1059828 | 198392 | PM | CMPLT | CMPLT | 2023-10-10 11:01 | 2023-10-10 11:01 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/10/2023 11:00:46 AM, Campbell, Connor) Disinfected, cleaned, rinsed and refilled
1058168 | 198393 | CM | CMPLT | CMPLT | 2023-10-10 11:44 | 2023-11-10 15:51 | "Battery 1 replacement required" | (TC 10/10/2023 11:44:14 AM, Cortez, John) Confirmed issue, SOH below 20% Will order replacement batteries and the charger(battery conditioner) (TC 10/10/2023 11:44:14 AM, Cortez, John) -battery part#: HMA-MSP369108 -battery charger part#: 369104 (TC 11/08/2023 04:19:55 PM, Cortez, John) Parts arrived, Will charge new batteries overnight, and condition faulty battery (TC 11/10/2023 03:50:13 PM, Cortez, John) Completed Preop tests successfully, returned to service
| 198394 | RDS | CMPLT | CMPLT | 2023-10-10 13:05 | 2023-10-10 13:11 | Weekly checks in CCU | (TC 10/10/2023 01:10:06 PM, Scero, Tomlyn) Weekly checks complete. Defib 30J tests complete.
67936 | 198395 | CM | CMPLT | CMPLT | 2023-10-10 13:23 | 2023-10-10 13:24 | south west latch is broken | (TC 10/10/2023 01:23:07 PM, Campbell, Connor) replaced latch with a spare performed a functional test functional test passed unit returned to service
1063746 | 198396 | PM | CMPLT | CMPLT | 2023-10-10 14:44 | 2023-10-10 14:45 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/10/2023 02:45:03 PM, Campbell, Connor) Drained, disinfected, rinsed and refilled with sterile water
67932 | 198397 | CM | CMPLT | CMPLT | 2023-10-11 07:27 | 2023-10-11 08:11 | Broken latch | (TC 10/11/2023 08:04:31 AM, Rhuland, Courtney) Replaced broken side door latch Functional check - OK
1067437 | 198398 | FC | CMPLT | CMPLT | 2023-10-11 07:32 | 2023-10-11 07:33 | QC checklist | (TC 10/11/2023 07:32:24 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
66839 | 198399 | CM | CMPLT | CMPLT | 2023-10-11 07:49 | 2023-10-11 07:51 | Reported that unable to measure NIBP | (TC 10/11/2023 07:48:50 AM, Baghirzade, Nara) Checked the NIBP settings in the X2; set to manual mode and NIBP window appears on the screen. Used allen key to open the NIBP connector; checked that it is in tact and clean. Took multiple NIBP measurements with simulator; each measurement was successfully taken. Checked ECG and SpO2, both are functional. No fault found.
| 198400 | CM | CMPLT | CMPLT | 2023-10-11 08:23 | 2023-10-19 09:27 | Order Power Strip and Holder for O.R | (TC 10/19/2023 09:27:31 AM, Chukri, Samir) Received parts and give it to Conner to assembly
64290 | 198401 | CM | CMPLT | CMPLT | 2023-10-11 09:45 | 2023-10-11 09:47 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 10/11/2023 09:44:52 AM, Rhuland, Courtney) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609340 Put monitor back on the arm. Confirmed that monitor is functional.
64278 | 198402 | CM | CMPLT | CMPLT | 2023-10-11 09:45 | 2023-10-11 09:47 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 10/11/2023 09:46:43 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609367 Put monitor back on the arm. Confirmed that monitor is functional.
67217 | 198403 | CM | CMPLT | CMPLT | 2023-10-11 09:47 | 2023-10-11 09:48 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 10/11/2023 09:47:53 AM, Rhuland, Courtney) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609370 Put monitor back on the arm. Confirmed that monitor is functional.
64268 | 198404 | CM | CMPLT | CMPLT | 2023-10-11 09:47 | 2023-10-11 09:48 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 10/11/2023 09:47:35 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609338 Put monitor back on the arm. Confirmed that monitor is functional.
67216 | 198405 | CM | CMPLT | CMPLT | 2023-10-11 09:48 | 2023-10-11 09:49 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 10/11/2023 09:48:22 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609319 Put monitor back on the arm. Confirmed that monitor is functional.
67211 | 198406 | CM | CMPLT | CMPLT | 2023-10-11 09:51 | 2023-10-11 09:51 | Project- J13 option ( RS232 board ) upgrade in monitor | (TC 10/11/2023 09:50:48 AM, Baghirzade, Nara) Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609347 Put monitor back on the arm. Confirmed that monitor is functional.
65967 | 198407 | CM | CMPLT | CMPLT | 2023-10-11 10:05 | 2023-10-20 07:30 | barrel clamp and syringe knob both broken (likely due to a drop) | (TC 10/11/2023 10:05:52 AM, Campbell, Connor) replaced the syringe knob with a spare when the housing was opened to replace barrel clamp, the internal frame was cracked as well due to the drop internal frame must be ordered, repairs can be completed when the frame arrives (TC 10/19/2023 02:57:46 PM, Campbell, Connor) syringe knob was not closing all the way, had to replace it with another spare, now it closes well replaced the internal frame performed a check-in, it passed started a functional test to be left overnight, will complete WO and return to service when the functional test is passed (TC 10/20/2023 07:29:35 AM, Campbell, Connor) functional test passed, returned to service
66132 | 198408 | CM | CMPLT | CMPLT | 2023-10-11 11:53 | 2023-10-11 11:53 | Broken bracket | Replaced 2 x broken holder braket
1059209 | 198409 | CM | CMPLT | CMPLT | 2023-10-11 12:17 | 2023-10-11 13:41 | Battery malfunction | (TC 10/11/2023 01:39:59 PM, Scero, Tomlyn) Battery replaced and X3 changed before returning
| 198410 | CM | CMPLT | CMPLT | 2023-10-11 12:22 | 2023-10-11 12:23 | testing various "broken" SpO2 cables from PACU and NICU | (TC 10/11/2023 12:22:09 PM, Campbell, Connor) Tested cables from PACU on the X3, all were determined to be functional Tested cables from NICU, all were determined to be functional Cables returned to service
| 198411 | CM | CMPLT | CMPLT | 2023-10-11 13:41 | 2023-10-11 13:41 | Emergency power transfer meeting | (TC 10/11/2023 01:41:03 PM, Baghirzade, Nara) Emergency power transfer meeting
| 198412 | CM | CMPLT | CMPLT | 2023-10-11 13:44 | 2023-10-11 13:49 | Testing MX800 monitor when there is no power; preparation for emergency power transfer | (TC 10/11/2023 01:43:56 PM, Baghirzade, Nara) Plugged in the simulator to the X2 (connected to MX800 using rack). MX800 is plugged in to the red outlet; after disconnecting from power, the rack loses power, monitor loses power, and high pitched beep comes from the monitor. X2 continued monitoring and alarming. Plugged the monitor back to wall outlet and the X2 automatically connected to it as well as the monitor booted up on its own without having to press the power button. Charging 4 X2 batteries just in case the batteries die during the emergency power transfer. Worked with Courtney.
1059778 | 198413 | FC | CMPLT | CMPLT | 2023-10-11 13:50 | 2023-10-11 13:50 | QC checklist | (TC 10/11/2023 01:49:44 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
| 198414 | CM | CMPLT | CMPLT | 2023-10-11 14:33 | 2024-02-27 10:12 | BIS module repairs | (TC 10/11/2023 02:33:54 PM, Scero, Tomlyn) Price quote requested from Philips. They called and provided me with the price. Ordered 3 on ereq (TC 01/11/2024 02:17:21 PM, Scero, Tomlyn) Followed up with Fan from Mohawk medbuy to get an ETA on arrival. (TC 02/21/2024 08:33:38 AM, Scero, Tomlyn) followed up with Fan and Philips for an ETA again. (TC 02/21/2024 01:59:22 PM, Scero, Tomlyn) repaired and tested 2 BIS module cables. returned them to OR (TC 02/27/2024 10:10:52 AM, Scero, Tomlyn) order arrived. brought to use in RT shop
62987 | 198415 | CM | CMPLT | CMPLT | 2023-10-11 14:41 | 2023-10-11 14:44 | blue light "dying", too yellow | (TC 10/11/2023 02:41:36 PM, Campbell, Connor) dissasembled light, cleaned all the dust replaced the light on timer replaced one of the cracked fan covers replaced the high/low intensity switch (this fixed the issue of the light being too yellow, the light would flicker between blue and yellow when moved, this was due to a faulty high/low switch) calibrated the light intensity to fall within ranges specified in manual for high/low intensity
1054521 | 198416 | CM | CMPLT | CMPLT | 2023-10-11 14:54 | 2023-10-11 15:03 | "V1 lead doesn't read" | (TC 10/11/2023 02:54:08 PM, Cortez, John) Confirmed issue, V6 and RL lead also faulty Replaced 2 sets of chest leads and one limb lead. Tested ok
32510 | 198417 | CM | CMPLT | CMPLT | 2023-10-11 14:57 | 2023-10-11 14:57 | Not reaching position | restart check positions
32520 | 198418 | CM | CMPLT | CMPLT | 2023-10-11 14:59 | 2023-10-11 14:59 | not reaching position
1047913 | 198419 | CM | CMPLT | CMPLT | 2023-10-11 15:17 | 2023-10-11 15:20 | "water pressure error when using F-ST" | (TC 10/11/2023 03:16:53 PM, Cortez, John) Unable to repeat error, but renewed water anyway. Asked Gabby (Dovigi manager) to contact me if error returns
1056665 | 198420 | CM | CMPLT | CMPLT | 2023-10-11 15:18 | 2023-10-11 15:19 | Video not working | (TC 10/11/2023 03:18:03 PM, Reid, Emmerson ) Video was not working on the monitors. I power cycled the ScaleOR and secured the video cables slightly. Video came back. Working now.
1055792 | 198421 | CM | CMPLT | CMPLT | 2023-10-11 15:20 | 2023-10-11 15:20 | No electrode connected | (TC 10/11/2023 03:19:59 PM, Reid, Emmerson ) Connected electrode and did 30J test. Working now.
1062032 | 198422 | CM | CMPLT | CMPLT | 2023-10-11 15:22 | 2023-10-23 15:07 | "not moving" | (TC 10/10/2023 03:22:21 PM, Cortez, John) Confirmed error, Fowler actuator detached/broken from frame Requested service from Stryker (TC 10/12/2023 03:06:14 PM, Cortez, John) Bed repaired by Mark (stryker)
32540 | 198423 | CM | CMPLT | CMPLT | 2023-10-11 15:54 | 2023-10-11 15:54 | Error in recon sequence | Images not being reconstructed Transfer not going over to PACS -Speicfic case was holding up the recons and transfers Prioritized images for recon and transfer -Restarted the system and techs were able to compete the recons from the raw data -Images were going over to Pacs
1065159 | 198424 | CM | CMPLT | CMPLT | 2023-10-12 07:15 | 2023-10-12 07:16 | Unit displaying "replace SPO2 sensor" | (TC 10/12/2023 07:15:05 AM, Rhuland, Courtney) Unit stated that the device was displaying message even after changing sensor Replaced sensor and SPO2 worked Functional check - OK
1054774 | 198425 | CM | CMPLT | NFF | 2023-10-12 08:26 | 2023-10-12 08:29 | "retry temperature" alarm | (TC 10/12/2023 08:27:04 AM, Rhuland, Courtney) Error message on screen when device powered on Cleared error and took multiple test temperature readings Functional check - OK No fault found once error was cleared
67950 | 198426 | PA | CMPLT | CMPLT | 2023-10-12 08:36 | 2023-10-18 14:52 | Ordering incubator hood | (TC 10/12/2023 08:37:39 AM, Baghirzade, Nara) SC 1003073958
1063793 | 198427 | CM | CMPLT | CMPLT | 2023-10-12 08:58 | 2023-12-08 08:23 | Lamp is getting too hot | (TC 10/12/2023 08:58:16 AM, Reid, Emmerson ) Gathered from OR and let it run. It does get hot but not unusually hot. The handle for adjusting the light is missing, I reached out to the vendor to order a replacement.
501361 | 198428 | CM | CMPLT | CMPLT | 2023-10-12 09:03 | 2023-10-13 10:22 | Bad crystals | (TC 10/12/2023 09:03:25 AM, Fan, Tak) Philips case #0122437690 for replacement. (TC 10/13/2023 10:21:40 AM, Fan, Tak) B0WH6R replaced F09PMV on Oct 13, 2023
63026 | 198429 | PA | CMPLT | CMPLT | 2023-10-12 09:26 | 2023-12-08 07:51 | Ordering parts for repair and stock | (TC 10/12/2023 09:30:35 AM, Baghirzade, Nara) SC 1003074047 (TC 10/23/2023 08:23:33 AM, Baghirzade, Nara) Received 2 enclosure casings: PO 7202299409 (TC 10/25/2023 07:44:49 AM, Baghirzade, Nara) Received 2 switches. PO 7202299409 (TC 12/08/2023 07:50:13 AM, Baghirzade, Nara) Received full order PO 7202299409
1053391 | 198430 | CM | CMPLT | CMPLT | 2023-10-12 09:40 | 2023-11-06 11:45 | Not charging | (TC 10/12/2023 09:44:14 AM, Baghirzade, Nara) AC light indicator intermittently switches on/off. When moving the power cord around, it causes it to disconnect. Opened the device and found that the power connector on the board is loose and not making proper contact; additionally the part is also damaged. Contacted Ifrah and a service tech from KCI; individual board components are not sold however the main board can be purchased. Main board is on back order. Alternatively, device can be sent to KCI for repair. Presented Sri with all the options and waiting for their decision. Also emailed Ifrah about getting a loaner for the duration of the repair. (TC 10/12/2023 10:18:37 AM, Baghirzade, Nara) Sending it in for repair. Requested a loaner. Pick up confirmation: 12897375 Loaner confirmation: 13128042 (TC 10/12/2023 10:40:58 AM, Baghirzade, Nara) Created no charge PO: SC 1003074227 (TC 10/13/2023 12:56:54 PM, Baghirzade, Nara) VAC picked up by 3M. Attached pickup receipt. (TC 11/06/2023 11:38:32 AM, Baghirzade, Nara) Sri could not locate the VAC since Thursday, Nov 2nd. Called 3M to find out who signed for the drop off at MSH - found that 3M used their courier to deliver the VAC and it was delivered to shipping/receiving. After trying to locate it there, finally found the VAC at materials management. Opened the case and found that 3M upgraded the carrying bag and fixed all the covers including the charging port cover. Attached the service report as well as the inspection card and drop off receipt to this WO. Dropped off the VAC with the 11 South ward clerk and sent Sri a message confirming the drop off at 11 South.
1061999 | 198431 | CM | CMPLT | CMPLT | 2023-10-12 10:06 | 2023-10-24 13:41 | No head board movement - maybe motor issue | (TC 10/12/2023 10:06:25 AM, Ferreira, Qwynn) Went to unit to troubleshoot the bed but no resolve connected with Mark to look into the repair
1056125 | 198432 | PM | CMPLT | CMPLT | 2023-10-12 10:18 | 2023-10-12 10:19 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/12/2023 10:18:33 AM, Campbell, Connor) Drained, disinfected, cleaned, rinsed and refilled
1056125 | 198433 | CM | CMPLT | CMPLT | 2023-10-12 10:57 | 2023-10-12 11:00 | Reported no water during use with code | (TC 10/12/2023 10:57:30 AM, Canjimavadivel, Kamallesh ) Went up with water bottles and found reservoir had adequate water level. Asked nurse what the issue was and was told they got meassage related to checking water. Tried connecting the ports and noticed error message keeps showing. Grabbed spare from ICu to use in meantime. Discovered hole in blanket they tried to use. Installed new blanekt and let ICU spare with them. Brought down unit to verify functionality. No errors came while running continuosly. Intital error most likelt due to leaking blanket with hole.
66455 | 198434 | CM | CMPLT | CMPLT | 2023-10-12 12:03 | 2023-10-12 12:03 | unit is alarming | Unit was alarming. RT shut it down from main power switch. Turn it on and run a device check. It is functional. Draeger: "Based on what I see above I would recommend the following 1. Check that the power cord it tight and pull the power cord and make sure it switches to battery and the battery lasts the correct amount of time. Plug it back in and make sure it is fully recharging the battery. 2. Run the Device and breathing circuit check a number of times to make sure it goes through without issue 3. If both of these checkout then it looks like the unit can go back into service. Did 3 x shutdown and device checks. Run the unit. Functional. Returned back to RT and informed Rex.
| 198435 | CM | CMPLT | CMPLT | 2023-10-12 12:25 | 2023-11-02 06:54 | Slit lamp not powering on | (TC 10/12/2023 12:24:59 PM, Baghirzade, Nara) Few weeks ago when Henna asked to check the slit lamp, I was able to isolate the issue to the power outlet located on the arm underneath the slit lamp. Innovamed sent over a fuse to replace on the main board; replaced the fuse on the main board however it did not mitigate the issue. When testing the switch box under the slit lamp, it only powers on when plugging it directly to the wall outlet. When its plugged to the outlet on the arm, it does not power on. The base stand itself is powered on, all lights and chair motion controls are functional however there is no power specifically to the outlet on the arm. Tested the original fuse that was on the board and it is shorted. Henna to send pictures of the outlet and switch box and I'll email Innovamed to describe all steps taken. (TC 11/01/2023 06:53:29 AM, Baghirzade, Nara) Followed up with Henna. Innovamed to come on site to do the repair as their initial solution did not resolve the problem.
| 198436 | CM | CMPLT | CMPLT | 2023-10-12 12:29 | 2023-10-13 12:58 | Oto-ophthalmoscope wall unit installation | (TC 10/12/2023 12:29:02 PM, Baghirzade, Nara) Lyudmila had called to have the oto-ophthalmscope unit installed in room 20. The unit requires building services to come in and mount it on the wall. Lyudmila put in the service request. Also, she is requesting an additional vitals signs monitor for room 10. Will follow up with Nima or Becky regarding this.
67354 | 198437 | CM | CMPLT | CMPLT | 2023-10-12 13:13 | 2023-10-12 13:20 | Not turning on | (TC 10/12/2023 01:17:06 PM, Scero, Tomlyn) Plugged unit in and cofirmed no power to the unit. I noticed a rattling coming from the plug and found when I opened it that the White neutral wire wasn't connected to the plug at all. I reconnected it and tested the electrical safety. LC 16uA SF 32uA GWR 150mOhm Ran unit and found that the orings needed to be regreased because air was getting in the tubing. After I regreased them I ran the unit to make sure it was running properly. Tests complete.
1061333 | 198438 | FC | CMPLT | CMPLT | 2023-10-12 14:09 | 2023-10-12 14:09 | QC checklist | (TC 10/12/2023 02:09:02 PM, Baghirzade, Nara) QC checklist completed. Returned to 18 South
| 198439 | CM | CMPLT | CMPLT | 2023-10-12 14:10 | 2023-10-12 14:10 | Dry run for emergency power transfer | (TC 10/12/2023 02:09:50 PM, Baghirzade, Nara) Dry run for emergency power transfer
68759 | 198440 | CM | CMPLT | CMPLT | 2023-10-12 14:50 | 2023-10-13 07:06 | System error code 120.4610 and 120.4540 | (TC 10/12/2023 02:49:46 PM, Baghirzade, Nara) System error code 120.4610 and 120.4540 pertaining to the power supply. Left device to charge however charge is not being retained in the battery. Replaced the power supply. Replaced front case as the AC light was flickering on/off and partially the keypad was not working. Left device to charge overnight. (TC 10/13/2023 07:05:57 AM, Baghirzade, Nara) Checked battery status; battery in good condition and fully charged.
1068035 | 198441 | CM | CMPLT | CMPLT | 2023-10-12 14:59 | 2023-10-12 15:01 | Not specified on tag what the issue was | (TC 10/12/2023 02:59:30 PM, Baghirzade, Nara) Checked battery level; battery is good. Received it with battery at 59%, left it to charge and charged to full. Brought device to Matthew from IT, he power cycled the computer, and re-seated a connector. Device booted up properly. Sent it back to 11 South
1062467 | 198442 | CM | CMPLT | CMPLT | 2023-10-13 08:24 | 2023-10-13 08:32 | "not working" | (TC 10/13/2023 08:23:47 AM, Cortez, John) Confirmed issue, same error as WO from acouple of weeks ago. 3009-1-28-802 Power cycling bed clears error for now, but will monitor to see if error returns (TC 10/12/2023 08:23:47 AM, Cortez, John) Confirmed issue, same error as WO from acouple of weeks ago. 3009-1-28-802 Power cycling bed clears error for now, but will monitor to see if error returns
67969 | 198443 | CM | CMPLT | CMPLT | 2023-10-13 08:38 | 2023-10-13 08:39 | Door latch was broken | (TC 10/13/2023 08:38:23 AM, Baghirzade, Nara) Replaced door latch. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C and 74.8% RH. Incubator is fully functional.
1067051 | 198444 | PM | CMPLT | CMPLT | 2023-10-13 08:44 | 2023-10-13 08:57 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/12/2023 08:54:30 AM, Cortez, John) Bed+Intellidrive batteries were replaced 60 mOhms, 22 uAmps Pendant missing clip PM passed
1062019 | 198445 | CM | CMPLT | CMPLT | 2023-10-13 09:04 | 2023-10-13 09:09 | Bed error-extended | (TC 10/13/2023 09:07:15 AM, Rhuland, Courtney) Checked and cleared bed error logs Power cycled bed Noted errors in extended description for future referral *bed locator bracket noted to be bent* | Error codes: 3009-3-5-1370 3009-5-7-1203 3009-3-12-701
1055575 | 198446 | CM | CMPLT | CMPLT | 2023-10-13 09:06 | 2023-10-25 08:06 | Ordering cuffs for 14S and 14 N | received cuffs. PO# 7202299578 (TC 10/25/2023 08:07:18 AM, Arslan, Olcay) received cuffs. PO# 7202299578
1064838 | 198447 | CM | CMPLT | CMPLT | 2023-10-13 09:24 | 2023-10-13 09:24 | Batt malfunction | (TC 10/12/2023 09:24:01 AM, Cortez, John) Conditioned battery, passed
66403 | 198448 | CM | CMPLT | CMPLT | 2023-10-13 09:35 | 2023-10-20 09:58 | ordering brackets for heaters | received brackets. PO#7202299575
1062999 | 198449 | PM | CMPLT | CMPLT | 2023-10-13 09:36 | 2023-10-13 09:38 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/13/2023 09:36:15 AM, Canjimavadivel, Kamallesh ) Cleaned ports and hose ends with alcohol to remove buildup. Port 3 return leaking a bit when running cleaning cycle. Cleaning cycle complete
32540 | 198450 | CM | CMPLT | CMPLT | 2023-10-13 10:13 | 2023-10-12 10:11 | recon errors reported | recon errors reported Service completed troubleshooting recon unit Error logs tested Serviced replaced IDD controller CT phantom tested OK
| 198451 | RDS | CMPLT | CMPLT | 2023-10-13 10:49 | 2023-10-13 10:51 | Weekly Rounds: POD A and 18 South | (TC 10/13/2023 10:49:12 AM, Baghirzade, Nara) Weekly rounds in POD A and 18 South completed. See procedure.
| 198452 | RDS | CMPLT | CMPLT | 2023-10-13 10:52 | 2023-10-13 10:53 | Weekly rounds: NICU and RESUS | (TC 10/13/2023 10:51:59 AM, Baghirzade, Nara) Weekly rounds in NICU and RESUS completed. See procedure.
| 198453 | RDS | CMPLT | CMPLT | 2023-10-13 10:54 | 2023-10-13 10:55 | Weekly Rounds: 11 South | (TC 10/13/2023 10:53:39 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure.
1047862 | 198454 | FC | CMPLT | CMPLT | 2023-10-13 11:16 | 2023-10-13 11:17 | Functional check | (TC 10/13/2023 11:16:30 AM, Baghirzade, Nara) All functional tests passed. Ran an infusion at various rates; occluded the syringe and occlusion alarm is functional. Infusion was successfully completed at different rates. Sending pump back to NICU.
| 198455 | CM | CMPLT | CMPLT | 2023-10-13 12:14 | 2023-11-22 10:46 | Tools for workshop | (TC 10/13/2023 12:13:46 PM, Rhuland, Courtney) Ordering tools for desk WO #7202301064
| 198456 | CM | CMPLT | CMPLT | 2023-10-13 12:19 | 2023-10-11 12:20 | Ontario Cybersecurity Conference
1055811 | 198457 | CM | CMPLT | CMPLT | 2023-10-13 13:38 | 2023-10-13 13:46 | Software upgrade to 02.36.21 | (TC 10/13/2023 01:45:17 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055812 | 198458 | CM | CMPLT | CMPLT | 2023-10-13 13:39 | 2023-10-13 13:47 | Software upgrade to 02.36.21 | (TC 10/13/2023 01:46:04 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055935 | 198459 | CM | CMPLT | CMPLT | 2023-10-13 14:17 | 2023-10-13 14:22 | 12S floor audit - Masimo | (TC 10/13/2023 02:17:30 PM, Rhuland, Courtney) Met with nurse manager to discuss issues floor was having with Masimos Replaced faulty probe on this device All other devices were fine
| 198460 | CM | CMPLT | CMPLT | 2023-10-13 14:23 | 2023-10-13 14:24 | SPO2 sensors
1068036 | 198461 | ACC | CMPLT | CMPLT | 2023-10-13 15:01 | 2023-10-13 15:02 | Incoming inspection | (TC 10/13/2023 03:00:56 PM, Cortez, John) Asset tagged, visual inspection passed
1068037 | 198462 | ACC | CMPLT | CMPLT | 2023-10-13 15:03 | 2023-10-13 15:03 | Incoming inspection | (TC 10/13/2023 03:02:57 PM, Cortez, John) Asset tagged, visual inspection passed
1068038 | 198463 | ACC | CMPLT | CMPLT | 2023-10-13 15:04 | 2023-10-13 15:06 | Incoming inspection | (TC 10/13/2023 03:04:09 PM, Cortez, John) Asset tagged, visual inspection passed
1068039 | 198464 | ACC | CMPLT | CMPLT | 2023-10-13 15:07 | 2023-10-13 15:08 | Incoming inspection | (TC 10/13/2023 03:07:32 PM, Cortez, John) Asset tagged, visual inspection passed
1068040 | 198465 | ACC | CMPLT | CMPLT | 2023-10-13 15:08 | 2023-10-13 15:09 | Incoming inspection | (TC 10/13/2023 03:08:20 PM, Cortez, John) Asset tagged, visual inspection passed
| 198466 | OC | CMPLT | CMPLT | 2023-10-13 16:36 | 2023-10-13 16:40 | Kamran Samanian On Call, OCT 13 - OCT 19; | (TC 10/13/2023 04:36:35 PM, Samanian, Kamran) Received a call from helpdesk that " user can not associate the patient to philips monitor in PACU POD B bed 8 , ticket#INC-29605772" called John to check with the unit. it seemed that user wanted to sreach for MRN and after few tryies it showed up on the monitor however it was not used to admit the patient. Regardless IBE was checked and everything looked normal and checked data IN and OUT for this bed. Noticed that data processsing for ADT info is a bit slow in IBE and that could have been a root cause. checked few sample patient search and all good. (TC 10/13/2023 04:50:00 PM, Samanian, Kamran) Received a call from helpdesk that " user can not associate the patient to philips monitor in PACU POD B bed 8 , ticket#INC-29605772" called John to check with the unit. it seemed that user wanted to sreach for MRN and after few tryies it showed up on the monitor however it was not used to admit the patient. Regardless IBE was checked and everything looked normal and checked data IN and OUT for this bed. Noticed that data processsing for ADT info is a bit slow in IBE and that could have been a root cause. checked few sample patient search and all good. (TC 10/13/2023 04:40:00 PM, Samanian, Kamran) Received a call from helpdesk that " user can not associate the patient to philips monitor in PACU POD B bed 8 , ticket#INC-29605772" called John to check with the unit. it seemed that user wanted to sreach for MRN and after few tryies it showed up on the monitor however it was not used to admit the patient. Regardless IBE was checked and everything looked normal and checked data IN and OUT for this bed. Noticed that data processsing for ADT info is a bit slow in IBE and that could have been a root cause. checked few sample patient search and all good. (TC 10/13/2023 04:40:00 PM, Samanian, Kamran) Received a call from helpdesk that " user can not associate the patient to philips monitor in PACU POD B bed 8 , ticket#INC-29605772" called John to check with the unit. it seemed that user wanted to sreach for MRN and after few tryies it showed up on the monitor however it was not used to admit the patient. Regardless IBE was checked and everything looked normal and checked data IN and OUT for this bed. Noticed that data processsing for ADT info is a bit slow in IBE and that could have been a root cause. checked few sample patient search and all good.
| 198467 | RDS | CMPLT | CMPLT | 2023-10-16 07:48 | 2023-10-16 14:00 | Weekly Checks in CCU | (TC 10/16/2023 01:58:07 PM, Scero, Tomlyn) Weekly checks complete. Defib 30J tests passed.
62696 | 198468 | CM | CMPLT | CMPLT | 2023-10-16 08:04 | 2023-10-16 08:21 | Leads V1 and V3 not displaying consistently | (TC 10/16/2023 08:19:15 AM, Canjimavadivel, Kamallesh ) Replaced chest lead V1-V3. Also replaced chest lead V4-V6
63343 | 198469 | CM | CMPLT | CMPLT | 2023-10-16 08:21 | 2023-10-31 12:51 | V1-V3 leads damaged | (TC 10/16/2023 09:52:35 AM, Campbell, Connor) Replaced V1-V3 chest lead Conditioning battery (TC 10/17/2023 11:23:55 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete but SOH of battery is 60, below the acceptable threshold. Order for battery placed. Cart sent back to unit for time being but battery second battery needs to be installed (TC 10/31/2023 12:48:16 PM, Campbell, Connor) new batteries installed
| 198470 | PA | CMPLT | CMPLT | 2023-10-16 08:24 | 2023-11-17 13:25 | Ordering ECG chest leads | (TC 10/16/2023 08:31:55 AM, Canjimavadivel, Kamallesh ) Placed order for chest leads set (TC 11/17/2023 01:24:18 PM, Canjimavadivel, Kamallesh ) Received leads, labelled and out on shelf
66434 | 198471 | CM | CMPLT | CMPLT | 2023-10-16 09:02 | 2023-10-16 09:08 | No signal, will not turn on | (TC 10/16/2023 09:07:10 AM, Scero, Tomlyn) Replaced battery adaptor. Simulation ECG complete. Tests passed. Returned to CCU
1060001 | 198472 | CM | CMPLT | CMPLT | 2023-10-16 09:10 | 2023-10-16 09:24 | V3 not working | (TC 10/16/2023 09:22:46 AM, Campbell, Connor) Replaced V1-V3, RL, V7, V5R leads Replaced keyboard cover
1054522 | 198473 | CM | CMPLT | CMPLT | 2023-10-16 09:27 | 2023-10-16 09:52 | V6 not working | (TC 10/16/2023 09:50:57 AM, Campbell, Connor) Replaced V4-V6 lead
67148 | 198474 | CM | CMPLT | CMPLT | 2023-10-16 09:55 | 2023-10-16 10:18 | C13 error code | (TC 10/16/2023 10:07:00 AM, Canjimavadivel, Kamallesh ) Error code points to battery failure, replaced battery
1057056 | 198475 | CM | CMPLT | CMPLT | 2023-10-16 10:13 | 2023-12-28 12:42 | MSH BI RAD REVIEW ROOM SECUREVIEW #4 DR. MECHANT REPORTS MONITOR BRIGHTNESS IS ERRATIC SN# 9780222D1056 | (TC 10/16/2023 10:29:21 AM, Whittington, Scott) ref 01379122 (TC 10/16/2023 10:29:21 AM, Whittington, Scott) ref 01379122 SN# 9780222D1056 | MSH BI RAD REVIEW ROOM SECUREVIEW #4 DR. MECHANT REPORTS MONITOR BRIGHTNESS IS ERRATIC
66461 | 198476 | CM | CMPLT | CMPLT | 2023-10-16 11:37 | 2023-10-16 11:37 | Philips test | worked with Kamran
| 198477 | RDS | CMPLT | CMPLT | 2023-10-16 11:38 | 2023-10-16 11:38 | Weekly check for 14S and 14N | Weekly check for 14S and 14N
1068043 | 198478 | ACC | CMPLT | CMPLT | 2023-10-16 12:48 | 2023-10-16 14:52 | Incoming inspection | (TC 10/16/2023 02:51:06 PM, Campbell, Connor) Asset tagged and inspected
1068044 | 198479 | CM | CMPLT | CMPLT | 2023-10-16 13:04 | 2023-10-16 14:54 | Incoming inspection | (TC 10/16/2023 02:53:34 PM, Campbell, Connor) Asset tagged and inspected
1068045 | 198480 | CM | CMPLT | CMPLT | 2023-10-16 13:07 | 2023-10-16 14:54 | Incoming Inspection | (TC 10/16/2023 02:54:09 PM, Campbell, Connor) Asset tagged and inspected
1068046 | 198481 | CM | CMPLT | CMPLT | 2023-10-16 13:08 | 2023-10-16 14:55 | Incoming inspection | (TC 10/16/2023 02:54:46 PM, Campbell, Connor) Asset tagged and inspected
1068047 | 198482 | CM | CMPLT | CMPLT | 2023-10-16 13:09 | 2023-10-16 14:56 | Incoming inspection | (TC 10/16/2023 02:55:41 PM, Campbell, Connor) Asset tagged and inspected
1068048 | 198483 | CM | CMPLT | CMPLT | 2023-10-16 13:10 | 2023-10-16 14:56 | Incoming Inspection | (TC 10/16/2023 02:56:18 PM, Campbell, Connor) Asset tagged and inspected
1068049 | 198484 | CM | CMPLT | CMPLT | 2023-10-16 13:11 | 2023-10-16 14:58 | Incoming inspection | (TC 10/16/2023 02:56:59 PM, Campbell, Connor) Asset tagged and inspected
1068050 | 198485 | CM | CMPLT | CMPLT | 2023-10-16 13:12 | 2023-10-16 14:59 | Incoming inspection | (TC 10/16/2023 02:58:20 PM, Campbell, Connor) Asset tagged and inspected
1068051 | 198486 | CM | CMPLT | CMPLT | 2023-10-16 13:13 | 2023-10-16 14:59 | Incoming inspection | (TC 10/16/2023 02:58:52 PM, Campbell, Connor) Asset tagged and inspected
1068052 | 198487 | CM | CMPLT | CMPLT | 2023-10-16 13:14 | 2023-10-16 15:00 | Incoming inspection | (TC 10/16/2023 02:59:24 PM, Campbell, Connor) Asset tagged and inspected
1068053 | 198488 | CM | CMPLT | CMPLT | 2023-10-16 13:15 | 2023-10-16 15:00 | Incoming inspection | (TC 10/16/2023 03:00:01 PM, Campbell, Connor) Asset tagged and inspected
1068054 | 198489 | CM | CMPLT | CMPLT | 2023-10-16 13:16 | 2023-10-16 15:01 | Incoming Inspection | (TC 10/16/2023 03:00:43 PM, Campbell, Connor) Asset tagged and inspected
1068055 | 198490 | CM | CMPLT | CMPLT | 2023-10-16 13:17 | 2023-10-16 15:02 | Incoming inspection | (TC 10/16/2023 03:01:18 PM, Campbell, Connor) Asset tagged and inspected
1068056 | 198491 | CM | CMPLT | CMPLT | 2023-10-16 13:18 | 2023-10-16 15:02 | Incoming inspection | (TC 10/16/2023 03:01:58 PM, Campbell, Connor) Asset tagged and inspected
1068057 | 198492 | CM | CMPLT | CMPLT | 2023-10-16 13:19 | 2023-10-16 15:03 | Incoming inspection | (TC 10/16/2023 03:02:43 PM, Campbell, Connor) Asset tagged and inspected
1068058 | 198493 | CM | CMPLT | CMPLT | 2023-10-16 13:20 | 2023-10-16 15:04 | Incoming Inspection | (TC 10/16/2023 03:03:28 PM, Campbell, Connor) Asset tagged and inspected
1055793 | 198494 | CM | CMPLT | CMPLT | 2023-10-16 13:21 | 2023-10-16 13:24 | Software upgrade to 02.36.21 | (TC 10/16/2023 01:23:19 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055797 | 198495 | CM | CMPLT | CMPLT | 2023-10-16 13:21 | 2023-10-16 13:24 | Software upgrade to 02.36.21 | (TC 10/16/2023 01:23:41 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055806 | 198496 | CM | CMPLT | CMPLT | 2023-10-16 13:22 | 2023-10-16 13:25 | Software upgrade to 02.36.21 | (TC 10/13/2023 01:24:31 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055794 | 198497 | CM | CMPLT | CMPLT | 2023-10-16 13:22 | 2023-10-16 13:25 | Software upgrade to 02.36.21 | (TC 10/13/2023 01:24:55 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055796 | 198498 | CM | CMPLT | CMPLT | 2023-10-16 13:22 | 2023-10-16 13:26 | Software upgrade to 02.36.21 | (TC 10/13/2023 01:25:23 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055810 | 198499 | CM | CMPLT | CMPLT | 2023-10-16 13:23 | 2023-10-16 13:26 | Software upgrade to 02.36.21 | (TC 10/13/2023 01:25:55 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1068059 | 198500 | CM | CMPLT | CMPLT | 2023-10-16 13:29 | 2023-10-16 15:04 | Incoming inspection | (TC 10/16/2023 03:04:04 PM, Campbell, Connor) Asset tagged and inspected
1068060 | 198501 | CM | CMPLT | CMPLT | 2023-10-16 13:32 | 2023-10-16 15:05 | Incoming inspection | (TC 10/16/2023 03:04:39 PM, Campbell, Connor) Asset tagged and inspected
1068061 | 198502 | CM | CMPLT | CMPLT | 2023-10-16 13:34 | 2023-10-16 15:06 | Incoming Inspection | (TC 10/16/2023 03:05:14 PM, Campbell, Connor) Asset tagged and inspected
1068062 | 198503 | CM | CMPLT | CMPLT | 2023-10-16 13:35 | 2023-10-16 15:06 | Incoming Inspection | (TC 10/16/2023 03:06:13 PM, Campbell, Connor) Asset tagged and inspected
1068064 | 198504 | CM | CMPLT | CMPLT | 2023-10-16 13:37 | 2023-10-16 15:08 | Incoming Inspection | (TC 10/16/2023 03:08:00 PM, Campbell, Connor) Asset tagged and inspected
1068065 | 198505 | CM | CANCL | CANCL | 2023-10-16 13:38 |  | Incoming Inspection
1068063 | 198506 | CM | CMPLT | CMPLT | 2023-10-16 14:10 | 2023-10-16 15:07 | Incoming inspection | (TC 10/16/2023 03:06:50 PM, Campbell, Connor) Asset tagged and inspected
1056415 | 198507 | CM | CMPLT | CMPLT | 2023-10-16 14:39 | 2023-10-16 15:02 | Low battery error | (TC 10/16/2023 02:57:43 PM, Canjimavadivel, Kamallesh ) Replaced batteries and replaced damaged probe. Cleared error codes
| 198508 | CM | CMPLT | CMPLT | 2023-10-16 15:30 | 2023-12-28 12:42 | MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 | (TC 10/17/2023 06:37:38 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 APPLY FOR NETWORK ACCESS (TC 10/25/2023 02:09:26 PM, Whittington, Scott) INC29755669 (TC 10/25/2023 02:09:26 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INC29755669 (TC 11/06/2023 12:23:59 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INC29755669 SYSTEM ON SITE, NETWORKED AND DICOM CONNECTED (TC 11/07/2023 11:49:58 AM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INC29755669 SYSTEM OPERATING (TC 11/08/2023 01:43:40 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INC29755669 MOVED SYSTEM TO ROOM E2 BESIDE BPP WILL NOT CONNECT HAD MIRZA RE-WHITELIST ORIGINAL MAC GIVEN AND ONE THAT WAS DISPLAYED ON SYSTEM SAW IN SETUP THAT DNS WASNT AUTO, ALTERED SYSTEM ON NETWORK (TC 11/16/2023 03:33:54 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INVESTIGATE AND CONFIRM STUDIES ON JDMIPACS INVESTIGATE CINE SPEED PLAYBACK ON PACS IS REALTIME TARA MENTIONED THERE WAS A CASE IN ALBERTA ON AN AGFA IMPAX PACS SYSTEM THAT PLAYED BACK FAST (TC 11/21/2023 01:02:34 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INC29755669 SYSTEM MOVED TO MSH 15 SPP TABLET COMM ISSUES (TC 11/22/2023 02:43:03 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INC29755669 SYSTEM MOVED TO MSH 15 SPP T (TC 11/23/2023 04:00:59 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 (TC 11/23/2023 04:01:09 PM, Whittington, Scott) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 (TC 11/23/2023 04:01:18 PM, Dolinar, John) MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 | MSH OPG SPP SAMSUNG HERA W10 ULTRASOUND TRIAL NOV 7- NOV 23 INC29755669
1049902 | 198509 | CM | CMPLT | CMPLT | 2023-10-16 15:44 | 2023-10-17 13:21 | System batteries low | (TC 10/16/2023 03:43:52 PM, Fan, Tak) Ordered batteries. (TC 10/17/2023 01:20:48 PM, Chau, Ken) System batteries replaced.
1060518 | 198510 | CM | CMPLT | CMPLT | 2023-10-16 19:07 | 2023-10-16 19:06 | No image on touch screen in the exam room. | Ordering the control board that interfaces the touch screen.
| 198511 | RDS | CMPLT | CMPLT | 2023-10-17 07:43 | 2023-10-17 11:29 | Weekly rounds for ED
| 198512 | RDS | CMPLT | CMPLT | 2023-10-17 07:43 | 2023-10-19 12:51 | Weekly rounds for ICU
1054269 | 198513 | CM | CMPLT | CMPLT | 2023-10-17 07:48 | 2023-10-17 07:52 | Not working | (TC 10/17/2023 07:49:56 AM, Canjimavadivel, Kamallesh ) Replaced temp probe Cleared logs
1062097 | 198514 | CM | CMPLT | CMPLT | 2023-10-17 08:38 | 2023-10-23 08:42 | ordering PS500 batteries | received batteries. PO#7202302015
1055801 | 198515 | CM | CMPLT | CMPLT | 2023-10-17 10:29 | 2023-10-17 10:30 | Software upgrade to 02.36.21 | (TC 10/17/2023 10:29:17 AM, Scero, Tomlyn) software upgraded from 02.32.06.00 to 02.36.21.00. new config uploaded.
1055809 | 198516 | CM | CMPLT | CMPLT | 2023-10-17 10:55 | 2023-10-17 11:15 | Software upgrade to 02.36.21 | (TC 10/17/2023 11:14:51 AM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055808 | 198517 | CM | CMPLT | CMPLT | 2023-10-17 10:55 | 2023-10-17 11:16 | Software upgrade to 02.36.21 | (TC 10/17/2023 11:15:42 AM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055805 | 198518 | CM | CMPLT | CMPLT | 2023-10-17 10:55 | 2023-10-17 11:17 | Software upgrade to 02.36.21 | (TC 10/17/2023 11:16:15 AM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055804 | 198519 | CM | CMPLT | CMPLT | 2023-10-17 10:56 | 2023-10-17 11:17 | Software upgrade to 02.36.21 | (TC 10/17/2023 11:16:45 AM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055802 | 198520 | CM | CMPLT | CMPLT | 2023-10-17 10:56 | 2023-10-17 11:18 | Software upgrade to 02.36.21 | (TC 10/17/2023 11:17:18 AM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
68966 | 198521 | CM | CMPLT | NFF | 2023-10-17 11:09 | 2023-10-17 11:11 | not charging/broken power cord | (TC 10/17/2023 11:09:39 AM, Campbell, Connor) performed a check in including ESA passed check in and ESA bent cord and played around with it in different outlets and positions, no fault or disconnections obersved returned to service
68899 | 198522 | CM | CMPLT | CMPLT | 2023-10-17 11:17 | 2023-10-17 11:20 | not charging | (TC 10/17/2023 11:17:28 AM, Campbell, Connor) confirmed power cord would lose connection at certain angles replaced power cord performed a check in and ESA check in and ESA passed refer to ASM and procedure for ESA returned to service
68878 | 198523 | CM | CMPLT | CMPLT | 2023-10-17 11:59 | 2023-10-17 12:05 | came down with tag, no description of problem | (TC 10/17/2023 11:59:50 AM, Campbell, Connor) unit would not detect it was plugged in tried multiple plugs and wiggling the cord determined power cord was the issue replaced the power cord performed a check in, refer to ASM performed ESA, refer to procedure all tests passed returned to service
69036 | 198524 | CM | CMPLT | CMPLT | 2023-10-17 14:17 | 2023-10-17 14:18 | broken front case clamp | (TC 10/17/2023 02:16:40 PM, Campbell, Connor) replaced front clamp with a spare, no charge performed a check in function check passed returned to service
1055795 | 198525 | CM | CMPLT | CMPLT | 2023-10-17 14:32 | 2023-10-17 14:33 | Software upgrade to 02.36.21 | (TC 10/17/2023 02:32:28 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1055774 | 198526 | CM | CMPLT | CMPLT | 2023-10-17 14:38 | 2023-10-19 13:35 | Software upgrade to 02.36.21 | (TC 10/18/2023 01:32:41 PM, Canjimavadivel, Kamallesh ) Trying to transfer and upload log (TC 10/19/2023 01:34:20 PM, Canjimavadivel, Kamallesh ) Software upgraded from 2.32.06.00 to version 2.36.21.00. Uploaded new config
1049608 | 198527 | CM | CMPLT | CMPLT | 2023-10-17 18:40 | 2023-11-20 08:39 | MSH OPG SPP GE V10 3-531 MAC DROPPED | (TC 10/17/2023 06:39:49 PM, Whittington, Scott) MSH OPG SPP GE V10 3-531 MAC DROPPED | MSH OPG SPP GE V10 3-531 MAC DROPPED
| 198528 | CM | CMPLT | CMPLT | 2023-10-17 19:02 | 2023-09-08 20:03 | MSH BI NETWORK | (TC 09/08/2023 08:02:00 PM, Whittington, Scott) MSH BI NETWORK | MSH BI NETWORK
62623 | 198529 | CM | CMPLT | CMPLT | 2023-10-17 19:14 | 2023-10-26 12:42 | cracked hood and not powering on | (TC 10/17/2023 07:15:59 PM, Ferreira, Qwynn) Hood cracked from over tightening of the screw into the latch Will need new hood potentially Cable in the back was not seated correctly and caused the incubator to fault Reseated cables in the back Will confirm with Samir for new hood (TC 10/26/2023 12:39:39 PM, Rhuland, Courtney) Repaired small crack in hood with UV10 MED resin Replaced side door latch Ran functional check - 35C and 75% Read 35C and 74.6%
32510 | 198530 | CM | CMPLT | CMPLT | 2023-10-17 19:24 | 2023-09-15 20:24 | MSH DR1 TUBE HEAD REPAIRS | (TC 09/15/2023 07:23:39 PM, Whittington, Scott) MSH DR1 TUBE HEAD REPAIRS | MSH DR1 TUBE HEAD REPAIRS
53076 | 198531 | CM | CMPLT | CMPLT | 2023-10-17 19:51 | 2023-10-20 14:09 | X2 died on transport, blender face plate fell off, and transport caster wheels missing breaks | (TC 10/17/2023 07:51:47 PM, Ferreira, Qwynn) cycled the x2 battery and displayed green light X2 charging now, might have left transport unplugged or let battery drain too much Will wait for Samirs opinion on the blender face plate and for help with replacing wheels Functional Check - OK (TC 10/20/2023 02:07:18 PM, Chukri, Samir) Replaced the O2/Air Blender. will send the defective one for repair. Found one IAD leg with brake in the storage room 18floor switched and return the transport backe to service.
| 198532 | CM | CMPLT | CMPLT | 2023-10-18 08:27 | 2024-01-23 14:39 | Batteries for biomed stock | (TC 10/18/2023 08:48:13 AM, Rhuland, Courtney) Ordered 40 PCU batteries SC #1003079009
1066939 | 198533 | CM | CMPLT | CMPLT | 2023-10-18 08:42 | 2023-10-18 08:43 | Recorder printer not working | (TC 10/18/2023 08:41:45 AM, Canjimavadivel, Kamallesh ) Flipped printer paper roll and restarted central. Test print successful
1067052 | 198534 | PM | CMPLT | CMPLT | 2023-10-18 08:54 | 2023-10-18 09:04 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/17/2023 09:01:46 AM, Cortez, John) Bed+Intellidrive batteries were replaced, pendant clip missing, wrench LED lit up, error code 1041, verified that spacers used for transport were removed 60mohms, 21uA PM passed
1046787 | 198535 | PM | CMPLT | CMPLT | 2023-10-18 09:15 | 2023-10-18 09:22 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/17/2023 09:15:37 AM, Cortez, John) Bed+Intellidrive batteries were replaced 50 mOhms, 25 uAmps Pendant missing Foot section missing 4 screws on underside, will look to replace PM passed
1046452 | 198536 | CM | CMPLT | CMPLT | 2023-10-18 10:11 | 2023-11-14 10:06 | Keeps turning on and off when we try and use the cautery pen | (TC 10/18/2023 10:13:37 AM, Rendon Lopez , Elizabeth) Removed from OR 3, testing, confirmed problem, Monopolar 1 will not activate intermittently when pen moved. To be sent out for repair, requested loaner (TC 10/24/2023 10:57:11 AM, Rendon Lopez , Elizabeth) Received loaner 1068078. Sent out device for repair (TC 11/07/2023 11:35:30 AM, Rendon Lopez , Elizabeth) Received repair qupte from Conmed, device no longer under consignment. Approved repaired, added cost (TC 11/13/2023 10:05:46 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK. Returned to L&D OR 1
1054996 | 198537 | CM | CMPLT | CMPLT | 2023-10-18 10:14 | 2023-10-18 10:17 | Software update | (TC 10/18/2023 10:14:28 AM, Rhuland, Courtney) Temp reading was only abble to read in continuous mode Updated software version and predictive mode set Installed filter fan Functional check - OK
502089 | 198538 | RQ | CMPLT | CMPLT | 2023-10-18 11:11 | 2023-10-25 14:18 | Gantry not functioning; service request message on screen |  | Gantry not functioning; service request message on screen
1066475 | 198539 | CM | CMPLT | CMPLT | 2023-10-18 13:45 | 2023-10-30 13:42 | Manifold gasket not installed correctly | (TC 10/18/2023 01:45:18 PM, Reid, Emmerson ) PM was completed last week and they OR had immediate problems with the manifold for the 4L canister. Reached out to Styker for a fix
1057950 | 198540 | CM | CMPLT | CMPLT | 2023-10-18 13:47 | 2024-01-09 13:13 | Table brake foot has broken off | (TC 10/18/2023 01:47:18 PM, Reid, Emmerson ) Table had a wobble. After inspection, one of the brakes had the plastic foot broken off. I moved this table into OR12 and brought in another table. (TC 11/28/2023 01:39:59 PM, Reid, Emmerson ) Brought the table down and placed and order for the foot to be replaced. (TC 01/09/2024 01:11:44 PM, Reid, Emmerson ) tipped over the table and was able to replace the foot. Working now.
67240 | 198541 | CM | CMPLT | CMPLT | 2023-10-18 14:10 | 2023-10-23 12:36 | battery not charging | (TC 10/18/2023 02:10:00 PM, Campbell, Connor) determined the power cord was faulty swapped it with a cord from a different unit that is being sent back for repair power cord works, battery now charges performed a functional test as per the service manual, pump flow rate is within tolerance functional check complete, returned to service
| 198542 | PA | CMPLT | CMPLT | 2023-10-18 14:21 | 2023-11-01 08:51 | Accessories | (TC 10/18/2023 02:53:18 PM, Canjimavadivel, Kamallesh ) Placed order for accessories (TC 11/01/2023 08:50:19 AM, Canjimavadivel, Kamallesh ) Order didn't get filled and no longer need items, canceling WO
| 198543 | CM | CMPLT | CMPLT | 2023-10-19 10:04 | 2023-10-19 10:05 | unpacking and preparing power bars for samir | (TC 10/19/2023 10:04:12 AM, Campbell, Connor) unpacked power bars and mounting brackets applied locktite to screws and mounted the power bars to the brackets disposed of the packaging
| 198544 | PA | CMPLT | CMPLT | 2023-10-19 10:35 | 2023-10-30 13:20 | Ordering Distilled Water | (TC 10/23/2023 01:44:55 PM, Reid, Emmerson ) Order was cancelled because of backorder and had to creat a new shopping cart 1003083335 (TC 10/30/2023 01:19:12 PM, Reid, Emmerson ) Recieved
| 198545 | CM | CMPLT | CMPLT | 2023-10-19 11:27 | 2023-11-16 11:37 | Order Wheel assembly for transport incubator. | (TC 11/16/2023 11:35:19 AM, Chukri, Samir) Received wheels with wrong stem size emailed to company Tirrie and will send the correct stem size kits.
69164 | 198546 | CM | CMPLT | CMPLT | 2023-10-19 11:43 | 2023-10-23 12:37 | unspecified error message | (TC 10/19/2023 11:43:34 AM, Campbell, Connor) using analog sensor test, top pressure sensor was determined to be faulty front case latch was also broken due to wear on the front case plastic replaced the top pressure sensor replaced the front case performed a check-in check-in passed, pump is now functional returned to service
| 198547 | CM | CMPLT | CMPLT | 2023-10-19 14:07 | 2023-10-24 08:04 | Oreder Bins from Uline | (TC 10/19/2023 02:08:07 PM, Chukri, Samir) Ordered 20 bins for the biomed shelf . (TC 10/24/2023 08:04:18 AM, Chukri, Samir) Received
1061214 | 198548 | CM | CMPLT | CMPLT | 2023-10-19 14:10 | 2023-10-19 14:11 | Light drifting up | (TC 10/19/2023 02:10:28 PM, Reid, Emmerson ) Had to come down to get my tools. I adjusted the weight on the arm and balanced it. Working now
1061202 | 198549 | CM | ACTIV | ISSUE | 2023-10-19 14:12 |  | Bluetooth not working | (TC 10/19/2023 02:51:40 PM, Reid, Emmerson ) Tried restarting VMC-3 multiple times and going through settings. Also checked the bluetoother reciever and everything seemed fine but the bluetooth was not appearing.
| 198550 | CM | CMPLT | CMPLT | 2023-10-19 14:21 | 2023-10-18 14:20 | worked with Kamran | Worked with kamran on VN500 server connection
1062097 | 198551 | CM | CMPLT | CMPLT | 2023-10-19 14:22 | 2023-10-19 14:22 | screen turned off. ventilator continued working. unable to power down | Collected logs and sent them to Draeger. Their response: Based on what I see above I would recommend the following 1. Check the gas supply in the room at the wall for proper gas supply pressure. Typically most hospitals are 50 to 55 PSI. 2. Go into the Hit software and check the O2 and Air flow at 5L/min. and 120L/min. Make sure it is Ok at the higher flows. 3. It appears the Cockpit restart occurred once the unit was powered back on due to the O2 measurement issues earlier when it was powered off. 4. Run the device and breathing circuit checks and make sure they pass without issue. 5. If everything checks out then it can probably go back into service. Restarted unit 3 times. Passed pre-use check. Run unit 3 days. it is functional. Returned back to RT and informed Rex.
| 198552 | RDS | CMPLT | CMPLT | 2023-10-19 14:55 | 2023-10-19 14:57 | Weekly Checks For OR | (TC 10/19/2023 02:56:10 PM, Reid, Emmerson ) All checks complete
| 198553 | RDS | CMPLT | CMPLT | 2023-10-19 14:58 | 2023-10-19 15:00 | Weekly Checks For PACU | (TC 10/19/2023 02:59:59 PM, Reid, Emmerson ) All checks complete.
| 198554 | RDS | ACTIV | ISSUE | 2023-10-19 15:01 |  | Weekly Checks For PACU
1049819 | 198555 | CM | CMPLT | CMPLT | 2023-10-19 15:29 | 2023-10-19 15:30 | Missing power cord | (TC 10/19/2023 03:29:11 PM, Rhuland, Courtney) Checked vital signs monitor and power cord was missing Replaced with spare from biomed Functional check - OK
1059982 | 198556 | PM | CMPLT | CMPLT | 2023-10-19 15:47 | 2023-10-19 15:52 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/18/2023 03:50:17 PM, Cortez, John) Bed+Intellidrive batteries were replaced 60mohms, 25uA no pendant 1-68 deg head section 90.4kg scale verification reading PM passed
1055898 | 198557 | CM | CMPLT | CMPLT | 2023-10-19 15:58 | 2023-10-19 16:00 | "not working" | (TC 10/19/2023 03:57:51 PM, Cortez, John) BP hose connector (end that connects to Root BP port) faulty, replaced All parameters tested ok
1055902 | 198558 | CM | CMPLT | CMPLT | 2023-10-19 16:06 | 2023-10-19 16:29 | "won't turn on, temp probe not working" | (TC 10/19/2023 04:18:43 PM, Cortez, John) Root seems to turn on normally, Added Radical battery as per recall, Calibrated BP, Replaced temp probe, All parameters tested ok
67093 | 198559 | CM | CMPLT | CMPLT | 2023-10-19 16:37 | 2023-10-19 16:38 | Broken? (send down to biomed without tag) | (TC 10/19/2023 04:37:04 PM, Cortez, John) Replaced power cord (ground pin broken off) Couldnt find any other issue with monitor All parameters tested ok
32540 | 198560 | CM | CMPLT | CMPLT | 2023-10-20 09:12 | 2023-10-19 09:12 | XC Error | XC error -HV Overcurrent and Over mA -Troubleshooting pointed to x-ray tube, ordered parts -X-ray tube replaced and calibrations completed -System passed all QC tests. (TC 10/20/2023 03:49:29 PM, Debebe, Adam) -Prepared parts for shipment at shipping/receiving
1061202 | 198561 | CM | CMPLT | CMPLT | 2023-10-20 09:39 | 2023-10-20 09:40 | Installing IV Pole Power Bars | (TC 10/20/2023 09:39:12 AM, Reid, Emmerson ) Installed power bars in OR1-4, 6, 8-11, 14
1056456 | 198562 | CM | CMPLT | CMPLT | 2023-10-20 09:52 | 2023-10-20 09:52 | BP not working | updated SW and installed a battery on Radical 7. tested unit. functiional.
| 198563 | CM | CMPLT | CMPLT | 2023-10-20 09:54 | 2023-10-20 09:54 | worked at 14S | helped Dovina to find and/or order cables for masimo units.
| 198564 | CM | CMPLT | CMPLT | 2023-10-20 09:56 | 2023-10-20 09:56 | meeting | meeting
| 198565 | CM | CMPLT | CMPLT | 2023-10-20 12:18 | 2023-11-08 14:46 | Incubator Chassis order
69327 | 198566 | CM | CMPLT | CMPLT | 2023-10-20 12:18 | 2023-10-20 12:24 | Created by Failed PM WO # 194450 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/20/2023 12:18:57 PM, Campbell, Connor) replaced upper presure sensor PM complete/passed, refer to ASM | (TC 10/20/2023 12:17:58 PM, Campbell, Connor) PM failed pressure calibration
1059208 | 198567 | CM | CMPLT | CMPLT | 2023-10-20 12:33 | 2023-10-20 12:49 | Battery needs to be changed | (TC 10/20/2023 12:48:35 PM, Scero, Tomlyn) battery replaced and left to charge overnight
65416 | 198568 | CM | CMPLT | CMPLT | 2023-10-20 13:16 | 2023-11-23 08:31 | System failure 8 | (TC 10/20/2023 01:15:57 PM, Rhuland, Courtney) Replaced both relay and control board to clear system failure Noticed humidifier was loose and when removing humidifier, chassis broke Chassis ordered (TC 11/17/2023 02:54:53 PM, Rhuland, Courtney) Began chassis replacement (TC 11/22/2023 09:15:03 AM, Baghirzade, Nara) Helped with chassis replacement. (TC 11/22/2023 02:54:54 PM, Rhuland, Courtney) Completed chassis replacement Replaced humidifier gasket (TC 11/23/2023 08:30:30 AM, Rhuland, Courtney) Functional check - OK Temp - 35C Humiditiy - 75.4%
68755 | 198569 | CM | CMPLT | CMPLT | 2023-10-20 14:14 | 2023-10-23 11:49 | not connecting to library, error 600.6835 | (TC 10/20/2023 02:14:14 PM, Campbell, Connor) worked with john to diagnose, updates the network config and reseated the board library is now updated and working no more error message performed a check-in, passed performing a test infusion over the weekend to test functionality, will update when complete (TC 10/23/2023 11:47:59 AM, Campbell, Connor) functional infusion test complete PCU returned to service
69327 | 198570 | CM | CMPLT | CMPLT | 2023-10-20 14:19 | 2023-10-20 14:20 | error, not transfusing | (TC 10/20/2023 02:18:53 PM, Campbell, Connor) replaced pressure sensor during the previous PM before working on this issue, refer to WO 198566 issue resulted from the pressure sensor, after it was replaced and a functional check was performed, pump is now functional returned to service
65880 | 198571 | CM | CMPLT | NFF | 2023-10-20 14:29 | 2023-10-20 14:29 | does not turn on | (TC 10/20/2023 02:28:41 PM, Campbell, Connor) performed a PM first, refer to WO 193866 performed a functional test infusion no issue found returned to service
1067437 | 198572 | FC | CMPLT | CMPLT | 2023-10-23 08:01 | 2023-10-23 08:02 | QC checklist | (TC 10/23/2023 08:01:16 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
| 198573 | RDS | CMPLT | CMPLT | 2023-10-23 08:03 | 2023-10-23 08:03 | weekly check for 14S and 16N | completed
1059828 | 198574 | PM | CMPLT | CMPLT | 2023-10-23 08:11 | 2023-11-13 10:36 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/23/2023 08:11:09 AM, Campbell, Connor) Disinfected, cleaned, rinsed and refilled
1060974 | 198575 | CM | CMPLT | CMPLT | 2023-10-23 09:12 | 2023-10-23 09:13 | Missing power cord | (TC 10/23/2023 09:11:50 AM, Reid, Emmerson ) Was missing power cord. Brought up another one. Working now
66466 | 198576 | CM | CMPLT | CMPLT | 2023-10-23 09:25 | 2023-10-31 12:03 | ordering C500 parts for VN500 vent | ordering C500 parts for VN500 vent. received parts PO#7202305305
1062414 | 198577 | CM | ACTIV | ISSUE | 2023-10-23 09:28 |  | not working | (TC 10/23/2023 09:28:03 AM, Samanian, Kamran) ticket#101246598 was created to building service checked the bed and all functions were OK when the bed was plugged in. noticed battery was dead. recommended to keep the bed plugged in to charge the battery.
1066639 | 198578 | CM | CMPLT | NFF | 2023-10-23 09:37 | 2023-10-23 09:39 | error code 33 | (TC 10/23/2023 09:37:18 AM, Campbell, Connor) performed a functional check of the accuracy of flow rate ran for 45 minutes flow rate was accurate no errors found pump is functional, returned to service
1054543 | 198579 | CM | CMPLT | CMPLT | 2023-10-23 10:13 | 2023-10-23 10:14 | not reading ECG | (TC 10/23/2023 10:12:47 AM, Samanian, Kamran) checked the machine without ECG clips on the simulator and all looked OK. NFF. re-attached the new clips and ask the unit to test it to make sure machin works on the patient as well.
66680 | 198580 | CM | CMPLT | NFF | 2023-10-23 10:53 | 2023-10-23 10:54 | Reported humidifier not working | (TC 10/23/2023 10:52:59 AM, Baghirzade, Nara) Ran a functional to check the humidifer; humidifier is working. Set the humidifer to 80% RH at 36C and measured 81% on the analyzer. Could not find fault with the incubator.
1063746 | 198581 | PM | CMPLT | CMPLT | 2023-10-23 12:35 | 2023-10-23 12:36 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/23/2023 12:35:24 PM, Campbell, Connor) Disinfected, cleaned, rinsed and refilled
1068077 | 198582 | ACC | CMPLT | CMPLT | 2023-10-23 13:58 | 2023-10-23 14:36 | Incoming inspection | (TC 10/23/2023 02:34:30 PM, Canjimavadivel, Kamallesh ) Asset tagged and all tests pass
64961 | 198583 | CM | CANCL | CANCL | 2023-10-23 15:10 |  | Wrong WO generated
64961 | 198584 | PM | CMPLT | CMPLT | 2023-10-23 15:13 | 2023-10-23 09:16 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/23/2023 09:09:00 AM, Baghirzade, Nara) PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. NIBP- Simulated: 120/80 Results: 122/83 Simulated: 80/40 Results: 79/43 Simulated: 190/120 Results: 192/121 ECG- cable is in good condition. Simulated 70 BPM ECG Signal Resulted in 70 BPM, clear QRS complex. Simulated 95 BPM ECG Signal Resulted in 95 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 95% Simulated: 85% Results: 85% (triggered low SpO2 alarm) Battery is in good condition. Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Device operates when unplugged from AC, on battery. Good physical condition. Screen in good condition. All buttons working and responsive. Electrical Safety: Current leakage = 33.0 uA Power cord resistance = 220 mohms ECG leads leakage: LL = 01.6 uA LA = 01.5 uA RA = 01.6 uA Updated date and time.
64216 | 198585 | PM | CMPLT | CMPLT | 2023-10-24 07:25 | 2023-10-23 10:26 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/23/2023 10:24:00 AM, Baghirzade, Nara) PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. NIBP- Simulated: 120/80 Results: 120/83 Simulated: 80/40 Results: 81/42 Simulated: 190/120 Results: 191/118 ECG- cable is in good condition. Simulated 70 BPM ECG Signal Resulted in 70 BPM, clear QRS complex. Simulated 95 BPM ECG Signal Resulted in 95 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 96% Simulated: 85% Results: 86% (triggered low SpO2 alarm) Battery is in good condition. Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Device operates when unplugged from AC, on battery. Good physical condition. Screen in good condition. All buttons working and responsive. Electrical Safety: Current leakage = 20.0 uA Power cord resistance = 137 mohms ECG leads leakage: LL = 01.8 uA LA = 01.8 uA RA = 01.5 uA Updated date and time.
64960 | 198586 | PM | CMPLT | CMPLT | 2023-10-24 07:29 | 2023-10-23 11:31 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/23/2023 11:29:00 AM, Baghirzade, Nara) PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. NIBP- Simulated: 120/80 Results: 124/83 Simulated: 80/40 Results: 81/38 Simulated: 190/120 Results: 193/122 ECG- cable is in good condition. Simulated 70 BPM ECG Signal Resulted in 70 BPM, clear QRS complex. Simulated 95 BPM ECG Signal Resulted in 95 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 96% Simulated: 85% Results: 85% (triggered low SpO2 alarm) Battery is in good condition. Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Device operates when unplugged from AC, on battery. Good physical condition. Screen in good condition. All buttons working and responsive. Electrical Safety: Current leakage = 48.0 uA Power cord resistance = 251 mohms ECG leads leakage: LL = 01.8 uA LA = 01.7 uA RA = 01.3 uA Updated date and time.
65591 | 198587 | PM | CMPLT | CMPLT | 2023-10-24 07:35 | 2023-10-23 13:36 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/23/2023 01:34:00 PM, Baghirzade, Nara) PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. NIBP- Simulated: 120/80 Results: 120/81 Simulated: 80/40 Results: 79/42 Simulated: 190/120 Results: 194/120 ECG- cable is in good condition. Simulated 70 BPM ECG Signal Resulted in 70 BPM, clear QRS complex. Simulated 95 BPM ECG Signal Resulted in 95 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 95% Simulated: 85% Results: 86% (triggered low SpO2 alarm) Battery is in good condition. Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Device operates when unplugged from AC, on battery. Good physical condition. Screen in good condition. All buttons working and responsive. Electrical Safety: Current leakage = 34.0 uA Power cord resistance = 166 mohms ECG leads leakage: LL = 01.9 uA LA = 01.6 uA RA = 01.6 uA Updated date and time.
| 198588 | CM | CMPLT | CMPLT | 2023-10-24 07:38 | 2023-10-23 14:37 | Meeting with Gaby | (TC 10/23/2023 02:37:00 PM, Baghirzade, Nara) Met with Gaby to go over the response email from Medisoft. Following up on the issues with the bodybox and hypair.
1055527 | 198589 | CM | CMPLT | CMPLT | 2023-10-24 07:54 | 2023-10-24 07:56 | Reported that thermometer is consistently reading low | (TC 10/24/2023 07:54:26 AM, Baghirzade, Nara) Probe had damage on the connector side; replaced probe. Used calibration slug to verify 37C; thermometer read 37C. Functional test passed.
1051713 | 198590 | CM | CMPLT | CMPLT | 2023-10-24 08:51 | 2023-10-24 08:53 | Repurposing thermometer | (TC 10/24/2023 08:51:02 AM, Baghirzade, Nara) Repurposing thermometer for NICU in place of the old fasttemp filac (no longer funcitonal). Verified functionality of thermometer; used calibration slug and confirmed that thermometer reads 37C. Took multiple measurements successfully. Thermometer is fully functional. Reprinted asset tag and updated all details on TMS.
65592 | 198591 | PM | CMPLT | CMPLT | 2023-10-24 08:57 | 2023-10-24 08:58 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/24/2023 08:56:48 AM, Baghirzade, Nara) PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. NIBP- Simulated: 120/80 Results: 120/82 Simulated: 80/40 Results: 84/44 Simulated: 190/120 Results: 193/121 ECG- cable is in good condition. Simulated 70 BPM ECG Signal Resulted in 70 BPM, clear QRS complex. Simulated 95 BPM ECG Signal Resulted in 95 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 96% Simulated: 85% Results: 85% (triggered low SpO2 alarm) Battery is in good condition. Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Device operates when unplugged from AC, on battery. Good physical condition. Screen in good condition. All buttons working and responsive. Electrical Safety: Current leakage = 33.0 uA Power cord resistance = 158 mohms ECG leads leakage: LL = 01.7 uA LA = 01.6 uA RA = 01.6 uA Updated date and time.
| 198592 | CM | CMPLT | CMPLT | 2023-10-24 09:23 | 2023-12-01 08:20 | Ordering HL-90 board parts | (TC 10/24/2023 02:17:28 PM, Reid, Emmerson ) Ordered parts on plexxus (TC 10/30/2023 08:10:35 AM, Reid, Emmerson ) Recieved the LEDs and the Piezos (TC 12/01/2023 08:15:25 AM, Reid, Emmerson ) All Recieved.
| 198593 | TEDU | CMPLT | CMPLT | 2023-10-24 09:31 | 2023-09-25 09:31 | EMG Training | EMG Training
| 198594 | TEDU | CMPLT | CMPLT | 2023-10-24 09:33 | 2023-10-24 09:34 | EEG Training | EEG Training.
500869 | 198595 | CM | CMPLT | CMPLT | 2023-10-24 11:44 | 2023-11-03 14:27 | MSH 5 US ROOM 6. APLIO XG RETURNED TO CLINICAL USE IN ROOM 6 AS i800 PRISM INSTALLED IN ER G7 | (TC 10/24/2023 11:44:12 AM, Whittington, Scott) MSH 5 US ROOM 6. APLIO XG RETURNED TO CLINICAL USE IN ROOM 6 AS i800 PRISM INSTALLED IN ER G7 | MSH 5 US ROOM 6. APLIO XG RETURNED TO CLINICAL USE IN ROOM 6 AS i800 PRISM INSTALLED IN ER G7
67361 | 198596 | PM | CMPLT | CMPLT | 2023-10-24 12:23 | 2023-10-24 13:57 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/24/2023 12:24:29 PM, Campbell, Connor) was waiting for a front case for a long time, so performed a PM to test all functionality adjusted LCD, over temp alarm and circulating temperature pots PM passed/complete, refer to procedure
1061964 | 198597 | CM | CMPLT | CMPLT | 2023-10-24 13:29 | 2023-10-24 13:30 | Left Side controls not working | (TC 10/24/2023 01:28:51 PM, Ferreira, Qwynn) Received a call for a bed repair. Noticed the left side was not working Checked pins and noticed they were dirty - cleaned Reset the bed and powered off/on All controls working as normal
| 198598 | CM | CMPLT | CMPLT | 2023-10-24 13:47 | 2023-10-24 13:49 | BP issues in ER Y44 | (TC 10/23/2023 01:48:13 PM, Cortez, John) Helped troubleshoot with nurse, BP hose seems to be at fault, replaced
1059778 | 198599 | FC | CMPLT | CMPLT | 2023-10-24 14:00 | 2023-10-24 14:01 | QC checklist | (TC 10/24/2023 02:00:35 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
| 198600 | CM | CMPLT | CMPLT | 2023-10-24 14:01 | 2023-10-24 14:01 | Meeting with Adeel | (TC 10/24/2023 02:01:10 PM, Baghirzade, Nara) Meeting with Adeel
| 198601 | CM | CMPLT | CMPLT | 2023-10-24 14:02 | 2023-10-24 14:05 | Meeting with PFT/Gaby | (TC 10/24/2023 02:01:51 PM, Baghirzade, Nara) Meeting regarding adding upper and lower limits to CPET reporting. Installing the review software on the receptionist's hospital laptop. Screenshots of DLCO/TLCO on Hypair. Went through the errors in linearity for DLCO. Any other issues that have occurred.
| 198602 | CM | CMPLT | CMPLT | 2023-10-24 14:06 | 2023-10-24 14:06 | Biomed huddle | (TC 10/24/2023 02:05:44 PM, Baghirzade, Nara) Biomed huddle
| 198603 | RDS | CMPLT | CMPLT | 2023-10-24 14:25 | 2023-10-24 14:27 | Weekly Checks For PACU | (TC 10/24/2023 02:26:40 PM, Reid, Emmerson ) Rounds completed
1049901 | 198604 | CM | CMPLT | CMPLT | 2023-10-24 14:38 | 2023-11-20 08:37 | MSH R61 REVOLUTION GENERIC SMALL DETECTOR SMALL BATTERY ORDER ON PO 7500153644 | (TC 10/24/2023 02:37:49 PM, Whittington, Scott) MSH R61 REVOLUTION GENERIC SMALL DETECTOR SMALL BATTERY ORDER ON PO 7500153644 (TC 11/20/2023 08:34:46 AM, Whittington, Scott) (TC 10/24/2023 02:37:49 PM, Whittington, Scott) MSH R61 REVOLUTION GENERIC SMALL DETECTOR SMALL BATTERY ORDER ON PO 7500153644 QTY 2 BATTERIES PN# SPAF1767 RECEIVED OCT 30 | MSH R61 REVOLUTION GENERIC SMALL DETECTOR SMALL BATTERY ORDER ON PO 7500153644
67355 | 198605 | PM | CMPLT | CMPLT | 2023-10-24 14:42 | 2023-10-24 15:14 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/24/2023 03:13:34 PM, Campbell, Connor) Adjusted the LCD, overtemp alarm and circulation temperature pots Pm complete/passed, refer to procedure
1067436 | 198606 | FC | CMPLT | CMPLT | 2023-10-24 15:45 | 2023-10-24 15:45 | QC check | (TC 10/24/2023 03:45:05 PM, Cortez, John) QC check passed
1055794 | 198607 | CM | CMPLT | CMPLT | 2023-10-24 16:16 | 2023-10-24 16:17 | Potential battery issue | (TC 10/24/2023 04:16:04 PM, Cortez, John) Dashboard states that unit is not "ready", Battery status unknown, swapped and calibrated battery
1047847 | 198608 | CM | CMPLT | CMPLT | 2023-10-25 07:51 | 2024-01-30 08:55 | Ordering parts for repair and stock | (TC 10/25/2023 07:55:21 AM, Baghirzade, Nara) SC 1003085074 (TC 01/26/2024 09:07:46 AM, Baghirzade, Nara) Received 1/2 items ordered: PO 7202306853 (TC 01/26/2024 09:07:46 AM, Baghirzade, Nara) PO 7202306853 (TC 01/30/2024 08:54:58 AM, Baghirzade, Nara) PO 7202306853
1046873 | 198609 | CM | CMPLT | CMPLT | 2023-10-25 08:00 | 2023-10-24 08:00 | alarming | tested unit. not functional. removed from service. informed Rex - RT.
| 198610 | CM | CMPLT | CMPLT | 2023-10-25 08:03 | 2023-10-24 08:02 | biomed meeting | biomed meeting
| 198611 | CM | CMPLT | CMPLT | 2023-10-25 08:04 | 2023-10-24 08:04 | worked at 16N | worked at 16N(14N)
| 198612 | RDS | CMPLT | CMPLT | 2023-10-25 08:27 | 2023-10-25 09:21 | Weekly check for ED
62244 | 198613 | CM | CMPLT | CMPLT | 2023-10-25 08:28 | 2023-10-25 08:29 | Repurposed thermometer | (TC 10/25/2023 08:27:53 AM, Baghirzade, Nara) Repurposed thermometer for NICU in place of the old fasttemp filac. Used calibration slug to verify that the thermometer reads 37C. Confirmed that it reads 37C. Reprinted asset tag and updated details on TMS. Thermometer functional tests passed, device is fully functional.
1051697 | 198614 | CM | CMPLT | CMPLT | 2023-10-25 08:30 | 2023-10-25 08:31 | Repurposed for NICU | (TC 10/25/2023 08:30:20 AM, Baghirzade, Nara) Repurposed thermometer for NICU in place of the old fasttemp filac. Used calibration slug to verify that the thermometer reads 37C. Confirmed that it reads 37C. Reprinted asset tag and updated details on TMS. Thermometer functional tests passed, device is fully functional.
| 198615 | RDS | CMPLT | CMPLT | 2023-10-25 08:32 | 2023-11-08 14:11 | Weekly check for ICU
1051702 | 198616 | CM | CMPLT | CMPLT | 2023-10-25 08:33 | 2023-10-25 08:33 | Repurposed for NICU | (TC 10/25/2023 08:33:08 AM, Baghirzade, Nara) Repurposed thermometer for NICU in place of the old fasttemp filac. Used calibration slug to verify that the thermometer reads 37C. Confirmed that it reads 37C. Reprinted asset tag and updated details on TMS. Thermometer functional tests passed, device is fully functional.
65291 | 198617 | CM | CMPLT | CMPLT | 2023-10-25 08:35 | 2023-10-25 08:37 | overheating and pump stops | (TC 10/25/2023 08:35:02 AM, Campbell, Connor) adjusted LCD, overtemp alarm and circulation temperature as per the manual Ran pump for a few hours after, pump is functional and not overheating or shutting off Pump front cover was also leaking, replaced the water chamber gasket and water chamber cover, pump is not leaking now will perform PM to test the rest of the functionality
65291 | 198618 | PM | CMPLT | CMPLT | 2023-10-25 08:38 | 2023-10-25 08:39 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/25/2023 08:38:38 AM, Campbell, Connor) PM complete/passed, refer to procedure
66434 | 198619 | CM | CMPLT | NFF | 2023-10-25 08:43 | 2023-10-25 09:05 | Nurse said it wasn't working intermittently | (TC 10/24/2023 09:03:48 AM, Scero, Tomlyn) Tested telepak in biomed and up in PACU pod A. Working fine, no loss of signal where it should be working
68372 | 198620 | CM | CMPLT | CMPLT | 2023-10-25 09:09 | 2023-10-31 13:00 | Vol agent gas sensor failure | (TC 10/27/2023 10:33:23 AM, Scero, Tomlyn) Observed AGM failure during case. SEVO was being delivered but no MAC % or SEV % was being displayed. Vol agent sensor failure alarm displayed. I first tried to use another SEVO module, no change. The gas machine was swapped and brought out of the OR. I was able to download the DSR and review the error codes. Found that the error was coming from a failed ILCA gas sensor in the PGA. Contacted Draeger for a quote and ordered the part. emailed peter to follow up. (TC 10/30/2023 12:46:16 PM, Scero, Tomlyn) I spoke with Salur in person about the issue. He suggested that I do the IPM-L for the PGA and check for leaks in the PATO. It's very rare for the ILCA2 to fail he said. (TC 10/31/2023 12:55:13 PM, Scero, Tomlyn) Turned the entire device off including the reset button, resetting everything and closing all valves. After restarting the self test passed and let unit run for 15 mins on man/spont to check the O2 values. O2 values passed. PGA was tested as per IPM-L, all tests passed and no leaks were found in the PATO. Allowed the AGM to run in Pressure Control mode for 1 hour with 2% SEVO and values were monitored. All values normal as expected and no errors or alarms were present. Vol. agent sensor alarm no longer able to replicate. Possible stuck valve or intermittent issues. Device can be used , but if error is present again the sensor will need to be changed.
1050733 | 198621 | CM | CMPLT | CMPLT | 2023-10-25 09:19 | 2023-10-25 09:20 | Smoke Evac working on continuous mode only | (TC 10/11/2023 09:19:05 AM, Rendon Lopez , Elizabeth) Found activation cable not fully plugged in. Tested several times-OK
1050729 | 198622 | CM | CMPLT | CMPLT | 2023-10-25 09:20 | 2023-10-25 09:21 | Malfunctioning activation cable | (TC 10/11/2023 09:20:28 AM, Rendon Lopez , Elizabeth) Replaced cable with spare from unit. Tested-OK
66730 | 198623 | CM | CMPLT | CMPLT | 2023-10-25 09:22 | 2023-10-25 10:36 | Call from IT to check connection to fetalink | (TC 10/16/2023 09:21:59 AM, Rendon Lopez , Elizabeth) Checked fetal monitor, not connecting, restarted CCE, changed USB plugs, unable to connect. Contacted IT to resolve issue
1059746 | 198624 | CM | CMPLT | CMPLT | 2023-10-25 10:39 | 2023-10-25 10:40 | Turning on by itself | (TC 10/16/2023 10:39:19 AM, Rendon Lopez , Elizabeth) Tested, NFF
1051390 | 198625 | CM | CMPLT | NFF | 2023-10-25 10:49 | 2023-10-25 10:49 | Not picking up contractions | (TC 10/17/2023 10:48:40 AM, Rendon Lopez , Elizabeth) Tested-OK
1054158 | 198626 | CM | CMPLT | NFF | 2023-10-25 10:49 | 2023-10-25 10:50 | Not picking up FHR | (TC 10/17/2023 10:49:28 AM, Rendon Lopez , Elizabeth) Tested-OK
1060134 | 198627 | CM | CMPLT | NFF | 2023-10-25 10:50 | 2023-10-25 10:51 | Not picking up FHR | (TC 10/17/2023 10:50:15 AM, Rendon Lopez , Elizabeth) Tested-OK
1059202 | 198628 | CM | CMPLT | CMPLT | 2023-10-25 10:51 | 2023-10-25 10:53 | Battery malfunction alarm | (TC 10/20/2023 10:51:52 AM, Rendon Lopez , Elizabeth) Device in L&D OR 1, alarm battery malfunction during case. Replaced battery, Functional tests completed, working properly.
1068078 | 198629 | ACC | CMPLT | CMPLT | 2023-10-25 10:55 | 2023-10-25 10:55 | Loaner incoming Inspection | (TC 10/24/2023 10:54:49 AM, Rendon Lopez , Elizabeth) Loaner incoming inspection, all tests passed. Placed loaner in OR 1
1046452 | 198630 | CM | CANCL | CANCL | 2023-10-25 10:56 |  | Cautery will turn on and off | Duplicate
1062419 | 198631 | CM | CMPLT | CMPLT | 2023-10-25 10:58 | 2023-10-25 10:59 | Bedside panel not working | (TC 10/24/2023 10:58:15 AM, Rendon Lopez , Elizabeth) Reset bed battery switch, erased error logs, tested all panels and functions, all working
67234 | 198632 | CM | CMPLT | CMPLT | 2023-10-25 10:59 | 2023-10-30 10:18 | Turns on but screen blank | (TC 10/24/2023 10:59:32 AM, Rendon Lopez , Elizabeth) Confirmed issue, removed from L&D, found malfunctioning Control board (TC 10/26/2023 10:16:46 AM, Rendon Lopez , Elizabeth) Replaced control board, cleaned dust. Testing (TC 10/27/2023 10:16:46 AM, Rendon Lopez , Elizabeth) Warmer working properly, all functions okay. Replaced blender, tested Resus sytem-OK. Returned to L&D as spare
1050474 | 198633 | CM | CMPLT | CMPLT | 2023-10-25 11:09 | 2023-10-25 11:11 | Replacement for IVF printer | (TC 10/25/2023 11:09:17 AM, Rhuland, Courtney) Tested functionality of printer in biomed to assure it would be compatible with IVF setup Brought printer to IVF t oreplace out of support, worn printer Functional check - OK
1054963 | 198634 | CM | CMPLT | CMPLT | 2023-10-25 11:28 | 2023-10-25 11:31 | "touch screen not accurate, not recording BP or temp" | (TC 10/25/2023 11:28:27 AM, Cortez, John) -Touch screen seems to be ok, unable to replicate issue -Replaced temp probe well -Upgraded software to 2.1.4.9 -BP a bit off, calibrated
32540 | 198635 | CM | CMPLT | CMPLT | 2023-10-25 13:02 | 2023-10-24 13:00 | Recon errors reported during clinical case | Recon errors reported during clinical case Service completed troubleshooting Service replaced 2 MHR-BP recon boards CT phantom tested OK
66042 | 198636 | CM | CMPLT | CMPLT | 2023-10-25 13:38 | 2023-10-25 13:38 | Broken cases | (TC 10/25/2023 01:37:44 PM, Rhuland, Courtney) Replaced cracked front and rear case (recall) Performed PM All tests passed Refer to ASM for details
1059778 | 198637 | FC | CMPLT | CMPLT | 2023-10-25 14:22 | 2023-10-25 14:45 | QC checklist | (TC 10/25/2023 02:44:44 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
502089 | 198638 | RQ | CMPLT | CMPLT | 2023-10-25 15:25 | 2023-10-27 14:33 | Symbia 2 persistent error | (TC 10/27/2023 02:32:01 PM, Green, Paul) Walked cusotmer through this problem and was able to get them going . | gantry can't be moved, message to call for service
502089 | 198639 | RQ | CMPLT | CMPLT | 2023-10-25 15:27 | 2023-10-27 14:32 | CT communication lost | (TC 10/26/2023 02:31:15 PM, Green, Paul) Had cusotmer power down and came back up ok. | error attempting SPECT CT acquisition
501974 | 198640 | RQ | CMPLT | CMPLT | 2023-10-25 15:30 | 2023-10-27 14:22 | probe and well qc error | (TC 10/27/2023 02:21:04 PM, Green, Paul) HV cable was coiled up as well as sealed sources were mixed up. Ran cal, ok. | probe & well qc failure
| 198641 | CM | CMPLT | CMPLT | 2023-10-26 08:36 | 2023-11-01 14:55 | Spare HL-90 parts | (TC 10/25/2023 08:47:42 AM, Rendon Lopez , Elizabeth) Spare fluid warmer parts. Received
| 198642 | PA | CMPLT | CMPLT | 2023-10-26 08:45 | 2023-11-20 09:20 | Warranty replacement SPO2 probes Masimo | (TC 10/26/2023 09:06:38 AM, Scero, Tomlyn) All serial numbers taken down and customer feedback form filled out online. Masimo informed which were still under warranty and requested a No charge PO before I send them in. PO requested on ereq. (TC 11/20/2023 09:19:05 AM, Scero, Tomlyn) Sensors and calbes replaced. Delivered to RT OR.
64868 | 198643 | PM | CMPLT | CMPLT | 2023-10-26 08:45 | 2023-10-26 08:47 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/26/2023 08:45:41 AM, Campbell, Connor) PM complete, refer to procedure
64868 | 198644 | CM | CMPLT | CMPLT | 2023-10-26 08:49 | 2023-10-26 08:50 | pump not working and overheating | (TC 10/26/2023 08:48:48 AM, Campbell, Connor) adjusted LCD, overheating alarm and circulation temperature pots ran unit for a few hours unit is functional
68224 | 198645 | CM | CMPLT | CMPLT | 2023-10-26 11:46 | 2023-10-26 13:08 | put into service with new battery | (TC 10/26/2023 01:05:24 PM, Scero, Tomlyn) unit taken from storoage in biomed to use in the CCU. X2 given to the unit with a new battery.
67888 | 198646 | CM | CMPLT | CMPLT | 2023-10-26 11:46 | 2024-01-19 09:52 | Stopped working multiple times | (TC 10/25/2023 11:48:07 AM, Rendon Lopez , Elizabeth) Pump alarming error code 41622. To be sent out for repair. Requested RMA. (TC 10/26/2023 11:48:07 AM, Rendon Lopez , Elizabeth) Received RMA# 317085, created e req (TC 01/18/2024 09:49:21 AM, Rendon Lopez , Elizabeth) incoming inspection-OK, PM completed, all tests passed. See attached report
67878 | 198647 | CM | CMPLT | CMPLT | 2023-10-26 11:50 | 2024-01-19 09:39 | PCEA button not working | (TC 10/25/2023 11:50:10 AM, Rendon Lopez , Elizabeth) PCEA cord attached not being detected, defective connector. Requested RMA to send out for repair (TC 10/26/2023 11:50:10 AM, Rendon Lopez , Elizabeth) Received RMA# 317085, created e req (TC 01/18/2024 09:38:48 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK. See attached report
53118 | 198648 | CM | CMPLT | NFF | 2023-10-26 12:37 | 2023-10-26 12:39 | Broken tray | (TC 10/26/2023 12:37:15 PM, Rhuland, Courtney) Printer paper tray was damaged Replaced with spare from retired printer Ink ribbon was depleted, replaced when testing printer, paper was unable to feed through Endo had ordered wrong ribbon size, had to use smaller paper Ran test print and no issues
1068092 | 198649 | CM | CMPLT | CMPLT | 2023-10-26 12:46 | 2023-11-06 14:57 | Not working | (TC 10/26/2023 12:46:14 PM, Rendon Lopez , Elizabeth) Confirmed issue. Added to TMS, sending out for repair (TC 11/06/2023 02:56:17 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1068093 | 198650 | CM | CMPLT | CMPLT | 2023-10-26 12:51 | 2023-11-06 14:53 | Not picking up FHR | (TC 10/26/2023 12:50:53 PM, Rendon Lopez , Elizabeth) Confirmed issue, Added to TMS, sending out for repair (TC 11/06/2023 02:52:42 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1055669 | 198651 | CM | CMPLT | CMPLT | 2023-10-26 12:54 | 2023-11-06 14:46 | not working | (TC 10/26/2023 02:30:00 PM, Rendon Lopez , Elizabeth) Confirmed issue, sending out to repair under warranty (TC 11/06/2023 02:45:58 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1062779 | 198652 | CM | CMPLT | CMPLT | 2023-10-26 12:55 | 2023-11-15 10:14 | Adding upper and lower limits to CPET reporting | (TC 10/26/2023 12:55:22 PM, Baghirzade, Nara) Respirology doctors requested to have upper and lower limits on the third page of the CPET reporting. Software limitations did not allow me to add the columns in, had to recreate a new report. Added upper and lower limits for the following parameters: -FVC (L) -FEV1(L) -FEV1/FVC(%) -PEF(L/S) -FEF25-75(L/S) -FEF50(L/S) -FEF75(L/S) -SVC(L) -IC(L) -ERV(L) Added the following parameter measurements to the table: -Ref.Val -Observed mean value -%Ref. Val pre -Observed post mean value -%Ref. Val post -Observed post mean value -Pre. dif% -Observed post mean value -Pre.dif% (post) Adjusted decimal points for all parameter values except for FEV1/FVC. Requested meeting with Gaby to go over the report before overwriting the original report. (TC 10/26/2023 03:10:00 PM, Baghirzade, Nara) Met with Gaby to go over the additions the CPET report. Gaby requested the following changes: - 2 decimal points for mean predicted val % for flow @ 50%VC and flow @ 25%VC; -2 decimal points for pre. observed values for FVC, FEV1, flow @ 50%VC and flow @ 25%V. -lower and upper limits to remain at 1 decimal point -FEV1/FVC ratio % to be at 0 decimal points. Tried to run the report on version 1.35 but the values would not populate on the table. Tried on local and network mode but none worked. Will recreate the report on version 1.35, and will keep the version 1.34 report as backup. (TC 11/14/2023 10:31:42 AM, Baghirzade, Nara) Re-made the report with upper and lower limits on the CPET station. Waiting for Gaby's approval to publish the report. (TC 11/15/2023 10:13:41 AM, Baghirzade, Nara) Gaby approved the CPET report. Requested same changes to be made to the CPET long report. Added the following parameter measurements to the table: -Ref.Val -Observed mean value -%Ref. Val pre -Observed post mean value -%Ref. Val post -Observed post mean value -Pre. dif% -Observed post mean value -Pre.dif% (post) Applied the changes to the CPET long report and made it available on the CPET station.
| 198653 | CM | CMPLT | CMPLT | 2023-10-26 13:35 | 2023-12-18 13:26 | Cleaning up the lasers information on TMS | (TC 10/26/2023 01:35:30 PM, Baghirzade, Nara) Looked through the inventory of lasers on TMS; noted a few discrepancies. Went with Jessica to the OR to locate and confirm that the lasers are there and their status. Was able to find 4/6 lasers we were looking for. An ultrapulse encore (loaner 1049012) laser is still active on TMS; followed up with Tomi to see if she had returned it so we can change the status of the asset to retired. Jessica to confirm about a CO2 laser that we werent able to locate (62125).
65419 | 198654 | CM | CMPLT | CMPLT | 2023-10-26 13:49 | 2023-10-26 13:50 | overheating | (TC 10/26/2023 01:49:38 PM, Campbell, Connor) adjusted LCD, over temp alarm and circulation temperature unit is not overheating now, is functional returned to service
65419 | 198655 | PM | CMPLT | CMPLT | 2023-10-26 13:52 | 2023-10-26 13:54 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 10/26/2023 01:51:46 PM, Campbell, Connor) Replaced water chamber front case, was cracked and leaking PM complete/passed, refer to procedure
65955 | 198656 | CM | CMPLT | CMPLT | 2023-10-26 14:54 | 2023-10-26 14:56 | Plunger head not moving | (TC 10/26/2023 02:54:04 PM, Rhuland, Courtney) Opened device and reseated leadscrew Replaced rear case PM done to assure functionality All tests passed Refer to ASM for details
1062171 | 198657 | CM | CMPLT | CMPLT | 2023-10-27 07:20 | 2023-11-06 08:19 | Auxiliary outlets on the bed are not working | (TC 10/26/2023 02:20:00 PM, Baghirzade, Nara) Iliana reported that the aux outlets are not working after they've tried to plug equipment to it. Used DMM to verify that there is no power to both outlets. Plugged the power cord to other outlets on the wall but still not working. Measured wall outlets and wall outlets are functional. There is a patient on the bed so once patient is moved, I will request a service visit from Stryker. Iliana to let me know when bed is empty and moved to the hallway. (TC 11/02/2023 02:18:16 PM, Baghirzade, Nara) Tested the bed again after it has been removed from room 13 on 18 South. All motion controls are functional but auxiliary outlets are not working. Auxiliary power cord is in good condition. Initiated a service request with Stryker. (TC 11/06/2023 08:17:25 AM, Baghirzade, Nara) Attached service report to this WO. Bed is fully functional. Informed 18 South.
| 198658 | CM | CMPLT | CMPLT | 2023-10-27 07:22 | 2023-10-27 07:23 | NICU huddle | (TC 10/26/2023 03:40:00 PM, Baghirzade, Nara) NICU huddle
1067591 | 198659 | CM | CMPLT | CMPLT | 2023-10-27 07:24 | 2023-10-27 07:25 | Moved new NICU warmer to storage | (TC 10/26/2023 02:00:00 PM, Baghirzade, Nara) John Balcom requested that we move the new warmer that is stored in biomed to the NICU 18th fl. storage. Moved the warmer to the storage area and sent John and May emails confirming that the warmer is now in storage.
| 198660 | CM | CMPLT | CMPLT | 2023-10-27 09:10 | 2023-11-23 16:18 | Panda warmer parts and repairs | (TC 10/27/2023 09:14:49 AM, Rendon Lopez , Elizabeth) Ordering panda warmer spare control board, and blender overhaul (TC 10/30/2023 09:47:45 AM, Rendon Lopez , Elizabeth) Shipped blenders for overhaul (TC 11/08/2023 01:34:58 PM, Rendon Lopez , Elizabeth) Received blenders (TC 11/23/2023 04:17:55 PM, Rendon Lopez , Elizabeth) Received control board
1054417 | 198661 | CM | CMPLT | CMPLT | 2023-10-27 09:16 | 2023-11-06 14:36 | Physical damage | (TC 10/26/2023 09:16:13 AM, Rendon Lopez , Elizabeth) Confirm issue, sending out for repair (TC 11/06/2023 02:35:50 PM, Rendon Lopez , Elizabeth) Broken beyond repair. Retired from service
1054416 | 198662 | CM | CMPLT | CMPLT | 2023-10-27 09:17 | 2023-11-09 09:59 | Not detecting contractions | (TC 10/26/2023 09:17:17 AM, Rendon Lopez , Elizabeth) Confirmed issue, no readings in fetal monitor. Sending out for repair (TC 11/09/2023 09:51:35 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
| 198663 | RDS | CMPLT | CMPLT | 2023-10-27 09:57 | 2023-10-27 10:21 | Weekly checks in CCU | (TC 10/27/2023 10:19:24 AM, Scero, Tomlyn) Weekly checks complete. 30J test done on all defibs
64494 | 198664 | PM | CMPLT | CMPLT | 2023-10-27 10:12 | 2023-10-27 10:14 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 10/27/2023 10:12:28 AM, Baghirzade, Nara) Replaced battery. All functional tests passed on the aspirator. (TC 10/27/2023 10:13:01 AM, Rhuland, Courtney) Replaced battery. All functional tests passed on the aspirator.
62191 | 198665 | PM | CMPLT | CMPLT | 2023-10-27 10:14 | 2023-10-27 10:15 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 10/27/2023 10:14:29 AM, Baghirzade, Nara) PM passed and completed. (TC 10/27/2023 10:14:44 AM, Rhuland, Courtney) PM passed and completed.
| 198666 | CM | CMPLT | CMPLT | 2023-10-27 10:15 | 2023-10-27 10:20 | Incubator/Omnibed exam lights | (TC 10/27/2023 10:15:38 AM, Baghirzade, Nara) Picked up 5 lights from May's office. Unboxed all the lights. Replaced 1 melted light in RESUS with 1 new one, 1 light was height-adjusted, 1 old light was replaced with a new one, and 1 omnibed was missing a light entirely. Removed the adjustment knobs and replaced them with screws. 2 new lights left. Courtney will install them on incubators in NICU rooms. (TC 10/27/2023 10:19:15 AM, Rhuland, Courtney) Picked up 5 lights from May's office. Unboxed all the lights. Replaced 1 melted light in RESUS with 1 new one, 1 light was height-adjusted, 1 old light was replaced with a new one, and 1 omnibed was missing a light entirely. Removed the adjustment knobs and replaced them with screws.
| 198667 | RDS | CMPLT | CMPLT | 2023-10-27 10:20 | 2023-10-27 10:22 | Weekly Rounds: 18 South and PACU POD A | (TC 10/27/2023 10:20:17 AM, Baghirzade, Nara) Weekly rounds in 18 South and PACU POD A completed. See procedure.
| 198668 | RDS | CMPLT | CMPLT | 2023-10-27 10:23 | 2023-10-27 10:24 | Weekly Rounds: NICU | (TC 10/27/2023 10:22:51 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure.
| 198669 | RDS | CMPLT | CMPLT | 2023-10-27 10:25 | 2023-10-27 10:26 | Weekly Rounds: RESUS | (TC 10/27/2023 10:24:37 AM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure.
| 198670 | RDS | CMPLT | CMPLT | 2023-10-27 10:27 | 2023-10-27 10:29 | Weekly Rounds: 11 South | (TC 10/27/2023 10:26:59 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure.
68279 | 198671 | CM | CMPLT | CMPLT | 2023-10-27 10:39 | 2023-10-27 10:42 | Vinod requested to have all electrode leads and lead cables checked | (TC 10/26/2023 10:39:27 AM, Baghirzade, Nara) Vinod reported that from time to time, the signal does not show up on the acquisition window; there is no baseline. Requested to have all electrode leads and leads cables checked. Used DMM to check for continuity for each lead and each cable. Manipulated the wires to see if movement affects the signal. All continuity tests passed, could not find fault with the leads and cables. Next step is to check the amplifiers and do a water test. Vinod to confirm a time when machines are accessible.
1048564 | 198672 | CM | CMPLT | CMPLT | 2023-10-27 12:56 | 2023-10-27 12:57 | Prong was bent at a 90-degree angle | (TC 10/27/2023 12:55:54 PM, Baghirzade, Nara) Opened up the charger and removed the pins out of the plastic holder. Was able to bend the pin almost to the same angle as the other 2 pins. Re-assembled the charger and tested it. AC light comes on when battery is placed there for charging. Charger is fully functional. E-mailed Sharmaine to let her know that I dropped off the charger at their charging station.
1066935 | 198673 | CM | CMPLT | CMPLT | 2023-10-27 13:07 | 2023-10-27 13:10 | Moving equipment | (TC 10/27/2023 01:07:32 PM, Rhuland, Courtney) Issue with processor, swapped with one from bronchoscopy procedure room (1064939) with Rob from Olympus Waiting on new processor to be delivered, will swap again and send out for repairs and new processor to be in bronch.
1061214 | 198674 | CM | CMPLT | CMPLT | 2023-10-27 13:13 | 2023-10-27 13:14 | OR table remote not working | (TC 10/27/2023 01:13:22 PM, Reid, Emmerson ) Went up to replace the remote, however they had already gotten a working one I just had to return the others to empty tables.
1064939 | 198675 | CM | CMPLT | CMPLT | 2023-10-27 13:15 | 2023-10-27 13:16 | Moved | (TC 10/27/2023 01:15:09 PM, Rhuland, Courtney) Moved from bronch procedure room to endo room 4 Swapped with other processor (see asset 1066935) Waiting on new processor
| 198676 | CM | CMPLT | CMPLT | 2023-10-27 13:22 | 2023-10-27 13:23 | Meeting | (TC 10/27/2023 01:22:19 PM, Rhuland, Courtney) 1 on 1 with Adeel
1067933 | 198677 | CM | CMPLT | CMPLT | 2023-10-27 13:31 | 2023-10-27 13:50 | Installing vitalbridge in sim room | (TC 10/27/2023 01:31:16 PM, Canjimavadivel, Kamallesh ) Worked with vendor to setup vitalbridge
1046788 | 198678 | PM | CMPLT | CMPLT | 2023-10-27 13:35 | 2023-10-27 14:01 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 10/27/2023 01:54:59 PM, Cortez, John) Bed+Intellidrive batteries were replaced 60 mOhms, 25 uAmps Pendant missing Missing 3 screws, foot section PM passed
1055940 | 198679 | CM | CMPLT | CMPLT | 2023-10-27 14:02 | 2023-10-26 14:01 | Thermometer not working. Cover is stucked inside. | removed probe cover. installed battery in Radicall 7 . Tested unit. Functional.
| 198680 | CM | CMPLT | CMPLT | 2023-10-27 14:04 | 2023-10-26 14:03 | biomed meeting | biomed meeting
66474 | 198681 | CM | CMPLT | CMPLT | 2023-10-27 14:05 | 2023-10-27 14:05 | not passing circuit test | tested unit. patient circuit had a leak. fixed it. informed Rex.
63874 | 198682 | CM | CMPLT | CMPLT | 2023-10-27 14:34 | 2023-10-27 14:35 | Light replacement | (TC 10/27/2023 02:34:37 PM, Rhuland, Courtney) Replaced missing light on incubator with new one Noted other rooms that were missing lights on incubators
1067570 | 198683 | CM | CMPLT | NFF | 2023-10-27 14:37 | 2023-11-02 14:00 | Foot pedal not activating pump | (TC 10/27/2023 01:58:42 PM, Rhuland, Courtney) Checked device and foot pedal was activating pump Verified with other devices No fault found
67762 | 198684 | CM | CMPLT | CMPLT | 2023-10-30 09:24 | 2023-10-30 09:27 | PCEA dose cord not working | (TC 10/30/2023 09:24:42 AM, Reid, Emmerson ) Pin was bent in the dose cord. I was able to bend it back and it is working now.
1058606 | 198685 | PM | CMPLT | CMPLT | 2023-10-30 09:56 | 2023-10-30 09:57 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 09:56:32 AM, Reid, Emmerson ) PM Complete
1061554 | 198686 | PM | CMPLT | CMPLT | 2023-10-30 09:59 | 2023-10-30 10:00 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 09:59:06 AM, Reid, Emmerson ) PM Complete, see attached PM Report
66520 | 198687 | CM | CMPLT | CMPLT | 2023-10-30 10:00 | 2023-10-30 10:02 | Fetal recorder not working | (TC 10/25/2023 10:01:04 AM, Rendon Lopez , Elizabeth) Tested, self test and on demo mode, NFF
66518 | 198688 | CM | CMPLT | CMPLT | 2023-10-30 10:02 | 2023-10-30 10:04 | Recorder not working | (TC 10/25/2023 10:02:09 AM, Rendon Lopez , Elizabeth) Tested, recorder not feeding properly. Removed from ODU for repair (TC 10/26/2023 10:02:09 AM, Rendon Lopez , Elizabeth) Replaced recorder with repaired one. Calibrated, working properly
66522 | 198689 | CM | CMPLT | CMPLT | 2023-10-30 10:04 | 2023-10-30 10:05 | Recorder not working | (TC 10/25/2023 10:04:17 AM, Rendon Lopez , Elizabeth) Confirmed issue, recorder not feeding properly. Removed from ODU, placed spare. Replaced recorder with repaired one. Calibrated, working properly
66521 | 198690 | CM | CMPLT | CMPLT | 2023-10-30 10:06 | 2023-10-30 10:07 | Recorder not working | (TC 10/25/2023 10:06:49 AM, Rendon Lopez , Elizabeth) Tested, self test-ok. Cleaned printerhead. NFF, working properly
66519 | 198691 | CM | CMPLT | CMPLT | 2023-10-30 10:08 | 2023-10-30 10:09 | Recorder not working | (TC 10/25/2023 10:08:12 AM, Rendon Lopez , Elizabeth) Confirmed issue, recorder not feeding properly. Removed from ODU and placed spare monitor. Replaced recorder with repaired one. Calibrated, working properly
66535 | 198692 | CM | CMPLT | CMPLT | 2023-10-30 10:09 | 2023-10-30 10:10 | Check recorder | (TC 10/26/2023 10:09:28 AM, Rendon Lopez , Elizabeth) Checked recorder, self-test Ok, cleaned printerhead. Working properly
| 198693 | RDS | CMPLT | CMPLT | 2023-10-30 10:10 | 2023-10-31 14:01 | Weekly check in CCU | (TC 10/31/2023 02:00:46 PM, Scero, Tomlyn) weekly checks complete.
| 198694 | CM | CMPLT | CMPLT | 2023-10-30 10:11 | 2023-12-12 14:37 | L&D no asset WOs | (TC 10/25/2023 10:10:37 AM, Rendon Lopez , Elizabeth) Cleaning and repairing fetal monitor recorders for replacement (TC 10/26/2023 10:10:37 AM, Rendon Lopez , Elizabeth) Cleaning and repairing fetal monitor recorders for replacement (TC 11/06/2023 04:37:55 PM, Rendon Lopez , Elizabeth) Rounds (TC 11/07/2023 08:47:16 AM, Rendon Lopez , Elizabeth) OR 1 camera not working. Troubleshooted, camera working, not displayed, unplugged cables, reseated-OK (TC 11/09/2023 08:43:42 AM, Rendon Lopez , Elizabeth) Fetal recorder repaired (TC 11/13/2023 10:07:24 AM, Rendon Lopez , Elizabeth) Rounds (TC 11/16/2023 02:02:42 PM, Rendon Lopez , Elizabeth) L&D OR checking OR table remote control and hoover devices.
1068090 | 198695 | CM | CMPLT | CMPLT | 2023-10-30 10:12 | 2023-10-30 10:12 | Missing button | (TC 10/26/2023 10:11:46 AM, Rendon Lopez , Elizabeth) Added to TMS. Replaced button with spare, tested function-OK
1065515 | 198696 | PM | CMPLT | CMPLT | 2023-10-30 10:13 | 2023-10-30 10:17 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/30/2023 10:16:36 AM, Reid, Emmerson ) PM Complete, see attached PM Report
1067437 | 198697 | FC | CMPLT | CMPLT | 2023-10-30 12:09 | 2023-10-30 12:09 | QC check | passed QC check
1059778 | 198698 | FC | CMPLT | CMPLT | 2023-10-30 12:11 | 2023-10-30 12:11 | QC check | passed QC check
| 198699 | RDS | CMPLT | CMPLT | 2023-10-30 12:21 | 2023-10-30 12:21 | Weekly check for 14S and 16N(14N) | Weekly check for 14S and 16N(14N)
1061217 | 198700 | CM | CMPLT | NFF | 2023-10-30 14:26 | 2023-10-30 14:28 | Nurse PC not turning on | (TC 10/30/2023 02:26:42 PM, Reid, Emmerson ) Called to check out why the computer was not turning on. I turned it on and it started up without any problems.
1065697 | 198701 | CM | CMPLT | CMPLT | 2023-10-30 15:51 | 2023-12-06 16:57 | Error code | (TC 10/30/2023 03:51:29 PM, Rendon Lopez , Elizabeth) Error 110.6021 upon turning on. Device under warranty, sending out RMA# 303144724 (TC 12/06/2023 04:56:00 PM, Rendon Lopez , Elizabeth) Incoming inspection. Keypad replaced by BD. Check in tests, all passed
| 198702 | RDS | CMPLT | CMPLT | 2023-10-31 08:33 | 2023-10-31 08:36 | Weekly Checks For PACU | (TC 10/31/2023 08:34:23 AM, Reid, Emmerson ) Had to turn on several monitors and also moved an X3 to the dock.
| 198703 | RDS | CMPLT | CMPLT | 2023-10-31 08:37 | 2023-10-31 08:39 | Weekly Checks For OR | (TC 10/31/2023 08:38:20 AM, Reid, Emmerson ) All Checks complete.
1061211 | 198704 | CM | CMPLT | CMPLT | 2023-10-31 08:40 | 2023-10-31 08:42 | Nurse PC is not turning on | (TC 10/31/2023 08:40:28 AM, Reid, Emmerson ) Went up to chck and everything seemed to be connected properly, however the PC would not turn on. I called help desk and created ticket# 29819628. I let the nurses know this was helpdesk and handed over the ticket #
69292 | 198705 | CM | CMPLT | CMPLT | 2023-10-31 09:06 | 2023-11-16 13:38 | Created by Failed PM WO # 195152 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE |  | (TC 10/31/2023 09:01:01 AM, Baghirzade, Nara) Calibrated pressure and rate. During calibration, both passed. During PM, the rate was reading low (11.505). Will recalibrate and will assess the pressure sensors.
1056935 | 198706 | CM | CMPLT | CMPLT | 2023-10-31 09:09 | 2024-02-09 15:49 | Created by Failed PM WO # 197348 : PM Major Annual: Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 10/31/2023 10:05:33 AM, Canjimavadivel, Kamallesh ) Rattling was heard when shaking the device. Opened up to find broken pieces from the main board. Also found signs of corrosion on the msl connectors. Created a case with Philips to request a quote (TC 10/31/2023 01:50:08 PM, Canjimavadivel, Kamallesh ) Received quote from Philips, placed order for parts (TC 11/24/2023 09:14:04 AM, Canjimavadivel, Kamallesh ) Replaced the power board and MSL flex cable (TC 11/24/2023 03:16:12 PM, Canjimavadivel, Kamallesh ) Attempted to test module, did not work. Opened SupportTool and saw SN needed to be transferred to the board. Contacted tech support for instructions. Transferred SN and options. Attempted to complete PM. Still did not complete calibration step. (TC 01/04/2024 02:50:37 PM, Canjimavadivel, Kamallesh ) Attempted to calibrate, did not complete (TC 01/05/2024 01:28:22 PM, Canjimavadivel, Kamallesh ) Tried upgrading software and calibrating, did not work. Tested with different microstream and X3 and calibration also failed. Suspect issue with test set. Reached out to vendor for quote for new calibration line and tube (TC 02/09/2024 03:45:35 PM, Cortez, John) CO2 calibration failed after several attempts. Attached X3+microstream to MX800 setup and cal passed. Attached microstream to charging dock again and cal passed this time. | (TC 10/25/2023 10:59:51 AM, Canjimavadivel, Kamallesh ) Started PM, final calibration did not complete (TC 10/31/2023 09:07:35 AM, Canjimavadivel, Kamallesh ) Attempted to calibrate again, fail calibration
1054679 | 198707 | CM | CMPLT | CMPLT | 2023-10-31 10:20 | 2023-10-31 10:32 | "clog downstream error" | (TC 10/31/2023 10:21:09 AM, Cortez, John) Door missing, which is causing tube to bend/occlude Replaced door with spare, tested ok
| 198708 | RDS | CMPLT | CMPLT | 2023-10-31 10:31 | 2023-10-31 10:40 | Weekly check for ED
1054675 | 198709 | CM | CMPLT | NFF | 2023-10-31 10:47 | 2023-10-31 10:49 | "not working" | (TC 10/31/2023 10:47:29 AM, Cortez, John) Functional, ran test infusion for 30 minutes no issues found
1059838 | 198710 | CM | CMPLT | CMPLT | 2023-10-31 11:15 | 2023-12-28 12:43 | OPTHAMOLOGY 60 MURRAY ST 1st Floor ELLEX EYE CUBED CLINICAL REQUIRES IN-SERVICE | (TC 10/31/2023 11:16:38 AM, Whittington, Scott) OPTHAMOLOGY 60 MURRAY ST 1st Floor ELLEX EYE CUBED CLINICAL REQUIRES IN-SERVICE (TC 10/31/2023 01:46:45 PM, Baghirzade, Nara) Connected with Scott to discuss the issues with the ultrasound. Met with Henna to discuss the state of the ultrasound device. Ultrasound is not fully functional: -does not archive any patient data and images. -does not do UBM (supposed to). -require in-service training as only 1 doctor knows how to use the US. Acquired the PO, sent it to Naz. Naz to reach out to Clarion for in-service. Connected with Kamran to help them with storage; possible solution would be NAS or connecting it to the hospital network to PACS. Will visit the site on Nov. 1 with Kamran to get a better sense of what's required for image storage. (TC 11/02/2023 06:42:15 AM, Baghirzade, Nara) Went with Kamran to meet with Henna and Zain (US tech). Had Zain show us the functionality of the device and how complicated the software is to use without proper training. UBM probe is also not functional, when activated it vibrates for a moment and then stops. The device needs to be configured by vendor to suit the clinic needs. Zain also mentioned that the images do not populate the reports properly. The ultrasound device needs to be functionally tested. Naz to contact Clarion for in-service training. Will contact JDMI for functional testing. (TC 11/15/2023 02:07:33 PM, Whittington, Scott) OPTHAMOLOGY 60 MURRAY ST 1st Floor ELLEX EYE CUBED CLINICAL REQUIRES IN-SERVICE NOV 2 INTRODUCED NAZ TO PMH OCULAR ONCOLOGY 16TH FLOOR STAFF WHO TRAIN COLLEAGUES IN THIS DISCIPLINE NOV 15 MEET WITH MSH BIOMED TO DISCUSS (TC 11/20/2023 11:41:22 AM, Whittington, Scott) OPTHAMOLOGY 60 MURRAY ST 1st Floor ELLEX EYE CUBED CLINICAL REQUIRES IN-SERVICE VISIT CLINIC WITH NARA & NAZ | OPTHAMOLOGY 60 MURRAY ST 1st Floor ELLEX EYE CUBED CLINICAL REQUIRES IN-SERVICE From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: October 31, 2023 10:12 AM To: Baghirzade, Nargiz - Sinai Health System <Nargiz.Baghirzade@sinaihealth.ca> Cc: Alam, Adeel - Sinai Health System <Adeel.Alam@sinaihealth.ca>; Esmailzadeh, Nima - Sinai Health System <Nima.Esmailzadeh@sinaihealth.ca>; Chukri, Samir - Sinai Health System <Samir.Chukri@sinaihealth.ca>; Rhuland, Courtney - Sinai Health System <Courtney.Rhuland@sinaihealth.ca> Subject: Re: [External] FW: Re: Ellex EyeCube Ultrasound Machine Hi Nara The Clarion system is MIE# 1059838 Ellex Eye-Cubed This system is also used in PMH 18A Ocular clinic. Milad Mosabebi from TGH/PMH BIOMED has some experience with the PMH system. Perhaps MSH Opthamology could reach out to their PMH clinical colleagues for clinical instruction? Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Baghirzade, Nargiz <Nargiz.Baghirzade@sinaihealth.ca> Sent: Tuesday, October 31, 2023 7:45:24 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Alam, Adeel - Sinai Health System <Adeel.Alam@sinaihealth.ca>; Esmailzadeh, Nima - Sinai Health System <Nima.Esmailzadeh@sinaihealth.ca>; Chukri, Samir - Sinai Health System <Samir.Chukri@sinaihealth.ca>; Rhuland, Courtney - Sinai Health System <Courtney.Rhuland@sinaihealth.ca> Subject: [External] FW: Re: Ellex EyeCube Ultrasound Machine Good Morning, I received the e-mail below from the admin at the Ophthalmology clinic at 60 Murray, however I am not sure how to proceed as this involves an ultrasound machine. Please advise. Thank you. Regards, Nara Baghirzade, MLSO Biomedical Engineering Technologist, Biomedical Engineering Department T 416-586-8888 x 3887 C 437-231-8625 nara.baghirzade@sinaihealth.ca Mount Sinai Hospital, Joseph & Wolf Lebovic Health Complex 139-600 University Avenue, Toronto, ON M5G 1X5 A patient care, teaching and research centre affiliated with University of Toronto. From: Mistry, Henna Sent: October 30, 2023 2:23 PM To: Baghirzade, Nargiz <Nargiz.Baghirzade@sinaihealth.ca>; Rhuland, Courtney <Courtney.Rhuland@sinaihealth.ca> Cc: Abidin, Zain-ul <Zain-ul.Abidin@sinaihealth.ca> Subject: Re: Ellex EyeCube Ultrasound Machine Importance: High Hi Nara and Courtney, I hope you both are doing well. I apologize in advance for the lengthy email. We have a Clarion Ellex EyeCube Ultrasound machine that was purchased by the hospital through capital funding back in 2020/2021. As this was before my time here at Sinai, I do not know much about this. All I have been told was that Dr. Rand Simpson was part of the process in picking the machine, we received it, and a few technicians were shown some Youtube videos on how to use it. Since then the machine has not really been used much as this was during Covid, Dr. Simpson retired and did not return after Covid, and we still had the Sonomed machine which was working. Now the Sonomed unit is not working properly and this company no longer services this device in Canada. The technicians and the physicians have now turned to using the Ellex Eyecube however, as there was no proper training they cannot use it properly. We tried contacting the Clarion company to see if we could arrange someone to come on site to show the physician and staff how to use the machine properly however, they mentioned that since we do not have a service agreement they cannot do that. As the staff were not provide a proper in-service training back in 2020/2021 due to Covid, shouldn’t Clarion provide one now for us at no charge? According to one of the technicians that were here at the time, he said Dr. Rand Simpson put a clause or stipulation in the agreement on purchase that we get a proper in service training on the machine so we know how to use it. The physician, Dr. Carla Lutchman, that normally the ultrasounds is away this week but will need the machine working next week when she is back for clinic. Additionally, they mentioned that this machine needs to be archived. Myself and the technicians are not sure on what program or if it’s an external hard drive of some sorts as that was also not shown to them at the time. Shouldn’t Clarion have set this up when they came to deliver the device? We also tried to get this device on the hospital network as it is currently not connected to any internet port. We contacted IT, filled in some forms, and then they closed the ticket without doing anything. I am wondering if you could reach out to Clarion and see if you can get someone on site to help with this ASAP. Your help will be much appreciated! Thank you, Henna Mistry Office Coordinator, Department Of Ophthalmology T 416-586-4800 ext. 2299 F 416-586-8789 henna.mistry@sinaihealth.ca Mount Sinai Hospital, Joseph & Wolf Lebovic Health Complex L1-001, 60 Murray St. Box #3 Toronto, ON M5T 3L9
63236 | 198711 | CM | CMPLT | NFF | 2023-10-31 11:16 | 2023-10-31 11:17 | "not working" | (TC 10/31/2023 11:16:35 AM, Cortez, John) Ran functional tests (accuracy, alarms) successfully No fault found
1068097 | 198712 | CM | CMPLT | CMPLT | 2023-10-31 12:33 | 2023-10-31 12:34 | Asset tagging and QC | (TC 10/31/2023 12:32:54 PM, Baghirzade, Nara) Loaner device while department-owned device is being repaired at 3M. Asset tagged as loaner and input all the device information on TMS. QC before patient use: Physical inspection good. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 123 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 124 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Device is functional. Returned to 11S.
68279 | 198713 | CM | CMPLT | CMPLT | 2023-10-31 12:41 | 2024-02-02 07:50 | Amplifier and leads testing | (TC 10/31/2023 12:41:35 PM, Baghirzade, Nara) Tested both data input connections for the amplifier with different connectors and electrode sets. Shorted the 3 leads to do testing for noise at 10 uV scale. Both DINs have good noise levels, less than 10 uV. DIN2 intermittently picks up some excessive noise. Initial issue: when doing EMG, the lines randomly go flat or excessive noise is picked up. Also when doing needle EMG, after the needle has been inserted to the patient, there is no activity on the monitor and no noise coming from the needle. Will contact Natus for remote support. (TC 10/31/2023 03:00:52 PM, Baghirzade, Nara) Called tech support, created help desk ticket, and coordinated a time for remote support. Friday, Nov. 3 at 10 AM with Michael Babineau (TC 11/03/2023 11:39:00 AM, Baghirzade, Nara) Remote session with Michael Babineau from Natus. He checked the software, all profiles, and specifications. Suggested to check for the internal solder joints on the amplifier boards. Samir and I took apart both amplifiers and reinforced the joints with soldering. Called Natus back to book another remote session for amplifier testing. If this does not resolve the flat waveforms testing, then both systems would require updates. Since the systems are from 2016, parts are no longer available and end of support. Followed up with Nima regarding this. If required, we will set up a meeting with Doug (sales rep) to discuss parts for the system. (TC 11/06/2023 08:34:28 AM, Baghirzade, Nara) Placed amplifier back on system. Ran EDX HW programmer successfully. Waiting on next Natus remote troubleshooting steps. (TC 01/08/2024 02:36:33 PM, Baghirzade, Nara) Vinod reported another wasted needle and additional noise on the amplifier. Concluded that the connectors are not the issue as they have be re-soldered. Had a meeting with Doug and Andrew; Doug to confirm if we can purchase the amplifier or base unit separately from entire system. Emailed Doug the part number for the EDX base unit, waiting on his response. (TC 02/02/2024 07:46:28 AM, Baghirzade, Nara) Received quote from Doug. Andrew to put the parts on capital list. At next meeting with Doug, we will confirm for how long we can keep ordering the individual parts: amplifier and base unit.
1059181 | 198714 | CM | CMPLT | CMPLT | 2023-10-31 12:54 | 2023-10-31 12:57 | SpO2 Malfunction | (TC 10/31/2023 12:53:57 PM, Reid, Emmerson ) Went up to check. Needed to restart the the X3 and then it was working.
63207 | 198715 | CM | CMPLT | NFF | 2023-10-31 13:24 | 2023-10-31 13:27 | "not working" | (TC 10/31/2023 01:24:20 PM, Cortez, John) Ran functional tests (accuracy, alarms) successfully Unable to replicate issue
| 198716 | CM | CMPLT | CMPLT | 2023-10-31 13:44 | 2023-11-02 06:41 | Ophthalmoscopes not working | (TC 10/31/2023 01:44:12 PM, Baghirzade, Nara) Elena requested to have 2 ophthalmoscopes looked at: 1. ophthalmscope required a battery change. Replaced the battery for them. Ophthalmoscope is fully functional. 2. second ophthalmscope did not switch on even after battery and bulb change. Measured the battery voltage with DMM and it was reading approximately 2.5V; put the ophthalmoscope to charge. Will check to see if it runs after charging overnight. (TC 11/01/2023 08:00:00 AM, Baghirzade, Nara) Charged ophthalmoscope battery overnight. Ran it for an hour on battery, light intensity did not change and device is functional. Returned to Ana at ENT.
| 198717 | PA | CMPLT | CMPLT | 2023-10-31 13:51 | 2023-11-08 08:28 | ECG clips for ED | (TC 10/31/2023 02:02:10 PM, Canjimavadivel, Kamallesh ) ED requested clips for ECG cart, placed order to add to their stock (TC 11/08/2023 08:27:28 AM, Canjimavadivel, Kamallesh ) Received clips and dropped off to ED
501064 | 198718 | CM | CMPLT | CMPLT | 2023-10-31 14:18 | 2023-11-16 15:33 | msh opg CEOU-3231(Rm6) IU22 MAIN MONITORT FAULTY | (TC 10/31/2023 02:18:32 PM, Whittington, Scott) msh opg CEOU-3231(Rm6) IU22 MAIN MONITORT FAULTY REF# #0122523137 (TC 11/06/2023 03:04:27 PM, Whittington, Scott) msh opg CEOU-3231(Rm6) IU22 MAIN MONITORT FAULTY REF# #0122523137 (TC 11/07/2023 11:46:35 AM, Whittington, Scott) msh opg CEOU-3231(Rm6) IU22 MAIN MONITORT FAULTY MONITOR REPLACED (TC 11/16/2023 03:32:52 PM, Whittington, Scott) msh opg CEOU-3231(Rm6) IU22 MAIN MONITORT FAULTY PUT BACK IN RM 7 | msh opg CEOU-3231(Rm6) IU22 MAIN MONITORT FAULTY REF# #0122523137
501596 | 198719 | CM | CMPLT | CMPLT | 2023-10-31 14:20 | 2023-11-06 08:33 | MSH 15 SPP IU22 FAULTY PROBE C5-1 SN# B1K79X | (TC 10/31/2023 02:19:53 PM, Whittington, Scott) MSH 15 SPP IU22 FAULTY PROBE (TC 11/01/2023 10:12:11 AM, Whittington, Scott) MSH 15 SPP IU22 FAULTY PROBE C5-1 SN# B1K79X | MSH 15 SPP IU22 FAULTY PROBE C5-1 SN# B1K79X SITE 556348 REF# CHRIS 0122530220 NOV 6 RECEIVED PROBE SN# F09Y9J
| 198720 | CM | CMPLT | CMPLT | 2023-10-31 14:22 | 2024-02-26 13:30 | MSH 15 L&D ANAESTHESIA SONOSITE EDGE II PROBE ORDER | (TC 10/31/2023 02:22:22 PM, Whittington, Scott) MSH 15 L&D ANAESTHESIA SONOSITE EDGE II PROBE ORDER | MSH 15 L&D ANAESTHESIA SONOSITE EDGE II ADDITIONAL PROBE ORDER NAZ FROM MSH BIOMED ENG GROUP INITIATED
63214 | 198721 | CM | CMPLT | CMPLT | 2023-10-31 16:10 | 2024-05-16 11:41 | "feed error" | (TC 10/31/2023 04:10:44 PM, Cortez, John) Confirmed issue, Ran test infusion and kept getting feed errors Will send to Cardinal Health for repair/replacement (TC 05/16/2024 11:39:19 AM, Cortez, John) Received replacement (sn C20298092) Unboxed and completed functinal check
63430 | 198722 | CM | CMPLT | CMPLT | 2023-10-31 16:13 | 2024-05-16 11:14 | "broken" | (TC 10/31/2023 04:13:04 PM, Cortez, John) Cracked/broken case near peristaltic pump Will send to Cardinal Health (TC 05/16/2024 11:11:29 AM, Cortez, John) Replacement arrived (sn C20298882) Unboxed and completed functional check
52033 | 198723 | PM | ACTIV | READY | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52252 | 198724 | PM | CMPLT | PMFAIL | 2023-11-01 03:18 | 2023-11-06 07:48 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/03/2023 02:47:00 PM, Baghirzade, Nara) PM failed during functional testing of RH and temp. 9V rechargeable battery was replaced. Flip door was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured very low (74%); calibrated RH sensor, measured around 76%. Looking at past WOs, the RH sensor has not been replaced, only re-calibrated. Will replace sensor and run a functional test.
62779 | 198725 | PM | CANCL | CANCL | 2023-11-01 03:18 | 2024-07-04 09:21 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport
62989 | 198726 | PM | CMPLT | CMPLT | 2023-11-01 03:18 | 2024-06-26 12:58 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/26/2024 12:57:40 PM, Rhuland, Courtney) PM completed All tests passed Irradiance check - OK Electrical safety - OK
63074 | 198727 | PM | CMPLT | CMPLT | 2023-11-01 03:18 | 2023-12-18 10:42 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 12/18/2023 10:41:42 AM, Rhuland, Courtney) Battery check - OK Functional check - OK
63177 | 198728 | PM | ACTIV | ISSUE | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Insufflators
63524 | 198729 | PM | ACTIV | ISSUE | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
65120 | 198730 | PM | CMPLT | CMPLT | 2023-11-01 03:18 | 2024-04-15 10:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/15/2024 10:35:41 AM, Reid, Emmerson ) Had to run the volume test twice. PM Complete, see procedure for details.
65894 | 198731 | PM | CNP | ENL | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65897 | 198732 | PM | CANCL | CANCL | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65910 | 198733 | PM | CANCL | CANCL | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65930 | 198734 | PM | CANCL | CANCL | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65932 | 198735 | PM | CNP | ENL | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65971 | 198736 | PM | CNP | ENL | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65975 | 198737 | PM | CMPLT | CMPLT | 2023-11-01 03:18 | 2024-03-26 14:20 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/26/2024 02:19:45 PM, Rhuland, Courtney) PM completed by BD
65997 | 198738 | PM | CNP | ENL | 2023-11-01 03:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66055 | 198739 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66065 | 198740 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66243 | 198741 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-07-05 12:10 | PM-MAJOR - Annual PM For: Simulators, Multiparameter | (TC 07/05/2024 12:10:00 PM, Chukri, Samir) Certificate of calibration performed by Clare instrument company. Pls see attched
66470 | 198742 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-11-01 12:58 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. Replaced PS500 batteries. Tested. did not communicate with C500. reset the C500. Functional. SW.2.51 Leakage: 77uA and 91 ohm.
66490 | 198743 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-01-30 10:06 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 01/25/2024 10:05:11 AM, Rendon Lopez , Elizabeth) Performance assurance, all tests passed
67202 | 198744 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | BP Check
67222 | 198745 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-06 16:34 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/01/2023 04:33:35 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67225 | 198746 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-11-28 09:51 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 11/24/2023 09:50:47 AM, Rendon Lopez , Elizabeth) Blender replaced. PM completed, all tests passed
67226 | 198747 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-06 16:49 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/06/2023 04:49:07 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67380 | 198748 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67392 | 198749 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-06 16:46 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/05/2023 04:45:48 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67415 | 198750 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-01-31 15:44 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/31/2024 03:43:46 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67418 | 198751 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-06 16:52 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/05/2023 04:52:08 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67422 | 198752 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-12 09:40 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/11/2023 09:39:54 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67878 | 198753 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-01-19 09:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/18/2024 09:39:34 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67881 | 198754 | PM | CMPLT | PMFAIL | 2023-11-01 03:19 | 2024-01-19 09:29 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/18/2024 09:27:26 AM, Rendon Lopez , Elizabeth) Testing device, ran rate accuracy several times with different cassettes and device over delivering. requested RMA to send out for repair
67882 | 198755 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-02-29 09:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 02/28/2024 09:43:37 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67905 | 198756 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Insufflators
68547 | 198757 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-11-20 09:18 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 11/02/2023 02:42:03 PM, Scero, Tomlyn) All PM parts replaced including batteries. PM tests started as per IPM-L (TC 11/16/2023 09:15:17 AM, Scero, Tomlyn) Continued with PM proceedures (TC 11/17/2023 09:15:48 AM, Scero, Tomlyn) All PM procedures complete and passed. IPM-L results sheet attcahed. returned AGM to the OR. (TC 11/14/2023 09:16:49 AM, Scero, Tomlyn) Continued with PM procedures.
68590 | 198758 | PM | CMPLT | PMFAIL | 2023-11-01 03:19 | 2024-03-13 09:12 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
68613 | 198759 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-03 13:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:25:12 PM, Rhuland, Courtney) PM completed by BD
68616 | 198760 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-03 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:10:59 PM, Rhuland, Courtney) PM completed by BD
68621 | 198761 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68639 | 198762 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68658 | 198763 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-24 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68662 | 198764 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-15 09:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/11/2024 09:01:19 AM, Huang, Jessica) replaced the IUI connector and completed PM
68683 | 198765 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68692 | 198766 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-03 13:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:54:02 PM, Rhuland, Courtney) PM completed by BD
68745 | 198767 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68748 | 198768 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-20 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68757 | 198769 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-01 09:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 12/01/2023 09:35:27 AM, Campbell, Connor) PM passed/complete, refer to ASM refer to procedure for ESA
68772 | 198770 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68845 | 198771 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68856 | 198772 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-24 14:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68860 | 198773 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-15 08:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/10/2024 08:56:34 AM, Huang, Jessica) replaced the power cord and completed PM
68876 | 198774 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-24 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68887 | 198775 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68903 | 198776 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68904 | 198777 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68913 | 198778 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-13 09:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/11/2024 09:11:35 AM, Rendon Lopez , Elizabeth) PM completed by BD
68930 | 198779 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68978 | 198780 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68997 | 198781 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69007 | 198782 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69035 | 198783 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69110 | 198784 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69120 | 198785 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69158 | 198786 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69185 | 198787 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69195 | 198788 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69217 | 198789 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-04-26 09:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69239 | 198790 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-12 10:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69245 | 198791 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69280 | 198792 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69425 | 198793 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69454 | 198794 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69465 | 198795 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-21 07:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 12/21/2023 07:57:01 AM, Campbell, Connor) PM complete/passed, refer to ASM
69510 | 198796 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-19 21:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69517 | 198797 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-25 11:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69551 | 198798 | PM | CMPLT | PMFAIL | 2023-11-01 03:19 | 2024-04-04 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69625 | 198799 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-12 10:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69642 | 198800 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69658 | 198801 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-21 18:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
67363 | 198802 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
67365 | 198803 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
32520 | 198804 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-11 16:30 | MSH DR2 PM - Semi-Annual | (TC 12/11/2023 04:29:55 PM, Whittington, Scott) MSH DR2 PM - Semi-Annual | MSH DR2 PM - Semi-Annual
32520 | 198805 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-11-24 15:40 | MSH DR2 HARP NOV 30 TRGT HARP - Semi Annual | (TC 11/23/2023 11:54:51 AM, Whittington, Scott) MSH DR2 HARP NOV 30 TRGT HARP - Semi Annual (TC 11/23/2023 11:55:03 AM, Whittington, Scott) MSH DR2 HARP NOV 30 TRGT HARP - Semi Annual FOLLOWUP (TC 11/23/2023 11:54:51 AM, Whittington, Scott) MSH DR2 HARP NOV 30 TRGT HARP - Semi Annual | MSH DR2 HARP NOV 30 TRGT HARP - Semi Annual
32540 | 198806 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-05 16:19 | PM - Monthly | PM Major done Slice count - 9575 TOS phantom a(-991.0, 10.4) b(330.8, 14.5) c(132.2, 13.7) d(99.1, 13.6) e(-105.8, 12.0) f(-0.6, 13.9) -System passed all tests
32560 | 198807 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-11-29 14:39 | PM - Monthly | PM Major done Slice count - 1315220 TOS phantom a(-990.2, 9.7) b(334.0, 14.3) c(132.8, 13.5) d(96.1, 13.5) e(-106.3, 11.7) f(-0.6, 13.3) -System passed all tests
32580 | 198808 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-11-28 15:13 | PM - Monthly | PM Major done Mechanical done Slice count - 147692 TOS phantom a(-990.1, 10.7) b(335.5, 15.2) c(133.7, 13.8) d(97.4, 14.0) e(-106.7, 12.7) f(0.8, 13.8) -System passed all tests
39505 | 198809 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-05 16:19 | PM - Annual | Injector PM done -Flex Ring calibrated -Potentiometer calibrated -Pressure calibrated -Touchscreen calibrated -System passed all functional tests
39506 | 198810 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-01 15:23 | PM - Annual | Injector PM done -Flex Ring calibrated -Potentiometer calibrated -Pressure calibrated -Touchscreen calibrated -System passed all functional tests
395000 | 198811 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-02-06 15:48 | PM - Annual | Injector PM done -Flex Ring calibrated -Potentiometer calibrated -Pressure calibrated -Touchscreen calibrated -System passed all functional tests
502452 | 198812 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-19 08:43 | PM - Annual | (TC 12/19/2023 08:42:40 AM, Fan, Tak) PM Completed as per checklist. Gnd R = 163 mohm Max A = 1.9 / 75.5 / 149.5 uA
502453 | 198813 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-15 08:34 | PM - Annual | (TC 12/14/2023 08:31:37 AM, Fan, Tak) PM completed. Gnd R = 197 mohm Max A = 1.9 / 80.2 / 156.6 uA C2-9-D 3067770YP9 = 0.5 / 20.8 uA (May need replacement) RAB6-D 474511WX3 = 0.4 / 2.8 uA C1-5-D 201813YP3 = 0.4 / 2.2 uA
1046062 | 198814 | PM | ACTIV | READY | 2023-11-01 03:19 |  | Annual PM
1046442 | 198815 | PM | ACTIV | READY | 2023-11-01 03:19 |  | Annual PM
1046452 | 198816 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-11-14 10:07 | Annual PM | (TC 11/13/2023 10:06:40 AM, Rendon Lopez , Elizabeth) PM completed, after repair. All tests passed
1046634 | 198817 | PM | ACTIV | READY | 2023-11-01 03:19 |  | Annual PM
1046635 | 198818 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-07-11 11:14 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 07/11/2024 11:14:25 AM, Reid, Emmerson ) PM Complete, see attached PM report.
1046655 | 198819 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046656 | 198820 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046657 | 198821 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046658 | 198822 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046659 | 198823 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046781 | 198824 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-09-22 20:45 | PM - Semi-Annual | PM
1046781 | 198825 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-09-20 22:54 | HARP - Semi Annual | HARP
1047095 | 198826 | PM | CNP | ENL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047358 | 198827 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-07-24 14:31 | PM-MAJOR- Annual PM for: High flow module | (TC 07/24/2024 02:31:00 PM, Chukri, Samir) Certificate Calibration done by TSI attached.
1047359 | 198828 | PM | ACTIV | PV | 2023-11-01 03:19 |  | PM-MAJOR- Annual PM for: Testers
1047360 | 198829 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-08-27 07:58 | PM-MAJOR- Annual PM for: low flow module | (TC 08/14/2024 01:56:52 PM, Chukri, Samir) completed the KJV and shipped to TSI (TC 08/27/2024 07:57:35 AM, Chukri, Samir) Received from TSI with Calibration Certificate attached .
1047361 | 198830 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-08-27 08:01 | PM-MAJOR- Annual PM for: High Flow module | (TC 08/14/2024 01:59:17 PM, Chukri, Samir) Completed KJV and PO shipped to TSI (TC 08/27/2024 08:01:28 AM, Chukri, Samir) Received from TSI with Calibration Certificate attached.
1048531 | 198831 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-05-06 08:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/06/2024 08:31:02 AM, Rhuland, Courtney) Calibrated and performed check-in All tests passed Refer to ASM for details
1049204 | 198832 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-22 08:39 | PM-MAJOR - Annual PM For: Central Monitor | (TC 12/21/2023 08:34:33 AM, Canjimavadivel, Kamallesh ) Rebooted after update, cleaned Note: dome lights in front of each room flashed red as central was off
1049215 | 198833 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-22 08:33 | PM-MAJOR - Annual PM For: Central Monitor | (TC 12/21/2023 08:31:16 AM, Canjimavadivel, Kamallesh ) Rebooted after update, cleaned
1049216 | 198834 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Central Monitor
1049217 | 198835 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-02-22 13:47 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/22/2024 01:45:42 PM, Scero, Tomlyn) PM complete during SW patch
1049218 | 198836 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-02-22 13:48 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/22/2024 01:47:19 PM, Scero, Tomlyn) PM complete during SW patch
1049219 | 198837 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-02-22 13:48 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/22/2024 01:47:46 PM, Scero, Tomlyn) PM complete during SW patch
1049220 | 198838 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-02-22 13:48 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/22/2024 01:48:12 PM, Scero, Tomlyn) PM complete during SW patch
1049221 | 198839 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-02-22 13:44 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/22/2024 01:42:37 PM, Scero, Tomlyn) PM complete during SW patch
1049222 | 198840 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-12 09:15 | PM-MAJOR - Annual PM For: Central Monitor | (TC 12/08/2023 09:14:10 AM, Rendon Lopez , Elizabeth) Vacuumed central. PM completed, all tests passed
1049255 | 198841 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-01-23 14:27 | PM-MAJOR - Annual PM For: Infant T-Piece Resuscitator | (TC 01/23/2024 02:26:47 PM, Baghirzade, Nara) Performance assurance passed and completed. All values within range. @10 = 9 cmH2O @20 = 20 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 73 cmH2O Less than 80 = 75 cmH2O (TC 01/23/2024 02:27:09 PM, Rhuland, Courtney) Performance assurance passed and completed. All values within range. @10 = 9 cmH2O @20 = 20 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 73 cmH2O Less than 80 = 75 cmH2O
1049608 | 198842 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-15 08:40 | PM - Annual | (TC 12/15/2023 07:39:00 AM, Fan, Tak) PM completed. Gnd R = 171 mohm Max A = 0.4 / 70.6 / 141.0 uA C2-9-D 426666YP4 = 0.5 / 1.6 uA eM6C G2 230933KR4 = 0.5 / 2.2 uA C1-5-D 3231202WX8 = 0.5 / 1.8 uA
1049901 | 198843 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-13 16:01 | HARP - Semi Annual |  | Harp performed.
1049901 | 198844 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-13 16:02 | PM |  | PM Minor
1049902 | 198845 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-19 11:23 | HARP - Semi Annual |  | HARP performed.
1049902 | 198846 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-19 11:24 | PM |  | PM performed.
1049903 | 198847 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-06 17:03 | HARP - Semi Annual |  | HARP completed.
1049903 | 198848 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-05 17:04 | PM |  | PM performed.
1049904 | 198849 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-14 11:47 | HARP - Semi Annual |  | Harp completed.
1049904 | 198850 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-14 11:49 | PM |  | PM Minor
1049905 | 198851 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-06 17:04 | HARP - Semi Annual |  | HARP completed
1049905 | 198852 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-06 17:05 | PM |  | PM Minor completed.
1049940 | 198853 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-28 13:05 | PM - Annual | (TC 12/27/2023 01:02:35 PM, Fan, Tak) Gnd R = 168 mohm Max A = 0.2 / 164.7 / 317 uA ECG = 0.5 / 3.0 uA ISO = 7.8 uA M5Sc-D (303529YP2) = 0.7 / 1.9 uA 4Vc-D (329824YP7) = 0.7 / 2.1 uA
1049941 | 198854 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-01-02 12:18 | PM - Annual | (TC 12/29/2023 12:16:55 PM, Fan, Tak) Gnd R = 76 mohm Max A = 0.2 / 158.7 / 310 uA ECG = 0.5 / 2.2 uA ISO = 8.4 uA M5S = 0.5 / 3.7 uA
1050056 | 198855 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-12 09:00 | PM, | (TC 12/12/2023 08:59:53 AM, Rendon Lopez , Elizabeth) PM completed after repair. Attached tests
1050256 | 198856 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-21 18:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050267 | 198857 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050904 | 198858 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-15 08:37 | PM - Annual | (TC 12/14/2023 08:35:03 AM, Fan, Tak) PM completed. Gnd R = 191 mohm Max A = 2.0 / 69.3 / 138.5 uA C2-9-D 346321YP3 = 0.5 / 4.6 uA (May need replacement) eM6C G2 258237KR7 = 0.4 / 2.2 uA C1-6-D 350629YP2 = 0.5 / 1.8 uA
1051377 | 198859 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051386 | 198860 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-05-31 11:19 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 05/31/2024 11:19:00 AM, Rhuland, Courtney) PM completed Hours - 3604.5 Irradiance - 51.1 uW/cm^2nm Electrical safety: Ground resistance - 0.132 uOhms Leakage current - 0.2 uA Replaced power cord retainer
1051830 | 198861 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PLANNED MAINTENANCE-VACUUM PC
1051831 | 198862 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Central Monitor
1051840 | 198863 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-12 14:25 | PM MAJOR - PM For Central Monitor PC | (TC 12/12/2023 02:24:52 PM, Reid, Emmerson ) Completed the patch
1053320 | 198864 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM-MAJOR - Annual PM For: Insufflators
1053330 | 198865 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2024-03-04 15:16 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 03/04/2024 03:13:25 PM, Canjimavadivel, Kamallesh ) Replaced PM kit parts All calibrations and tests pass Leakage current: 24uA
1053876 | 198866 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM MAJOR - ANNUAL PM for: ESU
1054227 | 198867 | PM | CMPLT | CMPLT | 2023-11-01 03:19 | 2023-12-20 11:51 | PM - Annual | (TC 12/20/2023 07:48:00 AM, Fan, Tak) PM completed. Gnd R = 69 mohm Max A = 2.0 / 78.1 / 153.7 uA RAB-6-D = (474512WX1) = 0.6 / 3.1 uA C1-5-D (286042YP7) = 0.6 / 1.9 uA C2-9-D (343599YP7) = 0.6 / 2.5 uA
1054591 | 198868 | PM | ACTIV | ISSUE | 2023-11-01 03:19 |  | PM MAJOR-Annual Philips Monitoring, Central Station
1054599 | 198869 | PM | CANCL | CANCL | 2023-11-01 03:19 |  | PM MAJOR-Annual Philips Monitoring, Central Station
1055106 | 198870 | PM | CANCL | CANCL | 2023-11-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055165 | 198871 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-03-15 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055177 | 198872 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-04-09 10:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/09/2024 10:35:09 AM, Reid, Emmerson ) Replaced the front cover. Ran ASM Calibration, and PM, all tests passed. PM Complete.
1055751 | 198873 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-04-10 14:53 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:53:10 PM, Rhuland, Courtney) PM completed by BD
1056190 | 198874 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-07 09:14 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/07/2023 09:13:39 AM, Canjimavadivel, Kamallesh ) Completed patch updates and vacuumed PICiX
1057511 | 198875 | PM | ACTIV | ISSUE | 2023-11-01 03:20 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057693 | 198876 | PM | CNP | ENL | 2023-11-01 03:20 | 2024-07-19 07:56 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1057699 | 198877 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-28 08:46 | Battery Cycling | (TC 11/28/2023 08:45:10 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete Capacity 3/5 LEDs
1057700 | 198878 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-01 10:17 | Battery Cycling | (TC 12/01/2023 10:17:05 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057701 | 198879 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-16 10:09 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 02/16/2024 10:08:28 AM, Canjimavadivel, Kamallesh ) PM complete
1057702 | 198880 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-04-02 10:14 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 04/02/2024 10:13:37 AM, Canjimavadivel, Kamallesh ) Calibration completed, PM pass
1057703 | 198881 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-04-02 09:54 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 04/02/2024 09:38:57 AM, Canjimavadivel, Kamallesh ) Calibration complete, PM pass
1057705 | 198882 | PM | CNP | ENL | 2023-11-01 03:20 | 2024-07-19 07:56 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1057707 | 198883 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-09 15:56 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 02/09/2024 03:54:22 PM, Cortez, John) Calibration complete, PM pass
1057710 | 198884 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-01 10:12 | Battery Cycling | (TC 12/01/2023 10:12:05 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1057711 | 198885 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-30 09:10 | Battery Cycling | (TC 11/30/2023 09:10:08 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057713 | 198886 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-29 14:14 | Battery Cycling | (TC 11/29/2023 02:13:17 PM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057714 | 198887 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-28 08:45 | Battery Cycling | (TC 11/28/2023 08:43:50 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete Capacity 3/5 LEDs
1057715 | 198888 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-30 09:10 | Battery Cycling | (TC 11/30/2023 09:08:55 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057716 | 198889 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-04 09:36 | Battery Cycling | (TC 12/04/2023 09:36:03 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057717 | 198890 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-28 08:48 | Battery Cycling | (TC 11/28/2023 08:47:25 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete Capacity 4/5 LEDs
1057718 | 198891 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-04 09:35 | Battery Cycling | (TC 12/04/2023 09:33:33 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057719 | 198892 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-05 09:03 | Battery Cycling | (TC 12/05/2023 09:02:34 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1057720 | 198893 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-01 10:14 | Battery Cycling | (TC 12/01/2023 10:13:39 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057721 | 198894 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-29 08:29 | Battery Cycling | (TC 11/29/2023 08:27:24 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057722 | 198895 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-01 10:12 | Battery Cycling | (TC 12/01/2023 10:10:50 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1057723 | 198896 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-01 09:27 | Battery Cycling | (TC 12/01/2023 09:23:07 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1057724 | 198897 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-01 10:16 | Battery Cycling | (TC 12/01/2023 10:15:57 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057726 | 198898 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-05 09:04 | Battery Cycling | (TC 12/05/2023 09:03:29 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057727 | 198899 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-01 10:15 | Battery Cycling | (TC 12/01/2023 10:14:18 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1057730 | 198900 | PM | CANCL | CANCL | 2023-11-01 03:20 | 2024-06-26 14:40 | Battery Cycling
1057732 | 198901 | PM | CANCL | CANCL | 2023-11-01 03:20 | 2024-06-26 14:37 | Battery Cycling
1057733 | 198902 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-09 10:11 | Battery Cycling | (TC 11/09/2023 10:10:48 AM, Reid, Emmerson ) Conditioning cycling complete.
1057734 | 198903 | PM | CANCL | CANCL | 2023-11-01 03:20 | 2024-06-26 14:39 | Battery Cycling
1057735 | 198904 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-09 10:10 | Battery Cycling | (TC 11/09/2023 10:09:26 AM, Reid, Emmerson ) Conditioning cycling complete.
1057736 | 198905 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-08 08:11 | Battery Cycling | (TC 11/08/2023 08:11:02 AM, Reid, Emmerson ) Conditioning complete.
1057737 | 198906 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-08 08:10 | Battery Cycling | (TC 11/08/2023 08:10:05 AM, Reid, Emmerson ) Conditioning complete.
1057738 | 198907 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-09 10:08 | Battery Cycling | (TC 11/09/2023 10:07:52 AM, Reid, Emmerson ) Conditioning cycling complete.
1057739 | 198908 | PM | CNP | EIU | 2023-11-01 03:20 | 2024-09-18 14:10 | Battery Cycling
1057740 | 198909 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-08 08:09 | Battery Cycling | (TC 11/08/2023 08:08:50 AM, Reid, Emmerson ) Conditioning complete
1057758 | 198910 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-07 09:19 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/07/2023 09:14:33 AM, Canjimavadivel, Kamallesh ) Completed patch updates and vacuumed PICiX
1057760 | 198911 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-07 09:25 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/07/2023 09:22:20 AM, Canjimavadivel, Kamallesh ) Completed patch updates and vacuumed PICiX
1057761 | 198912 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-07 09:27 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/07/2023 09:26:21 AM, Canjimavadivel, Kamallesh ) Completed patch updates and vacuumed PICiX
1057868 | 198913 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-09-17 15:03 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 09/17/2024 03:03:27 PM, Reid, Emmerson ) PM complete, see procedure for details.
1058237 | 198914 | PM | CANCL | CANCL | 2023-11-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059330 | 198915 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-01 10:13 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/01/2024 10:12:05 AM, Baghirzade, Nara) Checked network settings for the correct DNS primary and server - all good. Vacuumed the PICIX. Overview fully functional. (TC 02/01/2024 10:12:53 AM, Rhuland, Courtney) Checked network settings for the correct DNS primary and server - all good. Vacuumed the PICIX. Overview fully functional.
1059440 | 198916 | PM | ACTIV | ISSUE | 2023-11-01 03:20 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059441 | 198917 | PM | ACTIV | ISSUE | 2023-11-01 03:20 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059530 | 198918 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-12 14:28 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/12/2023 02:27:59 PM, Reid, Emmerson ) Completed the patch
1059531 | 198919 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-12 14:26 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/12/2023 02:26:13 PM, Reid, Emmerson ) Completed the patch
1059532 | 198920 | PM | ACTIV | ISSUE | 2023-11-01 03:20 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059533 | 198921 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-12 14:27 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/12/2023 02:27:08 PM, Reid, Emmerson ) Completed the patch
1059534 | 198922 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-12 14:29 | PM MAJOR-Monitors, Physiologic, Central Station
1059541 | 198923 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:25 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:24:10 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:24:37 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059542 | 198924 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:28 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:27:11 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:27:23 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059543 | 198925 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:36 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:35:37 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:35:44 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059544 | 198926 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:30 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:29:30 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:29:37 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059545 | 198927 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 14:04 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 02:02:38 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. Took longer to shut down and system tried to update but no update was logged Device is monitoring as expected
1059546 | 198928 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-01 10:17 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/01/2024 10:14:05 AM, Baghirzade, Nara) Checked network settings for the correct DNS primary and server - all good. However when switched to patient monitoring, the overview was running on local. Double checked the network settings and everything was correct. Ran the setup and it would fail to connect to mshphiixpri server. Called Kamran for support, he also tried to run the setup which would fail and had worked on the backend of the server which had made the overview connect successfully. Will monitor this overview. Vacuumed the PICIX. (TC 02/01/2024 10:16:37 AM, Rhuland, Courtney) Checked network settings for the correct DNS primary and server - all good. However when switched to patient monitoring, the overview was running on local. Double checked the network settings and everything was correct. Ran the setup and it would fail to connect to mshphiixpri server. Called Kamran for support, he also tried to run the setup which would fail and had worked on the backend of the server which had made the overview connect successfully. Will monitor this overview. Vacuumed the PICIX.
1059547 | 198929 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:31 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:31:09 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:31:16 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059548 | 198930 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:29 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:28:17 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:28:28 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059549 | 198931 | PM | ACTIV | ISSUE | 2023-11-01 03:20 |  | PM-MAJOR - Annual PM For: Central Monitor
1059550 | 198932 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-01 10:14 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/01/2024 10:13:25 AM, Baghirzade, Nara) Checked network settings for the correct DNS primary and server - all good. Vacuumed the PICIX. Overview fully functional. (TC 02/01/2024 10:13:32 AM, Rhuland, Courtney) Checked network settings for the correct DNS primary and server - all good. Vacuumed the PICIX. Overview fully functional.
1059551 | 198933 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-12 09:13 | PM-MAJOR - Annual PM For: Central Monitor | (TC 12/08/2023 09:12:53 AM, Rendon Lopez , Elizabeth) Vacuumed central. PM completed, all tests passed
1059552 | 198934 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:38 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:37:26 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:37:35 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059553 | 198935 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:38 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:38:03 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:38:10 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059554 | 198936 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:33 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:32:09 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:32:17 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059555 | 198937 | PM | CANCL | CANCL | 2023-11-01 03:20 |  | PM-MAJOR - Annual PM For: Central Monitor
1059556 | 198938 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:35 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:34:08 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:34:18 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059557 | 198939 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-02-14 12:34 | PM-MAJOR - Annual PM For: Central Monitor | (TC 02/14/2024 12:33:19 PM, Baghirzade, Nara) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring. (TC 02/14/2024 12:33:25 PM, Rhuland, Courtney) PM Completed. Vacuumed inside/outside PIICIX. Confirmed that PIICIX is back on monitoring.
1059873 | 198940 | PM | CNP | ENL | 2023-11-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1059941 | 198941 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-01-31 09:06 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/26/2024 09:05:30 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1060155 | 198942 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-04 09:38 | Battery Cycling | (TC 12/04/2023 09:37:41 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060156 | 198943 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-23 10:23 | Battery Cycling | (TC 11/23/2023 10:22:46 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete Capacity 4/5 LEDs
1060157 | 198944 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-04 09:37 | Battery Cycling | (TC 12/04/2023 09:36:47 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060195 | 198945 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-23 10:22 | Battery Cycling | (TC 11/23/2023 10:19:34 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete Capacity 4/5 LEDs
1060198 | 198946 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-30 09:11 | Battery Cycling | (TC 11/30/2023 09:10:49 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060336 | 198947 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-30 09:12 | Battery Cycling | (TC 11/30/2023 09:11:27 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1060337 | 198948 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-29 08:30 | Battery Cycling | (TC 11/29/2023 08:29:24 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1060338 | 198949 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-29 08:32 | Battery Cycling | (TC 11/29/2023 08:31:26 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1060901 | 198950 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-28 08:47 | Battery Cycling | (TC 11/28/2023 08:46:16 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete Capacity 4/5 LEDs
1060902 | 198951 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-29 08:29 | Battery Cycling | (TC 11/29/2023 08:28:57 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1061254 | 198952 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-23 10:24 | Battery Cycling | (TC 11/23/2023 10:23:48 AM, Canjimavadivel, Kamallesh ) Battery conditioning complete Capacity 3/5 LEDs
1062777 | 198953 | PM | ACTIV | ISSUE | 2023-11-01 03:20 |  | Planned Maintenance
1062796 | 198954 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-01-12 14:18 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 01/12/2024 02:17:24 PM, Rhuland, Courtney) PM complete See attached checklist for details
1063331 | 198955 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-20 11:48 | PM - Annual | (TC 12/20/2023 08:45:00 AM, Fan, Tak) PM completed. Gnd R = 87 mohm Max A = 1.7 / 82.5 / 163.5 uA RM6C (257046KR3) = 0.5 / 4.2 uA C1-6-D (410203YP4) = 0.4 / 2.3 uA C2-9-D (407509YP9) = 0.5 / 1.8 uA (TC 01/03/2024 03:08:44 PM, Fan, Tak) Replaced OLED monitor.
1063358 | 198956 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-11 15:07 | PM-MAJOR - Annual PM For: Central Monitor | (TC 12/11/2023 03:02:25 PM, Canjimavadivel, Kamallesh ) Completed patch updates and rebooted central
1063808 | 198957 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-20 15:13 | PM - Annual | (TC 12/20/2023 03:10:53 PM, Fan, Tak) PM completed as per checklist. Gnd R = 84 mohm Max A = 1.8 / 77.3 / 152.9 uA C2-9-D (417012YP2) = 0.4 / 87.4 uA (Needs replacement) RAB6 (1094320WX2) = 0.4 / 2.5 uA C1-6-D (418024YP6) = 0.4 / 2.1 uA C4-8-D (160728PD2) = 0.4 / 1.6 uA (TC 01/02/2024 10:45:40 AM, Fan, Tak) C2-9-D (417012YP2) replaced by 483791YP0 on Jan 2, 2023.
1063934 | 198958 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-11-29 14:36 | PM - Annual | Injector PM done -Flex Ring calibrated -Potentiometer calibrated -Pressure calibrated -Touchscreen calibrated -System passed all functional tests
1063935 | 198959 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2024-04-26 14:35 | PM - Annual | Injector PM done -Flex Ring calibrated -Potentiometer calibrated -Pressure calibrated -Touchscreen calibrated -System passed all functional tests -Injector had to be taken to PMH shop to complete the PM.
1064496 | 198960 | PM | ACTIV | READY | 2023-11-01 03:20 |  | Battery Cycling
1064497 | 198961 | PM | CNP | EIU | 2023-11-01 03:20 | 2024-09-19 11:27 | Battery Cycling
1065405 | 198962 | PM | ACTIV | ISSUE | 2023-11-01 03:20 |  | PM-MAJOR-Annual PM For: Rehab Chair
1066939 | 198963 | PM | CMPLT | CMPLT | 2023-11-01 03:20 | 2023-12-22 08:43 | PM-MAJOR - Annual PM For: Central Monitor | (TC 12/21/2023 08:39:50 AM, Canjimavadivel, Kamallesh ) Rebooted after update, cleaned Note: dome lights in front of each room flashed red as central was off
1057764 | 198964 | PM | CMPLT | CMPLT | 2023-11-01 08:20 | 2023-12-08 09:17 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/08/2023 09:16:29 AM, Canjimavadivel, Kamallesh ) Completed patch install and vacuumed PICiX
| 198965 | PA | CANCL | CANCL | 2023-11-01 08:21 |  | Realize Medical Elucis Software
1067968 | 198966 | PM | CMPLT | CMPLT | 2023-11-01 08:22 | 2023-12-18 11:25 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 12/11/2023 03:08:02 PM, Canjimavadivel, Kamallesh ) Rebooted central, some updates still pending after reboot. (TC 12/15/2023 11:24:38 AM, Cortez, John) Rebooted
1055788 | 198967 | CM | CMPLT | CMPLT | 2023-11-01 09:38 | 2023-11-01 09:39 | Electrodes not connected | (TC 11/01/2023 09:38:14 AM, Reid, Emmerson ) Went to check the defib and saw that it did not have any pads. I went to the storage room and could not find any so I took an extra pad from the other defib. Then, did the 30J test.
501173 | 198968 | CM | CMPLT | CMPLT | 2023-11-01 10:15 | 2023-11-01 10:18 | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 probe artifact c5-1 | (TC 11/01/2023 10:17:42 AM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 probe artifact c5-1 MSH 15 SPP IU22 FAULTY PROBE C5-1 SN# B1K79X SITE 556348 REF# CHRIS 0122530220 | MSH 15 SPP PROCEDURE ROOM Dorothy Large Procedure Room 501173 probe artifact c5-1 MSH 15 SPP IU22 FAULTY PROBE C5-1 SN# B1K79X SITE 556348 REF# CHRIS 0122530220
32510 | 198969 | CM | CMPLT | CMPLT | 2023-11-01 12:27 | 2023-11-07 11:48 | MSH DR1 SIEMENS DMLA LICENCE | (TC 11/01/2023 12:26:44 PM, Whittington, Scott) MSH DR1 SIEMENS DMLA LICENCE (TC 11/07/2023 11:48:01 AM, Whittington, Scott) MSH DR1 SIEMENS DMLA LICENCE INSTALLED | MSH DR1 SIEMENS DMLA LICENCE From: LAPINS, HARALD <harald.lapins@siemens-healthineers.com> Sent: November 1, 2023 1:13 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: NOSEWORTHY, GEORGE <george.noseworthy@siemens-healthineers.com>; Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Uptime.Ca <uptime.ca@siemens.com> Subject: [External] RE: MSH ARISTOS LICENCE KEYS SITE 6026 & 6027 Hi Scott Please see your updated keys. Thanks Harald host ID 0A6E9398 0F6627DA1BA847E8744C host ID 0A6E9791 5FD6E70AD3BFA7CAEC17 From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Wednesday, November 1, 2023 12:25 PM To: Lapins, Harald <harald.lapins@siemens-healthineers.com> Cc: Noseworthy, George <george.noseworthy@siemens-healthineers.com>; Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Uptime.Ca <uptime.ca@siemens.com> Subject: MSH ARISTOS LICENCE KEYS SITE 6026 & 6027 Hi Harald I confirmed with Matt that the licenses have been paid in May. Could you please forward licence keys for host ID 0A6E9398 host ID 0A6E9791 Thanks Scott
32520 | 198970 | CM | CMPLT | CMPLT | 2023-11-01 12:29 | 2023-11-20 08:30 | MSH DR2 SIEMENS DMLA LICENCE | (TC 11/01/2023 12:28:48 PM, Whittington, Scott) MSH DR2 SIEMENS DMLA LICENCE (TC 11/07/2023 11:49:13 AM, Whittington, Scott) MSH DR2 SIEMENS DMLA LICENCE INSTALLED | MSH DR2 SIEMENS DMLA LICENCE From: LAPINS, HARALD <harald.lapins@siemens-healthineers.com> Sent: November 1, 2023 1:13 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: NOSEWORTHY, GEORGE <george.noseworthy@siemens-healthineers.com>; Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Uptime.Ca <uptime.ca@siemens.com> Subject: [External] RE: MSH ARISTOS LICENCE KEYS SITE 6026 & 6027 Hi Scott Please see your updated keys. Thanks Harald host ID 0A6E9398 0F6627DA1BA847E8744C host ID 0A6E9791 5FD6E70AD3BFA7CAEC17 From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Wednesday, November 1, 2023 12:25 PM To: Lapins, Harald <harald.lapins@siemens-healthineers.com> Cc: Noseworthy, George <george.noseworthy@siemens-healthineers.com>; Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca>; Uptime.Ca <uptime.ca@siemens.com> Subject: MSH ARISTOS LICENCE KEYS SITE 6026 & 6027 Hi Harald I confirmed with Matt that the licenses have been paid in May. Could you please forward licence keys for host ID 0A6E9398 host ID 0A6E9791 Thanks Scott
| 198971 | CM | CMPLT | CMPLT | 2023-11-01 13:00 | 2023-11-01 13:00 | Biomed meetig | Biomed meeting
1066935 | 198972 | CM | CMPLT | CMPLT | 2023-11-01 13:22 | 2024-02-14 11:39 | system is sent for repair | (TC 11/01/2023 01:22:10 PM, Samanian, Kamran) device is sent for repair, KJV tracking# 398779
| 198973 | CM | CMPLT | CMPLT | 2023-11-01 13:26 | 2024-04-26 11:15 | BP reading high multiple NICU rooms | (TC 11/01/2023 01:26:02 PM, Rhuland, Courtney) Reports from NICU staff of BP reading abnormally high Attempted to recreate with simulator but could not Talked with multiple nurses and deducted that the cuffs used are the issue Contacted Philips to see if backordered cuffs are available Will contact materials management to order Philips cuffs if they are back in stock
| 198974 | CM | CMPLT | CMPLT | 2023-11-01 13:31 | 2023-11-01 13:32 | Meeting | (TC 11/01/2023 01:31:05 PM, Rhuland, Courtney) Huddle
67796 | 198975 | CM | CMPLT | CMPLT | 2023-11-01 13:38 | 2023-11-01 13:50 | MSL Cable loose | (TC 11/01/2023 01:38:21 PM, Reid, Emmerson ) Went up to PACU, MSL cable was loose. Working now.
| 198976 | CM | CMPLT | CMPLT | 2023-11-01 14:29 | 2023-11-01 14:59 | Accessories charge back | (TC 11/01/2023 02:29:08 PM, Canjimavadivel, Kamallesh ) ICU requested adapters as they went missing. Provided 2 adapters
| 198977 | CM | CMPLT | CMPLT | 2023-11-01 14:55 | 2023-11-01 14:57 | Leaking blood pressure cuff manual BP | (TC 11/01/2023 02:55:37 PM, Rhuland, Courtney) Checked device and found leak in cuff Brought a cuff back from biomed and issue resolved Functional check - OK
65417 | 198978 | CM | CMPLT | CMPLT | 2023-11-01 14:58 | 2023-11-01 15:02 | Wrong door | (TC 11/01/2023 02:58:08 PM, Rhuland, Courtney) Came down with the wrong side door Replaced door with new door - replaced recall latches Functional check - OK
| 198979 | CM | CMPLT | CMPLT | 2023-11-01 16:07 | 2023-12-04 09:36 | misc biomed - nov 2023 | (TC 11/01/2023 04:06:44 PM, Cortez, John) huddle (TC 11/01/2023 04:07:11 PM, Cortez, John) Returned UPSs and Masimo viewstation to emergency power outlets (TC 11/03/2023 03:57:49 PM, Cortez, John) Verified on the unit that there's enough clearance between procuity beds and secure connects, for Adeel/Samir (TC 11/06/2023 03:04:20 PM, Cortez, John) philips huddle (TC 11/06/2023 03:06:09 PM, Cortez, John) Draeger online modules (TC 11/20/2023 04:06:34 PM, Cortez, John) Delivered training AED to 10S (Tas, CNS) (TC 11/20/2023 04:07:14 PM, Cortez, John) huddle (TC 11/22/2023 03:42:40 PM, Cortez, John) systems commitee meeting (TC 11/27/2023 02:47:15 PM, Cortez, John) Helped Samir replace door on AED cabinet (tunnel to TGH) (TC 11/29/2023 04:13:10 PM, Cortez, John) huddle (TC 11/29/2023 04:13:32 PM, Cortez, John) Went with Kam and Kamran to meet with ICU regarding Cerner shutdown and manual admitting of patients to monitors (TC 11/29/2023 04:13:32 PM, Cortez, John) Went with Kam and Kamran to meet with ICU regarding Cerner shutdown and manual admitting of patients to monitors
| 198980 | CM | CMPLT | CMPLT | 2023-11-01 17:51 | 2023-11-01 17:51 | MSH IR BIOPSY GUNS ENGRAVE WITH REF# | (TC 11/01/2023 05:51:18 PM, Dolinar, John) MSH IR BIOPSY GUNS ENGRAVE WITH REF# | MSH IR BIOPSY GUNS ENGRAVE WITH REF#
1057056 | 198981 | CM | CMPLT | CMPLT | 2023-11-01 17:54 | 2023-12-14 10:21 | MSH BI RAD REVIEW ROOM 12th Floor - North Side WORKSTATION UPS'S BATTERIES REACHING END OF LIFE | (TC 11/01/2023 05:54:54 PM, Dolinar, John) MSH BI RAD REVIEW ROOM 12th Floor - North Side WORKSTATION UPS'S BATTERIES REACHING END OF LIFE SPEC OUT BATTERIES REQUIRED FROM APC AND PROVIDE IVY WITH INFO FOR HER TO ORDER (TC 12/14/2023 10:22:57 AM, Dolinar, John) Batteries replaced. | MSH BI RAD REVIEW ROOM 12th Floor - North Side WORKSTATION UPS'S BATTERIES REACHING END OF LIFE. Batteries replaced.
| 198982 | CM | CMPLT | CMPLT | 2023-11-02 06:41 | 2023-11-02 06:42 | Biomed huddle | (TC 11/01/2023 06:41:25 AM, Baghirzade, Nara) Biomed huddle.
1061333 | 198983 | CM | CMPLT | CMPLT | 2023-11-02 06:48 | 2023-11-13 11:25 | Screen randomly goes blank during use | (TC 11/01/2023 06:48:28 AM, Baghirzade, Nara) Sarabeth and another nurse reported that the VAC screen randomly goes blank and they are not able to change any settings however the VAC still continues to run and suction. They switched the VAC off/on, runs for a while and the screen goes blank again. Could not find anything pertaining to screen timeouts in the settings menu. Contacted KCI for technical support and troubleshooting but the VAC is on a patient. Will troubleshoot it with Sarabeth's supervision and permission. Confirmation # for troubleshooting: 12928713 (TC 11/02/2023 02:15:47 PM, Baghirzade, Nara) Met with Sarabeth to discuss the VAC. Overnight there were no complaints however she was not entirely sure. Expressed to her that KCI and I want to do troubleshooting on the device. Also let her know my concerns of the continuous use while the device is experiencing technical issues. Recommended that Sarabeth swaps out the VAC with another one and send this one down to biomed for troubleshooting. Still waiting for the VAC. (TC 11/13/2023 11:24:08 AM, Baghirzade, Nara) 3M was unable to troubleshoot over the phone and highly recommended to send the unit to 3M for inspection. Coordinated pickup. Pick up confirmation #: 12949325
| 198984 | CM | CMPLT | CMPLT | 2023-11-02 06:52 | 2023-11-02 14:08 | Broken monitor display | (TC 11/01/2023 06:51:55 AM, Baghirzade, Nara) Gaby found that the monitor display for the bodybox is broken off the mount. For now the monitor is sitting on an angle but still functional. Will look into repurposing a monitor display or purchasing a new one. (TC 11/02/2023 02:07:37 PM, Baghirzade, Nara) Assisted Gaby with helpdesk inquiry. As soon as a ticket number was generated, I was able to locate Jigar. He replaced the monitor. We tested, all functional.
| 198985 | CM | CMPLT | CMPLT | 2023-11-02 06:56 | 2023-11-02 06:57 | NICU NIBP cable | (TC 11/02/2023 06:56:12 AM, Baghirzade, Nara) Tested the NIBP cable from NICU. Cable is functional. No physical damage on the cable or the connector.
| 198986 | CM | CMPLT | CMPLT | 2023-11-02 09:07 | 2024-04-10 14:26 | Sterile water for NICU | (TC 11/03/2023 11:17:27 AM, Rhuland, Courtney) SC #1003094138 (TC 11/17/2023 07:57:24 AM, Rhuland, Courtney) PO #4150266898 (TC 12/06/2023 07:41:11 AM, Rhuland, Courtney) SC #1003123283
1065159 | 198987 | CM | CMPLT | CMPLT | 2023-11-02 09:19 | 2023-11-02 09:20 | Unable to ID clinician | (TC 11/02/2023 09:18:37 AM, Rhuland, Courtney) Scanned test barcode and scanner works Went up to unit and tested with nurse badge and could not ID any nurses Pushed conifguration again and issue resolved Functional check - OK
67947 | 198988 | CM | CMPLT | NFF | 2023-11-02 09:22 | 2023-11-02 09:24 | "Humidity not working" | (TC 11/02/2023 09:22:03 AM, Rhuland, Courtney) Ran incubator for an hour with analyzer and calibrated humidity Humidity was working No fault found
1061202 | 198989 | CM | CMPLT | CMPLT | 2023-11-02 10:16 | 2023-11-02 10:17 | ESU was not working | (TC 11/02/2023 10:16:12 AM, Reid, Emmerson ) Said ESU was not working. The power plug was loose. I pushed it back in and it was working. They said this was reaccurig, so I pull some extra slack to help prevent it in the future.
1061216 | 198990 | CM | CMPLT | CMPLT | 2023-11-02 10:18 | 2023-11-02 13:42 | Issue with the Camera | (TC 11/02/2023 10:20:41 AM, Reid, Emmerson ) Dr was complaining the scope was narrowing the FoV. We replaced the camera head and the Dr said it was better but not great. He made it though the case but wants me to help setup the next case. (TC 11/02/2023 01:35:43 PM, Reid, Emmerson ) Went up for the beginning of the case. It was working.
67964 | 198991 | CM | CMPLT | CMPLT | 2023-11-02 11:10 | 2023-11-02 11:13 | Broken parts | (TC 11/02/2023 11:11:37 AM, Rhuland, Courtney) Broken pieces came down inside omnibed, but no pieces missing from device. Replaced incorrect inside door latch Calibrated humidity Functional check - OK
66717 | 198992 | PM | CMPLT | CMPLT | 2023-11-02 11:22 | 2023-11-02 11:34 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 11/02/2023 11:22:32 AM, Rhuland, Courtney) Updated device description & manufacturer Replaced enclosure Replaced standby switch Light intensity - 76.5 uW/cm^2-nm Electrical safety - OK
1057292 | 198993 | CM | CMPLT | CMPLT | 2023-11-02 12:28 | 2023-11-03 10:58 | Repurposing from BH | (TC 11/03/2023 10:55:26 AM, Rhuland, Courtney) Changed asset tag for 60 Murray Performed functional check Replaced missing hose Brought it to 60 Murray family medicine clinic
| 198994 | CM | CMPLT | CMPLT | 2023-11-02 13:33 | 2023-11-02 13:34 | Bulb replacement on exam light | (TC 11/02/2023 01:33:00 PM, Rhuland, Courtney) Replaced bulb on Welch Allyn 48740 exam light @OPG
67933 | 198995 | CM | CMPLT | CMPLT | 2023-11-02 13:55 | 2023-11-02 13:58 | Broken latches on door | (TC 11/02/2023 01:55:05 PM, Baghirzade, Nara) Replaced 2 broken latches. Ran a functional test however the RH% failed, was reading too low. Calibrated the RH sensor, ran another functional for 40 minutes, and RH % failed again, this time it was reading too high. Looking at previous WOs, the RH sensor was previously calibrated with the salt vile. Replaced the RH sensor. (Already charged to NICU) Functional test after RH sensor replacement was conducted and completed by Courtney. (TC 11/02/2023 01:57:20 PM, Rhuland, Courtney) Functional tests passed and completed. RH sensor was calibrated. Set temp to 35C and humidity to 75%; measured 35C and 75%. Incubator is fully functional.
1054999 | 198996 | CM | CMPLT | CMPLT | 2023-11-02 14:03 | 2023-11-02 14:05 | Reported NIBP discepancies | (TC 11/02/2023 02:02:44 PM, Baghirzade, Nara) Calibrated zero point and span point. Passed calibration test. Passed functional tests. Followed up on an e-mail chain with NICU nursing staff and RT.
1061200 | 198997 | CM | CMPLT | CMPLT | 2023-11-02 14:05 | 2023-11-02 14:05 | Reported NIBP discrepancies | (TC 11/02/2023 02:05:12 PM, Baghirzade, Nara) Calibrated zero point and span point. Passed calibration test. Passed functional tests. Followed up on an e-mail chain with NICU nursing staff and RT.
501106 | 198998 | CM | CMPLT | CMPLT | 2023-11-02 14:06 | 2023-11-03 14:39 | MSH 5 US IU22 MAIN MONITOR NOT COMING ON CONSISTENTLY, WHEN ON, PURPLE IMAGE, THEN NO IMAGE | (TC 11/02/2023 02:24:16 PM, Whittington, Scott) MSH 5 US IU22 MAIN MONITOR NOT COMING ON CONSISTENTLY, WHEN ON, PURPLE IMAGE, THEN NO IMAGE msh 5 us iu22 monitor site 553758 REF# ANTHONY case #0122537687 shipping to TGH (TC 11/03/2023 02:28:33 PM, Whittington, Scott) TC 11/02/2023 02:24:16 PM, Whittington, Scott) MSH 5 US IU22 MAIN MONITOR NOT COMING ON CONSISTENTLY, WHEN ON, PURPLE IMAGE, THEN NO IMAGE msh 5 us iu22 monitor site 553758 REF# ANTHONY case #0122537687 shipping to TGH sn# AN040814007077 RMA R438998935 | MSH 5 US IU22 MAIN MONITOR NOT COMING ON CONSISTENTLY, WHEN ON, PURPLE IMAGE, THEN NO IMAGE msh 5 us iu22 monitor site 553758 REF# ANTHONY case #0122537687 shipping to TGH
1054998 | 198999 | CM | CMPLT | CMPLT | 2023-11-02 14:07 | 2023-11-02 14:07 | Reported NIBP discrepancies | (TC 11/02/2023 02:06:43 PM, Baghirzade, Nara) Calibrated zero point and span point. Passed calibration test. Passed functional tests. Followed up on an e-mail chain with NICU nursing staff and RT.
1055000 | 199000 | CM | CMPLT | CMPLT | 2023-11-02 14:07 | 2023-11-02 14:07 | Reported NIBP discrepancies | (TC 11/02/2023 02:07:16 PM, Baghirzade, Nara) Calibrated zero point and span point. Passed calibration test. Passed functional tests. Followed up on an e-mail chain with NICU nursing staff and RT.
1067050 | 199001 | PM | CMPLT | CMPLT | 2023-11-02 14:08 | 2023-11-02 14:18 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 11/02/2023 02:08:17 PM, Cortez, John) Bed+Intellidrive batteries were replaced 60 mOhms, 22 uAmps Pendant missing clip Connector/cable loose on right side of foot section loose PM passed
32510 | 199002 | CM | CMPLT | CMPLT | 2023-11-02 14:16 | 2023-11-03 09:47 | MSH DR1 GRID GUIDE (ANNUCIATOR ACTIVATOR) BROKEN | (TC 11/02/2023 02:16:34 PM, Whittington, Scott) MSH DR1 GRID GUIDE (ANNUCIATOR ACTIVATOR) BROKEN | MSH DR1 GRID GUIDE (ANNUCIATOR ACTIVATOR) BROKEN
1054984 | 199003 | CM | CMPLT | CMPLT | 2023-11-03 07:01 | 2023-11-03 07:03 | BP cuff keeps inflating non-stop | (TC 11/03/2023 07:01:59 AM, Baghirzade, Nara) NIBP hose was leaking. Replaced the NIBP hose with a spare I had. Functional tests of SpO2, NIBP, and temperature passed.
| 199004 | CM | CMPLT | CMPLT | 2023-11-03 07:45 | 2023-11-03 07:45 | biomed meeting | biomed meeting
| 199005 | CM | CMPLT | CMPLT | 2023-11-03 08:05 | 2023-11-02 08:05 | worked on TMS | worked on TMS
1064771 | 199006 | PM | CMPLT | CMPLT | 2023-11-03 09:03 | 2023-11-03 09:03 | PM MAJOR - ANNUAL PM for: Insuflator | (TC 11/03/2023 09:02:51 AM, Reid, Emmerson ) Turned on and tested flow and occlusion. PM Complete.
| 199007 | RDS | CMPLT | CMPLT | 2023-11-03 09:06 | 2023-11-03 09:08 | Weekly Rounds: 18 South and POD A | (TC 11/03/2023 09:05:57 AM, Baghirzade, Nara) Weekly rounds in 18 South and POD A completed. See procedure.
| 199008 | RDS | CMPLT | CMPLT | 2023-11-03 09:08 | 2023-11-03 09:09 | Weekly Rounds: NICU | (TC 11/03/2023 09:08:09 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure.
| 199009 | RDS | CMPLT | CMPLT | 2023-11-03 09:10 | 2023-11-03 09:11 | Weekly Rounds: RESUS | (TC 11/03/2023 09:09:57 AM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure.
1054294 | 199010 | CM | CMPLT | CMPLT | 2023-11-03 09:12 | 2023-11-07 10:40 | Probe artifact | (TC 11/02/2023 09:12:20 AM, Fan, Tak) Retrieved probe and tested. Damaged crystals. Ordered replacement. Order# 800491309, RMA: 50026255. (TC 11/07/2023 10:39:06 AM, Fan, Tak) Replaced probe. There seems to be 2 probes with the same ME#. Confirmed 2 probes exist at SPP. Split other probe into 1068132.
| 199011 | CM | CMPLT | CMPLT | 2023-11-03 09:13 | 2023-11-03 09:16 | Weekly Rounds: 11 South | (TC 11/03/2023 09:13:12 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure.
1056502 | 199012 | CM | CMPLT | CMPLT | 2023-11-03 09:29 | 2023-11-14 08:06 | Blank screen, not functional | (TC 11/03/2023 09:28:47 AM, Baghirzade, Nara) During rounds, I noticed a Masimo device with a note stating that the SpO2 module is not working. Brought it down to biomed and booted the Root + Radical; root booted up properly however the radical screen went blank with all buttons continously flashing. Tried to restart it but same outcome. Removed the battery, docked to Root, but still the same outcome. Checked the pin connectors, all good. Requested service from Masimo. (TC 11/03/2023 11:37:29 AM, Baghirzade, Nara) Contacted masimo for support; they requested a no-charge po. SC 1003094282 (TC 11/14/2023 08:02:41 AM, Baghirzade, Nara) Received Radical 7 from Masimo. PO 7500197413 (TC 11/14/2023 08:06:23 AM, Baghirzade, Nara) Sent radical to Masimo. RMA: 1042730 KJV courier #: 400918
1061209 | 199013 | CM | CMPLT | CMPLT | 2023-11-03 10:42 | 2023-11-03 10:44 | Annual Inspection | (TC 11/03/2023 10:42:34 AM, Reid, Emmerson ) Checked all the booms. Checked the monitors and needed to adjust one of the handles as it was loose. iPad was charging, lights were working. Inspection complete.
1061215 | 199014 | CM | CMPLT | CMPLT | 2023-11-03 10:45 | 2023-11-03 10:46 | Annual Inspection | (TC 11/03/2023 10:44:58 AM, Reid, Emmerson ) Checked all the booms. Checked the monitors and needed to adjust two of the handles as they were loose. iPad was charging, lights were working. Inspection complete.
| 199015 | CM | CMPLT | CMPLT | 2023-11-03 11:05 | 2023-11-03 11:07 | Changing connectors on BP cuff | (TC 11/03/2023 11:04:50 AM, Rhuland, Courtney) Went to diabetes clinic with connectors, but they needed to change double tube to single Got single tube from biomed and brought cuff and a few spare single tubes back to clinic
1047472 | 199016 | CM | CMPLT | CMPLT | 2023-11-03 11:14 | 2023-12-01 10:34 | ordering parts for VN500 | ordering VN500 parts for RT. Received parts. PO#7202312951
| 199017 | CM | CMPLT | CMPLT | 2023-11-03 11:37 | 2023-11-03 11:37 | worked on VN500 parts ordering | worked on VN500 parts ordering
| 199018 | CM | CMPLT | CMPLT | 2023-11-03 12:37 | 2023-11-03 12:43 | Reported that Bay A and Bay F have SpO2 connectivity issues | (TC 11/03/2023 12:40:29 PM, Baghirzade, Nara) May called about SpO2 connectivity issues at Bay A and F. Found that on the measurement selection, the SpO2 on X2 and on the standalone module had the same parameter name. With the same parameter name, it causes conflicts for measurements. Changed the parameter names to: X2 = SpO2 SpO2 module = SpO2pr Checked all 6 bays, all SpO2 modules and X2 SpO2s are functional. (TC 11/03/2023 12:42:22 PM, Rhuland, Courtney) May called about SpO2 connectivity issues at Bay A and F. Found that on the measurement selection, the SpO2 on X2 and on the standalone module had the same parameter name. With the same parameter name, it causes conflicts for measurements. Changed the parameter names to: X2 = SpO2 SpO2 module = SpO2pr Checked all 6 bays, all SpO2 modules and X2 SpO2s are functional.
502044 | 199019 | CM | CMPLT | CMPLT | 2023-11-03 15:53 | 2023-11-03 15:04 | Temperature if going up. | (TC 11/03/2023 03:52:43 PM, Green, Paul) MRT contacted us and temperature is going above 17C. Went to site and cleaned fans and filters. Water on software not indicating low, but opened covers anyway. Sensor for water stuck in "ON" postion, but resivour bone dry. Filled up water and allowed to circulate. Had module errors which required opening detector more than once. Several different modules would fail, requiring reseating. Now all ok, but humidity high. Will monitor as we do not have any dessicant onsite.
52252 | 199020 | CM | CMPLT | CMPLT | 2023-11-06 07:48 | 2023-11-06 11:37 | Created by Failed PM WO # 198724 : PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/06/2023 11:29:11 AM, Baghirzade, Nara) Replaced the RH sensor. PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 66.1 uA Power cord resistance = 58 mohms | (TC 11/03/2023 02:47:00 PM, Baghirzade, Nara) PM failed during functional testing of RH and temp. 9V rechargeable battery was replaced. Flip door was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured very low (74%); calibrated RH sensor, measured around 76%. Looking at past WOs, the RH sensor has not been replaced, only re-calibrated. Will replace sensor and run a functional test.
1059746 | 199021 | CM | CMPLT | CMPLT | 2023-11-06 08:47 | 2023-11-14 10:05 | Not turning on | (TC 10/30/2023 08:47:06 AM, Rendon Lopez , Elizabeth) Faulty main board. Inquired for repair, sending out for repair is more expensive than purchasing a new device. Notified Clinical staff, waiting for instructions to proceed (TC 11/08/2023 08:52:31 AM, Rendon Lopez , Elizabeth) Ordering new doppler for replacement (TC 11/13/2023 10:04:39 AM, Rendon Lopez , Elizabeth) Incoming inspection new device, changed SN, warranty updated to 2025
1058192 | 199022 | CM | CMPLT | CMPLT | 2023-11-06 08:55 | 2023-11-06 08:56 | Dose cord not working | (TC 11/06/2023 08:55:01 AM, Reid, Emmerson ) Used a flat head to bend the cables back into place. Turned it on and the dose cord is working now.
1051701 | 199023 | CM | CMPLT | CMPLT | 2023-11-06 08:56 | 2023-11-06 08:57 | Not working | (TC 10/31/2023 08:56:08 AM, Rendon Lopez , Elizabeth) Device malfunction. Retired, replaced with spare 1051709
1051709 | 199024 | CM | CMPLT | CMPLT | 2023-11-06 08:57 | 2023-11-06 08:58 | Performance check | (TC 10/31/2023 08:57:22 AM, Rendon Lopez , Elizabeth) Device back into service for 15M, tested-OK
1059238 | 199025 | CM | CMPLT | CMPLT | 2023-11-06 08:58 | 2023-11-06 08:59 | Check function | (TC 10/31/2023 08:58:23 AM, Rendon Lopez , Elizabeth) Check function-OK Returned to 15M
1051269 | 199026 | CM | CMPLT | CMPLT | 2023-11-06 08:59 | 2023-11-06 09:00 | Malfunction | (TC 10/31/2023 08:59:25 AM, Rendon Lopez , Elizabeth) Battery malfunction and physical damage. Unrepairable, retire
1055159 | 199027 | CM | CMPLT | CMPLT | 2023-11-06 09:04 | 2023-11-06 09:05 | Channel error | (TC 10/31/2023 09:04:29 AM, Rendon Lopez , Elizabeth) Downloaded error log. Fatal bottle pressure sensor failure multiple times. Replaced pressure sensor.
| 199028 | RDS | CMPLT | CMPLT | 2023-11-06 09:09 | 2023-11-06 09:50 | Weekly Checks For PACU | (TC 11/06/2023 09:49:12 AM, Reid, Emmerson ) Weekly check complete. Also was called to check the ECG on Philips monitors in Bay 36 and 37. Both were working properly
69276 | 199029 | CM | CMPLT | CMPLT | 2023-11-06 09:11 | 2023-11-06 09:12 | Channel error | (TC 11/01/2023 09:11:22 AM, Rendon Lopez , Elizabeth) Downloaded error log, confirmed, fatal bottle pressure sensor error. Replaced pressure sensor, calibrated, check-in tests, all passed
1068108 | 199030 | ACC | CMPLT | CMPLT | 2023-11-06 09:15 | 2024-01-19 15:52 | Incoming inspection | (TC 11/02/2023 09:14:53 AM, Rendon Lopez , Elizabeth) Unpacked warmer, entered in TMS. Found bed assembly and wheels had physical damaged, bed assembly will have to be replaced. Waiting for getinge response (TC 01/19/2024 03:50:59 PM, Rendon Lopez , Elizabeth) Received new warmer. Incoming inspection, all tests passed, changed SN in TMS, attached delivery documents
| 199031 | CM | CMPLT | CMPLT | 2023-11-06 09:20 | 2023-12-05 09:24 | Servo n batteries for PM | Received
1061211 | 199032 | CM | CMPLT | NFF | 2023-11-06 09:50 | 2023-11-06 09:51 | Need to logon | (TC 11/06/2023 09:50:40 AM, Reid, Emmerson ) Called because they couldn't logon to Olympus VMC. They just needed to press "logon"
66431 | 199033 | CM | CMPLT | CMPLT | 2023-11-06 10:31 | 2023-11-06 10:53 | not turning on | (TC 11/06/2023 10:43:22 AM, Scero, Tomlyn) Sally called to say the telepak wasn't turning on. I changed the battery adaptor and installed batteries. Ran the device for a bit on the simulator to confirm functionality. Returned the device to the CCU when complete. Battery adapter already charged to CCU.
| 199034 | RDS | CMPLT | CMPLT | 2023-11-06 10:31 | 2023-11-06 10:43 | Weekly checks in CCU | (TC 11/06/2023 10:42:18 AM, Scero, Tomlyn) Weekly checks complete. 30J tests done on all defibs
1050073 | 199035 | CM | CMPLT | CMPLT | 2023-11-06 10:36 | 2023-11-06 11:12 | Vitals tower smashed | (TC 11/06/2023 11:08:59 AM, Canjimavadivel, Kamallesh ) Replaced back housing and SpO2 module from spares. Replaced SpO2 sensor. Swapped roll stands | Housing was broken
1062324 | 199036 | CM | CMPLT | CMPLT | 2023-11-06 10:54 | 2023-11-07 15:11 | Spilled fluid and now no display | (TC 11/06/2023 10:54:03 AM, Ferreira, Qwynn) Reset the bed Cleaned fluid Made sure no remaining fluid near connections Checked the connections Functional check - OK Showed connor some of the smart bed features and troubleshooting tips
| 199037 | CM | CMPLT | CMPLT | 2023-11-06 11:38 | 2023-11-06 11:38 | Changed time on transport incubators | (TC 11/06/2023 11:37:47 AM, Baghirzade, Nara) Updated time on all transport incubators in NICU and NICU storage with Samir.
1062968 | 199038 | CM | CMPLT | CMPLT | 2023-11-06 12:06 | 2023-12-01 11:09 | not working, needs to be sent for repair | (TC 11/30/2023 11:06:54 AM, Huang, Jessica) Picked up the device in OR and replaced the lamp the device was still not working contacted the vendor and shipped it for warranty repair PSU and lamp were replaced by the vendor got it back and sent it back to OR (TC 12/01/2023 11:06:54 AM, Huang, Jessica) Picked up the device in OR and replaced the lamp the device was still not working contacted the vendor and shipped it for warranty repair PSU and lamp were replaced by the vendor got it back and sent it back to OR
| 199039 | CM | CMPLT | CMPLT | 2023-11-06 12:27 | 2023-11-20 08:28 | MSH OPG 3 SPP ASTRAIA DOWN | (TC 11/06/2023 12:28:08 PM, Whittington, Scott) MSH OPG 3 SPP ASTRAIA DOWN | MSH OPG 3 SPP ASTRAIA DOWN
| 199040 | CM | CMPLT | CMPLT | 2023-11-06 12:38 | 2023-11-13 11:24 | Sonomed Teamviewer installation for remote support | (TC 11/06/2023 12:38:08 PM, Baghirzade, Nara) Worked with Kamran. Opened services list -> found WLAN disabled -> enabled. Restarted the PC. Used my phone hotspot to connect to PC wifi; downloaded teamviewer. Kamran confirmed that application is working by remoting in. Asked Zain to let me know time/date when the vendor wants to remote in to check the system. Disabled the WLAN and will re-enable when the vendor needs to connect. Confirmed that there is no wifi connectivity before shutting the PC down. (TC 11/07/2023 01:55:46 PM, Baghirzade, Nara) Erik Cordova from sonomed remoted in to the sonomed system. Tested the A-probe and UBM; suspected board issue at the lemo connnector side but cannot completely confirm it. Sonomed requested the clinic to send the unit over to their facility for evaluation. Erik to send an email with updates and options for the clinic.
| 199041 | RDS | CMPLT | CMPLT | 2023-11-06 13:18 | 2023-11-06 13:20 | Weekly Checks For OR | (TC 11/06/2023 01:19:16 PM, Reid, Emmerson ) Rounds complete
| 199042 | RDS | CMPLT | CMPLT | 2023-11-06 13:18 | 2023-11-06 13:35 | Weekly check for ED
1054597 | 199043 | CM | CMPLT | CMPLT | 2023-11-06 13:33 | 2023-12-06 11:47 | MSH Echo - Charging Dock Broken | (TC 11/06/2023 01:33:00 PM, Chau, Ken) Charging dock broken and requires replacement, trying to contact GE to see if its replacable. 1-550129356741 not covered under any contract/plan, quote for repair is required. (TC 12/05/2023 11:45:15 AM, Chau, Ken) Returned unit to MSH Echo as requested. unit still functional and just broken charger dock. Department bought extra batteries and charging dock
501596 | 199044 | CM | CMPLT | CMPLT | 2023-11-06 14:09 | 2023-11-06 14:09 | MSH 15 SPP IU22 PROBE C5-1 SN# F09Y9J ACCEPTANCE | (TC 11/06/2023 02:09:10 PM, Whittington, Scott) MSH 15 SPP IU22 PROBE C5-1 SN# F09Y9J ACCEPTANCE | MSH 15 SPP IU22 PROBE C5-1 SN# F09Y9J ACCEPTANCE
64078 | 199045 | CM | CMPLT | CMPLT | 2023-11-06 14:23 | 2023-11-06 14:24 | Bed wont tilt | (TC 11/06/2023 02:22:46 PM, Baghirzade, Nara) Greased the tilt. Attached it to the bed. Moves smoothly and stays in place once affixed to a specific spot. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and RH 75%. Incubator is fully functional.
1050314 | 199046 | CM | CMPLT | NFF | 2023-11-06 14:25 | 2023-11-06 14:31 | Temp probe not working | (TC 11/06/2023 02:29:29 PM, Canjimavadivel, Kamallesh ) Tested probe with tester and calibration plug. Values within spec. No fault found
65923 | 199047 | CM | CMPLT | CMPLT | 2023-11-06 14:27 | 2023-11-06 14:28 | Channel error 354.6770 | (TC 11/06/2023 02:27:01 PM, Baghirzade, Nara) Re-seated pressure sensor. Check-in tests passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. Ran the pump at varying rates for over an hour; no issues or errors. Infusion completed successfully.
1057577 | 199048 | CM | CMPLT | CMPLT | 2023-11-06 14:51 | 2023-11-06 14:52 | Clear tape used as cover, infection control issue | (TC 11/06/2023 02:51:12 PM, Rendon Lopez , Elizabeth) Cleaned tape of back, tested functions, working properly.
1047061 | 199049 | CM | CMPLT | CMPLT | 2023-11-06 15:00 | 2023-11-06 15:09 | Not working | (TC 11/06/2023 03:01:13 PM, Rendon Lopez , Elizabeth) Tested, NFF
502044 | 199050 | CM | CMPLT | CMPLT | 2023-11-06 15:12 | 2023-11-06 15:15 | high humidity | 11/6/2023 - Replace the head and tube desicant on the device. Perform power up test and check humidity - passed. Device is in working condition.
1054418 | 199051 | CM | CMPLT | CMPLT | 2023-11-06 15:20 | 2023-11-23 16:28 | Physical damage | (TC 11/06/2023 03:20:24 PM, Rendon Lopez , Elizabeth) Send out for repair (TC 11/20/2023 04:27:20 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1049641 | 199052 | CM | CMPLT | CMPLT | 2023-11-06 15:23 | 2023-11-23 16:23 | Not working | (TC 11/06/2023 03:23:21 PM, Rendon Lopez , Elizabeth) Physical damage. Send out for repair (TC 11/20/2023 04:22:52 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1057477 | 199053 | CM | CMPLT | CMPLT | 2023-11-06 15:25 | 2023-11-23 16:24 | Doesn't show uterine activity | (TC 11/06/2023 03:26:19 PM, Rendon Lopez , Elizabeth) Confirmed issue. Sending out for repair (TC 11/20/2023 04:23:28 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK Serial number changed
69285 | 199054 | CM | CMPLT | CMPLT | 2023-11-06 16:30 | 2023-11-06 16:35 | Channel error | (TC 11/06/2023 04:34:25 PM, Rendon Lopez , Elizabeth) Downloaded error log, several fatal bottle pressure sensor failure. Replaced pressure sensor, calibrated.
1062081 | 199055 | CM | CMPLT | CMPLT | 2023-11-06 16:36 | 2023-11-06 16:37 | Daily checks error | (TC 11/06/2023 04:36:24 PM, Rendon Lopez , Elizabeth) Checked bed, cleared errors, working properly
1062070 | 199056 | CM | CMPLT | CMPLT | 2023-11-06 16:37 | 2023-11-06 16:37 | Daily checks errors | (TC 11/06/2023 04:37:09 PM, Rendon Lopez , Elizabeth) Checked bed, cleared errors, working properly
| 199057 | RDS | CMPLT | CMPLT | 2023-11-07 07:33 | 2023-11-07 07:33 | weekly check for 14S and 16N | weekly check for 14S and 16N
1061214 | 199058 | CM | CMPLT | CMPLT | 2023-11-07 08:56 | 2023-11-07 08:57 | Table not working | (TC 11/07/2023 08:56:06 AM, Reid, Emmerson ) Had to check the table. Remote was stuck on one screen. I unplugged and plugged back in the remote to reset it. Working now.
1059821 | 199059 | CM | CMPLT | CMPLT | 2023-11-07 08:58 | 2024-06-06 15:04 | Torn table top pad | (TC 11/07/2023 09:03:52 AM, Reid, Emmerson ) Ordering replacement pads (TC 05/01/2024 01:46:18 PM, Reid, Emmerson ) Followed up with Mohawk as to why this is not completed yet (TC 06/06/2024 03:03:40 PM, Reid, Emmerson ) Recieved, at last.
68788 | 199060 | CM | CMPLT | CMPLT | 2023-11-07 08:58 | 2023-11-07 09:02 | Reported that iAware was not connecting | (TC 11/07/2023 08:58:25 AM, Baghirzade, Nara) Went to network settings and checked: Status = associated SSID = mshmed Connected to mshmed. No network connectivity issues found. Possibly could have been an issue with the syringe pump, not PCU. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 33.0 mohms Leakage current: 74.7 uA
| 199061 | RDS | CMPLT | CMPLT | 2023-11-07 09:02 | 2023-11-07 09:04 | Weekly Rounds: 18 South and PACU POD A | (TC 11/07/2023 09:02:39 AM, Baghirzade, Nara) Weekly rounds in 18 South and PACU POD A completed. See procedure for results.
| 199062 | RDS | CMPLT | CMPLT | 2023-11-07 09:04 | 2023-11-07 09:06 | Weekly Rounds: 11 South | (TC 11/07/2023 09:04:25 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 199063 | RDS | CMPLT | CMPLT | 2023-11-07 09:06 | 2023-11-07 09:37 | Weekly Rounds: NICU | (TC 11/07/2023 09:14:20 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure for results.
63217 | 199064 | CM | CMPLT | NFF | 2023-11-07 09:15 | 2023-11-07 09:17 | pump saying clog in line | (TC 11/07/2023 09:15:35 AM, Campbell, Connor) primed and performed a pump flow rate check pump delivered expected amount no fault found returned to service
1049901 | 199065 | CM | CMPLT | CMPLT | 2023-11-07 09:22 | 2023-11-14 16:53 | MSH Revo R61 - System battery issues | (TC 11/07/2023 09:22:15 AM, Chau, Ken) parts ordered, to be replaced when arrives (TC 11/08/2023 05:20:54 PM, Whittington, Scott) MSH Revo R61 - System battery issues SYSTEM BATTERY REPLACEMENT (TC 11/08/2023 05:22:35 PM, Dolinar, John) MSH Revo R61 - System battery issues SYSTEM BATTERY REPLACEMENT (TC 11/08/2023 05:24:05 PM, Fan, Tak) MSH Revo R61 - System battery issues SYSTEM BATTERY REPLACEMENT | MSH Revo R61 - System battery issues
68224 | 199066 | CM | CMPLT | CMPLT | 2023-11-07 09:22 | 2023-11-07 09:49 | No V lead on monitor | (TC 11/07/2023 09:48:36 AM, Scero, Tomlyn) No V lead trace on monitor. Replaced both trunk and leads to fix the issue.
| 199067 | CM | CMPLT | CMPLT | 2023-11-07 11:52 | 2023-11-08 13:42 | MSH ANGIO LEAD APRON RACK ASSLY | (TC 11/07/2023 11:52:43 AM, Whittington, Scott) MSH ANGIO LEAD APRON RACK ASSLY (TC 11/08/2023 01:41:50 PM, Whittington, Scott) MSH ANGIO LEAD APRON RACK ASSLY (TC 11/08/2023 01:42:04 PM, Dolinar, John) MSH ANGIO LEAD APRON RACK ASSLY | MSH ANGIO LEAD APRON RACK ASSLY
68043 | 199068 | CM | CMPLT | CMPLT | 2023-11-07 13:01 | 2023-11-07 13:10 | Not working | (TC 11/07/2023 01:01:34 PM, Canjimavadivel, Kamallesh ) Cuff not inflating, found leak. Tested with different cuff, working normal.
67143 | 199069 | CM | CMPLT | CMPLT | 2023-11-07 13:44 | 2023-11-07 13:51 | Not working | (TC 11/07/2023 01:45:06 PM, Canjimavadivel, Kamallesh ) Cuff not inflating, found leak. Replaced cuff and tested ok
| 199070 | CM | CMPLT | CMPLT | 2023-11-07 13:51 | 2023-12-08 14:49 | Beds for Neurology | (TC 11/07/2023 01:51:04 PM, Baghirzade, Nara) Met with Naz to discuss the need for 2 beds in Neurology; the current beds in neuro are end of support, outdated, and not working. One is a potential risk to patients. Naz was able to acquire 2 beds from the repurposing plan they have; we located the two beds in storage, moved other beds around to make these 2 beds accessible. As soon as the work is coordinated with Vinod, the old beds will be removed from neuro and replaced with the new stryker beds. Waiting on coordination with Vinod. (TC 12/08/2023 02:47:35 PM, Baghirzade, Nara) Bed move was completed. Old beds were placed on 18 and managed by support services for disposal. Naz confirmed with Ula and Michael Badour.
| 199071 | CM | CMPLT | CMPLT | 2023-11-07 13:53 | 2023-11-08 09:28 | Oto-ophthalmoscopes not working | (TC 11/07/2023 01:53:35 PM, Baghirzade, Nara) 1/2 otoscopes required a new bulb. Replaced the bulb (U3100). The other looks like it has a very depleted battery or needs a new battery entirely. Placed it to charge on the charging station overnight. Will check tomorrow and proceed as needed. Will order more batteries for ENT. (TC 11/08/2023 09:27:12 AM, Baghirzade, Nara) Charged overnight. Otoscope is fully functional. Suggested to Elena and Ana to make sure that the doctors, fellow, and themselves dock the handles properly, ensuring that it makes contact with the charger.
| 199072 | RDS | CMPLT | CMPLT | 2023-11-07 13:57 | 2023-11-07 13:59 | Weekly Rounds: RESUS | (TC 11/07/2023 01:57:25 PM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure.
68862 | 199073 | CM | CMPLT | CMPLT | 2023-11-07 13:59 | 2023-11-07 14:05 | Reported that battery is not charging | (TC 11/07/2023 02:03:51 PM, Baghirzade, Nara) Battery was depleted; left it to charge for a couple of hours. Noted that at the beginning it had less than 5 mins of run time on battery. After charging, it was up to 10 hours of battery run time. Ran the PCU on battery with a syringe pump for basic infusion for 3 hours; unit is completely functional on battery. Connected to ASM to check battery status; all battery values are within specs. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 48.0 mohms Leakage current: 63.5 uA
1062172 | 199074 | CM | CMPLT | CMPLT | 2023-11-07 14:37 | 2023-11-20 16:10 | Noise when head section moves up/down | (TC 11/07/2023 02:37:31 PM, Cortez, John) Confirmed issue, Actuator is damaged, will request service with Stryker (TC 11/20/2023 04:08:17 PM, Cortez, John) Mark (Stryker) repaired bed I swapped broken headboard with one from spare bed on 6th floor
1062309 | 199075 | CM | CMPLT | CMPLT | 2023-11-07 14:47 | 2023-11-07 14:48 | Time change | (TC 11/07/2023 02:47:09 PM, Baghirzade, Nara) Updated the time due to DST for the vent on top of the transport incubator.
1061217 | 199076 | CM | CMPLT | CMPLT | 2023-11-07 15:03 | 2023-11-07 15:05 | Doppler batteries dead | (TC 11/07/2023 03:03:26 PM, Reid, Emmerson ) Called to check out a doppler. Replaced 8 AA batteries. Working now.
| 199077 | CM | CMPLT | CMPLT | 2023-11-07 15:33 | 2023-11-07 15:34 | NICU huddle | (TC 11/07/2023 03:33:24 PM, Baghirzade, Nara) NICU huddle
| 199078 | PA | CMPLT | CMPLT | 2023-11-08 06:24 | 2023-11-13 11:35 | Ordering NIMH rechargeable batteries | (TC 11/08/2023 06:25:41 AM, Baghirzade, Nara) SC 1003099409 (TC 11/13/2023 11:34:20 AM, Baghirzade, Nara) PO 7202315209
1060000 | 199079 | CM | CMPLT | CMPLT | 2023-11-08 09:25 | 2023-11-08 09:27 | Client view down | (TC 11/08/2023 09:25:06 AM, Baghirzade, Nara) Screen was blank, could not get in to task manager, and mouse/keyboard were unresponsive. Rebooted the PC however on the first reboot, it was frozen again. Rebooted the PC again, and safetynet loaded, but it took longer than usual to boot up. 11 South PSN is functional.
1062142 | 199080 | CM | CMPLT | CMPLT | 2023-11-08 09:43 | 2023-12-21 09:19 | error code 3009-6-1-502 limited bed functionality | (TC 11/08/2023 09:42:46 AM, Baghirzade, Nara) Error code 3009-6-1-502 limited bed functionality. Power cycled the bed, did not mitigate the issue. Checked the settings and configurations, all good. Checked the lock system, none of the motion controls are locked. Created service request with Stryker. (TC 12/21/2023 09:17:58 AM, Baghirzade, Nara) Attached service report from Mark. Bed has been returned to service.
1058058 | 199081 | CM | CMPLT | CMPLT | 2023-11-08 09:47 | 2023-11-28 14:22 | 130ml/min leak, but passing self test | (TC 11/16/2023 02:20:02 PM, Scero, Tomlyn) Worked with Draeger tech support and they told me how to download and check for the actual leakage in the system during the self test. When the logs were reviewed it showed a leak just over 100 ml/min and 500 is acceptable. Draeger said it was safe to put into use.
1066840 | 199082 | CM | CMPLT | CMPLT | 2023-11-08 09:48 | 2023-11-08 09:57 | Small leak | (TC 11/07/2023 09:55:00 AM, Scero, Tomlyn) Cleaned APL and replaced seal ring that was damaged. Ran leak tests 10ml/min passed. Self test info - 100 and 24. Returned to zeus 1058058 in OR 12.
1064980 | 199083 | CM | CMPLT | CMPLT | 2023-11-08 09:49 | 2023-11-08 09:54 | Small leak | (TC 11/07/2023 09:52:16 AM, Scero, Tomlyn) Cleaned APL and replaced 2 sealing rings that looked damaged. Tested the device on the macine in RT. Leak 10ml/min Test info - 100 and 68 Tests passed but cleaning the valves might reduce the leak even more. Put on zeus in RT that's currently being PM'd and I will look at it again when I return from vacation.
68372 | 199084 | CM | CMPLT | CMPLT | 2023-11-08 10:38 | 2023-11-10 08:46 | Diva module failure | (TC 11/08/2023 10:38:38 AM, Rendon Lopez , Elizabeth) Went to OR to check device. AA had completely restarted device and changed diva module to port 2. Device is working for now, will come back and check logs (TC 11/08/2023 08:47:42 AM, Rendon Lopez , Elizabeth) Downloaded logs. Left Diva supply module leaking. Device is working with DIVA plugged in right slot, clinical staff is continuing to use device as is
66505 | 199085 | CM | CMPLT | CMPLT | 2023-11-08 10:41 | 2023-11-10 08:43 | Recorder not feeding paper | (TC 11/08/2023 10:40:54 AM, Rendon Lopez , Elizabeth) Confirmed issue. Removed device from room to replace recorder (TC 11/09/2023 08:42:54 AM, Rendon Lopez , Elizabeth) Replaced recorder with repaired one. Cleaned, calibrated, tested all functions-OK
1046292 | 199086 | CM | CMPLT | CMPLT | 2023-11-08 10:41 | 2023-11-08 10:42 | Requires calibration | (TC 11/08/2023 10:41:39 AM, Baghirzade, Nara) Calibrated the thermometer. Low target = 32.2C Measured = 32.3C High target = 40.6C Measured = 40.5C Thermometer is fully functional.
1046291 | 199087 | CM | CMPLT | CMPLT | 2023-11-08 10:43 | 2023-11-08 10:43 | Requires calibration | (TC 11/08/2023 10:42:58 AM, Baghirzade, Nara) Calibrated the thermometer. Low target = 32.2C Measured = 32.1C High target = 40.6C Measured = 40.5C Thermometer is fully functional.
1048264 | 199088 | CM | CMPLT | CMPLT | 2023-11-08 10:50 | 2023-11-08 10:51 | Unable to calibrate, damaged lens | (TC 11/08/2023 10:50:19 AM, Baghirzade, Nara) Lens is damaged -possibly due to the cleaning procedures. Unable to calibrate, failed every calibration attempt. Spoke to Sarabeth and she agreed to retire it as they have already purchased new replacement thermometers. Will dispose of it in e-waste.
| 199089 | CM | CMPLT | CMPLT | 2023-11-08 12:27 | 2023-11-08 12:27 | ESU Unit Discussion and training meeting | (TC 11/08/2023 12:27:03 PM, Chukri, Samir) meeting
| 199090 | CM | CMPLT | CMPLT | 2023-11-08 12:42 | 2023-11-07 12:42 | worked on PS500 order from RT | located PS500 skids. arranged installation with Rex, Steve and Salur.
1057513 | 199091 | CM | CMPLT | CMPLT | 2023-11-08 13:35 | 2023-12-08 08:23 | Ligature not recognized | (TC 11/08/2023 01:35:07 PM, Reid, Emmerson ) Ligature was not recognized during a case. I took it off the boom and tried with a test ligature, but was getting the same result. I have reached out to tech support for next steps. (TC 11/09/2023 10:04:27 AM, Reid, Emmerson ) Called with an issue of the position of the mobile cart. I told them this is an issue in progress and they said they would be able to make it though the case. (TC 12/08/2023 08:22:04 AM, Reid, Emmerson ) Tested with a demo ligasure wand, and it was working after the upgrade. Returned to the OR.
1068097 | 199092 | CM | CMPLT | CMPLT | 2023-11-08 13:43 | 2023-11-08 13:44 | Loaner pickup | (TC 11/08/2023 01:42:48 PM, Baghirzade, Nara) Picked up loaner VAC from Sri. Called 3M and coordinated a pick up. Attached pickup receipt.
67958 | 199093 | CM | CMPLT | CMPLT | 2023-11-08 13:47 | 2023-11-13 09:22 | Broken door latch | (TC 11/08/2023 01:46:46 PM, Baghirzade, Nara) Replaced the door latch. Ran a functional check with 75% RH and 35C temp; failed the functional check. Calibrated the RH twice. Calibrated RH with salt vial. RH still failing. Will replace RH sensor. (TC 11/13/2023 09:21:24 AM, Baghirzade, Nara) Replaced the RH sensor. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and RH measured at 75%. Incubator is fully functional.
| 199094 | CM | CMPLT | NFF | 2023-11-08 14:46 | 2023-11-08 14:48 | Thermometer check | (TC 11/08/2023 02:46:28 PM, Rhuland, Courtney) Requested to check their thermometers - stated readings were high Checked thermometers againsts calibrated biomed ones and temp was the same No fault found, returned to clinic
1061903 | 199095 | CM | CMPLT | CMPLT | 2023-11-08 15:55 | 2024-01-12 09:59 | Bed motion functions arent working | (TC 11/08/2023 03:55:12 PM, Cortez, John) Confirmed issue, Power cycling doesn't resolve problem, Error 3009-6-1-501 in logs Requested service from Stryker (TC 11/20/2023 09:58:34 AM, Cortez, John) Repaired by Mark, located on 12floor by service elevator
| 199096 | PA | CMPLT | CMPLT | 2023-11-09 09:40 | 2023-11-22 07:27 | Batteries for ICU beds | (TC 11/09/2023 09:44:49 AM, Canjimavadivel, Kamallesh ) Placed order for PM batteries (TC 11/22/2023 07:26:34 AM, Canjimavadivel, Kamallesh ) Received batteries, unboxed and placed in storage bin
1061211 | 199097 | CM | CMPLT | CMPLT | 2023-11-09 10:03 | 2023-11-09 10:04 | Issue Recording | (TC 11/09/2023 10:03:17 AM, Reid, Emmerson ) They were getting a blue screen when trying to record. I changed the video input for the recorder, as they had it set incorrectly. Working now.
1055790 | 199098 | CM | CMPLT | CMPLT | 2023-11-09 10:12 | 2023-11-09 10:13 | Not plugged in | (TC 11/09/2023 10:12:03 AM, Reid, Emmerson ) Power cord was lose in the back. Working now.
67963 | 199099 | CM | CMPLT | CMPLT | 2023-11-09 10:23 | 2023-11-09 10:25 | Bed tilt broken | (TC 11/09/2023 10:23:32 AM, Rhuland, Courtney) Repaired bed tilt handle Replaced missing side wall grommet Functional check - OK
530510 | 199100 | CM | CMPLT | CMPLT | 2023-11-09 10:31 | 2023-11-09 10:33 | Metal nipple broken off N2O tank | (TC 11/08/2023 10:31:25 AM, Rendon Lopez , Elizabeth) Removed from L&D, replaced N2O tank and sealing rings for all tanks. Tested function-OK
65978 | 199101 | CM | CMPLT | CMPLT | 2023-11-09 13:15 | 2023-11-09 13:16 | "unable to associate pump - Missing A" | (TC 11/09/2023 01:15:19 PM, Campbell, Connor) tested the IUI conntectors performed a check-in no fault found pump is wokring fine and associating with PC unit and other pumps as normal returned to service
1058272 | 199102 | CM | CMPLT | CMPLT | 2023-11-09 13:38 | 2023-11-13 14:11 | "battery malfunction" | (TC 11/10/2023 08:24:48 AM, Campbell, Connor) performed battery optimization twice battery capacity is too low, needs new battery waiting for batteries from a previous order to arrive (TC 11/13/2023 02:09:58 PM, Campbell, Connor) Found a spare new battery and replaced it battery is good, returned to service
502044 | 199103 | CM | CMPLT | CMPLT | 2023-11-09 14:39 | 2023-11-09 14:41 | Arm lock needs adjusting. | (TC 11/09/2023 02:38:51 PM, Green, Paul) Adjusted arm lock as customer noted it was bouncing out of the lock position. Adjusted and now seats better when in the locked postion.
1064355 | 199104 | CM | CMPLT | CMPLT | 2023-11-09 14:39 | 2023-11-09 14:52 | Missing clips | (TC 11/09/2023 02:40:09 PM, Canjimavadivel, Kamallesh ) Installed clips, tested functions with simulator
67957 | 199105 | CM | CMPLT | CMPLT | 2023-11-09 15:03 | 2023-11-09 15:05 | Humidity not working | (TC 11/09/2023 03:03:05 PM, Rhuland, Courtney) Humidity would not increase past 60% Noticed fan motor was making loud noise Reseated motor and fan and issue resolved Replaced humidifier Calibrated humidity with salt Functional check - OK
65905 | 199106 | CM | CMPLT | CMPLT | 2023-11-10 07:48 | 2023-11-10 07:49 | flange detect not working | (TC 11/10/2023 07:47:47 AM, Campbell, Connor) replaced flange detect sensor with a spare part performed a check-in flange detect works, check-in passed returned to service
65899 | 199107 | CM | CMPLT | NFF | 2023-11-10 08:31 | 2023-11-10 08:32 | "unable to detect the correct syringe size" | (TC 11/10/2023 08:30:55 AM, Campbell, Connor) checked error logs, there was a runtime error for syringe plunger position but no fatal error performed a check-in, passed tested multiple syringe sizes, all were recognized performed a test infusion no faults found returned to service
1048168 | 199108 | CM | CMPLT | CMPLT | 2023-11-10 08:44 | 2023-11-10 08:45 | Not working | (TC 11/09/2023 08:44:47 AM, Rendon Lopez , Elizabeth) Error P06, faulty temp probe. Replaced, tested-OK
1055644 | 199109 | CM | CMPLT | CMPLT | 2023-11-10 10:39 | 2023-12-28 10:52 | Damaged pendant and head board | (TC 11/10/2023 10:39:19 AM, Canjimavadivel, Kamallesh ) Placed a service request with Baxter. (TC 12/28/2023 10:50:16 AM, Canjimavadivel, Kamallesh ) Repair completed by vendor Dec 20, 2023. See attached report
502044 | 199110 | RQ | CMPLT | CMPLT | 2023-11-10 11:14 | 2023-11-10 14:56 | brakes on ERGO wonky | (TC 11/10/2023 02:54:32 PM, Green, Paul) adjusted brakes by removing detector. reinstalled detector. detector now had modules out. opened detector and reseated module 29. retapped the modules as tape seemed to be coming off. continued to adjust brake until satisfied with lever control. | brakes on ERGO wonky
1068156 | 199111 | ACC | CMPLT | CMPLT | 2023-11-10 11:29 | 2023-11-20 08:28 | MSH 5 US PROBE mC12-3 SN# B3FMDF BABY HEAD ACCEPTANCE | (TC 11/10/2023 11:29:04 AM, Whittington, Scott) MSH 5 US PROBE mC12-3 SN# B3FMDF BABY HEAD ACCEPTANCE | MSH 5 US PROBE mC12-3 SN# B3FMDF BABY HEAD ACCEPTANCE
501132 | 199112 | CM | CMPLT | CMPLT | 2023-11-10 12:41 | 2023-11-20 08:27 | MSH FRMR TWH IU22 PM BY PHILIPS | (TC 11/10/2023 12:40:51 PM, Whittington, Scott) MSH FRMR TWH IU22 PM BY PHILIPS | MSH FRMR TWH IU22 PM BY PHILIPS
501133 | 199113 | CM | CMPLT | CMPLT | 2023-11-10 12:42 | 2023-11-20 08:27 | MSH FRMR TWH IU22 PM BY PHILIPS | (TC 11/10/2023 12:42:28 PM, Whittington, Scott) MSH FRMR TWH IU22 PM BY PHILIPS | MSH FRMR TWH IU22 PM BY PHILIPS
1064351 | 199114 | CM | CMPLT | CMPLT | 2023-11-10 12:45 | 2023-11-20 08:27 | MSH ICU SPARQ PM BY PHILIPS | (TC 11/10/2023 12:45:04 PM, Whittington, Scott) MSH ICU SPARQ PM BY PHILIPS | MSH ICU SPARQ PM BY PHILIPS
1064352 | 199115 | CM | CMPLT | CMPLT | 2023-11-10 12:46 | 2023-11-20 08:26 | MSH ICU SPARQ PM B PHILIPS | (TC 11/10/2023 12:46:04 PM, Whittington, Scott) MSH ICU SPARQ PM B PHILIPS | MSH ICU SPARQ PM B PHILIPS
1057687 | 199116 | CM | CMPLT | CMPLT | 2023-11-10 12:59 | 2023-11-10 13:00 | Updating membrane date | (TC 11/10/2023 12:59:26 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1057665 | 199117 | CM | CMPLT | CMPLT | 2023-11-10 13:01 | 2023-11-10 13:02 | Updating membrane date | (TC 11/10/2023 01:01:20 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1057678 | 199118 | CM | CMPLT | CMPLT | 2023-11-10 13:03 | 2023-11-10 13:03 | Update membrane change date | (TC 11/10/2023 01:02:38 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1054046 | 199119 | CM | CMPLT | CMPLT | 2023-11-10 13:03 | 2023-11-10 13:04 | Update membrane change date | (TC 11/10/2023 01:03:28 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1054043 | 199120 | CM | CMPLT | CMPLT | 2023-11-10 13:04 | 2023-11-10 13:05 | Update membrane change date | (TC 11/10/2023 01:04:24 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1057691 | 199121 | CM | CMPLT | CMPLT | 2023-11-10 13:05 | 2023-11-10 13:05 | Update membrane change date | (TC 11/10/2023 01:05:08 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1057670 | 199122 | CM | CMPLT | CMPLT | 2023-11-10 13:05 | 2023-11-10 13:06 | Update membrane change date | (TC 11/10/2023 01:05:38 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1057685 | 199123 | CM | CMPLT | CMPLT | 2023-11-10 13:06 | 2023-11-10 13:06 | Update membrane change date | (TC 11/10/2023 01:06:07 PM, Rhuland, Courtney) Updated membrane change date to 31 days
1057963 | 199124 | CM | CMPLT | CMPLT | 2023-11-10 13:06 | 2023-11-10 13:07 | Update membrane change date | (TC 11/10/2023 01:06:33 PM, Rhuland, Courtney) Updated membrane change date to 31 days
| 199125 | CM | CMPLT | CMPLT | 2023-11-10 13:38 | 2024-01-17 12:15 | Spare incubator parts
66468 | 199126 | CM | CMPLT | CMPLT | 2023-11-10 14:18 | 2023-11-10 14:18 | Nebulizer option did not pass pre-use check | tested unit 3 times with full shut down and boot up and pre-use check. Functional. Informed Rex.
| 199127 | CM | CMPLT | CMPLT | 2023-11-10 14:20 | 2023-11-10 14:20 | PS500 installation | Worked with Salur and Rex. PS500 units come with LI-ion batteries. Saltur will confirm that they can be used in Canada.
502044 | 199128 | CM | CMPLT | CMPLT | 2023-11-10 15:05 | 2023-11-10 15:06 | delete patient studies on Broker due to error | (TC 11/10/2023 03:05:12 PM, Green, Paul) deleted patient study on broker and WCH thinking system due to entry error as per MSH. they will resend when corrections made.
1048472 | 199129 | CM | CMPLT | CMPLT | 2023-11-13 07:25 | 2023-11-13 07:26 | Detector not acquiring images | (TC 11/10/2023 07:25:05 AM, Fan, Tak) Checked system. Connected. May have battery conact issue. Replaced contacts. Sent email to Matt.
| 199130 | RDS | CMPLT | CMPLT | 2023-11-13 07:42 | 2023-11-13 07:42 | Weekly checks for 14N and 16N | Weekly checks for 14N and 16N
62696 | 199131 | CM | CMPLT | NFF | 2023-11-13 07:55 | 2023-11-13 08:30 | V6 not working | (TC 11/13/2023 07:54:59 AM, Reid, Emmerson ) Tested on a simulator and found no issue. Tried moving the cables and it was still working. NFF
1067437 | 199132 | FC | CMPLT | CMPLT | 2023-11-13 09:20 | 2023-11-13 09:20 | QC checklist | (TC 11/13/2023 09:19:51 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1067436 | 199133 | FC | CMPLT | CMPLT | 2023-11-13 09:20 | 2023-11-13 09:21 | QC checklist | (TC 11/13/2023 09:20:30 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
67928 | 199134 | PA | CMPLT | CMPLT | 2023-11-13 09:30 | 2023-11-23 14:50 | Ordering incubator hood | (TC 11/13/2023 09:34:01 AM, Baghirzade, Nara) SC 1003104444 (TC 11/23/2023 02:49:48 PM, Rhuland, Courtney) Hood arrived - PO 7202317769
52181 | 199135 | PA | CMPLT | CMPLT | 2023-11-13 09:35 | 2023-11-23 07:22 | Ordering humidity sensors | (TC 11/13/2023 09:36:18 AM, Baghirzade, Nara) SC 1003104446 (TC 11/23/2023 07:20:39 AM, Baghirzade, Nara) PO 7202317769
1068077 | 199136 | CM | CMPLT | CMPLT | 2023-11-13 09:59 | 2023-11-30 08:09 | Cables Cut | (TC 11/13/2023 09:59:17 AM, Reid, Emmerson ) Went with Tommy, the cables from one of the opthalmoscope heads and the power cable have been cut somehow. Taking down to biomed to repair. (TC 11/17/2023 11:22:42 AM, Canjimavadivel, Kamallesh ) Contacted PCM for further details. Used extra parts from older device to replaced damaged components. Emailed vendor for quote for ophthalmoscope head. (TC 11/21/2023 10:29:35 AM, Canjimavadivel, Kamallesh ) Received quote, placed order for ophthalmoscope head (TC 11/30/2023 08:08:20 AM, Canjimavadivel, Kamallesh ) Installed new head and bulb, brought device up to B16 and installed on the headwall.
1054215 | 199137 | CM | CMPLT | CMPLT | 2023-11-13 10:09 | 2023-11-13 14:08 | not working | (TC 11/13/2023 02:06:56 PM, Scero, Tomlyn) Thermometer still had security mode activated. Deactivated SEC mode and returned the thermometer to CCU | SEC - message displayed
67914 | 199138 | CM | CMPLT | CMPLT | 2023-11-13 10:11 | 2023-11-13 10:13 | missing west door | (TC 11/13/2023 10:11:29 AM, Campbell, Connor) replaced door with a spare, spare door needed the latch recall done as well functional check complete returned to service Functional tests passed and comple (TC 11/13/2023 10:40:58 AM, Baghirzade, Nara) Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75% RH. Incubator is fully functional.
1057921 | 199139 | CM | CMPLT | CMPLT | 2023-11-13 10:29 | 2023-11-13 10:29 | not powering up | tested. not functional. removed from service. informed Rex.
1053380 | 199140 | CM | CMPLT | CMPLT | 2023-11-13 10:32 | 2023-11-13 10:32 | expr. casette valve experied | replaced expiratory casettte valve. updated date.
1059828 | 199141 | PM | CMPLT | CMPLT | 2023-11-13 10:35 | 2023-11-13 10:35 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/13/2023 10:35:02 AM, Campbell, Connor) Disinfected, Cleaned, rinsed and refilled
52181 | 199142 | CM | CMPLT | CMPLT | 2023-11-13 10:43 | 2023-11-13 10:44 | Reported not working | (TC 11/13/2023 10:42:39 AM, Baghirzade, Nara) Could not find fault with the incubator. Scale cable was stuck under the door and preventing the door from opening; got the cable out and checked it; no damage on the cable, looks good. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C on all temperature points and measured 77% as the highest recorded RH, still within the operating margin. Incubator is fully functional.
1059778 | 199143 | FC | CMPLT | CMPLT | 2023-11-13 11:13 | 2023-11-13 11:13 | QC checklist | (TC 11/13/2023 11:12:40 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
| 199144 | RDS | CMPLT | CMPLT | 2023-11-13 11:21 | 2023-11-13 11:22 | Weekly Checks For PACU | (TC 11/13/2023 11:21:57 AM, Reid, Emmerson ) Complete
1049204 | 199145 | CM | CMPLT | CMPLT | 2023-11-13 11:23 | 2023-11-13 11:24 | Central was not connecting for Pod B | (TC 11/13/2023 11:23:02 AM, Reid, Emmerson ) Possibly due to a temporary power outage, the central for Pod B disconnected from the rooms and the server. I reconnected it. Working now.
1061333 | 199146 | CM | CMPLT | CMPLT | 2023-11-13 11:26 | 2024-01-05 07:06 | Screen goes blank during use | (TC 11/13/2023 11:26:41 AM, Baghirzade, Nara) Accidently closed the previous WO: 198983 Issue is ongoing. (TC 12/20/2023 02:25:07 PM, Baghirzade, Nara) Repair estimate was approved by Katie (11N PCM). SC 1003134088 (TC 01/05/2024 07:05:18 AM, Baghirzade, Nara) Received VAC from KCI. Attached service/QC report. Returned VAC back to 11N and followed up with Sarabeth. | Accidently closed the previous WO: 198983 Issue is ongoing.
| 199147 | PA | CMPLT | CMPLT | 2023-11-13 12:04 | 2023-11-20 14:23 | Realize Medical Elucis VR Software | Completed and charged. Invoices are stored under T drive 3D printing
| 199148 | CM | CMPLT | CMPLT | 2023-11-13 12:41 | 2023-11-13 12:41 | Biomed meeting | Biomed meeting
| 199149 | CM | CMPLT | CMPLT | 2023-11-13 12:42 | 2023-11-13 12:42 | RT meeting | RT meeting
1064980 | 199150 | CM | CMPLT | CMPLT | 2023-11-13 14:30 | 2023-11-13 14:34 | 64 ml leak during selftest | (TC 11/13/2023 02:32:48 PM, Scero, Tomlyn) Took breathing system apart fully. cleaned all parts including all valves. was able to et the leak down to 44 ml/min during the self test which is more acceptable.
32560 | 199151 | CM | CMPLT | CMPLT | 2023-11-13 20:05 | 2023-11-13 20:03 | S/D CON connectivity/REC application problem reported | S/D CON connectivity/RECON application problem reported Service completed troubleshooting of S/D con Service replaced system drive/D-con/backup completed CT phantom tested OK
1048018 | 199152 | CM | CMPLT | CMPLT | 2023-11-14 07:52 | 2023-12-20 08:48 | Black out areas on the display | (TC 11/13/2023 02:52:00 PM, Baghirzade, Nara) Tested the #0 blade (SN: TN1326) by plugging it in to the CMAC; there are black out spots on the display. Suspected lens damage on the blade. Tested a different blade and no issues with the display or blade. Tested the cable and blades with a CMAC system in the OR and can confirm that #0 blade (SN: TN1326) is the issue. Contacted Kim to check if blade is still under warranty. Returned the CMAC to Rex and sent an update to Michelle and Rex regarding the blade issue. (TC 11/15/2023 10:38:15 AM, Baghirzade, Nara) Kim requested to have the CMAC blade sent in for inspection. Generated a no-charge PO. (TC 11/20/2023 07:21:05 AM, Baghirzade, Nara) Sending blade for inspection. KJV Waybill# 401921 (TC 11/20/2023 07:21:05 AM, Baghirzade, Nara) Sending blade for inspection. PO 7500198016 KJV Waybill# 401921 (TC 12/14/2023 08:50:30 AM, Baghirzade, Nara) Repair quote was sent by vendor; Steve Arai approved the repair. Attached approved repair quote to this WO. Sent e-mail to Mohawkmedbuy to update the cost on the PO. (TC 12/20/2023 08:46:47 AM, Baghirzade, Nara) Received blade from repair. Will drop it off with Rex in NICU RT (TC 12/20/2023 08:47:22 AM, Baghirzade, Nara) PO 7500198016
1055803 | 199153 | CM | CMPLT | CMPLT | 2023-11-14 08:10 | 2023-11-14 08:11 | No AC Power | (TC 11/14/2023 08:10:06 AM, Reid, Emmerson ) Checked the defib and it was plugged in. I ran 30J test and let is transmit
1056502 | 199154 | CM | CMPLT | CMPLT | 2023-11-14 08:21 | 2023-11-14 08:42 | Received replacement radical 7 | (TC 11/14/2023 08:40:39 AM, Baghirzade, Nara) Radical replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root/Radical via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed and placed asset tag. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
67928 | 199155 | CM | CMPLT | CMPLT | 2023-11-14 09:03 | 2023-11-14 09:04 | Broken porthole hinge | (TC 11/14/2023 09:02:55 AM, Baghirzade, Nara) Replaced porthole hinge (west door). Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75% RH. Incubator is fully functional.
64181 | 199156 | CM | CMPLT | CMPLT | 2023-11-14 10:05 | 2023-11-14 10:08 | Recorder not working | (TC 11/14/2023 10:05:21 AM, Reid, Emmerson ) Recorder was not working. Checked the paper and devices on central. Seemed to be working but was not, yellow light was still flashing on the front. I restarted the central and the printer connected and light turned green. Ran a test print. Working now.
| 199157 | RDS | CMPLT | CMPLT | 2023-11-14 10:26 | 2023-11-14 10:26 | Weekly checks in CCU | (TC 11/14/2023 10:25:46 AM, Scero, Tomlyn) Weekly checks complete in CCU. 30J tests passed on defibs.
1068196 | 199158 | CM | CMPLT | CMPLT | 2023-11-14 11:25 | 2024-04-03 09:50 | 2 YR PM | (TC 04/03/2024 09:49:35 AM, Scero, Tomlyn) PM parts replaced and tests complete. Passed with no leaks. <100 mmhg 19 mmhg total leak
| 199159 | CM | CMPLT | CMPLT | 2023-11-14 12:48 | 2023-11-14 12:50 | Stryker beds for Neurology | (TC 11/14/2023 12:48:19 PM, Baghirzade, Nara) Coordinated with Naz to have 2 beds for Neurology dept. The current beds are old, end of support, and not functional. Naz was able to acquire 2 beds from their repurposing plan. Coordinated with support services to have 2 porters move the 2 beds in and out. Old beds are being disposed of by support services. Showed Vinod how to use the bed and all the controls on it. Modified the bed configuration to remove wifi and secure connect as both are not applicable.
67936 | 199160 | CM | CMPLT | CMPLT | 2023-11-14 12:51 | 2023-11-14 12:52 | Reported that humidity was not working | (TC 11/14/2023 12:51:11 PM, Baghirzade, Nara) No fault found. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75% RH. Incubator is fully functional.
1049215 | 199161 | CM | CMPLT | CMPLT | 2023-11-14 12:54 | 2023-11-16 14:14 | Getting email alerts | (TC 11/14/2023 12:54:23 PM, Reid, Emmerson ) Kam let me know he was recieving dozens of email alerts saying the central was down. I went to check and it was working fine. He was still getting the alerts so I went up and rebooted the central
| 199162 | CM | CMPLT | CMPLT | 2023-11-14 14:01 | 2023-11-14 14:02 | Meeting with Andrew | (TC 11/14/2023 02:01:27 PM, Baghirzade, Nara) Met with Andrew to go over the clinics inventory: -dental -ent -ophthalmology -pft Andrew and I then went to ENT and PFT to check inventory and discuss which equipment the clinics want to request for capital.
| 199163 | CM | CMPLT | CMPLT | 2023-11-14 15:37 | 2023-11-14 15:37 | NICU huddle | (TC 11/14/2023 03:37:05 PM, Baghirzade, Nara) NICU huddle
1046381 | 199164 | CM | CMPLT | CMPLT | 2023-11-14 15:55 | 2024-01-08 09:34 | Sherlock not working | (TC 11/13/2023 01:55:00 PM, Fan, Tak) Checked sherlock system. USB seems OK, spoke to techs, issue was with ECG. Couldn't confirm but checked board and pins. Probably worn down over the time. Turned the pins to help with contact. Contacted BD. | Could not reproduce problem. Tested OK.
502044 | 199165 | CM | CMPLT | CMPLT | 2023-11-14 16:14 | 2023-11-15 13:30 | brake adjustment | (TC 11/15/2023 01:27:34 PM, Green, Paul) Talked to Micheal from Digirad. Went to site and waited for time on system. When time permitted, adjusted brake as per Digirad. Had MRTs confirm functionality of arm brake.
1049901 | 199166 | CM | CMPLT | CMPLT | 2023-11-14 16:57 | 2023-11-14 16:57 | Tube head covers melted, streched exposure cord, detector cal required. |  | Replaced tube head covers, performed detector cal and replaced exposure cable.
| 199167 | CM | CMPLT | CMPLT | 2023-11-14 16:58 | 2023-11-14 16:58 | Site coverage
1054045 | 199168 | CM | CMPLT | CMPLT | 2023-11-15 09:19 | 2023-11-15 09:21 | Setting the membrane change interval | (TC 11/15/2023 09:19:18 AM, Baghirzade, Nara) Downloaded V-STATS software on biomed laptop as it was initially unavailable. Connected device to the software; configured the membrane change interval to 31 days and set it to the specific NICU membrane change profile. Confirmed on device that the interval has been changed to 31 days. (TC 11/15/2023 09:21:02 AM, Rhuland, Courtney) Downloaded V-STATS software on biomed laptop as it was initially unavailable. Connected device to the software; configured the membrane change interval to 31 days and set it to the specific NICU membrane change profile. Confirmed on device that the interval has been changed to 31 days.
1057689 | 199169 | CM | CMPLT | CMPLT | 2023-11-15 09:22 | 2023-11-15 09:22 | Setting the membrane change interval | (TC 11/15/2023 09:21:38 AM, Baghirzade, Nara) Downloaded V-STATS software on biomed laptop as it was initially unavailable. Connected device to the software; configured the membrane change interval to 31 days and set it to the specific NICU membrane change profile. Confirmed on device that the interval has been changed to 31 days. (TC 11/15/2023 09:21:44 AM, Rhuland, Courtney) Downloaded V-STATS software on biomed laptop as it was initially unavailable. Connected device to the software; configured the membrane change interval to 31 days and set it to the specific NICU membrane change profile. Confirmed on device that the interval has been changed to 31 days.
| 199170 | CM | CANCL | CANCL | 2023-11-15 10:37 |  | Sending in the CMAC blade for inspection | Duplicated WO generated. Cancelled this WO.
63876 | 199171 | CM | CMPLT | NFF | 2023-11-15 13:01 | 2023-11-15 14:22 | Skin probe reading wrong temperature | (TC 11/15/2023 01:01:12 PM, Rhuland, Courtney) Tag stated probe was reading 1 degree difference between set temp and probe Ran functional check at 36 deg and 75% humidity with skin probe Replaced broken light Skin probe - 35.8 deg Incubator - 36 deg Analyzer - 35.9 deg Humidity - 75.3% No Fault Found
| 199172 | CM | CMPLT | CMPLT | 2023-11-15 13:51 | 2023-11-15 13:52 | Meeting | (TC 11/15/2023 01:51:10 PM, Rhuland, Courtney) Capital meeting
501618 | 199173 | CM | CMPLT | CMPLT | 2023-11-16 06:19 | 2023-11-20 08:21 | MSH OPG SPP PROBE C5-1 SN# B1KX93 SITE# 53198439 | (TC 11/16/2023 06:21:10 AM, Whittington, Scott) MSH OPG SPP PROBE C5-1 SN# B1KX93 SITE# 53198439 | msh opg spp probe c5-1 sn# B1KX93 site 53198439 REF# CHRIS 0122665852
502049 | 199174 | CM | CMPLT | CMPLT | 2023-11-16 06:20 | 2023-11-20 08:25 | MSH OPG SPP PROBE C5-1 SN# B1KX93 SITE# 53198439 | (TC 11/16/2023 06:20:54 AM, Whittington, Scott) MSH OPG SPP PROBE C5-1 SN# B1KX93 SITE# 53198439 | msh opg spp probe c5-1 sn# B1KX93 site 53198439 msh opg spp probe c5-1 sn# B1KX93 site 53198439 REF# CHRIS 0122665852
1055166 | 199175 | CM | CMPLT | CMPLT | 2023-11-16 07:44 | 2023-11-16 07:46 | Bent/damaged male IUI connector | (TC 11/16/2023 07:43:54 AM, Campbell, Connor) replaced male IUI connector with a spare, no charge performed a check in as per ASM pump is functional, returned to service
1056665 | 199176 | CM | CMPLT | CMPLT | 2023-11-16 08:23 | 2023-11-16 08:28 | DVI Connector Broken | (TC 11/16/2023 08:23:27 AM, Reid, Emmerson ) The female screw for the DVI cable connector on the DVI out on the microscope has come out on one side. The cable is also damaged. I changed the cable. I was not able to fix the connector but it was stll securely connected. Tested and it is working.
1061205 | 199177 | CM | CMPLT | CMPLT | 2023-11-16 08:33 | 2023-11-16 08:34 | DVI Port on boom missing the female pin | (TC 11/16/2023 08:33:11 AM, Reid, Emmerson ) Replaced the pin, and tested. Working now.
1050904 | 199178 | CM | CMPLT | CMPLT | 2023-11-16 09:27 | 2023-11-16 09:29 | System not connecting to DICOM nodes | (TC 11/14/2023 09:27:11 AM, Fan, Tak) Contacted Mirza, submitted request for network jack change. (TC 11/15/2023 09:27:11 AM, Fan, Tak) Retrieved network info for change. (TC 11/15/2023 09:27:11 AM, Fan, Tak) DICOM spooler not sending. Reset spooler, jobs transferring.
| 199179 | CM | CMPLT | CMPLT | 2023-11-16 09:33 | 2023-11-16 09:33 | Incubator parts and shelves organizing
65121 | 199180 | PM | CMPLT | CMPLT | 2023-11-16 09:50 | 2023-11-16 10:25 | PM MAJOR - ANNUAL PM for: Infusion Pumps | (TC 11/16/2023 10:25:01 AM, Reid, Emmerson ) Downloaded firmware V040201 and Pushed APS IV PCA MAR 2022. PM Complete, see procedure for details.
1063746 | 199181 | PM | CMPLT | CMPLT | 2023-11-16 10:41 | 2023-11-16 10:42 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/16/2023 10:41:22 AM, Campbell, Connor) Drained, disinfected, cleaned, refilled
| 199182 | CM | CMPLT | CMPLT | 2023-11-16 11:30 | 2023-11-15 11:30 | Biomed meeting | biomed meeting
| 199183 | CM | CMPLT | CMPLT | 2023-11-16 12:44 | 2024-01-17 09:30 | Bunnell PM parts | (TC 11/16/2023 01:13:34 PM, Rendon Lopez , Elizabeth) SC 1003108157 Bunnell HFV PM parts (TC 11/20/2023 08:23:07 AM, Rendon Lopez , Elizabeth) Received all excep gas filter element Received all
1056568 | 199184 | CM | CMPLT | CMPLT | 2023-11-16 14:08 | 2023-11-16 14:09 | Battery died | (TC 11/16/2023 02:08:05 PM, Reid, Emmerson ) Replaced the battery and did battery check. Showed 4/5 bars
63875 | 199185 | CM | CMPLT | CMPLT | 2023-11-16 14:23 | 2024-01-18 10:08 | Humidifier failure | (TC 11/16/2023 02:26:23 PM, Rhuland, Courtney) Humidifier failure error when turned on Found heater cartridge to be faulty when measuring resistance Replaced with spare cartridge Replaced humidifier bottom gasket Replaced reservoir switch button Error cleared - running functional check (TC 11/17/2023 08:27:54 AM, Rhuland, Courtney) Ran for over an hour and humidity could not reach set level - waiting for new humidifier (TC 12/06/2023 02:54:35 PM, Rhuland, Courtney) Attempted to replace humidifier, chassis broke Began chassis replacement (TC 01/16/2024 03:13:38 PM, Rhuland, Courtney) Completed chassis replacement and replaced humidifier (TC 01/17/2024 02:31:56 PM, Rhuland, Courtney) Replaced hood Bed tilt would not lock - replaced brake pad assembly and ball bearing Replaced power cord Replaced door latch (TC 01/18/2024 10:03:50 AM, Rhuland, Courtney) Completed rebuild Humidifier test failed after calibration Replaced humidifier sensor Replaced broken light Replaced chipped inside door Functional check - OK
| 199186 | CM | CMPLT | CMPLT | 2023-11-16 14:26 | 2023-11-16 14:28 | Meeting with Andrew | (TC 11/16/2023 02:26:41 PM, Baghirzade, Nara) Met with Andrew to reitre some of the assets in ENT that we confirmed are no longer there. Went over the inventory list for Ophthalmology and Urology clinics. Went to the clinics physically to confirm if devices are there and which to retire/replace. (TC 11/16/2023 02:27:32 PM, Rhuland, Courtney) Went over the inventory list for Ophthalmology and Urology clinics. Went to the clinics physically to confirm if devices are there and which to retire/replace.
1059839 | 199187 | CM | CMPLT | CMPLT | 2023-11-16 15:20 | 2023-11-16 15:20 | not working | (TC 09/01/2023 03:20:00 PM, Huang, Jessica) got the call from OR and saying the device wasn't working replaced the tube, tested it, working fine
1060974 | 199188 | CM | CMPLT | CMPLT | 2023-11-16 15:26 | 2023-11-16 15:26 | bipolar foot pedal | (TC 09/04/2023 03:26:13 PM, Huang, Jessica) replaced the bipolar foot pedal
62485 | 199189 | CM | CMPLT | CMPLT | 2023-11-16 15:27 | 2023-11-16 15:27 | bipolar foot pedal | (TC 10/11/2023 03:33:12 PM, Huang, Jessica) replaced the bipolar foot pedal
501071 | 199190 | CM | CMPLT | CMPLT | 2023-11-16 15:29 | 2023-12-07 08:31 | CEOU RM 12 turns off spontaneously site 553550 | (TC 11/16/2023 03:28:44 PM, Whittington, Scott) CEOU RM 12 turns off spontaneously site 553550 | CEOU RM 12 turns off spontaneously site 553550
501072 | 199191 | CM | CMPLT | CMPLT | 2023-11-16 15:30 | 2024-02-26 13:30 | CEOU RM 5 study txfr issue | (TC 11/16/2023 03:30:23 PM, Whittington, Scott) CEOU RM 5 study txfr issue | CEOU RM 5 study txfr issue
501066 | 199192 | CM | CMPLT | CMPLT | 2023-11-16 15:31 | 2024-02-26 13:30 | CEOU RM 4 burner | (TC 11/16/2023 03:31:14 PM, Whittington, Scott) CEOU RM 4 burner | CEOU RM 4 burner
1054421 | 199193 | CM | CMPLT | CMPLT | 2023-11-16 15:34 | 2023-11-16 15:34 | replaced the SPO2 sensor
1046531 | 199194 | CM | CMPLT | CMPLT | 2023-11-16 15:35 | 2023-11-16 15:35 | replaced the battery
541265 | 199195 | CM | CMPLT | CMPLT | 2023-11-16 15:36 | 2023-11-16 15:36 | replaced the battery
| 199196 | CM | CMPLT | CMPLT | 2023-11-16 15:37 | 2023-11-16 15:37 | remote control sensor defect | (TC 10/09/2023 03:37:14 PM, Huang, Jessica) got a call in OR regarding the table remote swapped it in OR 12 and ordered the new one
1054070 | 199197 | CM | CMPLT | CMPLT | 2023-11-16 15:41 | 2023-11-16 15:41 | internal hardware failure | (TC 10/09/2023 03:41:27 PM, Huang, Jessica) contacted Baxter for tech support disconnected the battery and reconnected it back
| 199198 | CM | CMPLT | CMPLT | 2023-11-16 15:47 | 2023-11-16 15:47 | helping Nara to repair the photo therapy lights
| 199199 | CM | CMPLT | CMPLT | 2023-11-16 15:48 | 2023-11-16 15:48 | checking port information in the OR room
| 199200 | CM | CMPLT | CMPLT | 2023-11-16 15:49 | 2023-11-16 15:49 | checking the ECG device | (TC 09/05/2023 03:49:21 PM, Huang, Jessica) foroozan reported the ECG device wasnt working. checked it, working fine and inform them to keep the electrodes in the dry and close bag
1064170 | 199201 | CM | CMPLT | CMPLT | 2023-11-16 15:52 | 2023-11-16 15:52 | new incoming inspection
1064173 | 199202 | CM | CMPLT | CMPLT | 2023-11-16 15:53 | 2023-11-16 15:53 | new incoming inspection
1064169 | 199203 | CM | CMPLT | CMPLT | 2023-11-16 15:54 | 2023-11-16 15:54 | new incoming inspection
1064172 | 199204 | CM | CMPLT | CMPLT | 2023-11-16 15:56 | 2023-11-16 15:56 | new incoming inspection
1064171 | 199205 | CM | CMPLT | CMPLT | 2023-11-16 15:56 | 2023-11-16 15:56 | new incoming inspection
1064168 | 199206 | CM | CMPLT | CMPLT | 2023-11-16 16:03 | 2023-11-16 16:03 | new incoming inspection
1054071 | 199207 | CM | CMPLT | CMPLT | 2023-11-16 16:04 | 2023-11-16 16:04 | replaced the temperature probe
| 199208 | CM | CMPLT | CMPLT | 2023-11-16 16:06 | 2023-11-16 16:06 | repaired the OR table accessories
1067803 | 199209 | CM | CMPLT | CMPLT | 2023-11-16 16:10 | 2023-11-16 16:10 | new incoming inspection
1067804 | 199210 | CM | CMPLT | CMPLT | 2023-11-16 16:11 | 2023-11-16 16:11 | new incoming inspection
1067805 | 199211 | CM | CMPLT | CMPLT | 2023-11-16 16:12 | 2023-11-16 16:12 | new incoming inspection
1067806 | 199212 | CM | CMPLT | CMPLT | 2023-11-16 16:12 | 2023-11-16 16:12 | new incoming inspection
1067807 | 199213 | CM | CMPLT | CMPLT | 2023-11-16 16:13 | 2023-11-16 16:13 | new incoming inspection
1067809 | 199214 | CM | CMPLT | CMPLT | 2023-11-16 16:14 | 2023-11-16 16:14 | new incoming inspection
1067810 | 199215 | CM | CMPLT | CMPLT | 2023-11-16 16:14 | 2023-11-16 16:14 | new incoming inspection
1067811 | 199216 | CM | CMPLT | CMPLT | 2023-11-16 16:15 | 2023-11-16 16:15 | new incoming inspection
1067812 | 199217 | CM | CMPLT | CMPLT | 2023-11-16 16:16 | 2023-11-16 16:16 | new incoming inspection
1067808 | 199218 | CM | CMPLT | CMPLT | 2023-11-16 16:16 | 2023-11-16 16:16 | new incoming inspection
1067813 | 199219 | CM | CMPLT | CMPLT | 2023-11-16 16:17 | 2023-11-16 16:17 | new incoming inspection
| 199220 | CM | CMPLT | CMPLT | 2023-11-16 16:20 | 2023-11-16 16:20 | ordered the new hose for the hovermatt air supply
1059519 | 199221 | CM | CMPLT | CMPLT | 2023-11-16 16:21 | 2023-11-16 16:21 | not working | (TC 09/14/2023 04:21:08 PM, Huang, Jessica) picked it up in the OR room and tested it, working fine sent it back to OR
541283 | 199222 | CM | CMPLT | CMPLT | 2023-11-16 16:24 | 2023-11-16 16:24 | broken probe head | (TC 09/19/2023 04:23:51 PM, Huang, Jessica) replaced the probe cable, noticed the probe head is defectived. contacted roxon for the quote and sent it for approval. shipped the probe and didnt get the approval for a few months. the defectived probe was sent back without repair. repair quote got approved finally, sent the probe for repair again and received it back passed calibration
541304 | 199223 | CM | CMPLT | CMPLT | 2023-11-16 16:27 | 2023-11-16 16:27 | not working
1048636 | 199224 | CM | CMPLT | CMPLT | 2023-11-16 16:28 | 2023-11-16 16:28 | defective battery | (TC 09/19/2023 04:28:24 PM, Huang, Jessica) replaced the battery
| 199225 | CM | CMPLT | CMPLT | 2023-11-16 16:29 | 2023-11-16 16:29 | broken suction tree
1061385 | 199226 | CM | CMPLT | CMPLT | 2023-11-16 16:31 | 2023-11-16 16:31 | not showing right image on the side monitors | (TC 09/22/2023 04:31:32 PM, Huang, Jessica) got the call in the OR room, the image was not displayed properly on the side monitors, tried the different cable and couldnt fix the issue. swapped with the other machine, working fine. took the device into the different OR room, working fine
| 199227 | CM | CMPLT | CMPLT | 2023-11-16 17:00 | 2023-11-16 17:00 | getting the cart ready for OR
| 199228 | CM | CMPLT | CMPLT | 2023-11-16 17:08 | 2023-11-16 17:08 | requesting electrical checks in the OR room | (TC 09/15/2023 05:07:44 PM, Huang, Jessica) the side monitors were off while the ESU was operating. removed the power cable from the boom and connected to the wall outlet, it was working for awhile and happened it again during the case. arranged the building service and the boom specialist from olympus to have electrical checks. nothing was found reported to the OR leadership (TC 09/15/2023 05:07:44 PM, Huang, Jessica) the side monitors were off while the ESU was operating. removed the power cable from the boom and connected to the wall outlet, it was working for awhile and happened it again during the case. arranged the building service and the boom specialist from olympus to have electrical checks. nothing was found reported to the OR leadership (TC 09/18/2023 05:11:29 PM, Huang, Jessica) contacted Olympus to have integration team to check the room nothing was found
120017 | 199229 | CM | CMPLT | CMPLT | 2023-11-16 17:14 | 2023-11-16 17:14 | not working | (TC 10/03/2023 05:14:20 PM, Huang, Jessica) replaced the battery
| 199230 | CM | CMPLT | CMPLT | 2023-11-16 17:19 | 2023-11-16 17:19 | 6 food pumps not working | (TC 10/03/2023 05:18:45 PM, Huang, Jessica) 7th floor sent 6 food pumps, tested them and working fine informed the nurse educator
120014 | 199231 | CM | CMPLT | CMPLT | 2023-11-16 17:20 | 2023-11-16 17:20 | not working | (TC 10/03/2023 05:20:29 PM, Huang, Jessica) replaced the battery
| 199232 | CM | CMPLT | CMPLT | 2023-11-16 17:22 | 2023-11-16 17:22 | sending the genius calibrator for annual calibration
| 199233 | CM | CMPLT | CMPLT | 2023-11-16 17:23 | 2023-11-16 17:23 | the recorder isnt working
1047216 | 199234 | CM | CMPLT | CMPLT | 2023-11-16 17:24 | 2023-11-16 17:24 | not working | (TC 10/10/2023 05:23:56 PM, Huang, Jessica) sent the thermometer for warranty repair
541268 | 199235 | CM | CMPLT | CMPLT | 2023-11-16 17:25 | 2023-11-16 17:25 | not working | (TC 10/10/2023 05:25:27 PM, Huang, Jessica) replaced the battery
541042 | 199236 | CM | CMPLT | CMPLT | 2023-11-16 17:26 | 2023-11-16 17:26 | not working
541311 | 199237 | CM | CMPLT | CMPLT | 2023-11-16 17:27 | 2023-11-16 17:27 | not working | (TC 10/10/2023 05:27:33 PM, Huang, Jessica) replaced the battery
541049 | 199238 | CM | CMPLT | CMPLT | 2023-11-16 17:28 | 2023-11-16 17:28 | not working | (TC 10/10/2023 05:28:22 PM, Huang, Jessica) replaced the BP cuff
1063375 | 199239 | CM | CMPLT | CMPLT | 2023-11-16 17:29 | 2023-11-16 17:29 | not working | (TC 10/10/2023 05:29:31 PM, Huang, Jessica) sent it to the vendor for warranty repair
1047215 | 199240 | CM | CMPLT | CMPLT | 2023-11-16 17:30 | 2023-11-16 17:30 | not working | (TC 10/10/2023 05:30:34 PM, Huang, Jessica) sent it for warranty repair
1053411 | 199241 | CM | CMPLT | CMPLT | 2023-11-16 17:31 | 2023-11-16 17:31 | not working | (TC 10/10/2023 05:31:28 PM, Huang, Jessica) sent it for warranty repair
1047217 | 199242 | CM | CMPLT | CMPLT | 2023-11-16 17:34 | 2023-11-16 17:34 | not working | (TC 10/10/2023 05:34:28 PM, Huang, Jessica) sent it for warranty repair
1047203 | 199243 | CM | CMPLT | CMPLT | 2023-11-16 17:35 | 2023-11-16 17:35 | not working | (TC 10/10/2023 05:35:11 PM, Huang, Jessica) sent it for warranty repair
1054886 | 199244 | CM | CMPLT | CMPLT | 2023-11-16 17:36 | 2023-11-16 17:36 | not working | (TC 10/10/2023 05:36:01 PM, Huang, Jessica) sent it for warranty repair
1057124 | 199245 | CM | CMPLT | CMPLT | 2023-11-16 17:37 | 2023-11-16 17:37 | replaced the temperature probe
1064423 | 199246 | CM | CMPLT | CMPLT | 2023-11-16 17:40 | 2023-11-16 17:40 | not working | (TC 10/16/2023 05:40:00 PM, Huang, Jessica) tested it, working fine problem not found
| 199247 | CM | CMPLT | CMPLT | 2023-11-16 17:43 | 2023-11-16 17:43 | requesting quote for the joystick on the moible headlight
120012 | 199248 | CM | CMPLT | CMPLT | 2023-11-16 17:44 | 2023-11-16 17:44 | removed from service
541268 | 199249 | CM | CMPLT | CMPLT | 2023-11-16 17:45 | 2023-11-16 17:45 | not working
1046463 | 199250 | CM | CMPLT | CMPLT | 2023-11-16 17:47 | 2023-11-16 17:47 | not working | (TC 10/17/2023 05:46:49 PM, Huang, Jessica) replaced the SPO2 extension cable
541311 | 199251 | CM | CMPLT | CMPLT | 2023-11-16 17:48 | 2023-11-16 17:48 | replaced the temperature probe
1046633 | 199252 | CM | CMPLT | CMPLT | 2023-11-16 17:49 | 2023-11-16 17:49 | replaced the battery
| 199253 | CM | CMPLT | CMPLT | 2023-11-16 17:53 | 2023-11-16 17:53 | checking neptune device in OR
1064441 | 199254 | CM | CMPLT | CMPLT | 2023-11-16 17:54 | 2023-11-16 17:54 | not working | (TC 10/24/2023 05:54:16 PM, Huang, Jessica) sent it for warranty repair
1063364 | 199255 | CM | CMPLT | CMPLT | 2023-11-16 17:55 | 2023-11-16 17:55 | not working | (TC 10/24/2023 05:54:59 PM, Huang, Jessica) sent it for warranty repair
| 199256 | CM | CMPLT | CMPLT | 2023-11-16 17:56 | 2023-11-16 17:56 | arranging to install airseal in the OR room
541263 | 199257 | CM | CMPLT | CMPLT | 2023-11-16 17:58 | 2023-11-16 17:58 | not working | (TC 10/31/2023 05:57:45 PM, Huang, Jessica) replaced the SPO2 sensor
1060181 | 199258 | CM | CMPLT | CMPLT | 2023-11-16 17:59 | 2023-11-16 17:59 | not working | (TC 10/31/2023 05:58:46 PM, Huang, Jessica) sent it for warranty repair
1060184 | 199259 | CM | CMPLT | CMPLT | 2023-11-16 17:59 | 2023-11-16 17:59 | not working | (TC 10/31/2023 05:59:39 PM, Huang, Jessica) sent it for warranty repair
1064442 | 199260 | CM | CMPLT | CMPLT | 2023-11-16 18:01 | 2023-11-16 18:01 | not working | (TC 10/31/2023 06:00:47 PM, Huang, Jessica) sent it for warranty repair
1051899 | 199261 | CM | CMPLT | CMPLT | 2023-11-16 18:11 | 2023-11-16 18:11 | missing the rear case screws | (TC 11/16/2023 06:11:22 PM, Huang, Jessica) missing screws and nuts, contacted mother's choice for PN and quote, was infromed the parts are not available. the vendor sent two spare parts, installed them on the pump and tested it working fine
| 199262 | CM | CMPLT | CMPLT | 2023-11-16 18:22 | 2023-11-16 18:22 | received a call from OR
67956 | 199263 | CM | CMPLT | CMPLT | 2023-11-17 08:45 | 2023-11-17 09:41 | Missing power cord | (TC 11/17/2023 09:37:26 AM, Rhuland, Courtney) Replaced power cord and light with spare from retired unit Funcitonal check - OK Electrical safety - OK
1066460 | 199264 | FC | CMPLT | CMPLT | 2023-11-17 09:07 | 2023-11-17 09:07 | QC checklist | (TC 11/17/2023 09:06:54 AM, Baghirzade, Nara) Completed QC as per Sarabeth's request. QC completed. Returned to 18 South.
| 199265 | RDS | CMPLT | CMPLT | 2023-11-17 09:08 | 2023-11-17 09:12 | Weekly Rounds: NICU | (TC 11/17/2023 09:08:00 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure.
| 199266 | RDS | CMPLT | CMPLT | 2023-11-17 09:12 | 2023-11-17 09:14 | Weekly Rounds: 11 South | (TC 11/17/2023 09:12:23 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 199267 | RDS | CMPLT | CMPLT | 2023-11-17 09:15 | 2023-11-17 09:16 | Weekly Rounds: 18 South | (TC 11/17/2023 09:14:50 AM, Baghirzade, Nara) Weekly rounds on 18 South completed. See procedure for results.
1055587 | 199268 | CM | CMPLT | CMPLT | 2023-11-17 10:36 | 2023-11-17 10:37 | temp probe broken | (TC 11/17/2023 10:35:47 AM, Campbell, Connor) Temp probe had a cover wedged into the storage port removed the temp probe cover inspected the thermometer probe reseated the probe tested the probe unit functional, returned to service
| 199269 | CM | CMPLT | CMPLT | 2023-11-17 10:48 | 2023-11-17 10:48 | Full team meeting | (TC 11/17/2023 10:47:56 AM, Baghirzade, Nara) Full team meeting
| 199270 | RDS | CMPLT | CMPLT | 2023-11-17 11:34 | 2023-11-17 11:43 | Weekly Rounds: RESUS | (TC 11/17/2023 11:34:18 AM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure for results.
| 199271 | RDS | CMPLT | CMPLT | 2023-11-17 11:43 | 2023-11-17 11:44 | Weekly Rounds: PACU POD A | (TC 11/17/2023 11:43:09 AM, Baghirzade, Nara) Weekly rounds in PACU POD A completed. See procedure for results.
| 199272 | MEP | CMPLT | CMPLT | 2023-11-17 13:40 | 2024-05-01 14:11 | Alaris Pumps Upgrade Projet | (TC 11/17/2023 01:52:32 PM, Reid, Emmerson ) SC 1003109300 (TC 04/11/2024 02:59:55 PM, Reid, Emmerson ) Tried upgrading the syringe pump with ASM software, was not have luck but it seems like there may have been other issues with it.
1057758 | 199273 | CM | CMPLT | CMPLT | 2023-11-17 13:49 | 2023-11-17 13:50 | High GDI Resource Usage Alert (Focal Point Message) | (TC 11/17/2023 01:48:55 PM, Canjimavadivel, Kamallesh ) Received multiple focal point alerts for same issue. Recommended action was to restart central. Went to unit and restarted central.
62633 | 199274 | PM | CMPLT | CMPLT | 2023-11-17 14:00 | 2023-11-17 14:01 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 11/14/2023 02:00:04 PM, Rendon Lopez , Elizabeth) Replaced observation light assembly, after troubleshooting intermittent issue. PM completed, all tests passed.
| 199275 | CM | CMPLT | CMPLT | 2023-11-17 14:04 | 2023-11-17 14:04 | biomed meeting | biomed meeting
| 199276 | CM | CMPLT | CMPLT | 2023-11-17 14:05 | 2023-11-17 14:05 | PS500 installation | worked with Salur and Rex.
1061215 | 199277 | CM | CMPLT | CMPLT | 2023-11-17 14:17 | 2023-11-17 14:18 | Table Remote now working | (TC 11/17/2023 02:17:04 PM, Reid, Emmerson ) Changed the remote with a new one and it worked. Tested the remote and it was working but it was missing screws so I brought it down to repair.
65841 | 199278 | CM | CMPLT | CMPLT | 2023-11-17 14:18 | 2024-03-12 14:09 | Hose damage | (TC 11/16/2023 02:21:41 PM, Rendon Lopez , Elizabeth) Device in OR 3 L&D, confirmed issue. Received part information. ordering replacement and spare Received
1061214 | 199279 | CM | CMPLT | CMPLT | 2023-11-17 14:18 | 2023-11-17 14:19 | PACS computer not showing on monitor | (TC 11/17/2023 02:18:14 PM, Reid, Emmerson ) Display settings were wrong. I told them it was help desk not us but I fixed it.
1050244 | 199280 | CM | ACTIV | ISSUE | 2023-11-17 14:21 |  | Alarming non-stop | (TC 11/17/2023 02:21:34 PM, Reid, Emmerson ) Tommy brought it down and it was alarming. Front panel was not working so I had to replace it in order to get the alarm to stop. LCD dispaly also seems to be broken and needs more investigation.
1047913 | 199281 | CM | CMPLT | CMPLT | 2023-11-17 15:34 | 2023-11-17 15:41 | "water pressure low" | (TC 11/17/2023 03:34:07 PM, Cortez, John) Same error as the one from a month back. This time however, I was able to repeat the issue by firing the F-SW handpiece. Error appeared within a minute. Nothing visually/audibly off. Reset device, which seems to resolve issue. On startup, device restarts water cycle. Perhaps leaving unit on idle is causing issues. Renewed water anyways and informed manager
1067437 | 199282 | FC | CMPLT | CMPLT | 2023-11-20 07:24 | 2023-11-20 07:24 | QC checklist | (TC 11/20/2023 07:23:45 AM, Baghirzade, Nara) As per Sri's request; QC checklist completed. Returned to 11 South.
| 199283 | RDS | CMPLT | CMPLT | 2023-11-20 08:07 | 2023-11-20 10:28 | Weekly check for ED
| 199284 | RDS | CMPLT | CMPLT | 2023-11-20 08:07 | 2023-11-30 09:36 | Weekly check for ICU
502049 | 199285 | ACC | CMPLT | CMPLT | 2023-11-20 08:23 | 2023-11-27 13:39 | MSH OPG SPP PROBE C5-1 ARTIFACT REPLACEMENT ACCEPTANCE | (TC 11/20/2023 08:23:39 AM, Whittington, Scott) MSH OPG SPP PROBE C5-1 ARTIFACT REPLACEMENT ACCEPTANCE (TC 11/23/2023 04:05:20 PM, Dolinar, John) MSH OPG SPP PROBE C5-1 ARTIFACT REPLACEMENT ACCEPTANCE msh opg spp probe c5-1 sn# B1KX93 site 53198439 REF# CHRIS 0122665852 SHIPPING TO TGH 1NU160 PROBE NOT YET DELIVERED/FOUND? TEMP SUBSTITUTE MIE 502225 FROM TGH IU 22SPARES (TC 11/23/2023 04:05:20 PM, Dolinar, John) MSH OPG SPP PROBE C5-1 ARTIFACT REPLACEMENT ACCEPTANCE msh opg spp probe c5-1 sn# B1KX93 site 53198439 REF# CHRIS 0122665852 SHIPPING TO TGH 1NU160 PROBE NOT YET DELIVERED/FOUND? TEMP SUBSTITUTE MIE 502225 FROM TGH IU 22 SPARES (TC 11/27/2023 01:38:11 PM, Whittington, Scott) MSH OPG SPP PROBE C5-1 ARTIFACT REPLACEMENT ACCEPTANCE SN# F09ZWZ | MSH OPG SPP PROBE C5-1 ARTIFACT REPLACEMENT ACCEPTANCE msh opg spp probe c5-1 sn# B1KX93 site 53198439 REF# CHRIS 0122665852 SHIPPING TO TGH 1NU160 RETURN FAULTY PROBE VIA TGH FEDEX 7866 5712 2188 RMA #R439015887
| 199286 | RDS | CMPLT | CMPLT | 2023-11-20 08:34 | 2023-11-20 08:34 | weekly check for 14 S and 16 N | weekly check for 14 S and 16 N
1068220 | 199287 | CM | CMPLT | CMPLT | 2023-11-20 08:51 | 2023-12-28 12:44 | MSH US i800 PRISM ACCEPTANCE | (TC 11/28/2023 04:02:35 PM, Whittington, Scott) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH (TC 11/28/2023 04:06:31 PM, Dolinar, John) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH (TC 11/28/2023 04:06:49 PM, Chau, Ken) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH (TC 11/28/2023 04:07:02 PM, Fan, Tak) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH | MSH US i800 PRISM ACCEPTANCE 55 mOHM 275 uA
1068221 | 199288 | CM | CMPLT | CMPLT | 2023-11-20 08:52 | 2023-12-28 12:44 | MSH US i800 PRISM ACCEPTANCE | (TC 11/28/2023 04:07:45 PM, Whittington, Scott) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH (TC 11/28/2023 04:07:55 PM, Dolinar, John) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH (TC 11/28/2023 04:08:11 PM, Chau, Ken) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH (TC 11/28/2023 04:08:28 PM, Fan, Tak) MSH US i800 PRISM ACCEPTANCE SYSTEM ARRIVED ON SITE PROGRAMMED MOVED TO MSH | MSH US i800 PRISM ACCEPTANCE 160mOHM 280 uA
| 199289 | RDS | CMPLT | CMPLT | 2023-11-20 09:40 | 2023-11-20 09:42 | Weekly Checks For PACU | (TC 11/20/2023 09:41:34 AM, Reid, Emmerson ) Weekly round complete
| 199290 | RDS | CMPLT | CMPLT | 2023-11-20 09:42 | 2023-11-20 09:44 | Weekly Checks For OR | (TC 11/20/2023 09:43:17 AM, Reid, Emmerson ) Weekly rounds complete
1053330 | 199291 | CM | CMPLT | CMPLT | 2023-11-20 11:02 | 2023-11-20 11:19 | Battery 1 needs replacement | (TC 11/20/2023 11:13:03 AM, Cortez, John) Swapped with spare battery that was conditioned recently. Will monitor if error returns 6455mAhs 20 cycles 1%error 35.6% SOH
1055797 | 199292 | CM | CMPLT | CMPLT | 2023-11-20 11:24 | 2023-11-20 11:35 | Dashboard mentions battery fault | (TC 11/20/2023 11:24:05 AM, Cortez, John) No battery alerts on actual defib, Battery LED lit up Swapped with spare anyway, conditioned faulty battery
1066421 | 199293 | CM | CMPLT | CMPLT | 2023-11-20 12:45 | 2023-11-20 12:45 | power plug broken | repaired power plug
| 199294 | CM | CMPLT | CMPLT | 2023-11-20 12:46 | 2023-11-20 12:46 | biomed meeting | biomed meeting
| 199295 | CM | CMPLT | CMPLT | 2023-11-20 12:47 | 2023-11-20 12:47 | PS500 installation | worked with Salur from Draeger
1068226 | 199296 | CM | CMPLT | CMPLT | 2023-11-20 13:51 | 2023-11-22 14:18 | MSH ER VIVID 1Q PROBE ACCEPTANCE | (TC 11/20/2023 07:13:00 AM, Whittington, Scott) MSH ER VIVID 1Q PROBE ACCEPTANCE CAT# H48062AF TYPE: TV DESC: E8Cs-RS SERVICE PART# 5670375 SN# 1268200WX1 | MSH ER VIVID 1Q PROBE ACCEPTANCE CAT# H48062AF TYPE: TV DESC: E8Cs-RS SERVICE PART# 5670375 SN# 1268200WX1
1068227 | 199297 | CM | CMPLT | CMPLT | 2023-11-20 13:52 | 2023-11-22 14:18 | MSH ER VIVID 1Q PROBE ACCEPTANCE | (TC 11/21/2023 07:13:27 AM, Whittington, Scott) MSH ER VIVID 1Q PROBE ACCEPTANCE MSH ER VIVID 1Q PROBE ACCEPTANCE CAT# H48062AF TYPE: TV DESC: E8Cs-RS SERVICE PART# 5670375 SN# 1268223WX3 (TC 11/20/2023 07:13:27 AM, Whittington, Scott) MSH ER VIVID 1Q PROBE ACCEPTANCE MSH ER VIVID 1Q PROBE ACCEPTANCE CAT# H48062AF TYPE: TV DESC: E8Cs-RS SERVICE PART# 5670375 SN# 1268223WX3 | MSH ER VIVID 1Q PROBE ACCEPTANCE MSH ER VIVID 1Q PROBE ACCEPTANCE CAT# H48062AF TYPE: TV DESC: E8Cs-RS SERVICE PART# 5670375 SN# 1268223WX3
1068228 | 199298 | CM | CMPLT | CMPLT | 2023-11-20 13:53 | 2023-11-22 14:19 | MSH ER VIVID 1Q PROBE ACCEPTANCE | (TC 11/20/2023 07:14:01 AM, Whittington, Scott) MSH ER VIVID 1Q PROBE ACCEPTANCE H45551ZE SN#NT105505 SERVICE PN# 5729431 MSH ER VIVID 1Q PROBE ACCEPTANCE CAT# H45551ZE DESC: 6tC-RS SN#NT105505 SERVICE PN# 5729431 | MSH ER VIVID 1Q PROBE ACCEPTANCE H45551ZE SN#NT105505 SERVICE PN# 5729431 MSH ER VIVID 1Q PROBE ACCEPTANCE CAT# H45551ZE DESC: 6tC-RS SN#NT105505 SERVICE PN# 5729431
1068229 | 199299 | CM | CMPLT | CMPLT | 2023-11-20 13:54 | 2023-11-22 14:19 | MSH ER VIVID 1Q PROBE ACCEPTANCE | (TC 11/21/2023 07:15:15 AM, Whittington, Scott) MSH ER VIVID 1Q PROBE ACCEPTANCE MSH ER VIVID 1Q PROBE ACCEPTANCE H45551ZE SN#NT105527 SERVICE PN# 5729431 | MSH ER VIVID 1Q PROBE ACCEPTANCE MSH ER VIVID 1Q PROBE ACCEPTANCE H45551ZE SN#NT105527 SERVICE PN# 5729431
| 199300 | PA | CMPLT | CMPLT | 2023-11-20 14:37 | 2023-11-28 13:43 | Ordering DVI-D cables for OR | (TC 11/20/2023 03:12:59 PM, Reid, Emmerson ) Ordered DVI cables from digikey for OR
1060463 | 199301 | CM | CMPLT | CMPLT | 2023-11-20 15:14 | 2023-12-08 08:22 | Cable is broken | (TC 11/20/2023 03:14:15 PM, Reid, Emmerson ) Went with Samir to investigate. Cable was broken at the end connecting to the scopeguide and I was not able to repair it. Handed off to Kamran.
1054995 | 199302 | CM | CMPLT | CMPLT | 2023-11-20 15:21 | 2023-11-20 15:22 | Thermometer not working | (TC 11/20/2023 03:21:01 PM, Ferreira, Qwynn) replaced with spare probe and was working fine - previous probe was not making any reading.
1068233 | 199303 | CM | CMPLT | CMPLT | 2023-11-20 15:22 | 2023-11-22 14:20 | MSH ER GE VIVID IQ TEE PROBE BITE HOLE INDICATOR ACCEPTANCE | (TC 11/21/2023 07:15:36 AM, Whittington, Scott) MSH ER GE VIVID IQ TEE PROBE BITE HOLE INDICATOR ACCEPTANCE | MSH ER GE VIVID IQ TEE PROBE BITE HOLE INDICATOR ACCEPTANCE
1061959 | 199304 | CM | CMPLT | CMPLT | 2023-11-20 16:02 | 2023-11-20 16:03 | Scale issue, "call service" prompt | (TC 11/20/2023 04:02:36 PM, Cortez, John) Confirmed issue, power cycling bed clears error
1061458 | 199305 | CM | CMPLT | CMPLT | 2023-11-20 16:05 | 2023-11-20 16:06 | Motion functions not working | (TC 11/20/2023 04:05:31 PM, Cortez, John) Confirmed issue, power cycling bed clears error
1054978 | 199306 | CM | CMPLT | CMPLT | 2023-11-21 07:04 | 2023-11-21 07:06 | Oximetry not working | (TC 11/21/2023 07:04:01 AM, Baghirzade, Nara) SpO2 was not working; picking up interference and not displaying SpO2%. Tested trunk cable; cable ok. SpO2 finger sensor was defective. Replaced SpO2 finger sensor. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055918 | 199307 | CM | CMPLT | CMPLT | 2023-11-21 07:06 | 2023-11-21 07:06 | Oximetry not working | (TC 11/21/2023 07:05:52 AM, Baghirzade, Nara) SpO2 was not working; picking up interference and not displaying SpO2%. Tested trunk cable; cable ok. SpO2 finger sensor was defective. Replaced SpO2 finger sensor. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 199308 | CM | CMPLT | CMPLT | 2023-11-21 07:07 | 2023-11-21 07:07 | Meeting with Scott and Naz | (TC 11/20/2023 11:06:00 AM, Baghirzade, Nara) Met with Scott and Naz to discuss ophthalmic ultrasounds/support/current issues.
| 199309 | CM | CMPLT | CMPLT | 2023-11-21 07:08 | 2023-11-21 07:08 | PM huddle | (TC 11/20/2023 02:07:00 PM, Baghirzade, Nara) Biomed PM huddle
65922 | 199310 | CM | CANCL | CANCL | 2023-11-21 07:42 |  | Cancelled WO. | Created WO by mistake.
69437 | 199311 | CM | CMPLT | CMPLT | 2023-11-21 08:44 | 2023-11-21 09:57 | Buttons are unresponsive | (TC 11/21/2023 08:44:07 AM, Baghirzade, Nara) Connected pump to ASM and ran the keyboard test; all buttons are unresponsive. Replaced the front case. Check-in tests passed and completed. IUI connectors and device cleaned. Device is fully functional.
| 199312 | CM | ACTIV | ISSUE | 2023-11-21 10:35 |  | ED miscellaneous | (TC 11/21/2023 10:35:01 AM, Canjimavadivel, Kamallesh ) Fixed laminator machine (TC 11/27/2023 03:16:04 PM, Canjimavadivel, Kamallesh ) Functional test of the Autopulse (TC 11/30/2023 09:59:21 AM, Canjimavadivel, Kamallesh ) Validating mounting bracket for specula holders (TC 11/30/2023 01:17:30 PM, Canjimavadivel, Kamallesh ) Testing cables (TC 12/01/2023 02:30:26 PM, Canjimavadivel, Kamallesh ) Added diagnostic stations to TMS and unboxed (TC 12/08/2023 10:51:37 AM, Canjimavadivel, Kamallesh ) Verifying patch updates on centrals. Install not completed on PC end, need to validate issue (TC 12/19/2023 03:30:00 PM, Canjimavadivel, Kamallesh ) Working with Stryker rep to get access to B14 in order to troubleshoot issue with VGA port (TC 02/07/2024 08:27:33 AM, Canjimavadivel, Kamallesh ) Checking diagnostic stations in various rooms as per request from unit (TC 02/09/2024 08:08:52 AM, Canjimavadivel, Kamallesh ) Rover crash cart had broken hinge. Need to replaced rivets. Drilled out old rivets and installed new rivets. Returned to unit and informed PCM (TC 02/08/2024 03:00:00 PM, Canjimavadivel, Kamallesh ) Rover crash cart had broken hinge. Need to replaced rivets. Drilled out old rivets and installed new rivets. Returned to unit and informed PCM (TC 02/14/2024 09:00:00 AM, Canjimavadivel, Kamallesh ) Looking into consumable items requested by ED team that need to be update on SAP. Spoke to Becky for assistance. (TC 02/16/2024 01:07:48 PM, Canjimavadivel, Kamallesh ) Monthly touch base meeting (TC 02/27/2024 01:46:52 PM, Canjimavadivel, Kamallesh ) Updated asset info for ED Annex after move (TC 02/27/2024 02:58:20 PM, Canjimavadivel, Kamallesh ) BD recall discussion with ED team (TC 02/28/2024 02:31:25 PM, Canjimavadivel, Kamallesh ) Checked in defib that had readiness issue and corrected locations on dashboard (TC 03/01/2024 02:53:15 PM, Canjimavadivel, Kamallesh ) Testing SpO2 cables that were sent down (TC 03/08/2024 02:31:41 PM, Canjimavadivel, Kamallesh ) Prepped defib for ED education next week (TC 03/12/2024 10:18:34 AM, Canjimavadivel, Kamallesh ) Looking into material orders (TC 03/15/2024 08:19:20 AM, Canjimavadivel, Kamallesh ) Follow-up on consumable ordering info on e-req (TC 03/25/2024 02:50:31 PM, Canjimavadivel, Kamallesh ) Go through cable inventory in the ED storage rooms and got part numbers for missing accessories. Provided to unit (TC 04/11/2024 02:42:16 PM, Canjimavadivel, Kamallesh ) Go over plan for ED final equipment list (TC 05/10/2024 10:38:48 AM, Canjimavadivel, Kamallesh ) Dropped off vent to storage room (TC 05/16/2024 09:51:48 AM, Canjimavadivel, Kamallesh ) Sorting amico delivery for ED Final items (TC 06/03/2024 10:29:00 AM, Canjimavadivel, Kamallesh ) Testing SpO2 cables. Sent part ordering info to unit (TC 06/06/2024 01:45:00 PM, Canjimavadivel, Kamallesh ) Meeting with Becky and Nima to discuss standardizing SpO2 cables in ED (TC 06/14/2024 01:32:50 PM, Canjimavadivel, Kamallesh ) Transition planning discussion (TC 06/14/2024 01:33:22 PM, Canjimavadivel, Kamallesh ) Validated inventory and requested quote for RD SpO2 sensors+cables (TC 07/10/2024 01:46:58 PM, Canjimavadivel, Kamallesh ) Asset tagging equipment installed during go-live (TC 08/13/2024 10:26:06 AM, Canjimavadivel, Kamallesh ) Inquiry regarding temp probe trunk. Followed and provided recommendation to connect with CNS and PCM (TC 08/16/2024 11:23:21 AM, Canjimavadivel, Kamallesh ) Monthly touchbase with PCM (TC 08/21/2024 02:42:37 PM, Canjimavadivel, Kamallesh ) VSM repurposing from ED (TC 10/09/2024 03:07:07 PM, Canjimavadivel, Kamallesh ) Working on ED exam lights inventory plan and added PM schedules (TC 10/24/2024 02:48:21 PM, Canjimavadivel, Kamallesh ) Philips upgrade touch base meeting (TC 11/26/2024 03:45:33 PM, Canjimavadivel, Kamallesh ) Swapped out X3s for PM (TC 11/28/2024 11:01:53 AM, Canjimavadivel, Kamallesh ) Swapping X3s for PMs. Checking VSM as per request from staff
| 199313 | CM | CMPLT | CMPLT | 2023-11-21 11:09 | 2023-11-21 11:10 | Meeting with ENT (Soojin) and Andrew | (TC 11/21/2023 11:09:18 AM, Baghirzade, Nara) Meeting with ENT (Soojin) and Andrew
| 199314 | CM | CMPLT | CMPLT | 2023-11-21 11:10 | 2023-11-21 11:11 | PICiX HW in NICU confirmation | (TC 11/21/2023 11:09:56 AM, Baghirzade, Nara) As per Kamran's request, I checked each central and overview station for the PICiX HW model. All NICU PICiX HW are model: HP Engage Flex Pro. Confirmed with Kamran.
1067012 | 199315 | CM | CMPLT | CMPLT | 2023-11-21 12:00 | 2023-11-21 13:04 | Broken 8pin patient cable | (TC 11/21/2023 01:02:59 PM, Scero, Tomlyn) Unit brought down by ECG tech. Replaced USB cable and gave them clear tabs.
1054466 | 199316 | CM | CMPLT | CMPLT | 2023-11-21 13:25 | 2023-12-28 10:25 | Created by Failed PM WO # 193445 : PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 12/28/2023 10:03:44 AM, Canjimavadivel, Kamallesh ) PM complete, see procedure (TC 12/28/2023 10:18:58 AM, Canjimavadivel, Kamallesh ) Repair completed by Vince Dec 20, 2023. See attached service report | (TC 11/21/2023 01:12:40 PM, Cortez, John) Bed+Intellidrive batteries were replaced 70 mOhms, 25 uAmps Pendant missing clip PM failed, buttons on bottom right rail not working during chair/flat test
1059745 | 199317 | CM | CMPLT | CMPLT | 2023-11-21 13:32 | 2023-11-21 14:57 | Broken | (TC 11/21/2023 02:53:45 PM, Rendon Lopez , Elizabeth) Damaged probe. Replaced probe and cable, tested-OK. Placed back in stand, returned to L&D
1047867 | 199318 | CM | CMPLT | CMPLT | 2023-11-21 13:33 | 2023-11-22 14:10 | System failure: check plunger sensor and force sensor test | (TC 11/21/2023 01:33:19 PM, Baghirzade, Nara) Upon boot up, there are two system failure errors: 1. check plunger sensor 2. force sensor test Opened the plunger head, first replaced the force sensor however the force sensor test kept failing. Placed the old force sensor back and replaced the small PCB in the plunger head. Kept failing upon boot up. (TC 11/21/2023 01:34:41 PM, Baghirzade, Nara) Opened the device completely to check the plunger cabling. Cabling is ok. Took the data cable off another medfusion pump however the cable is still not the issue. Checked the clutch. Checked all components including motor. All ok. (TC 11/21/2023 01:36:19 PM, Baghirzade, Nara) Kept the data cable from the another medufsion in this pump. Aseembled device back together. Upon boot up, the force sensor test failed again. Went through the spare parts in biomed and consistently the force sensor kept failing until was able to find a PCB and force sensor that passed the self test. Tested all the spare parts for the force sensor and plunger head and found that many of the spare parts were defective. Gave to Samir for further testing and decision making. Since all parts were disconnected/connected to the PCB boards, will have to calibrate each sensor and run for a prolonged period of time before handing it over to NICU. (TC 11/22/2023 11:55:24 AM, Baghirzade, Nara) After doing a full calibration, proceeded to run the pump to check for occlusion alarm functionality. A reoccuring error during syringe load "check syringe plunger sensor". Ran a digital sensor test and found that the syringe plunger sensors are defective. Replaced entire plungerhead. (TC 11/22/2023 11:56:39 AM, Baghirzade, Nara) Performed a full calibration of: -syringe size and position -force sensor -plunger position sensor -syringe size sensor See procedure for results. Ran the pump for an extended period of time with various rates. Completed successfully with no errors.
1054221 | 199319 | CM | CMPLT | CMPLT | 2023-11-22 08:14 | 2023-11-22 08:28 | Dispaly "SEC" and "0" on display | (TC 11/22/2023 08:23:24 AM, Canjimavadivel, Kamallesh ) Redocked thermometer to reset security counter and increased counter
65992 | 199320 | CM | CMPLT | CMPLT | 2023-11-22 10:46 | 2023-11-22 10:47 | Reported channel error | (TC 11/21/2023 10:46:05 AM, Baghirzade, Nara) Downloaded error logs - noted a couple of errors pertaining to the pressure sensor and linear sensor however log only. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. Existing SN barcode was tattered; reprinted and replaced with a new SN barcode sticker. Will run the syringe pump for a couple of hours to check for any issues/errors during infusion. (TC 11/22/2023 10:46:36 AM, Baghirzade, Nara) Ran pump for a couple of hours with varying rates. Changed rates in the middle of the infusion as well. Basic infusion completed successfully with no issues or errors. Pump is fully functional.
| 199321 | CM | ACTIV | ISSUE | 2023-11-22 10:55 |  | Cord Protector
1059828 | 199322 | PM | CMPLT | CMPLT | 2023-11-22 11:43 | 2023-11-22 11:44 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/22/2023 11:43:18 AM, Campbell, Connor) Disinfected, cleaned, rinsed and refilled
| 199323 | CM | CMPLT | CMPLT | 2023-11-22 11:52 | 2023-11-22 11:53 | Capital equipment meeting with dentistry and Andrew | (TC 11/22/2023 11:52:29 AM, Baghirzade, Nara) Capital equipment meeting with dentistry and Andrew
1053883 | 199324 | CM | CMPLT | CMPLT | 2023-11-22 12:02 | 2023-11-22 12:03 | Leaking CO2 tank | (TC 11/22/2023 12:02:01 PM, Rhuland, Courtney) Connector from hose to tank was flush with yoke Worked with Samir to Adjust the height to allow better connection with tank Replaced green seal with rubber seal Tested and issue resolved.
| 199325 | CM | CMPLT | CMPLT | 2023-11-22 12:04 | 2023-11-22 12:05 | Dental equipment meeting
1057302 | 199326 | CM | CMPLT | CMPLT | 2023-11-22 14:03 | 2023-11-22 14:15 | MSH OPG SPP OLD SOAK ROOM VOLUSON 10 RECONNECT AFTER SAMSUNG DEMO | (TC 11/22/2023 02:03:38 PM, Whittington, Scott) MSH OPG SPP OLD SOAK ROOM VOLUSON 10 RECONNECT AFTER SAMSUNG DEMO | MSH OPG SPP OLD SOAK ROOM VOLUSON 10 RECONNECT AFTER SAMSUNG DEMO msh opg spp soak rm ge vol 10 hook up after demo
1049608 | 199327 | CM | CMPLT | CMPLT | 2023-11-22 14:05 | 2023-11-22 14:14 | MSH OPG SPP D6 CLEAN MOUSE | (TC 11/22/2023 02:05:02 PM, Whittington, Scott) MSH OPG SPP D6 CLEAN MOUSE | MSH OPG SPP D6 CLEAN MOUSE msh opg spp vol10 small rooom mouse clean
502453 | 199328 | CM | CMPLT | CMPLT | 2023-11-22 14:09 | 2024-02-26 13:29 | MSH OPG SPP-3525(Rm D1) AFTER MOULD REMEDIATION, RECONNECT ULTRASOUND TO NETWORK | (TC 11/22/2023 02:09:26 PM, Whittington, Scott) MSH OPG SPP-3525(Rm D1) AFTER MOULD REMEDIATION, RECONNECT ULTRASOUND TO NETWORK (TC 11/23/2023 11:47:29 AM, Fan, Tak) MSH OPG SPP-3525(Rm D1) AFTER MOULD REMEDIATION, RECONNECT ULTRASOUND TO NETWORK SYSTEM OFFLINE SYSTEM CANNOT REACH HSC TAK RESTORED WIRELESS RESOLVED LAN CONNECTION TO MSH NETWORK (TC 11/23/2023 11:48:20 AM, Dolinar, John) MSH OPG SPP-3525(Rm D1) AFTER MOULD REMEDIATION, RECONNECT ULTRASOUND TO NETWORK SYSTEM OFFLINE SYSTEM CANNOT REACH HSC TAK RESTORED WIRELESS RESOLVED LAN CONNECTION TO MSH NETWORK (TC 11/23/2023 11:48:33 AM, Whittington, Scott) MSH OPG SPP-3525(Rm D1) AFTER MOULD REMEDIATION, RECONNECT ULTRASOUND TO NETWORK SYSTEM OFFLINE SYSTEM CANNOT REACH HSC TAK RESTORED WIRELESS RESOLVED LAN CONNECTION TO MSH NETWORK REQUIRES RECONNECTION TO HSC | MSH OPG SPP-3525(Rm D1) AFTER MOULD REMEDIATION, RECONNECT ULTRASOUND TO NETWORK SYSTEM OFFLINE SYSTEM CANNOT REACH HSC
501585 | 199329 | CM | CMPLT | CMPLT | 2023-11-22 14:11 | 2023-11-22 14:13 | MSH OPG SPP E2 RECONNECT SYSTEM AFTER SAMSUNG DEMO | (TC 11/22/2023 02:11:07 PM, Whittington, Scott) MSH OPG SPP E2 RECONNECT SYSTEM AFTER SAMSUNG DEMO | MSH OPG SPP E2 RECONNECT SYSTEM AFTER SAMSUNG DEMO
| 199330 | CM | CMPLT | CMPLT | 2023-11-22 14:13 | 2023-11-22 14:13 | ENT inventory with Andrew and Courtney | (TC 11/22/2023 02:12:45 PM, Baghirzade, Nara) ENT inventory with Andrew and Courtney
| 199331 | CM | CMPLT | CMPLT | 2023-11-22 14:14 | 2023-11-22 14:15 | Looking for 2 incubators to recall | (TC 11/22/2023 02:13:41 PM, Baghirzade, Nara) Looking for the remaining 2 incubators that requires the latch recall. Found 1 incubator in MBU. Could not locate the second incubator. Looked in L&D OR hallway, RESUS storage, NICU rooms and storage, MBU, and soiled rooms as well. Could not locate.
| 199332 | CM | CMPLT | CMPLT | 2023-11-22 15:02 | 2023-11-22 15:02 | ENT inventory | (TC 11/22/2023 03:01:37 PM, Rhuland, Courtney) Checked procedure rooms for equipment to replace under capitol
| 199333 | CM | CMPLT | CMPLT | 2023-11-22 15:03 | 2023-11-22 15:03 | Incubator inventory | (TC 11/22/2023 03:02:42 PM, Rhuland, Courtney) Searched for incubators left for recall - found 2 could not locate 1
1053391 | 199334 | CM | CMPLT | CMPLT | 2023-11-23 07:47 | 2023-11-23 07:48 | QC after/before patient use | (TC 11/23/2023 07:46:53 AM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 124 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by an hour. Device is functional. Returned to 11S.
| 199335 | PA | CMPLT | CMPLT | 2023-11-23 07:49 | 2024-01-08 13:58 | Ordering activac filters & battery | (TC 11/23/2023 07:56:00 AM, Baghirzade, Nara) SC 1003113790 (TC 01/08/2024 01:56:53 PM, Baghirzade, Nara) PO 7202326038
1059983 | 199336 | PM | CMPLT | PMFAIL | 2023-11-23 09:26 | 2023-12-27 15:03 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 12/27/2023 03:01:09 PM, Canjimavadivel, Kamallesh ) Found loose wires on left side near foot board. Swapped out with repaired bed on 18.
52114 | 199337 | CM | CANCL | CANCL | 2023-11-23 11:07 |  | Rebuild
32510 | 199338 | CM | CMPLT | CMPLT | 2023-11-23 12:03 | 2023-11-23 15:55 | MSH DR1 NO OGP MOVEMENT | (TC 11/23/2023 12:02:46 PM, Whittington, Scott) MSH DR1 NO OGP MOVEMENT REPAIR TUBE GUARD CAR, CABLING (TC 11/23/2023 12:02:46 PM, Whittington, Scott) MSH DR1 NO OGP MOVEMENT REPAIR TUBE GUARD CAR, CABLING (TC 11/27/2023 02:10:41 PM, Whittington, Scott) MSH DR1 NO OGP MOVEMENT | MSH DR1 NO OGP MOVEMENT
| 199339 | CM | CMPLT | CMPLT | 2023-11-23 12:44 | 2023-11-23 12:44 | PS500 installation | worked with Salur and Rex
| 199340 | CM | CMPLT | CMPLT | 2023-11-23 13:53 | 2023-11-23 13:54 | NICU Capital meeting
1061957 | 199341 | CM | CNP | EIU | 2023-11-23 13:55 |  | No functions working | (TC 11/23/2023 01:55:26 PM, Canjimavadivel, Kamallesh ) Went up to check bed, patient still in bed 1033A. Was not able to confirm issue as nurse was not around to assist. Power cycled bed and looked up error codes, found multiple error codes from today and multiple from the past. Was not able to test if functions were working. Need to verify functions | Error codes found: 3009-5-10-2000 3009-5-6-1213 3009-5-7-1203
1056030 | 199342 | CM | CMPLT | CMPLT | 2023-11-23 14:20 | 2024-02-07 16:07 | requested the replacement | (TC 11/23/2023 04:05:29 PM, Huang, Jessica) picked up the device from the OR desk requested the replacement unit from the vendor SN 190267719 got replaced sent the device back to OR
67928 | 199343 | CM | CMPLT | CMPLT | 2023-11-23 14:51 | 2023-11-24 10:21 | Hood replacement | (TC 11/23/2023 02:50:38 PM, Rhuland, Courtney) Replaced cracked hood (TC 11/24/2023 10:19:51 AM, Rhuland, Courtney) Replaced broken humidity sensor Functional check - OK Humidity - 75.6% Temp - 35C
62700 | 199344 | CM | CMPLT | CMPLT | 2023-11-23 15:10 | 2023-11-23 15:18 | Keyboard cover ripped | (TC 11/23/2023 03:10:15 PM, Canjimavadivel, Kamallesh ) Replaced keyboard cover
1064341 | 199345 | PM | CMPLT | CMPLT | 2023-11-23 15:28 | 2023-11-23 15:29 | Annual PM | (TC 11/23/2023 03:29:06 PM, Rendon Lopez , Elizabeth) PM completed, see attached report | PM completed by Medaris medical Technician. See attached Report
1064341 | 199346 | CM | CMPLT | CMPLT | 2023-11-23 15:32 | 2023-11-23 15:33 | Table back upper part makes cracking noises | (TC 11/21/2023 03:31:46 PM, Rendon Lopez , Elizabeth) Found issue when testing defective remote. Contacted medaris for repair. (TC 11/23/2023 03:31:46 PM, Rendon Lopez , Elizabeth) Medaris tech repaired table. See attached report
1064593 | 199347 | PM | CMPLT | CMPLT | 2023-11-23 15:48 | 2023-11-23 15:51 | Annual PM | (TC 11/23/2023 03:50:05 PM, Rendon Lopez , Elizabeth) PM completed, see attached report | PM performed by Medaris Medical Technician, See attached report
1049633 | 199348 | PM | CMPLT | CMPLT | 2023-11-23 15:57 | 2023-11-23 15:58 | Annual PM | (TC 11/23/2023 03:58:09 PM, Rendon Lopez , Elizabeth) PM completed, see attached report. | PM performed by Medaris Medical Technician. See attached report.
1069256 | 199349 | PM | CMPLT | CMPLT | 2023-11-23 16:00 | 2023-11-23 16:02 | Annual PM | (TC 11/23/2023 04:01:53 PM, Rendon Lopez , Elizabeth) Added device to TMS. PM completed, see attached report | PM completed by Medaris medical Technician. See attached Report
1069257 | 199350 | PM | CMPLT | CMPLT | 2023-11-23 16:04 | 2023-11-23 16:05 | Annual PM | (TC 11/23/2023 04:04:45 PM, Rendon Lopez , Elizabeth) Added device to TMS. PM completed, see attached report | PM completed by Medaris medical Technician. See attached Report
1067510 | 199351 | CM | CMPLT | CMPLT | 2023-11-23 16:10 | 2023-11-23 16:12 | Device not shutting off | (TC 11/21/2023 04:10:37 PM, Rendon Lopez , Elizabeth) Confirmed issue. Contacted supplier to check for parts. Device needs to be send out for repair or retire. notified unit (TC 11/23/2023 04:10:37 PM, Rendon Lopez , Elizabeth) L&D decided to purchase new unit, retired device from service.
1056128 | 199352 | CM | CMPLT | CMPLT | 2023-11-23 16:28 | 2023-11-23 16:28 | Not working | (TC 11/20/2023 04:28:08 PM, Rendon Lopez , Elizabeth) Tested-NFF
1056055 | 199353 | CM | CMPLT | CMPLT | 2023-11-23 18:23 | 2023-11-23 18:22 | Red line/streaks on the right side f the large scale monitor. | Troubleshoot the problem to a faulty adoptor in the B-Cabinet, ordered parts.
| 199354 | PA | CMPLT | CMPLT | 2023-11-24 07:47 | 2023-12-18 09:00 | TC70 Keyboard covers for stock | (TC 11/24/2023 07:52:36 AM, Canjimavadivel, Kamallesh ) Placed order for keyboard covers (TC 12/07/2023 10:56:04 AM, Canjimavadivel, Kamallesh ) Followed up on order, price update needed to be confirmed. (TC 12/18/2023 08:59:09 AM, Canjimavadivel, Kamallesh ) Received covers, labelled and put on stock shelf | To be charged back as used
| 199355 | CM | CMPLT | CMPLT | 2023-11-24 10:23 | 2023-11-24 10:24 | Dental equipment walk through | (TC 11/24/2023 09:22:00 AM, Rhuland, Courtney) Walk through of dental clinic equipment list for capital planning
1053883 | 199356 | CM | CMPLT | CMPLT | 2023-11-24 11:13 | 2023-11-29 08:31 | Ordering O-rings | (TC 11/24/2023 11:15:52 AM, Rhuland, Courtney) Ordered O-ring for CO2 tank
1069262 | 199357 | CM | CMPLT | CMPLT | 2023-11-24 12:57 | 2023-11-30 10:25 | Incoming Inspection | (TC 11/30/2023 10:24:12 AM, Ferreira, Qwynn) Functional Test for the sunction was done and passed Handed over to steve once it passed electrical safety
| 199358 | CM | CMPLT | CMPLT | 2023-11-24 13:11 | 2023-12-15 13:10 | Order for Stryker Procuity bed batteries, biomed stock | (TC 11/24/2023 01:24:11 PM, Cortez, John) Went with Samir to verify what batteries the Procuity beds use Placed order for batteries (TC 12/13/2023 01:01:45 PM, Cortez, John) Reached out to plexxus+BBM battery regarding order. Was informed that order already arrived (TC 12/15/2023 01:07:30 PM, Cortez, John) Couldn't locate in receiving. Item sent to wrong location, RM 979A (TC 12/14/2023 01:07:30 PM, Cortez, John) Couldn't locate in receiving. Item sent to wrong location, RM 979A (TC 12/15/2023 01:09:11 PM, Cortez, John) Picked up package in LTRI room 979A
1062127 | 199359 | CM | CMPLT | CMPLT | 2023-11-24 13:29 | 2023-11-24 13:32 | "bed stuck in elevated position" | (TC 11/24/2023 01:29:11 PM, Cortez, John) Power cycling bed resolves issue
| 199360 | CM | CMPLT | CMPLT | 2023-11-24 14:53 | 2023-12-01 11:25 | spare parts for zeus | (TC 11/24/2023 02:55:18 PM, Scero, Tomlyn) Received email from draeger with part info to order. (TC 12/01/2023 11:24:33 AM, Scero, Tomlyn) Parts received
| 199361 | CM | CMPLT | CMPLT | 2023-11-24 15:19 | 2023-11-24 15:20 | Move Cavitron from OR19 to OR2 | (TC 11/24/2023 03:19:11 PM, Canjimavadivel, Kamallesh ) Disconnected hose and plugs from boom in OR19 and moved device to OR2 and reconnected
1049047 | 199362 | CM | CMPLT | NFF | 2023-11-24 15:58 | 2023-11-24 15:59 | "doesn't get hot" | (TC 11/24/2023 03:58:07 PM, Cortez, John) Measured output with thermocouple, and temperature was accurate to set value. Not able to replicate issue.
| 199363 | CM | CMPLT | CMPLT | 2023-11-24 17:06 | 2023-12-07 17:52 | panda blender overhaul | (TC 11/24/2023 05:07:19 PM, Rendon Lopez , Elizabeth) Panda blender for overhaul (TC 12/07/2023 05:52:08 PM, Rendon Lopez , Elizabeth) Received
1063808 | 199364 | CM | CMPLT | CMPLT | 2023-11-24 17:14 | 2023-11-24 17:14 | Satellite monitor has no image. |  | Replaced s video cables and ordered new ones.
| 199365 | CM | CMPLT | CMPLT | 2023-11-24 17:16 | 2023-11-24 17:16 | Site coverage
| 199366 | RDS | CMPLT | CMPLT | 2023-11-27 07:45 | 2023-11-27 07:45 | weekly check for 14S and 16N | weekly check for 14S and 16N
62500 | 199367 | CM | CMPLT | CMPLT | 2023-11-27 08:41 | 2023-11-27 08:46 | Can't reach the setting temperature and not going beyond 27C alarming every 10 mins. | (TC 11/27/2023 08:42:01 AM, Chukri, Samir) Found plastic object inside on the fan parallel, Cleaned and replaced two wheels with brakes for IAD
1054522 | 199368 | CM | CMPLT | CMPLT | 2023-11-27 08:56 | 2023-11-27 08:59 | V3 not working | (TC 11/27/2023 08:56:40 AM, Canjimavadivel, Kamallesh ) Confirmed issue. Replaced chest lead set V1-V3. Tested all leads with simulator
1060002 | 199369 | CM | CMPLT | CMPLT | 2023-11-27 09:20 | 2023-11-27 10:43 | Broken leads | (TC 11/27/2023 10:38:22 AM, Canjimavadivel, Kamallesh ) Multiple leads were found damaged and taped up. Replaced leads LA, LL, RA, RL, V1-V3. Tested with simulator.
501071 | 199370 | CM | CMPLT | CMPLT | 2023-11-27 12:55 | 2023-12-07 08:31 | Unit shuts down during procedure. | (TC 11/30/2023 10:06:00 AM, Fan, Tak) Parts delayed. Informed techs. Checked system but system was off in the morning. Started up but crashed/powered-off after speaking to techs again. Could not boot system, keeps crashing after 3 seconds. Blue light on PS not flashing though. Informed Jillian. (TC 12/04/2023 01:12:00 PM, Fan, Tak) PS replaced issue persists. (TC 12/05/2023 02:13:05 PM, Fan, Tak) UMB replaced, issue persists. Jillian unsure about parts. Swapped PS with other system. Issue "resolved". Left PS in. Will swap with new one when arrives. Whitelisted new MAC. (TC 12/05/2023 02:15:44 PM, Fan, Tak) Return parts. (TC 12/07/2023 08:30:30 AM, Fan, Tak) Swapped PS back to spare and installed replacement PS. System OK. Returned part. | Called Philips. Ref# 0122696450
| 199371 | CM | CMPLT | CMPLT | 2023-11-27 14:02 | 2023-11-27 14:02 | OPG 8th floor overhead light | OPG 8th floor overhead light bulb replacement. Did not work. They will order new light.
| 199372 | CM | CMPLT | CMPLT | 2023-11-27 14:05 | 2023-11-27 14:05 | PS500 installation | worked with Kam and Rex. Salur is coming after 2 weeks. We need to find the last unit.
| 199373 | CM | CMPLT | CMPLT | 2023-11-27 14:06 | 2023-11-27 14:06 | created daily-check list | created daily-check list
| 199374 | RDS | CMPLT | CMPLT | 2023-11-27 14:24 | 2023-11-27 14:26 | Weekly Checks For PACU | (TC 11/27/2023 02:25:06 PM, Reid, Emmerson ) Defib needed to be plugged in and did 30J test.
| 199375 | CM | CMPLT | CMPLT | 2023-11-27 14:24 | 2023-11-27 14:24 | Cerner Downtime meeting. | (TC 11/27/2023 02:24:16 PM, Chukri, Samir) meeting with Nima and Adeel.
| 199376 | CM | CMPLT | CMPLT | 2023-11-27 14:27 | 2023-11-27 14:27 | Replaced borken door for the AED box on the tunnel between Sick kits and MSH | (TC 11/27/2023 02:27:07 PM, Chukri, Samir) work with John and replaced the broken door
1060121 | 199377 | CM | CMPLT | CMPLT | 2023-11-27 14:35 | 2023-11-27 14:37 | Function verification | (TC 11/27/2023 02:34:52 PM, Canjimavadivel, Kamallesh ) Visual check pass, no issues found. Opened up LifeBand and found one side was jammed in the hinge clip. Cleared jam and reseated band. Turned on device and tested functionality. All functions working as intended. Returned to ICU. | Request from ED to verifiy if all functions were working as previously it did not work during a case
1063808 | 199378 | CM | CMPLT | CMPLT | 2023-11-27 14:39 | 2023-11-27 14:55 | MSH 15 SPP PROCEDURE ROOM REMOTE VIDEO | (TC 11/27/2023 02:54:48 PM, Whittington, Scott) MSH 15 SPP PROCEDURE ROOM REMOTE VIDEO (TC 11/27/2023 02:54:56 PM, Dolinar, John) MSH 15 SPP PROCEDURE ROOM REMOTE VIDEO | MSH 15 SPP PROCEDURE ROOM REMOTE VIDEO
1060518 | 199379 | CM | CMPLT | CMPLT | 2023-11-27 20:59 | 2023-11-27 20:58 | No Table motion. | Table movement was calibrated.
| 199380 | CM | CMPLT | CMPLT | 2023-11-28 07:39 | 2023-11-29 14:46 | Ordering porthole seals for incubators
68800 | 199381 | CM | CMPLT | NFF | 2023-11-28 08:32 | 2023-11-28 08:33 | making buzzing sounds | (TC 11/28/2023 08:32:19 AM, Campbell, Connor) performed a check in on the PCU and it was fine performed a test with and IVP unit and it was fine as well no fault found
69483 | 199382 | CM | CMPLT | NFF | 2023-11-28 08:33 | 2023-11-28 08:34 | broken making buzzing sound | (TC 11/28/2023 08:33:30 AM, Campbell, Connor) performed a check in, passed and there was no buzzing sound at all during infusion or alarm no fault found
1066872 | 199383 | CM | CMPLT | CMPLT | 2023-11-28 08:36 | 2023-11-28 08:37 | Ophthalmoscope not working | (TC 11/28/2023 08:35:56 AM, Canjimavadivel, Kamallesh ) Removed head and found bulb was missing. Installed bulb and tested. Working ok.
68672 | 199384 | CM | CMPLT | CMPLT | 2023-11-28 09:47 | 2023-11-28 09:48 | Not working | (TC 11/23/2023 09:47:08 AM, Rendon Lopez , Elizabeth) Charging (TC 11/27/2023 09:47:08 AM, Rendon Lopez , Elizabeth) Downloaded error logs, no errors. Testing on battery-OK. Check in tests, all passed
1058281 | 199385 | CM | CMPLT | CMPLT | 2023-11-28 09:49 | 2023-11-28 09:50 | Will not turn on | (TC 11/23/2023 09:49:01 AM, Rendon Lopez , Elizabeth) Confirmed issue, no light when charging. Defective front case/keypad (TC 11/27/2023 09:49:01 AM, Rendon Lopez , Elizabeth) Replaced front case. Check in tests, all passed
67393 | 199386 | PM | CMPLT | CMPLT | 2023-11-28 09:52 | 2023-11-28 09:52 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 11/24/2023 09:52:11 AM, Rendon Lopez , Elizabeth) Blender replaced. PM completed, all tests passed
69065 | 199387 | CM | CMPLT | CMPLT | 2023-11-28 09:53 | 2023-11-28 09:55 | Channel error | (TC 11/27/2023 09:53:44 AM, Rendon Lopez , Elizabeth) Downloaded error log. E 240.4150.0 - Bottle pressure sensor fatal failure. Replaced bottle pressure sensor, calibrated
69311 | 199388 | CM | CMPLT | CMPLT | 2023-11-28 09:56 | 2023-11-28 09:58 | Door broken | (TC 11/27/2023 09:57:27 AM, Rendon Lopez , Elizabeth) Broken front door cover. Replaced, performed check in tests, all passed
| 199389 | CM | CMPLT | CMPLT | 2023-11-28 10:06 | 2023-11-28 10:06 | Hekp Carla with the right Vacuume Regulator configurator for 9S | (TC 11/28/2023 10:06:39 AM, Chukri, Samir) visiting 9S and meet with Carla, to get the right vacuum configurtor and connectors. completed the order form and emailed it to Carla.
502088 | 199390 | RQ | CMPLT | CMPLT | 2023-11-28 10:36 | 2023-11-28 10:59 | After completing patient and about to change from High energy collimator to low energy, P-scope showed "call service". With this message, only service | (TC 11/28/2023 10:57:42 AM, Green, Paul) Uncommanded motion after installing heavy collimators. SLD1 persistant errror occurred. Was able to clear. | Service was called. Tino answered and promptly came to the department to clear the log. Thanks for the great service!
1056473 | 199391 | CM | CMPLT | CMPLT | 2023-11-28 13:48 | 2023-11-28 13:48 | thermometer not working | replaced thermometer probe. Tested temp and BP. Functional.
1061209 | 199392 | CM | CMPLT | CMPLT | 2023-11-28 13:48 | 2023-11-28 13:50 | Lights not connecting to iPad | (TC 11/28/2023 01:48:30 PM, Reid, Emmerson ) iPad was coming in and out of connectiong with the lights. I reset the wifi on light one and it was working better but still had intermittent issues. I came back after the case and the connection issues had resolved themselves.
1066460 | 199393 | FC | CMPLT | CMPLT | 2023-11-28 13:49 | 2023-11-28 13:49 | QC check | passed QC check
68372 | 199394 | CM | CMPLT | CMPLT | 2023-11-28 14:22 | 2024-01-23 12:04 | Left side DIVA not working | (TC 11/23/2023 02:22:19 PM, Scero, Tomlyn) Downloading logs and reviewing with Draeger tech support. Worked with Salur on site to troubleshoot all of the valves in the Diva supply system. We found 1 possible place for a leak so I ordered that valve to replace. (TC 12/06/2023 12:02:11 PM, Scero, Tomlyn) Completely disassembled the zeus gas machine to remove the diva supply module. Replaced PM parts as I went along. (TC 12/07/2023 12:20:37 PM, Scero, Tomlyn) Replaced NO valve 2.31. Reassembled diva supply module into agm (TC 12/11/2023 03:03:49 PM, Scero, Tomlyn) Ran self test after reassmebling. Passed the self test. when I ran the unit on a vent mode it showed Air + N2O delivery cutoff error. (TC 01/10/2024 02:13:15 PM, Scero, Tomlyn) Replaced NC Suplly valve 2.26. Passing self test with no issues but still getting errors when trying to use the Left side DIVA port. Ordered another NC and NO valve and I'll replace the relief valve 2.62 and AGS valve next 2.60 (TC 01/19/2024 02:59:29 PM, Scero, Tomlyn) Replaced NC valve 2.62, test valve. Self test complete and passed with DIVA in left side 1. Ran with test lung, no errors. (TC 01/22/2024 12:37:47 PM, Scero, Tomlyn) finished replacing PM parts and putting the device back together. A leak was found and device wouldn't pass the self test. Troubleshooting found that one of the internal hoses wasn't fully seated and causing a leak. reseated and performed selftest with no errors. continued with PM procedures. (TC 01/23/2024 11:56:50 AM, Scero, Tomlyn) Device selftest passed with both DIVAs installed. No leaks or errors. Replacing the Test Valve (2.62) fixed the issue. Ready to continue with the PM procedures.
| 199395 | CM | CMPLT | CMPLT | 2023-11-28 14:23 | 2024-01-24 11:25 | working with Naz for transport incubator project in NICU | (TC 11/28/2023 02:22:50 PM, Chukri, Samir) working with Naz and checked all transport incubators in NICU (TC 11/30/2023 10:36:57 AM, Chukri, Samir) Prepairing the parts used on these incubator from 2018 to 2023 list and emailed to Naz up to her requested. (TC 12/07/2023 02:34:22 PM, Chukri, Samir) talking to Naz asking her to send me the list of the order. to review and check. (TC 12/19/2023 12:44:03 PM, Chukri, Samir) Meeting with Michelle and Olcay to add cylinder holder on the transport incubator. (TC 01/02/2024 11:45:10 AM, Chukri, Samir) quick chat with Naz Re: transport incubator.
| 199396 | RDS | CMPLT | CMPLT | 2023-11-28 14:27 | 2023-11-28 14:29 | Weekly checks in CCU | (TC 11/28/2023 02:26:54 PM, Scero, Tomlyn) Weekly checks in CCu complete. 30J test done on defibs
1046196 | 199397 | CM | CMPLT | CMPLT | 2023-11-28 15:38 | 2023-11-28 15:36 | X-ray tube replacement | X-ray tube replacement/ due to excessive tube arcing reported Service completed calibrations CT phantom tested OK
| 199398 | CM | CMPLT | CMPLT | 2023-11-28 15:52 | 2023-11-28 15:52 | NICU huddle | (TC 11/28/2023 03:51:39 PM, Rhuland, Courtney) Huddle regarding Cerner downtime
32520 | 199399 | CM | CMPLT | CMPLT | 2023-11-28 16:10 | 2023-11-29 08:53 | MSH DR2 SWAP OUT COLLIMATOR | (TC 11/28/2023 04:09:49 PM, Whittington, Scott) MSH DR2 SWAP OUT COLLIMATOR (TC 11/28/2023 04:10:05 PM, Dolinar, John) MSH DR2 SWAP OUT COLLIMATOR (TC 11/28/2023 05:52:00 PM, Whittington, Scott) MSH DR2 SWAP OUT COLLIMATOR | MSH DR2 SWAP OUT COLLIMATOR
1054537 | 199400 | PM | CMPLT | CMPLT | 2023-11-28 16:54 | 2023-11-28 16:55 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/28/2023 04:54:21 PM, Cortez, John) Completed cleaning cycle
1069271 | 199401 | ACC | CMPLT | CMPLT | 2023-11-28 18:18 | 2023-11-29 15:03 | MSH 5 US PVT-781VTE SN# FSD23X6345 acc | (TC 11/29/2023 03:03:42 PM, Whittington, Scott) MSH 5 US PVT-781VTE SN# FSD23X6345 ACCEPTANCE | MSH 5 US PVT-781VTE SN# FSD23X6345 ACCEPTANCE
1069272 | 199402 | ACC | CMPLT | CMPLT | 2023-11-28 18:20 | 2023-11-29 15:04 | MSH 5 US PVT781VTE SN# FSD23X6346FSD23X6346 acc | (TC 11/29/2023 03:04:20 PM, Whittington, Scott) MSH 5 US PVT781VTE SN# FSD23X6346FSD23X6346 acc | MSH 5 US PVT781VTE SN# FSD23X6346FSD23X6346 acc
1069273 | 199403 | ACC | CMPLT | CMPLT | 2023-11-28 18:22 | 2023-11-29 15:05 | MSH 5 US PVT-781 VTE SN# FSD2396074 ACC | (TC 11/29/2023 03:04:49 PM, Whittington, Scott) MSH 5 US PVT-781 VTE SN# FSD2396074 ACC | MSH 5 US PVT-781 VTE SN# FSD2396074 ACC
1069274 | 199404 | ACC | CMPLT | CMPLT | 2023-11-28 18:27 | 2023-11-29 15:05 | MSH 5 US PVT-781 VTE SN FSD2396070 ACC | (TC 11/29/2023 03:05:18 PM, Whittington, Scott) MSH 5 US PVT-781 VTE SN FSD2396070 ACC | MSH 5 US PVT-781 VTE SN FSD2396070 ACC
1069275 | 199405 | CM | CMPLT | CMPLT | 2023-11-28 18:28 | 2023-11-29 15:06 | MSH 5 US PVT 781VTE SN# FSD2396071 ACC | (TC 11/29/2023 03:05:55 PM, Whittington, Scott) MSH 5 US PVT 781VTE SN# FSD2396071 ACC | MSH 5 US PVT 781VTE SN# FSD2396071 ACC
1069276 | 199406 | CM | CMPLT | CMPLT | 2023-11-28 18:29 | 2023-11-29 15:06 | MSH 5 US PVT 781 VTE SN# FSD2396072 ACC | (TC 11/29/2023 03:06:34 PM, Whittington, Scott) MSH 5 US PVT 781 VTE SN# FSD2396072 ACC | MSH 5 US PVT 781 VTE SN# FSD2396072 ACC
1069277 | 199407 | ACC | CMPLT | CMPLT | 2023-11-28 18:31 | 2023-11-30 11:25 | MSH 5 US PVT 781VTE SN# FSD2396073 ACC | (TC 11/30/2023 11:25:03 AM, Whittington, Scott) MSH 5 US PVT 781VTE SN# FSD2396073 ACC | MSH 5 US PVT 781VTE SN# FSD2396073 ACC
1069278 | 199408 | CM | CMPLT | CMPLT | 2023-11-28 18:35 | 2023-11-30 11:26 | MSH 5 US PLI-2004BX SN# FSC2344214 ACC | (TC 11/30/2023 11:25:50 AM, Whittington, Scott) MSH 5 US PLI-2004BX SN# FSC2344214 ACC | MSH 5 US PLI-2004BX SN# FSC2344214 ACC
1069279 | 199409 | CM | CMPLT | CMPLT | 2023-11-28 18:37 | 2023-11-30 11:27 | MSH 5 US PLI-705BX SN# FSD2396114 ACC | (TC 11/30/2023 11:26:24 AM, Whittington, Scott) MSH 5 US PLI-705BX SN# FSD2396114 ACC | MSH 5 US PLI-705BX SN# FSD2396114 ACC
1069280 | 199410 | CM | CMPLT | CMPLT | 2023-11-28 18:39 | 2023-11-30 11:26 | MSH 5 US PLT-1202BT SN# FSC2355690 ACC | (TC 11/30/2023 11:26:47 AM, Whittington, Scott) MSH 5 US PLT-1202BT SN# FSC2355690 ACC | MSH 5 US PLT-1202BT SN# FSC2355690 ACC
1069281 | 199411 | CM | CMPLT | CMPLT | 2023-11-28 18:41 | 2023-11-30 12:16 | MSH 5 US PVI-475BX SN# FSD2383631 ACC | (TC 11/30/2023 12:15:30 PM, Whittington, Scott) MSH 5 US PVI-475BX SN# FSD2383631 ACC | MSH 5 US PVI-475BX SN# FSD2383631 ACC
1069282 | 199412 | ACC | CMPLT | CMPLT | 2023-11-28 20:13 | 2023-11-30 13:12 | MSH 5 US PROBE PLI-705BX SN# FSD2396088 ACC | (TC 11/30/2023 01:11:54 PM, Whittington, Scott) MSH 5 US PROBE PLI-705BX SN# FSD2396088 ACC | MSH 5 US PROBE PLI-705BX SN# FSD2396088 ACC
1069283 | 199413 | ACC | CMPLT | CMPLT | 2023-11-28 20:17 | 2023-11-30 13:12 | MSH 5 US PROBE PVI-475BX SN# FSD2383630 ACC | (TC 11/30/2023 01:12:30 PM, Whittington, Scott) MSH 5 USD PROBE PVI-475BX SN# FSD2383630 ACC | MSH 5 USD PROBE PVI-475BX SN# FSD2383630 ACC
32520 | 199414 | CM | CMPLT | CMPLT | 2023-11-28 21:17 | 2023-11-29 08:55 | MSH DR2 TUBE LONGITUDINAL MOVEMENT | (TC 11/28/2023 11:54:00 PM, Whittington, Scott) MSH DR2 TUBE LONGITUDINAL MOVEMENT | MSH DR2 TUBE LONGITUDINAL MOVEMENT
| 199415 | CM | CMPLT | CMPLT | 2023-11-29 08:15 | 2023-12-06 08:01 | O-rings for biomed shop | (TC 11/29/2023 08:29:21 AM, Rhuland, Courtney) PO #7202328398
68788 | 199416 | CM | CMPLT | NFF | 2023-11-29 08:27 | 2023-11-29 08:28 | not able to associate on powerchart or i aware | (TC 11/29/2023 08:26:48 AM, Campbell, Connor) checked error logs, nothing pertaining to network was found tested network connectivity, connected to the network fine no fault found returned to service
66471 | 199417 | CM | CMPLT | CMPLT | 2023-11-29 09:22 | 2023-11-29 09:22 | Needs O2 cell calibration(set to FiO2 0.21 read 0.18) | tested unit FiO2 at 21 read 21. calibrated PATO(O2). tested over night. put 50 and read 50. Functional. Returned unit back to RT and informed Rex.
1046572 | 199418 | CM | CMPLT | CMPLT | 2023-11-29 09:55 | 2024-02-07 16:08 | not working | (TC 11/27/2023 04:08:50 PM, Huang, Jessica) picked up the device from the OR desk shipped the device for warranty repair followed up with the vendor and got the device back tested it, sent it back to OR
1054543 | 199419 | CM | CMPLT | CMPLT | 2023-11-29 09:56 | 2023-11-30 11:04 | Not turning on | (TC 11/30/2023 10:58:59 AM, Baghirzade, Nara) Power cord was not seated properly; according to Courtney, the battery was depleted entirely. Charged overnight and battery level is full. Operated the ECG on battery to do functional tests. PIM test passed. Printer test passed. Connected ECG to simulator, 60 BPM, all waveforms are good, and printed the waveforms. ECG is fully functional.
1069284 | 199420 | ACC | CMPLT | CMPLT | 2023-11-29 11:17 | 2023-12-01 16:44 | MSH US ER G7 PROBE PVT 781 VTE SN# FSD2375604 ACCEPTANCE | (TC 12/01/2023 04:44:46 PM, Whittington, Scott) MSH US ER G7 PVT 781 VTE SN# FSD2375604 ACCEPTANCE | MSH US ER G7 PVT 781 VTE SN# FSD2375604 ACCEPTANCE
1069285 | 199421 | ACC | CMPLT | CMPLT | 2023-11-29 11:19 | 2023-12-01 16:45 | MSH US ER G7 PROBE PVT 781 VTE SN# FSD2375605 ACCEPTANCE | (TC 12/01/2023 04:45:20 PM, Whittington, Scott) MSH US ER G7 PROBE PVT 781 VTE SN# FSD2375605 ACCEPTANCE | MSH US ER G7 PROBE PVT 781 VTE SN# FSD2375605 ACCEPTANCE
1069286 | 199422 | ACC | CMPLT | CMPLT | 2023-11-29 11:23 | 2023-12-01 16:46 | MSH US ER G7 PROBE PVT781VTE SN#FSD2375606 ACCEPTANCE | (TC 12/01/2023 04:45:51 PM, Whittington, Scott) MSH US ER G7 PROBE PVT781VTE SN#FSD2375606 ACCEPTANCE | MSH US ER G7 PROBE PVT781VTE SN#FSD2375606 ACCEPTANCE
1069288 | 199423 | ACC | CMPLT | CMPLT | 2023-11-29 11:26 | 2023-12-07 11:38 | MSH US ER G7 PROBE PVT781VTE SN# FSC2334611 ACCEPTANCE | (TC 12/07/2023 11:38:22 AM, Whittington, Scott) MSH US ER G7 PROBE PVT781VTE SN# FSC2334611 ACCEPTANCE | MSH US ER G7 PROBE PVT781VTE SN# FSC2334611 ACCEPTANCE
1069290 | 199424 | ACC | CMPLT | CMPLT | 2023-11-29 11:32 | 2023-12-07 23:03 | MSH US ER G7 PROBE PVT781VTE SN# FSD2375655 ACCEPTANCE | (TC 12/07/2023 11:03:21 PM, Whittington, Scott) MSH US ER G7 PROBE PVT781VTE SN# FSD2375655 ACCEPTANCE | MSH US ER G7 PROBE PVT781VTE SN# FSD2375655 ACCEPTANCE
1069291 | 199425 | ACC | CMPLT | CMPLT | 2023-11-29 11:35 | 2023-12-07 23:04 | MSH US ER G7 PROBE PVT781VTE SN# FSD2375656 ACCEPTANCE | (TC 12/07/2023 11:04:20 PM, Whittington, Scott) MSH US ER G7 PROBE PVT781VTE SN# FSD2375656 ACCEPTANCE | MSH US ER G7 PROBE PVT781VTE SN# FSD2375656 ACCEPTANCE
1069292 | 199426 | CM | CMPLT | CMPLT | 2023-11-29 11:42 | 2023-12-07 23:05 | MSH US ER G7 PROBE PVT781VTE SN# FSD2375657 ACCEPTANCE | (TC 12/07/2023 11:05:25 PM, Whittington, Scott) MSH US ER G7 PROBE PVT781VTE SN# FSD2375657 ACCEPTANCE | MSH US ER G7 PROBE PVT781VTE SN# FSD2375657 ACCEPTANCE
1069293 | 199427 | ACC | CMPLT | CMPLT | 2023-11-29 11:43 | 2023-12-08 13:45 | MSH US ER G7 PROBE PLI-2002BT SN# FSC2343246 ACCEPTANCE | (TC 12/08/2023 01:45:02 PM, Whittington, Scott) MSH US ER G7 PROBE PVT781VTE SN# FSD2375655 ACCEPTANCE | MSH US ER G7 PROBE PVT781VTE SN# FSD2375655 ACCEPTANCE
1069294 | 199428 | CM | CMPLT | CMPLT | 2023-11-29 12:24 | 2023-12-08 13:46 | MSH US ER G7 PROBE PLI-705BX SN# FSC2335593 ACCEPTANCE | (TC 12/08/2023 01:45:59 PM, Whittington, Scott) MSH US ER G7 PROBE PLI-705BX SN# FSC2335593 ACCEPTANCE | MSH US ER G7 PROBE PLI-705BX SN# FSC2335593 ACCEPTANCE
1069295 | 199429 | ACC | CMPLT | CMPLT | 2023-11-29 12:27 | 2023-12-11 16:38 | MSH US ER G7 PROBE PLI-1205BX SN# FSD2385127 ACCEPTANCE | (TC 12/11/2023 04:37:51 PM, Whittington, Scott) MSH US ER G7 PROBE PLI-1205BX SN# FSD2385127 ACCEPTANCE (TC 08/28/2024 11:44:44 AM, Dolinar, John) MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2472163 REPLACES SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL AUG 28 2024 | MSH US ER G7 PROBE PLI-1205BX SN# FSD2385127 ACCEPTANCE
1069296 | 199430 | ACC | CMPLT | CMPLT | 2023-11-29 12:36 | 2023-12-14 14:52 | MSH US ER G7 PROBE PV-475BX SN# FSC2322331 ACCEPTANCE | (TC 12/14/2023 02:52:26 PM, Whittington, Scott) MSH US ER G7 PROBE PV-475BX SN# FSC2322331 ACCEPTANCE | MSH US ER G7 PROBE PV-475BX SN# FSC2322331 ACCEPTANCE
1069297 | 199431 | ACC | CMPLT | CMPLT | 2023-11-29 12:39 | 2023-12-14 14:53 | MSH US ER G7 PROBE PVI-574BX SN# FSB2383975 ACCEPTANCE | (TC 12/14/2023 02:53:25 PM, Whittington, Scott) MSH US ER G7 PROBE PVI-574BX SN# FSB2383975 ACCEPTANCE | MSH US ER G7 PROBE PVI-574BX SN# FSB2383975 ACCEPTANCE
| 199432 | CM | CMPLT | CMPLT | 2023-11-29 13:34 | 2023-11-29 13:34 | biomed meeting | biomed meeting
| 199433 | CM | CMPLT | CMPLT | 2023-11-29 14:24 | 2023-12-07 09:42 | L&D spare parts | (TC 11/29/2023 02:29:00 PM, Rendon Lopez , Elizabeth) L&D panda warmer and doppler spare parts Received
1049904 | 199434 | CM | CMPLT | CMPLT | 2023-11-29 14:39 | 2023-11-30 15:32 | MSH ER R64 - Broken Battery Contacts | (TC 11/29/2023 02:44:46 PM, Chau, Ken) Replaced contacts, detector calibrated Ordering system batteries to be replaced (TC 11/30/2023 03:31:03 PM, Fan, Tak) Replaced batteries.
| 199435 | CM | CMPLT | CMPLT | 2023-11-29 14:44 | 2023-11-29 14:44 | Huddle | (TC 11/29/2023 02:43:52 PM, Rhuland, Courtney) Weekly huddle
500859 | 199436 | CM | CMPLT | CMPLT | 2023-11-29 15:09 | 2023-11-29 15:09 | MSH US APLIO XG REMOVED FROM ROOM AND SORED AT MSH AS CANON PRISM INSTALLED | (TC 11/29/2023 03:09:19 PM, Whittington, Scott) MSH US APLIO XG REMOVED FROM ROOM AND SORED AT MSH AS CANON PRISM INSTALLED | MSH US APLIO XG REMOVED FROM ROOM AND SORED AT MSH AS CANON PRISM INSTALLED
500868 | 199437 | CM | CMPLT | CMPLT | 2023-11-29 15:10 | 2023-11-29 15:10 | MSH US APLIO XG REMOVED FROM ROOM AND SORED AT MSH AS CANON PRISM INSTALLED | (TC 11/29/2023 03:10:34 PM, Whittington, Scott) MSH US APLIO XG REMOVED FROM ROOM AND SORED AT MSH AS CANON PRISM INSTALLED | MSH US APLIO XG REMOVED FROM ROOM AND SORED AT MSH AS CANON PRISM INSTALLED
1063000 | 199438 | PM | CMPLT | CMPLT | 2023-11-29 16:10 | 2023-11-29 16:12 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/29/2023 04:10:13 PM, Cortez, John) Completed cleaning cycle. During rinse, temperature overshot to 26.1 deg C, then settled to set temperature. Port 2 also took awhile to show up as connected. Will monitor.
1060518 | 199439 | CM | CMPLT | CMPLT | 2023-11-29 20:47 | 2023-11-29 20:46 | Intermittant failure of table height motion. | Replaced AMC and potentiometer for this mption. Fixed the problem.
67943 | 199440 | CM | CMPLT | CMPLT | 2023-11-30 09:30 | 2023-11-30 10:28 | Scale not showing up | (TC 11/30/2023 10:26:22 AM, Rhuland, Courtney) Scale not registering with incubator Tried with different scale and same issue Opened back and noticed one of the pins were out of place Reseated pin in proper place and issue was resolved Calibrated and verified scale Functional check - OK
65900 | 199441 | CM | CMPLT | NFF | 2023-11-30 09:46 | 2023-11-30 09:46 | channel error | (TC 11/30/2023 09:45:52 AM, Campbell, Connor) preformed a check in, passed no fault found
65956 | 199442 | CM | CMPLT | NFF | 2023-11-30 10:09 | 2023-11-30 10:10 | unspecified channel error | (TC 11/30/2023 10:09:15 AM, Campbell, Connor) performed a check in, passed performed a test infusion, no error found
| 199443 | CM | CMPLT | CMPLT | 2023-11-30 10:29 | 2023-11-30 10:30 | Meeting | (TC 11/30/2023 10:28:38 AM, Rhuland, Courtney) Meeting with Andrew and Nara to discuss strategy to tag untagged devices in ENT and dental clinic
1061210 | 199444 | CM | CMPLT | CMPLT | 2023-11-30 11:12 | 2023-11-30 11:14 | Vitals not showing on the monitors | (TC 11/30/2023 11:12:44 AM, Reid, Emmerson ) Went to OR 9 to check the issue. Monitor was coming in and out of connection very easily. I got it settled and left, but the issue came back after 30 minutes. I went back up with a new cable and connector, I noticed the connector was damaged so I replaced it. Working now.
501067 | 199445 | CM | CMPLT | CMPLT | 2023-11-30 12:03 | 2023-12-06 15:04 | System noisy | (TC 11/30/2023 12:02:39 PM, Fan, Tak) Traced source of noise. GPU fan. Cleaned, system less nosiy after but still noisy. Placed call for new GPU. (TC 12/06/2023 08:03:00 AM, Fan, Tak) Replaced GPU. | GPU fan is loud
| 199446 | CM | CMPLT | CMPLT | 2023-11-30 13:37 | 2024-02-26 13:29 | MSH IR ANGIO PHILPS IVUS FROM TWH IR MOVED TO MSH IR ANGIO SETUP FOR MSH MWL | (TC 11/30/2023 01:37:08 PM, Whittington, Scott) MSH IR ANGIO PHILPS IVUS FROM TWH IR MOVED TO MSH IR ANGIO SETUP FOR MSH MWL (TC 12/19/2023 12:56:37 PM, Whittington, Scott) MSH IR ANGIO PHILPS IVUS FROM TWH IR MOVED TO MSH IR ANGIO SETUP FOR MSH MWL (TC 12/19/2023 12:56:47 PM, Whittington, Scott) MSH IR ANGIO PHILPS IVUS FROM TWH IR MOVED TO MSH IR ANGIO SETUP FOR MSH MWL (TC 12/19/2023 05:56:00 PM, Whittington, Scott) MSH IR ANGIO PHILPS IVUS FROM TWH IR MOVED TO MSH IR ANGIO SETUP FOR MSH MWL | MSH IR ANGIO PHILPS IVUS FROM TWH IR MOVED TO MSH IR ANGIO SETUP FOR MSH MWL
| 199447 | CM | CMPLT | CMPLT | 2023-11-30 14:29 | 2023-11-30 14:37 | Ophthalmology clinic inventory and walkthrough | (TC 11/30/2023 02:28:56 PM, Baghirzade, Nara) Inventory and walkthrough (TC 11/30/2023 02:36:14 PM, Rhuland, Courtney) Inventory and walkthrough
1049362 | 199448 | CM | CMPLT | CMPLT | 2023-11-30 20:44 | 2023-12-01 15:26 | Injector piston end pins stuck | Moved the spare injector to ER but failed due to a bad head cable Brought the injector back to the shop and replaced the head cable Swapped main injector with the spare and brought it to PMH shop to work on it Took apart the piston end and cleaned all the contrast Tested system and passed all functionality tests
63289 | 199449 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
22691 | 199450 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26100 | 199451 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26101 | 199452 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26104 | 199453 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26110 | 199454 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26112 | 199455 | PM | CMPLT | CMPLT | 2023-12-01 02:17 | 2024-11-08 12:20 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 11/08/2024 12:19:46 PM, Rhuland, Courtney) PM completed by Spectramedx Returned and functional check completed 21% O2 - 20.9% O2 40% O2 - 37.5% O2 60% O2 - 58.2% O2 80% O2 - 79.1% O2 100% O2 - 98.3% O2 Returned to RT
26113 | 199456 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26117 | 199457 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26146 | 199458 | PM | ACTIV | ISSUE | 2023-12-01 02:17 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26798 | 199459 | PM | CMPLT | CMPLT | 2023-12-01 02:17 | 2024-01-12 12:34 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 01/12/2024 12:31:46 PM, Rhuland, Courtney) PM complete See attached checklist for details
50605 | 199460 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
62905 | 199461 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
62913 | 199462 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
62916 | 199463 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63059 | 199464 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63125 | 199465 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63135 | 199466 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63239 | 199467 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Intra-Operative Heated Chemotherapy Pump
63246 | 199468 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Intra-Operative Heated Chemotherapy Pump
63303 | 199469 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-12 09:28 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/11/2023 09:27:05 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
63813 | 199470 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infant T-Piece Resuscitator | (TC 03/27/2024 07:58:01 AM, Rhuland, Courtney) unable to locate
64993 | 199471 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-12 09:07 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 12/07/2023 09:07:07 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
64994 | 199472 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-03 10:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/03/2024 10:21:02 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
65124 | 199473 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-15 09:42 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/15/2024 09:40:44 AM, Reid, Emmerson ) PM Complete, see procedure for details.
65127 | 199474 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65158 | 199475 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65167 | 199476 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65169 | 199477 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65389 | 199478 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
65754 | 199479 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM CO2 Module
65849 | 199480 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
65861 | 199481 | PM | CNP | ENL | 2023-12-01 02:18 | 2024-09-05 11:25 | PM-MAJOR - Annual PM For: Aspirators, Emergency
65867 | 199482 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
65875 | 199483 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65886 | 199484 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65901 | 199485 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65906 | 199486 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65934 | 199487 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-10 14:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:41:04 PM, Rhuland, Courtney) PM completed by BD
65957 | 199488 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65988 | 199489 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-21 08:30 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 08:29:16 AM, Rhuland, Courtney) PM completed by BD
66072 | 199490 | PM | CNP | ENL | 2023-12-01 02:18 | 2024-08-14 11:58 | PM-MAJOR - Annual PM For: Universal Mobile Lift | Could not locate at 17 Murray floor. Department did not know its location either. Nazanin(CE) is still looking for this portable patient lift.
66107 | 199491 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-20 09:03 | PM-MAJOR- Annual PM for: Pulmo Vista | (TC 12/20/2023 08:58:20 AM, Canjimavadivel, Kamallesh ) PM completed by vendor on Oct 29, 2023. See attached report
66188 | 199492 | PM | CNP | ENL | 2023-12-01 02:18 | 2024-08-14 12:54 | PM-MAJOR - Annual PM For: Aspirators, Emergency
66737 | 199493 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-24 11:55 | PM-MAJOR - Annual PM For: Giraffe Shuttle | (TC 01/24/2024 11:54:44 AM, Baghirzade, Nara) PM passed and completed. Fan is functional. Alarms work when foot pedal is activated. Battery health is good; green LED illuminated. Electrical safety: power cord resistance = 250 mohms leakage current = 82.9 uA (TC 01/24/2024 11:55:12 AM, Rhuland, Courtney) PM passed and completed. Fan is functional. Alarms work when foot pedal is activated. Battery health is good; green LED illuminated. Electrical safety: power cord resistance = 250 mohms leakage current = 82.9 uA
67224 | 199494 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:06 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/26/2024 09:06:14 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67228 | 199495 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-12 09:12 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/08/2023 09:11:55 AM, Rendon Lopez , Elizabeth) Replaced blender. PM completed, all tests passed
67229 | 199496 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:04 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/26/2024 09:03:06 AM, Rendon Lopez , Elizabeth) Replaced fuses on power board. Replaced blender. PM completed, all tests passed
67230 | 199497 | PM | CMPLT | PMFAIL | 2023-12-01 02:18 | 2023-12-12 09:09 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/07/2023 09:08:03 AM, Rendon Lopez , Elizabeth) Found Resus O2 knob loose during PM, removed from room
67231 | 199498 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-06 16:33 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/01/2023 04:32:11 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67232 | 199499 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:15 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/29/2024 09:15:05 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67233 | 199500 | PM | CMPLT | PMFAIL | 2023-12-01 02:18 | 2023-12-12 09:25 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/11/2023 09:23:59 AM, Rendon Lopez , Elizabeth) Resus unit leaking when in off position. Removed warmer from room for repair
67234 | 199501 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-12 09:11 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/07/2023 09:10:49 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. In OR5
67235 | 199502 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-24 11:54 | PM-MAJOR - Annual PM For: Infant Resuscitation System | (TC 01/24/2024 11:53:10 AM, Baghirzade, Nara) PM completed. All values are within manufacturer specifications. With flow adjusted to 0 L/min, checked that no flow is going through. Passed. Set flow to 7 L/min - @0 O2% = 0 @20 O2% = 20.9 @40 O2% = 38 @60 O2% = 59 @80 O2% = 79.4 @100 O2% = 99.3 Checked flow @ 5 L/min = 4.7 L/min @ 7 L/min = 6.8 L/min @ 10 L/min = 10 L/min (TC 01/24/2024 11:54:01 AM, Rhuland, Courtney) PM completed. All values are within manufacturer specifications. With flow adjusted to 0 L/min, checked that no flow is going through. Passed. Set flow to 7 L/min - @0 O2% = 0 @20 O2% = 20.9 @40 O2% = 38 @60 O2% = 59 @80 O2% = 79.4 @100 O2% = 99.3 Checked flow @ 5 L/min = 4.7 L/min @ 7 L/min = 6.8 L/min @ 10 L/min = 10 L/min
67375 | 199503 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67379 | 199504 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-10-31 10:36 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 10/31/2024 10:36:22 AM, Rhuland, Courtney) PM completed by Spectramedx Alarm needed to be replaced Returned and functional check completed 21% O2 - 20.9% O2 40% O2 - 38.8% O2 60% O2 - 60.8% O2 80% O2 - 80.7% O2 100% O2 - 99.9% O2 Returned to RT
67381 | 199505 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-13 11:18 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 02/13/2024 11:16:03 AM, Baghirzade, Nara) Functional testing completed. 21% = 20.9% 40% = 39.8% 60% = 61.2% 80% = 79.8% 100% = 99.1% When disconnecting air OR oxygen, blender alarm is functional. Attached service report.
67383 | 199506 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67390 | 199507 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-06 16:33 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/01/2023 04:32:55 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67391 | 199508 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-05 17:09 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 11/29/2023 05:08:21 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67394 | 199509 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 15:38 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/31/2024 03:37:37 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67395 | 199510 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-06 16:49 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/05/2023 04:48:23 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67396 | 199511 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:11 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/29/2024 09:10:59 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67410 | 199512 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:07 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/26/2024 09:06:48 AM, Rendon Lopez , Elizabeth) Missing diaphgram, replaced. PM completed, all tests passed
67412 | 199513 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-12 09:39 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/11/2023 09:39:10 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67414 | 199514 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:10 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/29/2024 09:09:56 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67419 | 199515 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 15:39 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/31/2024 03:38:34 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67420 | 199516 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-16 08:39 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 02/15/2024 08:38:52 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67421 | 199517 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-16 08:37 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 02/15/2024 08:36:45 AM, Rendon Lopez , Elizabeth) Missing diaphgram, replaced. PM completed, all tests passed
67423 | 199518 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-06 16:50 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/05/2023 04:49:59 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67424 | 199519 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-06 16:53 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/06/2023 04:52:46 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67427 | 199520 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:12 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/29/2024 09:12:16 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67873 | 199521 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-27 13:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 02/26/2024 01:39:54 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67875 | 199522 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-05-22 15:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/22/2024 03:15:12 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67884 | 199523 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-03 11:35 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/03/2024 11:34:37 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67886 | 199524 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-29 09:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 02/28/2024 09:47:42 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67889 | 199525 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-26 12:02 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | Refer to WO 201086
67891 | 199526 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-12 09:03 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 12/07/2023 09:02:55 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
68375 | 199527 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68460 | 199528 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Hysteroscopic Fluid Management System
68537 | 199529 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68589 | 199530 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
68597 | 199531 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
68604 | 199532 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
68617 | 199533 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68678 | 199534 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-25 10:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68707 | 199535 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-04 13:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/18/2024 01:55:56 PM, Huang, Jessica) replaced the power cord and completed PM
68725 | 199536 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68731 | 199537 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68749 | 199538 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-13 09:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:23:39 AM, Rendon Lopez , Elizabeth) PM completed by BD
68767 | 199539 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68801 | 199540 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-21 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68831 | 199541 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-21 09:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68843 | 199542 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68895 | 199543 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68909 | 199544 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68977 | 199545 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69017 | 199546 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69056 | 199547 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69063 | 199548 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-12 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 01:54:52 PM, Reid, Emmerson ) needed to fix the latch. PM Complete.
69069 | 199549 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-26 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69074 | 199550 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-20 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69099 | 199551 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69124 | 199552 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-12 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69137 | 199553 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69153 | 199554 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-19 19:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69168 | 199555 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69187 | 199556 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69208 | 199557 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-15 13:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69301 | 199558 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-25 11:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69335 | 199559 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-19 20:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69336 | 199560 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-19 20:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69370 | 199561 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-21 20:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69394 | 199562 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-03 14:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:28:42 PM, Rhuland, Courtney) PM completed by BD
69432 | 199563 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69484 | 199564 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69521 | 199565 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-26 09:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69544 | 199566 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69565 | 199567 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69569 | 199568 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-25 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69581 | 199569 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-14 07:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/14/2024 07:19:36 AM, Rhuland, Courtney) PM completed by BD
69595 | 199570 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69597 | 199571 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69601 | 199572 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69604 | 199573 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69615 | 199574 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69645 | 199575 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69651 | 199576 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69675 | 199577 | PM | CMPLT | PMFAIL | 2023-12-01 02:18 | 2024-03-14 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
530510 | 199578 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:09 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/26/2024 09:08:20 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67358 | 199579 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-05-01 11:12 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 04/05/2024 10:54:46 AM, Reid, Emmerson ) Unit was not turning on. I opened it up and reseated connections and it started working, but not innitially. I found a break in the power cord that was likely causing the intermittent issue. I reached out to the vendor to order a new cord and removed the old one. (TC 04/15/2024 10:52:36 AM, Reid, Emmerson ) Put in a PO for a new power cord (TC 04/24/2024 02:58:59 PM, Reid, Emmerson ) Replaced the power cord and it worked to turn on. Still need to PM. (TC 05/01/2024 11:11:37 AM, Reid, Emmerson ) Completed PM, see procedure for details.
67366 | 199580 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-01 10:06 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 12/01/2023 10:06:05 AM, Campbell, Connor) PM passed/complete, refer to procedure
63962 | 199581 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2023-12-05 17:07 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 11/29/2023 05:06:34 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
63963 | 199582 | PM | CMPLT | PMFAIL | 2023-12-01 02:18 | 2023-12-12 14:34 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/12/2023 02:33:20 PM, Rendon Lopez , Elizabeth) Resus gas outlet is loose. Removed device from LDR 18
30507 | 199583 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM - Annual
32540 | 199584 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM - Monthly
32560 | 199585 | PM | CANCL | CANCL | 2023-12-01 02:18 |  | PM - Monthly | PM canceled due to service calls
32580 | 199586 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-04 15:03 | PM - Monthly | PM Major done NRA - LL and L done Air cal - LL and L done Slice count - 166009 TOS phantom a(-997.7, 10.3) b(335.4, 14.1) c(134.2, 14.1) d(98.1, 14.7) e(-105.1, 13.4) f(1.7, 14.3) -System passed all tests
541281 | 199587 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | BLADDER SCANNER
1046398 | 199588 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1046654 | 199589 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-04-02 16:00 | PM-MAJOR - Annual PM For: Bladder Scanner | (TC 03/26/2024 03:59:46 PM, Huang, Jessica) PM completed
1047035 | 199590 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-19 13:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047847 | 199591 | PM | CANCL | CANCL | 2023-12-01 02:18 | 2024-06-18 12:32 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe
1048199 | 199592 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Insufflators
1048525 | 199593 | PM | CMPLT | PMFAIL | 2023-12-01 02:18 | 2024-03-22 11:59 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 11:58:48 AM, Baghirzade, Nara) Broken rear case and syringe requires full calibration
1048528 | 199594 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-12 13:56 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 01:55:47 PM, Baghirzade, Nara) Replaced rear case. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1048568 | 199595 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-09 15:54 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | (TC 01/31/2024 03:47:38 PM, Rendon Lopez , Elizabeth) Preparing test equipment for in-house PM (TC 02/06/2024 09:53:07 AM, Rendon Lopez , Elizabeth) PM tasks started, internal inspection, cleaning, lubrication, done. (TC 02/07/2024 09:53:07 AM, Rendon Lopez , Elizabeth) Calibration procedures done. Test equipment set up and adjustments (TC 02/08/2024 09:55:33 AM, Rendon Lopez , Elizabeth) Performance verification tests (TC 02/09/2024 03:52:20 PM, Rendon Lopez , Elizabeth) Final inspection done. Sent service forms to Bunnell for record (TC 02/09/2024 03:52:20 PM, Rendon Lopez , Elizabeth) Final inspection done. PM completed, all tests passed. Sent service forms to Bunnell for record
1048571 | 199596 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-09 09:52 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 02/01/2024 01:04:25 PM, Rendon Lopez , Elizabeth) Removed from NICU RT. PM started (TC 02/02/2024 01:04:25 PM, Rendon Lopez , Elizabeth) Patient box inspected, lubriicated, PM Parts replaced. (TC 02/05/2024 03:56:52 PM, Rendon Lopez , Elizabeth) Working on calibration, waiting for load tester set up (TC 02/07/2024 09:51:35 AM, Rendon Lopez , Elizabeth) Calibration procedures done (TC 02/08/2024 09:51:35 AM, Rendon Lopez , Elizabeth) Performance verification tests done. PM completed, all tests passed
1048572 | 199597 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-02-09 09:51 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 02/02/2024 03:29:48 PM, Rendon Lopez , Elizabeth) Patient box inspected, calibrated, PM parts replaced. Performed patient box calibration, having issues with set up- resolved. Calibration done. (TC 02/05/2024 03:55:44 PM, Rendon Lopez , Elizabeth) Working on PM procedures, having issues with load tester. Reached out to Bunnell for tester set up. (TC 02/07/2024 09:49:39 AM, Rendon Lopez , Elizabeth) Calibration procedures done. (TC 02/08/2024 09:49:39 AM, Rendon Lopez , Elizabeth) Performance verification tests done. PM completed, all tests passed
1048662 | 199598 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048701 | 199599 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049956 | 199600 | PM | ACTIV | PV | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Testers
1050087 | 199601 | PM | ACTIV | ISSUE | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
1050091 | 199602 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-01-31 09:08 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 01/26/2024 09:07:43 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1050253 | 199603 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050270 | 199604 | PM | CMPLT | CMPLT | 2023-12-01 02:18 | 2024-03-12 11:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050272 | 199605 | PM | CNP | ENL | 2023-12-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050478 | 199606 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-01 11:07 | PM-MAJOR- Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | SW upgraded to 4.4. Replaced door hinge. Performed PM. Passed pre-use check. Functional. Membrane %30. Next one 5000 hours
1050481 | 199607 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-04-08 09:24 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | Completed annual PM. Replaced expiratory cassette's membrane. Passed device check. Tested. Functional. leakage 21 uA. Resistance 110 mOhm
1050482 | 199608 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-01-10 11:50 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | Performed PM Pre check pass
1050619 | 199609 | PM | CMPLT | PMFAIL | 2023-12-01 02:19 | 2024-03-13 09:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1051370 | 199610 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-15 14:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051384 | 199611 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
1053339 | 199612 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-07-25 13:22 | PM-MAJOR- Annual PM for: Testers | (TC 07/25/2024 01:22:08 PM, Chukri, Samir) Received from TSI with Calibration certificate attached.
1053342 | 199613 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-07-24 14:13 | PM-MAJOR- Annual PM for: Testers | (TC 07/24/2024 02:13:01 PM, Chukri, Samir) Completed by TSI certificate calibration. attached.
1053345 | 199614 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-08-14 07:49 | PM-MAJOR- Annual PM for: Testers | (TC 07/31/2024 10:41:44 AM, Chukri, Samir) shipping the unit to TSI for certifier calibration. (TC 08/14/2024 07:49:19 AM, Chukri, Samir) Received from TSI certificate Calibration by TSI
1053346 | 199615 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-08-14 07:53 | PM-MAJOR- Annual PM for: Testers | (TC 07/31/2024 10:45:23 AM, Chukri, Samir) shipping Unit to TSI for Accredited Calibration (TC 08/14/2024 07:53:07 AM, Chukri, Samir) Received unit from TSI. Certificate Calibration by Vendor TSI.
1053734 | 199616 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-02-27 13:27 | PM-MAJOR- Annual PM for: Pacemakers, Cardiac, External | (TC 02/27/2024 01:26:35 PM, Scero, Tomlyn) PM tests complete. Functional and physical checks complete and passed.
1054044 | 199617 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM FOR: TRANSCUTANEOUS CO2 DIO
1055745 | 199618 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1055791 | 199619 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-05-02 15:05 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 05/02/2024 03:03:51 PM, Scero, Tomlyn) Defib PM complete. All tests passed. Results sheet attached to asset.
1056979 | 199620 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-01-23 14:26 | PM-MAJOR - Annual PM For: Infant T-Piece Resuscitator | (TC 01/23/2024 02:25:34 PM, Baghirzade, Nara) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 19 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 68 cmH2O Less than 80 = 75 cmH2O (TC 01/23/2024 02:25:49 PM, Rhuland, Courtney) Performance assurance passed and completed. All values within range. @10 = 10 cmH2O @20 = 19 cmH2O @40 = 39 cmH2O Greater than 60 = passed - 68 cmH2O Less than 80 = 75 cmH2O
1057575 | 199621 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057665 | 199622 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057670 | 199623 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057675 | 199624 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057678 | 199625 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057683 | 199626 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057685 | 199627 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057687 | 199628 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057689 | 199629 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057691 | 199630 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057743 | 199631 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM, SEMI-ANNUAL, ULTRASOUND GENERAL |  | PM, SEMI-ANNUAL, ULTRASOUND GENERAL CANCELLED NO AGREEMENT IN PLACE REPAIRS ONLY
1057845 | 199632 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-10-21 11:59 | PM-MAJOR - Annual PM For: Insufflators | (TC 10/17/2024 11:59:31 AM, Huang, Jessica) shipped the device out for PM and installed it in OR 13
1057853 | 199633 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-01-25 13:08 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 01/25/2024 01:07:20 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure for details
1057865 | 199634 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2023-12-18 08:27 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 12/18/2023 08:26:51 AM, Reid, Emmerson ) Replaced the battery. PM Complete
1057866 | 199635 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2023-12-08 09:15 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 12/08/2023 09:08:54 AM, Canjimavadivel, Kamallesh ) Replaced battery
1057872 | 199636 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Forced-Air
1057873 | 199637 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Forced-Air
1057881 | 199638 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Forced-Air
1057882 | 199639 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Forced-Air
1057885 | 199640 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Forced-Air
1057886 | 199641 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Forced-Air
1057887 | 199642 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Forced-Air
1057938 | 199643 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR-Annual PM for : Table, Operating
1057939 | 199644 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM For : Table, Operating
1057940 | 199645 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM For : Table, Operating
1057941 | 199646 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM for : Table, Operating
1057943 | 199647 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-25 10:47 | PM-MAJOR- Annual PM for : Table, Operating | (TC 03/22/2024 09:57:34 AM, Reid, Emmerson ) Bottom two brake pistons were not engaging. brought down to biomed to fix and PM (TC 03/25/2024 10:45:48 AM, Reid, Emmerson ) Tables needed to be deep cleaned with clorox. PM Complete, see proceedure.
1057945 | 199648 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | MP-MAJOR-Annual PM for : Table, Operating
1057946 | 199649 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR-Annual PM for : Table, Operating
1057947 | 199650 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR-Annual PM for : Table, Operating
1057948 | 199651 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM for : Table, Operating
1057950 | 199652 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-01-10 15:21 | PM-MAJOR- Annual PM for : Table, Operating | (TC 01/10/2024 03:19:41 PM, Reid, Emmerson ) Needed to do a deep clean of several surfaces. Also had to find the grease gun. Completed PM, see proceedure for details.
1057951 | 199653 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057954 | 199654 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057956 | 199655 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057961 | 199656 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057962 | 199657 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057963 | 199658 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057964 | 199659 | PM | ACTIV | READY | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057968 | 199660 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM for : Table, Operating
1057969 | 199661 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM for : Table, Operating
1057970 | 199662 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR-Annual PM for : Table, Operating
1057972 | 199663 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM for : Table, Operating
1057973 | 199664 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR- Annual PM for : Table, Operating
1058163 | 199665 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058164 | 199666 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058165 | 199667 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-02-28 10:54 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Adult/Pediatric | (TC 02/27/2024 10:46:00 AM, Cortez, John) Visual inspection passed Type of HEPA filter that came with PM kit different, Blower flow test fails as a result. Vacuumed original filter and re-installed. Other PM parts replaced. Calibration/general tests passed Preop tests passed Electrical safety: 23 uA Battery SOH >30%, charged overnight Blower 0%, vent at 58 operating hrs PM passed
1058166 | 199668 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-12 10:17 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 03/06/2024 04:24:43 PM, Cortez, John) Swapped out from OR Battery SOH >30% Operating hrs 486 Blower 0% (TC 03/11/2024 08:47:40 AM, Cortez, John) Electrical safety- 23 uA enclosure leakage PM parts replaced Adjustments/calibration tasks complete Component tests passed System tests passed General tasks completed PM complete
1058167 | 199669 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-05-09 11:55 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 05/09/2024 11:53:44 AM, Canjimavadivel, Kamallesh ) Replaced PM Kit parts Started battery calibration All calibrations and tests complete Enclosure Leakage: 24 uA PM complete
1058168 | 199670 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-02-29 11:03 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 02/28/2024 10:57:44 AM, Cortez, John) PM parts replaced, except HEPA filter. Ordered filter different Battery SOH >20% Blower module 0% Visual Inspection passed (TC 02/29/2024 11:02:36 AM, Cortez, John) Cal/component/system tests passed General tasks passed Preoperational check passed Electrical safety: 22 uA enclosure leakage
1058176 | 199671 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2023-12-29 10:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 12/29/2023 10:06:15 AM, Reid, Emmerson ) Pushed APS EPIDURAL - OCT 2023 library
1058178 | 199672 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-06-04 10:04 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/04/2024 10:03:48 AM, Reid, Emmerson ) PM Complete, see procedure for details.
1058182 | 199673 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058185 | 199674 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058189 | 199675 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-07-12 09:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/12/2024 09:22:21 AM, DeGuzman, Leila) PM completed. See procedure for details.
1058194 | 199676 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058221 | 199677 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058222 | 199678 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058223 | 199679 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058225 | 199680 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-12 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058226 | 199681 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-12 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058231 | 199682 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-15 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058232 | 199683 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-21 18:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058233 | 199684 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-20 10:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058234 | 199685 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-20 10:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058240 | 199686 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-20 11:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058244 | 199687 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-12 10:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058251 | 199688 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058254 | 199689 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058256 | 199690 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-14 08:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058258 | 199691 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058260 | 199692 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058263 | 199693 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-27 14:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058264 | 199694 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-12 09:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:32:04 AM, Rendon Lopez , Elizabeth) PM completed by BD
1058265 | 199695 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058268 | 199696 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058271 | 199697 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058280 | 199698 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058282 | 199699 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058283 | 199700 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058284 | 199701 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-12 13:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058288 | 199702 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058526 | 199703 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-02-22 14:36 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 02/20/2024 02:25:36 PM, Cortez, John) Visual test passed Replaced PM parts Obtained spare adult set from Nigel Operating hours: 10658 hrs Blower module at 1% Batteries SOH >20% Had issues with flow sensore test, eventually found out test set up was incorrect. Had filter installed on expiratory port Had issue with blower flow test. Flow 2100 kept failing, but other flow speeds passed. 21820 (tolerance: 21896-35978). Decided to keep going with PM, but will make note of fail and contact Bomimed. All other tests passed. Electrical safety 23 uA enclosure leakage. Preop test completed successfully. Completed PM
1058528 | 199704 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-05 16:21 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 03/04/2024 04:11:33 PM, Cortez, John) Swapped out from ICU pod B Preop tests passed Replaced PM parts (TC 03/05/2024 04:12:56 PM, Cortez, John) Charged battery, Battery SOH >30% 1465 operating hours, blower timer at 0% Visual inspection passed General tasks passed Electrical safety: 25 uA enclosure leakage Adjustement/calibration tasks complete Component tests passed System test>pressure fail. Paw sensor slightly low After troubleshooting and contacting Bomimed, turns out reseating expiratory valve resolves issue. Membrane wasn't fully seated All other tests passed, PM complete
1058529 | 199705 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-04-30 10:40 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 04/29/2024 11:25:49 AM, Canjimavadivel, Kamallesh ) Brought vent down to biomed. Started battery calibration cycle (TC 04/30/2024 10:38:18 AM, Canjimavadivel, Kamallesh ) Replaced PM parts All tests and calibrations pass Leakage current: 23uA
1058530 | 199706 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-06 11:55 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 03/05/2024 04:31:14 PM, Cortez, John) Swapped out unit from ICU pod A Preop tests passed Blower at 0%, battery SOH >30% Replaced PM parts (TC 03/05/2024 04:56:03 PM, Cortez, John) 1795 operating hrs (TC 03/06/2024 11:45:32 AM, Cortez, John) Adjustment/calibration tasks completed Component test>air entry failed. Loosening filter allows it to pass. All other component tests passed System tests passed Electrical safety: 25uAs enclosure leakage PM passed.
1059220 | 199707 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059224 | 199708 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059225 | 199709 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059419 | 199710 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059421 | 199711 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059424 | 199712 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059432 | 199713 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059443 | 199714 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059445 | 199715 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059451 | 199716 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059454 | 199717 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059834 | 199718 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-19 14:51 | PM-MAJOR - Annual PM For: Tourniquets, Pneumatic | (TC 03/05/2024 10:09:05 AM, Canjimavadivel, Kamallesh ) Grabbed device from unit for PM. Was not able to find calibration hose required. Jessica reaching out to vendor (TC 03/19/2024 01:48:38 PM, Canjimavadivel, Kamallesh ) Found calibration hoses. Went through manual to see if all required test equipment was available. Calibration Test: Pass Leak Test: Pass Touch Test: Pass See procedure for electrical safety
1059835 | 199719 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Tourniquets, Pneumatic
1059836 | 199720 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Tourniquets, Pneumatic
1059837 | 199721 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Tourniquets, Pneumatic
1059839 | 199722 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Tourniquets, Pneumatic
1059880 | 199723 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1059942 | 199724 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-02-16 08:39 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/15/2024 08:37:48 AM, Rendon Lopez , Elizabeth) Adjusted resuscitation system flow. PM completed, all tests passed
1059943 | 199725 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2023-12-06 16:45 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/05/2023 04:45:02 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1061392 | 199726 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: Insufflators
1062123 | 199727 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-01 15:16 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 03/01/2024 03:16:05 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1063000 | 199728 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-01-11 10:56 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 01/10/2024 10:55:21 AM, Cortez, John) Completed PM, and cleaning cycle
1064398 | 199729 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-06-26 11:43 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | (TC 01/29/2024 08:56:56 AM, Chukri, Samir) emaild Perry for PM quote. and loaner (TC 01/29/2024 02:32:38 PM, Chukri, Samir) Completed service request form from mmsi company and emailed to Perry. waiting for loaner ventilator to be switch for PM (TC 02/01/2024 01:22:54 PM, Chukri, Samir) Replaced PO for PM and received the Loaner ventilator (TC 02/05/2024 09:23:43 AM, Chukri, Samir) Swapped the loaner SN# 12059 with our MRI ventilators tested by Sandy (R.T) (TC 02/06/2024 09:28:53 AM, Chukri, Samir) going to shipping our ventilator for PM , repaired the courier decoument . RGA/RMA # LS0129 (TC 02/08/2024 10:59:58 AM, Chukri, Samir) Shipped the ventilator today by Purolator Tracking # 334547070444 (TC 06/20/2024 01:30:40 PM, Chukri, Samir) Received ventilator calibrated by Airon comany Please see attached PM certificate (TC 06/26/2024 11:43:07 AM, Chukri, Samir) Complete installation on the MRI incubator and swap it with the loaner retun the loaner to Perry .
1065320 | 199730 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1065321 | 199731 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1065322 | 199732 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1065323 | 199733 | PM | ACTIV | ISSUE | 2023-12-01 02:19 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1065686 | 199734 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065687 | 199735 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065690 | 199736 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065691 | 199737 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-25 10:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065692 | 199738 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065693 | 199739 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065694 | 199740 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065695 | 199741 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065696 | 199742 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065697 | 199743 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 13:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065698 | 199744 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065699 | 199745 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-25 11:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065700 | 199746 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065701 | 199747 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 15:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065702 | 199748 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-25 11:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065703 | 199749 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065704 | 199750 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 15:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065705 | 199751 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065706 | 199752 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065707 | 199753 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065708 | 199754 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065709 | 199755 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065710 | 199756 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065711 | 199757 | PM | CNP | ENL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065712 | 199758 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 15:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065713 | 199759 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 13:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065714 | 199760 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-18 15:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065715 | 199761 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-27 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065716 | 199762 | PM | CANCL | CANCL | 2023-12-01 02:19 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065717 | 199763 | PM | CMPLT | CMPLT | 2023-12-01 02:19 | 2024-03-27 14:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065718 | 199764 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065719 | 199765 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065720 | 199766 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065721 | 199767 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-04-26 08:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065722 | 199768 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-25 11:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065723 | 199769 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 15:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065724 | 199770 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 13:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065725 | 199771 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065726 | 199772 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065727 | 199773 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065728 | 199774 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-27 14:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065729 | 199775 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065730 | 199776 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-27 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065731 | 199777 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065732 | 199778 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065733 | 199779 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065734 | 199780 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 13:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065735 | 199781 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065739 | 199782 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 15:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065740 | 199783 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-21 18:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065741 | 199784 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 09:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065742 | 199785 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065743 | 199786 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 15:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065744 | 199787 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065745 | 199788 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065746 | 199789 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065747 | 199790 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-21 18:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065748 | 199791 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065749 | 199792 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 09:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065750 | 199793 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065751 | 199794 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065752 | 199795 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 08:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065753 | 199796 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065754 | 199797 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 10:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065755 | 199798 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 15:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065756 | 199799 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 15:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065757 | 199800 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065758 | 199801 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 08:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065759 | 199802 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065760 | 199803 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-21 18:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065761 | 199804 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 11:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065762 | 199805 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065763 | 199806 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065764 | 199807 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065765 | 199808 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065766 | 199809 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 15:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065767 | 199810 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065768 | 199811 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065769 | 199812 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065770 | 199813 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065771 | 199814 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 08:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065772 | 199815 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065773 | 199816 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065774 | 199817 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-04-26 10:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065775 | 199818 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 09:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065776 | 199819 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065777 | 199820 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 08:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065778 | 199821 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065779 | 199822 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 11:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065780 | 199823 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065781 | 199824 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065782 | 199825 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 09:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065783 | 199826 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 10:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065784 | 199827 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 11:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065785 | 199828 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 09:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065786 | 199829 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065787 | 199830 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065788 | 199831 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065789 | 199832 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 08:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065790 | 199833 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065791 | 199834 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 10:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065792 | 199835 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 08:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065793 | 199836 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 09:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065794 | 199837 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-11 19:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065795 | 199838 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 10:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065796 | 199839 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065797 | 199840 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 11:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065798 | 199841 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 08:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065799 | 199842 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065800 | 199843 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065801 | 199844 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065802 | 199845 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-14 09:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065803 | 199846 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 11:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065804 | 199847 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065805 | 199848 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065806 | 199849 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065807 | 199850 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065808 | 199851 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065809 | 199852 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065810 | 199853 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 10:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065811 | 199854 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 10:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065812 | 199855 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065813 | 199856 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 13:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065814 | 199857 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065815 | 199858 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065816 | 199859 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-07-03 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:43:06 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered
1065817 | 199860 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065818 | 199861 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065819 | 199862 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065820 | 199863 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 10:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065821 | 199864 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-15 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065822 | 199865 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065823 | 199866 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065824 | 199867 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065825 | 199868 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 10:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065826 | 199869 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065827 | 199870 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065828 | 199871 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065829 | 199872 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-21 18:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065830 | 199873 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-04-26 10:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065831 | 199874 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065832 | 199875 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065833 | 199876 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065834 | 199877 | PM | CNP | ENL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065835 | 199878 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-18 10:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065836 | 199879 | PM | CANCL | CANCL | 2023-12-01 02:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065837 | 199880 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065838 | 199881 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 14:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065839 | 199882 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065840 | 199883 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065841 | 199884 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-19 14:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065842 | 199885 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-22 08:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065843 | 199886 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-20 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065844 | 199887 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-12 11:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065845 | 199888 | PM | CMPLT | CMPLT | 2023-12-01 02:20 | 2024-03-21 18:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065846 | 199889 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065847 | 199890 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065848 | 199891 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-22 11:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065849 | 199892 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065850 | 199893 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065851 | 199894 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-14 09:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065852 | 199895 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065853 | 199896 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065854 | 199897 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 10:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065855 | 199898 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065856 | 199899 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 10:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065857 | 199900 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065858 | 199901 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-13 08:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 08:20:06 AM, Reid, Emmerson ) PM Complete.
1065859 | 199902 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 10:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065860 | 199903 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-20 14:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065861 | 199904 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065862 | 199905 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-04-04 13:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:28:53 PM, Rhuland, Courtney) PM completed by BD
1065863 | 199906 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065864 | 199907 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065865 | 199908 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065866 | 199909 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-20 14:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065867 | 199910 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065868 | 199911 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065869 | 199912 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-20 14:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065870 | 199913 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 18:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065871 | 199914 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065872 | 199915 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-04-24 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065873 | 199916 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065874 | 199917 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-22 13:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065875 | 199918 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065876 | 199919 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065877 | 199920 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065878 | 199921 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 11:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065879 | 199922 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065880 | 199923 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065881 | 199924 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065882 | 199925 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065883 | 199926 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065884 | 199927 | PM | CNP | ENL | 2023-12-01 02:21 | 2023-08-09 16:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065885 | 199928 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065886 | 199929 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065887 | 199930 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065888 | 199931 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065889 | 199932 | PM | CMPLT | PMFAIL | 2023-12-01 02:21 | 2024-03-14 14:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1065890 | 199933 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065891 | 199934 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065892 | 199935 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065893 | 199936 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065894 | 199937 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 18:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065895 | 199938 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065896 | 199939 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-04-26 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065897 | 199940 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065898 | 199941 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065899 | 199942 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065900 | 199943 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065901 | 199944 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-04-24 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065902 | 199945 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065903 | 199946 | PM | CANCL | CANCL | 2023-12-01 02:21 | 2024-04-02 16:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065904 | 199947 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-04-26 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065905 | 199948 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065906 | 199949 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065907 | 199950 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065908 | 199951 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065909 | 199952 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065910 | 199953 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065911 | 199954 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065912 | 199955 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065913 | 199956 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065914 | 199957 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065915 | 199958 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065916 | 199959 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-19 14:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065917 | 199960 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065918 | 199961 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065919 | 199962 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065920 | 199963 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065921 | 199964 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065922 | 199965 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065923 | 199966 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065924 | 199967 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 11:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065925 | 199968 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065926 | 199969 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 18:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065927 | 199970 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 11:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065928 | 199971 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065929 | 199972 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065930 | 199973 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065931 | 199974 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065932 | 199975 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065933 | 199976 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-22 13:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065934 | 199977 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-22 13:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065935 | 199978 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065936 | 199979 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-19 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065937 | 199980 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065938 | 199981 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 18:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065939 | 199982 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065940 | 199983 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065941 | 199984 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065942 | 199985 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065943 | 199986 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065944 | 199987 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065945 | 199988 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065946 | 199989 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 11:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065947 | 199990 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065948 | 199991 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065949 | 199992 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065950 | 199993 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-12 11:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065951 | 199994 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065952 | 199995 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-15 11:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065953 | 199996 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-22 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065954 | 199997 | PM | CNP | ENL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065955 | 199998 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065956 | 199999 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065957 | 200000 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-18 10:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065958 | 200001 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-22 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065959 | 200002 | PM | CANCL | CANCL | 2023-12-01 02:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065960 | 200003 | PM | CMPLT | CMPLT | 2023-12-01 02:21 | 2024-03-21 08:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065961 | 200004 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 18:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065962 | 200005 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 13:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065963 | 200006 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-18 10:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065964 | 200007 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065965 | 200008 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 14:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065966 | 200009 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-18 10:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065967 | 200010 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 11:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065968 | 200011 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065969 | 200012 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065970 | 200013 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065971 | 200014 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-18 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065972 | 200015 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-18 10:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065973 | 200016 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 14:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065974 | 200017 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065975 | 200018 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065976 | 200019 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065977 | 200020 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065978 | 200021 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-18 11:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065979 | 200022 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-19 14:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065980 | 200023 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 14:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065981 | 200024 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-18 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065982 | 200025 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065983 | 200026 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065984 | 200027 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 08:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065985 | 200028 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 09:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065986 | 200029 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065987 | 200030 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065988 | 200031 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 11:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065989 | 200032 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065990 | 200033 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065991 | 200034 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065992 | 200035 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065993 | 200036 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065994 | 200037 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065995 | 200038 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065996 | 200039 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065997 | 200040 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065998 | 200041 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065999 | 200042 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 11:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066000 | 200043 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066001 | 200044 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066002 | 200045 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 11:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066003 | 200046 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066004 | 200047 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066005 | 200048 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066006 | 200049 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-18 11:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066007 | 200050 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-12 10:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066008 | 200051 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-04-26 10:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066009 | 200052 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-04-26 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066010 | 200053 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 18:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066011 | 200054 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 11:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066012 | 200055 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066013 | 200056 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-12 10:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066014 | 200057 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066015 | 200058 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066016 | 200059 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 11:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066017 | 200060 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066018 | 200061 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 14:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066019 | 200062 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 11:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066020 | 200063 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 09:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066021 | 200064 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-04-04 13:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:31:25 PM, Rhuland, Courtney) PM completed by BD
1066022 | 200065 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066023 | 200066 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066024 | 200067 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 09:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066025 | 200068 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066026 | 200069 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066027 | 200070 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 09:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066028 | 200071 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066029 | 200072 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066030 | 200073 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-04-26 10:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066031 | 200074 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066032 | 200075 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 12:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066033 | 200076 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 14:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066034 | 200077 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066035 | 200078 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 12:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066036 | 200079 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066037 | 200080 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 12:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066038 | 200081 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066039 | 200082 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066040 | 200083 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066041 | 200084 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-04-26 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066042 | 200085 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 12:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066043 | 200086 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-12 10:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066044 | 200087 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-12 11:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066045 | 200088 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-19 14:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066046 | 200089 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-15 12:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066047 | 200090 | PM | CANCL | CANCL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066048 | 200091 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 15:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066049 | 200092 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 09:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066050 | 200093 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066051 | 200094 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 15:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066052 | 200095 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066053 | 200096 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 15:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066054 | 200097 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066055 | 200098 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-22 15:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066056 | 200099 | PM | CMPLT | CMPLT | 2023-12-01 02:22 | 2024-03-21 08:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066057 | 200100 | PM | CNP | ENL | 2023-12-01 02:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68680 | 200101 | CM | CMPLT | CMPLT | 2023-12-01 09:28 | 2023-12-01 09:30 | update network config | (TC 12/01/2023 09:28:50 AM, Campbell, Connor) update network config to BH
68684 | 200102 | CM | CMPLT | CMPLT | 2023-12-01 09:31 | 2023-12-01 09:31 | update network config | (TC 12/01/2023 09:30:40 AM, Campbell, Connor) update network config to BH
68757 | 200103 | CM | CMPLT | CMPLT | 2023-12-01 09:37 | 2023-12-01 09:37 | update network config | (TC 12/01/2023 09:36:52 AM, Campbell, Connor) update network config to BH
68836 | 200104 | CM | CMPLT | CMPLT | 2023-12-01 09:39 | 2023-12-01 09:39 | update network config | (TC 12/01/2023 09:38:46 AM, Campbell, Connor) update network config to BH
1069320 | 200105 | ACC | CMPLT | CMPLT | 2023-12-01 09:57 | 2023-12-01 09:59 | Incoming inspection and asset tagging | (TC 12/01/2023 09:56:44 AM, Baghirzade, Nara) Picked up the INVOS monitor from Dr. Deshpande's office for assembly. Assembled the monitor, attached the battery and docking station. Attached eletrical safety from the vendor. Created asset on TMS. Device is purchased by Dr. Deshpande and PO# was 759583 (requisition dated: October 25. 2023). The cost centre is Dr. Deshpande’s research IO 410018529. Device will be used in the NICU for research when participants are consented. (TC 12/01/2023 09:56:44 AM, Baghirzade, Nara) Picked up the INVOS monitor from Dr. Deshpande's office for assembly. Assembled the monitor, attached the battery and docking station. Attached eletrical safety from the vendor. Created asset on TMS. Device is purchased by Dr. Deshpande and PO# was 759583 (requisition dated: October 25. 2023). The cost centre is Dr. Deshpande’s research IO 410018529. Device will be used in the NICU for research when participants are consented. Dropped off the device back to Dr. Deshpande's office.
1069319 | 200106 | ACC | CMPLT | CMPLT | 2023-12-01 09:59 | 2023-12-01 10:02 | Incoming inspection and asset tagging | (TC 12/01/2023 09:59:23 AM, Baghirzade, Nara) Picked up the INVOS monitor from Dr. Deshpande's office for assembly. Assembled the monitor, attached the battery and docking station. Attached eletrical safety from the vendor. Created asset on TMS. Device is purchased by Dr. Deshpande and PO# was 759583 (requisition dated: October 25. 2023). The cost centre is Dr. Deshpande’s research IO 410018529. Device will be used in the NICU for research when participants are consented. Dropped off the device back to Dr. Deshpande's office.
| 200107 | OC | CMPLT | CMPLT | 2023-12-01 10:25 | 2023-12-08 12:20 | On Call - Dec 1-Dec 8, 2023 | (TC 12/01/2023 10:25:15 AM, Scero, Tomlyn) Daily checks complete (TC 12/03/2023 11:39:02 AM, Scero, Tomlyn) Cerner upgrade command centre and VSM change (TC 12/04/2023 02:18:53 PM, Scero, Tomlyn) Daily checks complete (TC 12/05/2023 10:18:14 AM, Scero, Tomlyn) daily checks (TC 12/06/2023 09:12:59 AM, Scero, Tomlyn) Daily checks (TC 12/07/2023 09:13:14 AM, Scero, Tomlyn) daily checks
| 200108 | CM | CMPLT | CMPLT | 2023-12-01 10:35 | 2023-11-30 10:35 | Biomed meeting | Biomed PM meeting
| 200109 | CM | CMPLT | CMPLT | 2023-12-01 10:36 | 2023-12-01 10:36 | worked on TMS | worked on TMS
| 200110 | CM | CMPLT | CMPLT | 2023-12-01 10:37 | 2023-11-30 10:37 | daily checks | daily checks
69562 | 200111 | CM | CMPLT | CMPLT | 2023-12-01 10:42 | 2023-12-01 11:47 | "keeps saying air in line after changing a new set" | (TC 12/01/2023 10:41:40 AM, Campbell, Connor) performed a check in, passed performed a test infusion, no fault found
65981 | 200112 | CM | CMPLT | CMPLT | 2023-12-01 10:42 | 2023-12-01 10:45 | Channel error reported | (TC 12/01/2023 10:41:43 AM, Baghirzade, Nara) Downloaded the error log; noted log only errors pertaining to the pressure sensor disc. The rear case had a few cracks/chips in it. Replaced the rear case and transferred the handle and latch on to the new rear case. Reprinted asset tag and inspection sticker. Reseated the pressure sensor disc. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
| 200113 | CM | CMPLT | CMPLT | 2023-12-01 11:18 | 2024-02-05 07:52 | Working with Adeel and Nima to HR Biomed and assignment update. | (TC 12/01/2023 11:18:35 AM, Chukri, Samir) working with Adeel via Slack and Phone review the HR and assginment list for back up Sr.technologist cover . (TC 12/01/2023 03:14:34 PM, Chukri, Samir) Zoom Meeting with Adeel (HR and WO-time ) (TC 02/05/2024 07:52:39 AM, Chukri, Samir) completed .
65956 | 200114 | CM | CMPLT | CMPLT | 2023-12-01 11:28 | 2023-12-01 11:30 | Channel error not specified | (TC 12/01/2023 11:28:08 AM, Baghirzade, Nara) Downloaded the error log; noted repeated errors pertaining to pressure disc failure. When disassembling the device, noted that the gasket on the rear case was completely damaged and torn. Replaced the gasket on the inside of the rear case. Replaced the pressure sensor disc. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65934 | 200115 | CM | CMPLT | CMPLT | 2023-12-01 12:36 | 2023-12-01 12:40 | Channel error when connecting to PCU | (TC 12/01/2023 12:36:42 PM, Baghirzade, Nara) Downloaded the error log; noted a couple of log only errors pertaining to the pressure sensor disc. Pressure sensor disc was previously replaced and has not been re-seated since. Re-seated the pressure sensor. Upon disassembly, noted that both the IUIs had damaged pins. Also noted that the left handle and rear case was cracked. Replaced left and right IUIs. Replaced left handle under recall. Replaced rear case under recall. Pressure sensor disc was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1069325 | 200116 | ACC | CMPLT | CMPLT | 2023-12-01 13:52 | 2023-12-12 09:58 | Incoming inspection | (TC 12/12/2023 09:56:55 AM, Canjimavadivel, Kamallesh ) Installed on boom in G4. Tested functions
1069326 | 200117 | ACC | CMPLT | CMPLT | 2023-12-01 13:55 | 2023-12-13 08:35 | Incoming inspection | (TC 12/13/2023 08:33:53 AM, Canjimavadivel, Kamallesh ) Installed on boom in B14. Tested all functions, working ok.
1069327 | 200118 | ACC | CMPLT | CMPLT | 2023-12-01 13:56 | 2023-12-08 14:07 | Incoming inspection | (TC 12/08/2023 02:06:31 PM, Canjimavadivel, Kamallesh ) Installed on boom in room B15. All functions pass
62696 | 200119 | CM | CMPLT | NFF | 2023-12-01 14:27 | 2023-12-01 14:28 | Lead II not working | (TC 12/01/2023 02:27:49 PM, Canjimavadivel, Kamallesh ) Tested all leads with simulator. Unable to recreate reported issue. Returned cart back to unit.
| 200120 | CM | CMPLT | CMPLT | 2023-12-01 15:18 | 2023-12-01 15:20 | working and showing Nara to resoldering the EEG interface connectors. | (TC 12/01/2023 03:18:48 PM, Chukri, Samir) Showing Nar to resoldering the EEG interface connectors.
1049362 | 200121 | CM | CMPLT | CMPLT | 2023-12-01 15:27 | 2023-12-01 15:27 | PM - Yearly | Injector PM done -Flex Ring calibrated -Potentiometer calibrated -Pressure calibrated -Touchscreen calibrated -System passed all functional tests
502089 | 200122 | RQ | CMPLT | CMPLT | 2023-12-01 15:34 | 2023-12-08 15:49 | screw on CT phantom (Symbia 2) needs to be replaced | (TC 12/08/2023 03:48:04 PM, Green, Paul) Ordered part. Replaced and notified customer. | screw on CT phantom needs to be replaced
32520 | 200123 | CM | CMPLT | CMPLT | 2023-12-04 08:02 | 2023-12-04 08:58 | MSH DR2 COLLIMATOR NOT COMMUNICATING BLADE OPEING INFO WITH SYSTEM | (TC 12/01/2023 08:01:00 PM, Whittington, Scott) MSH DR2 COLLIMATOR NOT COMMUNICATING BLADE OPEING INFO WITH SYSTEM | MSH DR2 COLLIMATOR NOT COMMUNICATING BLADE OPEING INFO WITH SYSTEM
501066 | 200124 | CM | CMPLT | CMPLT | 2023-12-04 08:19 | 2023-12-20 11:02 | msh ceou 4 site 553547 c5-1 probe artifact sn B3KJDR and disable constant send to burner | (TC 12/04/2023 08:58:33 AM, Whittington, Scott) msh ceou 4 site mie 501066 site 553547 c5-1 probe artifact sn B3KJDR and disable constatnt send to burner ref# belonged to site 553550 0122734112 probe ordered | msh ceou 4 site mie 501066 site 553547 c5-1 probe artifact sn B3KJDR and disable constatnt send to burner ref# belonged to site 553550 0122734112 letter rma Return Material Document Outbound DN: 862246377 / RMD DN: 846966396 ref# site#
501070 | 200125 | CM | CMPLT | CMPLT | 2023-12-04 08:26 | 2023-12-29 13:21 | msh ceou 10 qwerty kybd site 553549 | (TC 12/04/2023 08:56:49 AM, Whittington, Scott) msh ceou 10 qwerty kybd site 553549 msh ceou 10 qwerty kybd & cable site mie 501070 site 553549 ref# 0122734158 anthony white PARTS ORDERED (TC 12/29/2023 01:20:50 PM, Dolinar, John) QWERTY KYBD AND CABLE WERE FAULTY, WHICH WOULD INDUCE OPPANEL ERRORS QWERTY KYBD & CABLE REPLACED DECEMBER 29 | msh ceou 10 qwerty kybd site 553549 msh ceou 10 qwerty kybd & cable site mie 501070 site 553549 ref# 0122734158 anthony white
| 200126 | RDS | CMPLT | CMPLT | 2023-12-04 08:28 | 2023-12-04 14:19 | Weekly checks in CCU | (TC 12/04/2023 02:17:50 PM, Scero, Tomlyn) weekly checks complete. 30J test done
| 200127 | PA | CANCL | CANCL | 2023-12-04 08:30 |  | BP cuffs for stock
1059828 | 200128 | PM | CMPLT | CMPLT | 2023-12-04 09:16 | 2023-12-04 09:17 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/04/2023 09:16:12 AM, Campbell, Connor) Disinfected, cleaned, rinsed and refilled
1059778 | 200129 | FC | CMPLT | CMPLT | 2023-12-04 10:01 | 2023-12-04 10:01 | QC checklist | (TC 12/04/2023 10:01:01 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1067437 | 200130 | FC | CMPLT | CMPLT | 2023-12-04 10:02 | 2023-12-04 10:03 | QC checklist | (TC 12/04/2023 10:01:55 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1066460 | 200131 | FC | CMPLT | CMPLT | 2023-12-04 10:03 | 2023-12-04 10:04 | QC checklist | (TC 12/04/2023 10:03:32 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
1066462 | 200132 | FC | CMPLT | CMPLT | 2023-12-04 10:04 | 2023-12-04 10:05 | QC checklist | (TC 12/04/2023 10:04:27 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
1056125 | 200133 | PM | CMPLT | CMPLT | 2023-12-04 10:37 | 2023-12-04 10:41 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/04/2023 10:40:35 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
52178 | 200134 | CM | CMPLT | CMPLT | 2023-12-04 10:51 | 2023-12-04 10:53 | Broken latch (SW) | (TC 12/04/2023 10:52:01 AM, Baghirzade, Nara) Replaced SW latch. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 76% RH. Incubator is fully functional.
| 200135 | CM | CMPLT | CMPLT | 2023-12-04 10:53 | 2023-12-04 10:53 | 11 South inventory and walkthrough | (TC 12/04/2023 10:53:02 AM, Baghirzade, Nara) 11 South inventory and walkthrough with Andrew.
| 200136 | RDS | CMPLT | CMPLT | 2023-12-04 10:57 | 2023-12-04 11:00 | Weekly Checks For PACU | (TC 12/04/2023 10:59:25 AM, Reid, Emmerson ) Rounds complete. Also showed Eden how to work a patient simulator for a future clinical simulation
1054221 | 200137 | CM | CMPLT | CMPLT | 2023-12-04 13:49 | 2023-12-04 13:59 | "SEC" displaying on screen | (TC 12/04/2023 01:48:51 PM, Canjimavadivel, Kamallesh ) Removed security counter as there isn't an available dock to redock and clear counter as per request from staff
| 200138 | CM | ACTIV | ISSUE | 2023-12-04 14:08 |  | Discussion of Philips's new NBP ISO compliant inlet | (TC 12/04/2023 02:08:43 PM, Chukri, Samir) Meeting with Adeel, tomlyn and Maisha . wating to get a quotation from Philips
502088 | 200139 | RQ | CMPLT | CMPLT | 2023-12-04 14:11 | 2023-12-05 14:12 | Symbia 1 Detector 1 error -limit error | (TC 12/05/2023 02:12:50 PM, Green, Paul) Light curtain end cap was inserted incorrectly and a small piece broke off and caught inside. This caused light curtain to fail and could not be used. Found this during troubleshooting. Replaced light curtain glass and end cap. Tested, ok. | Symbia 1 Detector 1 error -limit error ; unable to start TB acqusition
| 200140 | CM | CMPLT | CMPLT | 2023-12-04 14:13 | 2023-12-01 14:33 | Cerner downtime | TC 12/01/2023 02:13:28 PM, Chukri, Samir) working during cerner dwontime at 1900 till about midnight.
1057660 | 200141 | CM | CMPLT | CMPLT | 2023-12-04 14:18 | 2023-12-04 14:19 | Low battery | (TC 12/04/2023 02:18:50 PM, Canjimavadivel, Kamallesh ) Replaced batteries and probe Cleared logs
1054225 | 200142 | CM | CMPLT | CMPLT | 2023-12-04 14:21 | 2023-12-04 14:22 | Low battery | (TC 12/04/2023 02:21:43 PM, Canjimavadivel, Kamallesh ) Repalced batteries
| 200143 | RDS | CMPLT | CMPLT | 2023-12-04 14:23 | 2023-12-04 14:23 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
| 200144 | CM | CMPLT | CMPLT | 2023-12-04 14:26 | 2023-12-04 14:26 | PS500 installation | worked with Salur, Steve and Rex
1057842 | 200145 | CM | CMPLT | CMPLT | 2023-12-04 15:11 | 2023-12-07 14:16 | Front panel not working | (TC 12/04/2023 03:13:14 PM, Reid, Emmerson ) Was called to OR6. Had to swtich the ESU with one from a cart. Control panel bipolar power down button was not working. (TC 12/05/2023 01:04:11 PM, Reid, Emmerson ) Tried reseasting the control panel, however this did not work. Will have to wait for the part to come in. (TC 12/07/2023 02:14:59 PM, Reid, Emmerson ) Replaced the display board. Testing will be done on the uncompleted PM WO.
1062999 | 200146 | PM | CMPLT | CMPLT | 2023-12-04 15:55 | 2023-12-04 15:56 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/04/2023 03:54:48 PM, Cortez, John) Completed cleaning cycle
| 200147 | RDS | CMPLT | CMPLT | 2023-12-05 07:56 | 2023-12-05 08:52 | Weekly check for ED
| 200148 | RDS | CMPLT | CMPLT | 2023-12-05 08:02 | 2023-12-05 15:06 | Weekly check for ICU
1049898 | 200149 | CM | CMPLT | CMPLT | 2023-12-05 09:54 | 2024-02-26 13:28 | MSH-5-RPCU Probe PROVIDED BY ENDOCH cable requires taping | (TC 12/05/2023 10:03:06 AM, Whittington, Scott) MSH-5-RPCU Probe PROVIDED BY ENDOCH cable requires taping | MSH-5-RPCU Probe PROVIDED BY ENDOCH cable requires taping
501442 | 200150 | CM | CMPLT | CMPLT | 2023-12-05 10:01 | 2024-02-26 13:27 | MSH RPCU LOAN PVT 375 FROM MSH GEN US SPARES | (TC 12/05/2023 10:01:54 AM, Whittington, Scott) MSH RPCU LOAN PVT 375 FROM MSH GEN US SPARES | MSH RPCU LOAN PVT 375 FROM MSH GEN US SPARES
62700 | 200151 | CM | CMPLT | CMPLT | 2023-12-05 10:38 | 2023-12-05 11:17 | V1-V3 chest lead replaced | (TC 12/05/2023 10:53:46 AM, Scero, Tomlyn) ECG tech brought down the machine and asked for a new V1-V3 chest lead because the current one was damaged. Chest lead taken from ICU extra stock.
1061215 | 200152 | CM | CMPLT | CMPLT | 2023-12-05 10:53 | 2023-12-05 10:54 | Vitals not showing on the monitors | (TC 12/05/2023 10:53:23 AM, Reid, Emmerson ) Cable came lose from the back of the MX800. Working now.
1061211 | 200153 | CM | CMPLT | CMPLT | 2023-12-05 10:54 | 2023-12-05 10:55 | Monitor frozen | (TC 12/05/2023 10:54:22 AM, Reid, Emmerson ) VMC monitor was frozen. I reseated the cables and it worked. Also had to move the SDI cable off the composite onto the proper port. Working now.
62701 | 200154 | CM | CMPLT | CMPLT | 2023-12-05 11:17 | 2023-12-05 11:26 | Keyboard cover damaged | (TC 12/05/2023 11:25:10 AM, Scero, Tomlyn) keyboard cover replaced. returned unit to the ECG department
| 200155 | PA | CMPLT | CMPLT | 2023-12-05 13:07 | 2023-12-28 11:03 | Ordering brackets for ICU | (TC 12/05/2023 01:15:14 PM, Canjimavadivel, Kamallesh ) Placed order for brackets (TC 12/08/2023 01:32:53 PM, Canjimavadivel, Kamallesh ) Received brackets, one of them was wrong one. Reached out to vendor for resolution. (TC 12/12/2023 02:30:00 PM, Canjimavadivel, Kamallesh ) Received confirmation from vendor to return and reship correct part. Initiated return using vendor provided account (TC 12/18/2023 11:18:25 AM, Canjimavadivel, Kamallesh ) Having difficulties creating label for return, reached out to vendor to provide shipping label. (TC 12/19/2023 08:03:19 AM, Canjimavadivel, Kamallesh ) Replacement was received, waiting on vendor to provide next steps for returning wrong part. (TC 12/27/2023 07:44:51 AM, Canjimavadivel, Kamallesh ) Received shipping label from vendor. Printed label and prepped box for shipping. (TC 12/28/2023 11:02:33 AM, Canjimavadivel, Kamallesh ) Dropped off package with loading dock for pickup
65880 | 200156 | CM | CMPLT | NFF | 2023-12-05 13:35 | 2023-12-05 13:37 | "unable to associate pump channel" | (TC 12/05/2023 01:34:40 PM, Campbell, Connor) Performed check in, passed plugged and unplugged multiple times to test iui connectors, no problem cleaned iui connectors performed a test infusion no fault found
65878 | 200157 | CM | CMPLT | NFF | 2023-12-05 13:40 | 2023-12-05 13:41 | "unable to associate pump channel" | (TC 12/05/2023 01:39:42 PM, Campbell, Connor) Pm was complete just before this test plugged and unplugged multiple times to test iui connectors, no problem cleaned iui connectors performed a test infusion no fault found
1064212 | 200158 | PM | CMPLT | CMPLT | 2023-12-05 14:25 | 2023-12-05 14:31 | PM-MAJOR - Annual PM For: Ventilators | (TC 12/05/2023 02:28:00 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 6989 hrs Total vent time: 4181 hrs
1049362 | 200159 | CM | CMPLT | CMPLT | 2023-12-05 16:24 | 2023-12-05 16:24 | Power 3026 Error | (TC 12/04/2023 04:24:31 PM, Debebe, Adam) Injector power 3026 Error -Troubleshooting done on site -Multiple errors suspected -Moved injector to shop - Faulty head cable, faulty power board (burnt) and faulty Base power supply. (TC 12/05/2023 04:24:31 PM, Debebe, Adam) Replaced Head cable, power board and and base power supply. -Also noted that there was spill on the injector head, took it apart to clean it and calibrated the flex rings. (TC 12/06/2023 02:36:25 PM, Debebe, Adam) Injector testing completed -Returned back to ER CT -Spare injector brought back to shop for inspection and erturned to CR room on the 5th floor.
67311 | 200160 | CM | CMPLT | CMPLT | 2023-12-05 16:54 | 2024-02-22 08:52 | Top heater not warming up | (TC 11/28/2023 04:59:34 PM, Rendon Lopez , Elizabeth) Received bag sensor under warranty. Replaced and found connector on main board broken off (not covered under warranty) Inquired for part pricing (TC 11/28/2023 05:01:05 PM, Rendon Lopez , Elizabeth) Returned device to OR3 while waiting for part, with note DON'T USE (TC 12/05/2023 05:01:05 PM, Rendon Lopez , Elizabeth) Received part information, ordered part (TC 01/17/2024 08:28:55 AM, Rendon Lopez , Elizabeth) Followed up with part delivery. priced changed and adjusted part number, will be delivered in February (TC 02/01/2024 03:36:15 PM, Rendon Lopez , Elizabeth) Middle heater having same issue, contacted service for warranty. Restarted warmer, seems to be working (TC 02/20/2024 02:47:31 PM, Rendon Lopez , Elizabeth) Received control board, replaced. Testing overnight (TC 02/21/2024 08:51:43 AM, Rendon Lopez , Elizabeth) Fluid warmer working properly through the night. Assembled back in OR 3.
1049903 | 200161 | CM | CMPLT | CMPLT | 2023-12-05 17:06 | 2023-12-05 17:06 | System will not charge. |  | replaced reel cord mechanism.
| 200162 | CM | CMPLT | CMPLT | 2023-12-05 17:09 | 2023-12-05 17:09 | Site coverage.
1062695 | 200163 | CM | CMPLT | CMPLT | 2023-12-05 17:10 | 2023-12-05 17:11 | Blue light malfunction | (TC 11/29/2023 05:10:00 PM, Rendon Lopez , Elizabeth) Confirmed issue, replaced device, updated Sn, tested-OK
1051307 | 200164 | CM | CMPLT | CMPLT | 2023-12-05 17:11 | 2023-12-05 17:11 | Blue light malfunction | (TC 11/29/2023 05:10:57 PM, Rendon Lopez , Elizabeth) Confirmed issue, replaced device, updated Sn, tested-OK
1069318 | 200165 | ACC | CMPLT | CMPLT | 2023-12-05 17:13 | 2023-12-05 17:13 | Incoming inspection | (TC 11/30/2023 05:12:44 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069305 | 200166 | ACC | CMPLT | CMPLT | 2023-12-05 17:13 | 2023-12-05 17:14 | Incoming inspection | (TC 11/30/2023 05:13:21 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069306 | 200167 | ACC | CMPLT | CMPLT | 2023-12-05 17:14 | 2023-12-05 17:14 | Incoming inspection | (TC 11/30/2023 05:13:56 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069307 | 200168 | ACC | CMPLT | CMPLT | 2023-12-05 17:17 | 2023-12-05 17:17 | Incoming inspection | (TC 11/30/2023 05:17:09 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069308 | 200169 | ACC | CMPLT | CMPLT | 2023-12-05 17:18 | 2023-12-05 17:18 | Incoming inspection | (TC 11/30/2023 05:17:59 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069309 | 200170 | ACC | CMPLT | CMPLT | 2023-12-05 17:18 | 2023-12-05 17:22 | Incoming inspection | (TC 11/30/2023 05:22:13 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069310 | 200171 | ACC | CMPLT | CMPLT | 2023-12-05 17:23 | 2023-12-05 17:23 | Incoming inspection | (TC 11/30/2023 05:22:50 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069311 | 200172 | ACC | CMPLT | CMPLT | 2023-12-05 17:24 | 2023-12-05 17:24 | Incoming inspection | (TC 11/30/2023 05:23:48 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069312 | 200173 | ACC | CMPLT | CMPLT | 2023-12-05 17:24 | 2023-12-05 17:26 | Incoming inspection | (TC 11/30/2023 05:24:25 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069313 | 200174 | ACC | CMPLT | CMPLT | 2023-12-05 17:26 | 2023-12-05 17:26 | Incoming inspection | (TC 11/30/2023 05:26:14 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069314 | 200175 | ACC | CMPLT | CMPLT | 2023-12-05 17:27 | 2023-12-05 17:27 | Incoming inspection | (TC 11/30/2023 05:26:51 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069316 | 200176 | ACC | CMPLT | CMPLT | 2023-12-05 17:28 | 2023-12-05 17:28 | Incoming inspection | (TC 11/30/2023 05:27:41 PM, Rendon Lopez , Elizabeth) Incoming inspection
1069317 | 200177 | ACC | CMPLT | CMPLT | 2023-12-05 17:28 | 2023-12-05 17:28 | Incoming inspection | (TC 11/30/2023 05:28:15 PM, Rendon Lopez , Elizabeth) Incoming inspection
1067353 | 200178 | CM | CMPLT | CMPLT | 2023-12-06 08:03 | 2023-12-06 08:04 | Cable replacement | (TC 12/06/2023 08:03:18 AM, Rhuland, Courtney) Replaced shorter cable with longer one Functional check - OK
1069362 | 200179 | CM | CMPLT | CMPLT | 2023-12-06 08:23 | 2023-12-06 08:26 | Incoming inspection | (TC 12/06/2023 08:23:41 AM, Rhuland, Courtney) Assembled and performed functional & electrical safety check
| 200180 | CM | CMPLT | CMPLT | 2023-12-06 08:25 | 2023-12-06 08:26 | Sr.technologist interview | (TC 12/05/2023 02:24:00 PM, Chukri, Samir) Sr. Technologist interview
1055937 | 200181 | CM | CMPLT | NFF | 2023-12-06 08:46 | 2023-12-06 08:46 | "Bp not working" | (TC 12/06/2023 08:45:45 AM, Campbell, Connor) Tested BP with two different cuffs functional check on spo2 and temp no faults found
1057115 | 200182 | CM | CMPLT | CMPLT | 2023-12-06 08:57 | 2023-12-06 08:57 | power cord missing | (TC 12/06/2023 08:56:40 AM, Campbell, Connor) replaced power cord with a spare from the office, no charge performed a functional check on temp, bp and spo2, all passed returned to service
68548 | 200183 | CM | CMPLT | CMPLT | 2023-12-06 09:03 | 2023-12-06 09:04 | Hose ripped | (TC 12/05/2023 09:03:07 AM, Cortez, John) Replaced cuff
1057033 | 200184 | CM | CMPLT | NFF | 2023-12-06 10:03 | 2023-12-06 10:06 | "broken not inflating" | (TC 12/06/2023 10:03:37 AM, Campbell, Connor) Tested BP, inflated fine, no fault tested temp, temp is fine SPO2 finger clip cable was damaged, wires exposed, replaced with a new one and charged no fault found, returned to service
53076 | 200185 | CM | CMPLT | CMPLT | 2023-12-06 10:28 | 2023-12-06 10:30 | Needed for transport but no stretcher | (TC 12/05/2023 10:28:09 AM, Ferreira, Qwynn) Worked with connor to replace the stretcher battery Charger pack was damaged, but NICU team needed ASAP Charged enough for the transport. USed manual mode to lift stretcher and incubator Informed Staff in the NICU about the transport and charging the stretcher battery pack Will need to swap over tomorrow
1054546 | 200186 | CM | CMPLT | NFF | 2023-12-06 10:33 | 2023-12-06 10:34 | Not working | (TC 12/06/2023 10:33:27 AM, Rhuland, Courtney) Reported that it was not turning on Ran functional check with 5kg test weight and no fault found Instructed unit to leave scale plugged in
| 200187 | CM | CMPLT | CMPLT | 2023-12-06 11:15 | 2023-12-06 11:15 | worked on patient beds at 16N and 14S | worked on patient beds at 16N and 14S
62779 | 200188 | CM | CMPLT | CMPLT | 2023-12-06 11:48 | 2023-12-06 11:51 | Broken power cord | (TC 12/06/2023 11:48:10 AM, Rhuland, Courtney) Switched from stretcher to transport legs Ground was missing from power bar power cord Replaced plug head with spare from biomed Electrical safety - OK Functional check - OK
1069363 | 200189 | ACC | CMPLT | CMPLT | 2023-12-06 12:01 | 2024-01-04 14:23 | MSH SPP GE VOLUSON EXPERT 22 MSH_US_SPP_VOLI ACCEPTANCE | (TC 12/06/2023 12:18:24 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE (TC 12/06/2023 12:18:24 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I TO BE INSTALLED IN ROOM D5 3-529 (TC 12/08/2023 01:34:34 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 0 MSH_US_SPP_VOLI ACCEPTANCE (TC 12/14/2023 12:36:18 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I CR1798 TO BE INSTALLED IN ROOM D5 3-529 (TC 12/15/2023 12:36:33 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I CR1798 TO BE INSTALLED IN ROOM D5 3-529 (TC 12/22/2023 01:17:22 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I CR1798 TO BE INSTALLED IN ROOM D5 3-529 SYSTEM HAD ARRIVED LST WEEK, FOUND IN AMY'S OFFICE GLEANED MAC & SN# SETUP PROBES IN TMS (TC 12/29/2023 01:00:31 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I CR1798 TO BE INSTALLED IN ROOM D5 3-529 SYSTEM HAD NOT BEEN RECEIVED ON DELIVERY OBTAINED PACKING SLIPS AND HAD RECEIVED IN RECEIVING DEPT ASTRAIA DETAILS TO BE DEALT WITH MIRZA PROVIDED STATIC IP From: Riaz, Mirza <mirza.riaz@dxc.com> Sent: December 28, 2023 10:39 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] RE: MSH OPG SPP provide a static IP on VLAN 172.19.72.1 for an additional GE Voluson Ultrasound with MAC 00:13:95:5A:F3:31, which will require NAC Whitelistng Hi Scott, See the IP address below, I also have whitelisted the MAC address and should be good now IP Address: 172.19.72.5 Mask: 255.255.255.0 Default Gateway: 172.19.72.1 Thanks, Mirza From: "HELP Desk (MSH)" <helpdesk.MSH@sinaihealth.ca> Date: December 28, 2023 at 2:54:10?PM EST To: "Whittington, Scott - UHN - University Health Network" <Scott.Whittington@uhn.ca> Cc: "Riaz, Mirza" <Mirza.Riaz@sinaihealth.ca> Subject: RE: MSH OPG SPP provide a static IP on VLAN 172.19.72.1 for an additional GE Voluson Ultrasound with MAC 00:13:95:5A:F3:31, which will require NAC Whitelistng Hello, Incident INC30462882 has been assigned to onsite to set up the static IP on VLAN. Please let us know if you need anything else. Best regards, Simon Pilon, MSH Service Desk Phone : 416 586 4800 x 4357 Toll Free:1-866-244-9909(MSH) Email: helpdesk.MSH@sinaihealth.ca or toronto.carecentre@dxc.com From: Whittington, Scott [mailto:Scott.Whittington@uhn.ca] Sent: Thursday, December 28, 2023 12:27 PM To: HELP Desk (MSH) <helpdesk.MSH@sinaihealth.ca> Cc: Riaz, Mirza <Mirza.Riaz@sinaihealth.ca> Subject: MSH OPG SPP provide a static IP on VLAN 172.19.72.1 for an additional GE Voluson Ultrasound with MAC 00:13:95:5A:F3:31, which will require NAC Whitelistng CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Good Afternooon Please provide a static IP on VLAN 172.19.72.1 for an additional GE Voluson Ultrasound with MAC 00:13:95:5A:F3:31, which will require NAC Whitelistng. Thanks, Scott 416-402-2462 (TC 01/02/2024 04:25:18 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I CR1798 TO BE INSTALLED IN ROOM D5 3-529 From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: January 2, 2024 1:45 PM To: Wong, Stephanie - Sinai Health System (Stephanie.Wong4@sinaihealthsystem.ca) <stephanie.wong4@sinaihealthsystem.ca> Cc: Ultrasound Support <Ultrasound.Support@uhn.ca>; Hollingshead, Andrea - Sinai Health System <Andrea.Hollingshead@sinaihealth.ca>; Catunto, Antonio (GE HealthCare) <ANTONIO.CATUNTO@ge.com>; Biro, Tiffany - Sinai Health System (Tiffany.Biro@sinaihealthsystem.ca) <tiffany.biro@sinaihealthsystem.ca> Subject: Re: [External] RE: MSH OPG SPP provide a static IP on VLAN 172.19.72.1 for an additional GE Voluson Ultrasound with MAC 00:13:95:5A:F3:31, which will require NAC Whitelistng Hi Stephanie Here is the info for the latest GE VOLUSON US for Astraia setup Room 3-529 D5 AE: MSH_US_SPP_VOLI (CAPITAL I) IP Address: 172.19.72.5 Mask: 255.255.255.0 Default Gateway: 172.19.72.1 (TC 01/04/2024 02:21:57 PM, Whittington, Scott) MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I CR1798 TO BE INSTALLED IN ROOM D5 3-529 PROGRAMMED SYSTEM GE TONY CONTUTO STATIC IP PROVIDED DID NOT WORK DHCP FOR NOW UNTIL HSC CONNECTION IS ESTABLISHED | MSH SPP GE VOLUSON 10 MSH_US_SPP_VOLI ACCEPTANCE , USING I CAPITAL I CR1798 TO BE INSTALLED IN ROOM D5 3-529 200 mOHM 235 uA
65959 | 200190 | CM | CMPLT | CMPLT | 2023-12-06 13:24 | 2023-12-06 13:25 | channel error | (TC 12/06/2023 01:24:13 PM, Campbell, Connor) looked through logs, it was a pressure sensor error reseated the pressure sensor flashed the pump performed a check in performed a test infusion pump is functional, no longer displays channel error returned to service
| 200191 | CM | CMPLT | CMPLT | 2023-12-06 14:45 | 2023-12-19 14:31 | Masimo accessories order, for biomed stock | (TC 12/06/2023 03:05:56 PM, Cortez, John) placed order for 5x spo2 sensors 1x box of adult cuffs (TC 12/12/2023 03:58:36 PM, Cortez, John) box of cuffs came in (TC 12/19/2023 02:30:44 PM, Cortez, John) SPO2 sensors came in
67411 | 200192 | PM | CMPLT | CMPLT | 2023-12-06 16:34 | 2023-12-06 16:35 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/01/2023 04:34:41 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67413 | 200193 | PM | CMPLT | CMPLT | 2023-12-06 16:37 | 2023-12-06 16:38 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/01/2023 04:37:24 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67428 | 200194 | PM | CMPLT | CMPLT | 2023-12-06 16:39 | 2023-12-06 16:39 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/01/2023 04:39:00 PM, Rendon Lopez , Elizabeth) Replaced damaged collection ring. PM completed, all tests passed
1049878 | 200195 | CM | CMPLT | CMPLT | 2023-12-06 16:41 | 2024-04-03 14:56 | Created by Failed PM WO # 198225 : PM-MAJOR - Annual PM For: YAG laser, gyne | (TC 12/05/2023 04:41:32 PM, Rendon Lopez , Elizabeth) Received loaner device (1069354) Prepared device for shipphing to Dornier (TC 12/05/2023 04:41:32 PM, Rendon Lopez , Elizabeth) Shipped device for repair (TC 04/03/2024 02:49:45 PM, Rendon Lopez , Elizabeth) Received device back from repair. Returned to L&D OR1, tested with Tasha to make sure it turns on no issues. | Device needs to be sent out for repair
1069354 | 200196 | ACC | CMPLT | CMPLT | 2023-12-06 16:43 | 2023-12-06 16:44 | Incoming inspection | (TC 12/05/2023 04:43:12 PM, Rendon Lopez , Elizabeth) Loaner device incoming inspection. Placed on cart, attached all labels,c leaned. Returned to L&D OR1. See attached report
67416 | 200197 | PM | CMPLT | CMPLT | 2023-12-06 16:51 | 2023-12-06 16:51 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/05/2023 04:51:00 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
502048 | 200198 | CM | CMPLT | CMPLT | 2023-12-06 16:57 | 2024-02-26 13:27 | MSH OPG SPP C2 iU22 DOPPLER ERRORS SN# B0MTHZ | (TC 12/07/2023 11:26:30 AM, Whittington, Scott) MSH OPG SPP C2 iU22 DOPPLER ERRORS SN# B0MTHZ | MSH OPG SPP C2 iU22 DOPPLER ERRORS SN# B0MTHZ
501070 | 200199 | CM | CMPLT | CMPLT | 2023-12-06 16:59 | 2023-12-29 13:19 | MSH CEOU RM 10 OPERATOR CONTROL SLOW? OR STICKY | (TC 12/07/2023 11:26:02 AM, Whittington, Scott) MSH CEOU RM 10 OPERATOR CONTROL SLOW? OR STICKY DELETE PT DB, WAS AT 13% USED | MSH CEOU RM 10 OPERATOR CONTROL SLOW? OR STICKY DELETE PT DB, WAS AT 13% USED QWERTY KYBD AND CABLE WERE FAULTY, WHICH WOULD INDUCE OPPANEL ERRORS QWERTY KYBD & CABLE REPLACED DECEMBER 29
1053381 | 200200 | CM | CMPLT | CMPLT | 2023-12-06 17:02 | 2023-12-06 17:05 | Expiratory cassette membrane needs replacement | (TC 12/06/2023 05:02:16 PM, Rendon Lopez , Elizabeth) Replaced membrane, reset hours. Device failing self tests on internal leak. Troubleshooting, found leak in inspiratory side, removed complete inspiratory channel, cleaned, reseated, device passing. Run on user mode, no issues. Returned to RT NICU
1067441 | 200201 | CM | CMPLT | CMPLT | 2023-12-06 17:09 | 2023-12-06 17:14 | Blue light malfunction | (TC 12/06/2023 05:13:57 PM, Rendon Lopez , Elizabeth) Replaced device, updated TMS SN. Charged, tested-OK
1059741 | 200202 | CM | CMPLT | CMPLT | 2023-12-06 17:15 | 2023-12-06 17:19 | Top button broken off | (TC 12/06/2023 05:15:14 PM, Rendon Lopez , Elizabeth) Replaced top cover from spare. Charged, tested-OK.
1057279 | 200203 | CM | CMPLT | CMPLT | 2023-12-07 09:04 | 2024-02-02 10:19 | Battery not charging | (TC 12/07/2023 09:03:45 AM, Campbell, Connor) battery not charging, tuns off when unplugged from AC, but works when plugged in displaying red charging light after being plugged in for a while, needs service removed unit accesories and removed it from the stand, will call masimo to send in for service and update this WO with the PO when done (TC 02/02/2024 10:15:25 AM, Cortez, John) Replacement unit received, unboxed electrical sfety passed- 130 mohms, 68 uAs Updated asset# on TMS Added config Added spare SPO2 sensor All parameters tested ok (TC 02/02/2024 10:25:44 AM, Cortez, John) Added radical battery
1056476 | 200204 | CM | CMPLT | CMPLT | 2023-12-07 09:06 | 2024-02-02 13:43 | battery not charging, does not turn on at all | (TC 12/07/2023 09:05:47 AM, Campbell, Connor) tried to charge battery, battery charging light indicator is red unit will not turn on when plugged in or when unplugged continual beeping and power light flashes red after displaying masimo boot screen unit needs service, will call masimo and update the WO with the PO once i have it (TC 02/02/2024 01:38:31 PM, Cortez, John) Replacement received, updated SN# on TMS Unboxed, added configuration Electrical safety passed- 150 mohms, 70 uAs Added Radical battery Missing mountinhg screw, added replacement screw All parameters tested ok
| 200205 | CM | CMPLT | CMPLT | 2023-12-07 09:13 | 2024-05-02 10:26 | Sending CMAC blades for inspection/repair | (TC 12/07/2023 09:21:01 AM, Baghirzade, Nara) SC 1003124288 (TC 12/08/2023 10:14:29 AM, Baghirzade, Nara) Steve Arai agreed to the exchange program for CMAC blade SN: YR30854. Will update purchasing department to change the cost to the quoted cost. (TC 12/08/2023 10:34:17 AM, Baghirzade, Nara) Shipping both CMAC blades to Storz. RMA # 200584007 Waybill # 405657 (TC 12/14/2023 07:54:28 AM, Baghirzade, Nara) Received CMAC blade from Storz as part of the exchange program. New SN: VM3051 po 7202332606 (TC 12/14/2023 11:56:21 AM, Baghirzade, Nara) Emailed Fan from mohawkmedbuy to change the cost of the PO. Indicated that the cost is only for the item with serial number YR30854. (TC 05/02/2024 10:25:14 AM, Baghirzade, Nara) Steve Arai does not want to repair any further CMAC blades as new ones have been ordered and more requests will be placed for next year's capital. | Sending 2 blades for inspection. 1. SN: TN1327 – image is very dark and grainy. Blade to be assessed for warranty replacement as it is less than two years old. 2. SN: YR30854 – image is grainy, light is not working. The 8403XX blades have reached end-of-service and are being in exchanged with the 8404XX’s. Forwarded quote to Steve Arai.
| 200206 | PA | CMPLT | CMPLT | 2023-12-07 10:03 | 2023-12-22 08:03 | Ordering PM Kits for Servo U | (TC 12/07/2023 01:09:23 PM, Canjimavadivel, Kamallesh ) Placed order for PM kits (TC 12/20/2023 08:02:25 AM, Canjimavadivel, Kamallesh ) Received PM kits
67922 | 200207 | CM | CMPLT | CMPLT | 2023-12-07 12:34 | 2023-12-07 12:35 | Humidity not working; can only go up to 10% humidity | (TC 12/07/2023 12:33:51 PM, Baghirzade, Nara) Replaced the humidity sensor. Calibrated the RH sensor. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.4% RH. Incubator is fully functional.
1059838 | 200208 | CM | CMPLT | CMPLT | 2023-12-07 12:37 | 2023-12-07 12:37 | Clarion Eyecubed In-service | (TC 12/07/2023 12:37:10 PM, Baghirzade, Nara) Clarion Eyecubed In-service
| 200209 | CM | CMPLT | CMPLT | 2023-12-07 12:40 | 2023-12-07 12:40 | Biomed meeting | Biomed meeting
1054070 | 200210 | CM | CMPLT | CMPLT | 2023-12-07 13:17 | 2023-12-07 13:18 | spare unit that need to be sent to 17th floor | (TC 12/07/2023 01:17:12 PM, Campbell, Connor) retrieved unit from storage tested all functions, passed charged battery brought to 17th floor (TC 12/07/2023 01:17:12 PM, Campbell, Connor) retrieved unit from storage tested all functions, passed charged battery brought to 17th floor
1055773 | 200211 | CM | CMPLT | CMPLT | 2023-12-07 13:52 | 2023-12-07 13:54 | Plug in AC - did not check in on dashboard | (TC 12/07/2023 01:52:15 PM, Baghirzade, Nara) Completed 30J for transmittance test. Transmitted successfully. Updated Tomi (on-call) & biomed team about this as she flagged it as part of her on-call checks.
1055521 | 200212 | CM | CMPLT | CMPLT | 2023-12-07 14:04 | 2023-12-07 14:05 | thermomter probe faulty, taking too long to read | (TC 12/07/2023 02:03:59 PM, Campbell, Connor) couldnt find any new probes, replaced with a spare probe from broken asset 66161 performed functionality test, thermomter now functional returned to service
1059828 | 200213 | CM | CMPLT | CMPLT | 2023-12-07 14:37 | 2023-12-08 09:01 | spraying water when turned on out of the 3rd top port | (TC 12/07/2023 02:37:30 PM, Campbell, Connor) replicated issue, cleaned all the ports and removed front cover will take apart more with nara to determine the exact issue and repair it (TC 12/08/2023 09:00:08 AM, Campbell, Connor) replaced the top 3rd port valve with the corresponding valve from the broken ED unit that was in the storage room added teflon tape on the threads tested functionality, valvle no longer leaking or spraying confirmed by nara returned to service | due to the seal getting stuck on the port
1064418 | 200214 | CM | CMPLT | CMPLT | 2023-12-07 14:39 | 2024-02-05 08:43 | Fan not working | (TC 12/07/2023 02:38:42 PM, Chukri, Samir) Checked the MRI incubator yesterday found the fan free but not operating (running ) need more investigation (TC 12/08/2023 02:06:39 PM, Chukri, Samir) went with Qwynn to video the fan problem and Qwynn will email it to Michael. (TC 02/05/2024 07:51:02 AM, Chukri, Samir) completed and replaced by vendor tech.
66161 | 200215 | CM | CMPLT | CMPLT | 2023-12-07 14:43 | 2023-12-07 14:44 | not working | (TC 12/07/2023 02:42:52 PM, Campbell, Connor) tried changing the batteries tried changing the probe tried changing the top part unit will not function, and is out of warratny will retire and ewaste
1069377 | 200216 | ACC | CMPLT | CMPLT | 2023-12-07 14:48 | 2023-12-13 10:58 | Incoming inspection | (TC 12/11/2023 10:23:09 AM, Cortez, John) passed
1069378 | 200217 | ACC | CMPLT | CMPLT | 2023-12-07 14:49 | 2023-12-13 11:00 | Incoming inspection | (TC 12/11/2023 10:59:44 AM, Cortez, John) Visual inspection and electrical safety passed
1069379 | 200218 | ACC | CMPLT | CMPLT | 2023-12-07 14:51 | 2023-12-13 11:02 | Incoming inspection | (TC 12/11/2023 11:02:00 AM, Cortez, John) Visual inspection and electrical safety passed
1069380 | 200219 | ACC | CMPLT | CMPLT | 2023-12-07 14:52 | 2023-12-13 11:04 | Incoming inspection | (TC 12/11/2023 11:03:57 AM, Cortez, John) Visual inspection and electrical safety passed
502089 | 200220 | CM | ACTIV | ISSUE | 2023-12-07 14:53 |  | detector calibrations | (TC 12/08/2023 03:46:09 PM, Green, Paul) Same problem on Symbia2 detector2. June has opened a call with Support and going to contact factory for input. Running a monthly but sources are weak. Will see if it completes by Monday.
1069381 | 200221 | ACC | CMPLT | CMPLT | 2023-12-07 14:55 | 2023-12-13 11:06 | Incoming inspection | (TC 12/11/2023 11:06:01 AM, Cortez, John) Visual inspection and electrical safety passed
1069382 | 200222 | ACC | CMPLT | CMPLT | 2023-12-07 14:56 | 2023-12-13 11:09 | Incoming inspection | (TC 12/11/2023 11:08:14 AM, Cortez, John) Visual inspection and electrical safety passed
1069383 | 200223 | ACC | CMPLT | CMPLT | 2023-12-07 14:57 | 2023-12-13 11:11 | Incoming inspection | (TC 12/11/2023 11:10:23 AM, Cortez, John) Visual inspection and electrical safety passed
1069384 | 200224 | ACC | CMPLT | CMPLT | 2023-12-07 15:03 | 2023-12-13 11:13 | Incoming inspection | (TC 12/11/2023 11:12:34 AM, Cortez, John) Visual inspection and electrical safety passed
32520 | 200225 | CM | CMPLT | CMPLT | 2023-12-07 22:39 | 2023-12-08 09:08 | MSH DR2 TABLE CONTROLLER FAULTY; REPAIR | (TC 12/07/2023 10:39:20 PM, Whittington, Scott) MSH DR2 TABLE CONTROLLER FAULTY; REPAIR (TC 12/07/2023 10:39:20 PM, Whittington, Scott) MSH DR2 TABLE CONTROLLER FAULTY; REPAIR | MSH DR2 TABLE CONTROLLER FAULTY; REPAIR
1062389 | 200226 | CM | CMPLT | NFF | 2023-12-08 08:55 | 2023-12-08 08:57 | Reported not working | (TC 12/08/2023 08:55:11 AM, Baghirzade, Nara) Tested all bed functions. No errors found. Bed is functional. No fault found. (TC 12/08/2023 08:56:16 AM, Rhuland, Courtney) Tested all bed functions. No errors found. Bed is functional. No fault found.
1063709 | 200227 | CM | CMPLT | CMPLT | 2023-12-08 09:24 | 2024-02-26 13:25 | MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN | (TC 12/08/2023 09:23:39 AM, Whittington, Scott) MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN (TC 12/08/2023 09:23:50 AM, Whittington, Scott) MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN (TC 12/08/2023 08:00:00 AM, Whittington, Scott) MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN (TC 12/20/2023 11:09:19 AM, Whittington, Scott) MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN MEET WITH REEBA & NAZ (TC 12/20/2023 05:20:29 PM, Dolinar, John) MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN INVESTIGATE PROBE DISTRIBUTION INVESTIGATE VIDEO O/P (TC 12/08/2023 08:00:00 AM, Whittington, Scott) MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN | MSH 7 OR SPARQ BATTERIES DRAINED, SYSTEM NOT PLUGGED IN
| 200228 | RDS | CMPLT | CMPLT | 2023-12-08 09:34 | 2023-12-08 09:39 | Weekly rounds: 18 South | (TC 12/08/2023 09:33:44 AM, Baghirzade, Nara) Weekly rounds at 18 South completed. See procedure for results.
1069386 | 200229 | CM | CMPLT | CMPLT | 2023-12-08 09:39 | 2024-01-30 11:40 | Incoming inspection
| 200230 | RDS | CMPLT | CMPLT | 2023-12-08 09:40 | 2023-12-08 09:50 | Weekly rounds: 7 PACU POD A | (TC 12/08/2023 09:39:37 AM, Baghirzade, Nara) Weekly rounds in 7 PACU POD A completed. See procedure for results.
| 200231 | RDS | CMPLT | CMPLT | 2023-12-08 09:51 | 2023-12-08 09:58 | Weekly rounds: 11 South | (TC 12/08/2023 09:55:18 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
1066474 | 200232 | CM | CMPLT | CMPLT | 2023-12-08 09:55 | 2023-12-20 13:49 | Filter needs to be changed | (TC 12/08/2023 09:54:56 AM, Reid, Emmerson ) Checked the filters this morning and found this one needs to be replaced. This should have been flagged by the SAs, so I followed up with Andrew who asked if I could reach out to Sean to find out why this happened and to get a Safer report filed. (TC 12/11/2023 08:04:24 AM, Reid, Emmerson ) Unit was flagged by Support Services. I pulled it from service and left it in OR12 with a sign not to use it. Waiting for a quote for new filters. (TC 12/20/2023 01:34:13 PM, Reid, Emmerson ) the three filters all came in. I changed the filter on this neptune and cleaned inside of the cavity with clorox before inserting the new filter. The old filter was disposed through biohazardous waste in a yellow bag. I ran the neptune on the "normal" cleaning cycle and let the nursing station know it can be use again. Working now.
| 200233 | RDS | CMPLT | CMPLT | 2023-12-08 09:58 | 2023-12-08 10:00 | Weekly rounds: NICU | (TC 12/08/2023 09:58:25 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure for results.
| 200234 | RDS | CMPLT | CMPLT | 2023-12-08 10:00 | 2023-12-08 10:01 | Weekly rounds: RESUS | (TC 12/08/2023 10:00:11 AM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure for results.
1055520 | 200235 | CM | CMPLT | CMPLT | 2023-12-08 10:02 | 2023-12-08 10:03 | not recording reading | (TC 12/08/2023 10:02:33 AM, Campbell, Connor) tested thermometer, when taking reading it just says OFF instead of a number determined that these thermometers are too old to be worth repairing, NICU is ordering new ones
1056444 | 200236 | CM | CMPLT | CMPLT | 2023-12-08 10:03 | 2023-12-08 10:06 | Reported not working; no issue stated | (TC 12/08/2023 10:03:00 AM, Baghirzade, Nara) During my rounds, i found a masimo with a paper on it that read "not working", asked the staff at the nursing station but no one knew what the issue with it was. It was also missing a power cord. Searched the unit and found a power cord that belongs to a Masimo unit, it was tucked near the window sill at the end of the hall. Plugged the Masimo to AC; battery charging ok. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1063746 | 200237 | PM | CMPLT | CMPLT | 2023-12-08 11:07 | 2023-12-08 11:08 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/08/2023 11:07:23 AM, Campbell, Connor) Drained, disinfected, cleaned, rinsed, refilled
1066935 | 200238 | CM | CMPLT | CMPLT | 2023-12-08 11:15 | 2024-01-30 09:35 | Device to be sent to Olympus for investigation | (TC 12/08/2023 01:57:59 PM, Cortez, John) Requested PO for repair and for the loaner (TC 12/11/2023 08:14:37 AM, Samanian, Kamran) arrnaged with KJV to ship the box to Olympus for repair, waybill# 405957 (TC 12/12/2023 04:17:33 PM, Cortez, John) Moved CLV put in room 4 (Ambulatory CLV) to bronch room Put loaner CLV (SN 7670841) in room 4. Tested both light sources with scope; both ok (TC 01/25/2024 10:29:34 AM, Cortez, John) Received back from repair. Loose paper stuck in between outer housing and light source frame. Removed. Incoming inspection otherwise passed. 5 mohms, 73 uAs Waiting on instruction from Anita for installation location (TC 01/29/2024 09:32:21 AM, Cortez, John) Installed light source back in room 4, tested ok with a scope
1064939 | 200239 | CM | CMPLT | CMPLT | 2023-12-08 11:57 | 2024-01-30 09:30 | Needs to be sent to Olympus for investigation | (TC 12/08/2023 01:59:01 PM, Cortez, John) Requested PO for repair and loaner (TC 12/11/2023 08:16:26 AM, Samanian, Kamran) arranged with KJV to ship the box to Olympus Waybill#405957 (TC 01/24/2024 10:01:29 AM, Cortez, John) Received back from Olympus. Incoming inspection passed 50 mohms, 30 uAs Waiting on instruction from Anita for installation location (TC 01/29/2024 09:27:44 AM, Cortez, John) Installed light source in bronch room, tested ok with scope
| 200240 | CM | ACTIV | ISSUE | 2023-12-08 12:55 |  | working with John and Kamran for Endo light source CLV-190 water leaking issues | (TC 12/08/2023 12:55:05 PM, Chukri, Samir) working with Johan and Kamran to investigating the water leaking into Olympus CLV-190 light source replaced PO for repair and loaner .
| 200241 | CM | CMPLT | CMPLT | 2023-12-08 14:00 | 2024-01-02 10:38 | misc biomed - dec 2023 | (TC 12/08/2023 02:00:41 PM, Cortez, John) Swapped processor from room 1 with spare from OR (TC 12/07/2023 02:01:35 PM, Cortez, John) Swapped processor from room 4 with one from bronch room (TC 12/07/2023 02:02:55 PM, Cortez, John) Replaced hose on manual BP in room 10-106 (TC 12/06/2023 02:03:45 PM, Cortez, John) Helped find spare temp cables and rectal probes (for use on Philips monitors) for Emerg. Borrowed from ICU, will charge back (TC 12/06/2023 02:07:58 PM, Cortez, John) Picked up lamp from OR. Asset was retired a few weeks ago. Handed to Jessica. (TC 12/08/2023 09:51:31 AM, Cortez, John) On-call daily check (TC 12/11/2023 09:52:06 AM, Cortez, John) On-call daily check (TC 12/12/2023 09:52:21 AM, Cortez, John) On-call daily check (TC 12/11/2023 09:53:02 AM, Cortez, John) Assisted endo with moving light sources between bronch room and room 1 Had an issue after installation into room 1. E315 popped on screen with "incompatible scope" message Worked after reseating cables (TC 12/12/2023 04:23:16 PM, Cortez, John) TMS work (TC 12/13/2023 11:17:02 AM, Cortez, John) on-call daily check (TC 12/13/2023 03:57:47 PM, Cortez, John) huddle (TC 12/13/2023 11:17:21 AM, Cortez, John) TMS work (TC 12/15/2023 12:56:09 PM, Cortez, John) Resetted 4S PICIX in emerg (TC 12/15/2023 01:11:44 PM, Cortez, John) IS meeting (TC 12/15/2023 03:22:17 PM, Cortez, John) IBE solution review (TC 12/19/2023 01:30:44 PM, Cortez, John) Met with Anna, Vincy and Becky regarding TruTurn feature of Isotour pumps Observed that angle of head section may affect speed of mattress inflation. Delivered second TruTurn capable mattress to 10N with Connor (TC 12/19/2023 01:35:25 PM, Cortez, John) Compiled spreadsheet of Endo CLV-190 locations/dates of arrival for Anita (TC 12/18/2023 01:58:34 PM, Cortez, John) Anita from Endo wanted biomed to inspect light source in room 4 for sign of fluid in air port tubing Disassembled and found no signs of water (TC 12/18/2023 02:34:43 PM, Cortez, John) Helped replace UPS battery in old MCR (TC 12/20/2023 08:33:38 AM, Cortez, John) ED Annex needed pumps, sent 2 PCUs and 4 LVPs (TC 12/22/2023 10:39:30 AM, Cortez, John) Helpled Nara with PICIX issue (TC 12/20/2023 11:54:48 AM, Cortez, John) Department meeting (TC 12/22/2023 11:55:20 AM, Cortez, John) systems committee (TC 12/22/2023 01:35:10 PM, Cortez, John) Helped Rob from Oympus swap out demo lightsource with original in endo room 3 (TC 12/27/2023 10:34:15 AM, Cortez, John) PFT troubleshooting with Nara and Samir
1054460 | 200242 | CM | CMPLT | CMPLT | 2023-12-08 14:06 | 2023-12-08 14:07 | "alarming, low solution level" | (TC 12/08/2023 02:06:19 PM, Cortez, John) Set was installed incorrectly. Reseated and errors cleared. Showed nurse correct set installation. (TC 12/06/2023 02:06:19 PM, Cortez, John) Set was installed incorrectly. Reseated and errors cleared. Showed nurse correct set installation.
1069387 | 200243 | ACC | CMPLT | CMPLT | 2023-12-08 14:18 | 2023-12-13 10:08 | Incoming Inspection | (TC 12/11/2023 10:06:26 AM, Cortez, John) Visual inspection and electrical safety passed
1069388 | 200244 | ACC | CMPLT | CMPLT | 2023-12-08 14:38 | 2023-12-13 10:15 | Incoming Inspection | (TC 12/11/2023 10:14:57 AM, Cortez, John) Visual inspection and electrical safety passed
1069389 | 200245 | ACC | CMPLT | CMPLT | 2023-12-08 14:43 | 2023-12-13 10:17 | Incoming Inspection | (TC 12/11/2023 10:16:32 AM, Cortez, John) Visual inspection and electrical safety passed
1069390 | 200246 | ACC | CMPLT | CMPLT | 2023-12-08 14:45 | 2023-12-13 10:35 | Incoming Inspection | (TC 12/11/2023 10:22:05 AM, Cortez, John) Visual inspection and electrical safety passed
1069391 | 200247 | ACC | CMPLT | CMPLT | 2023-12-08 14:47 | 2023-12-13 10:47 | Incoming Inspection | (TC 12/11/2023 10:45:44 AM, Cortez, John) Visual inspection and electrical safety passed
1069392 | 200248 | ACC | CMPLT | CMPLT | 2023-12-08 14:52 | 2023-12-13 10:51 | Incoming Inspection | (TC 12/11/2023 10:50:10 AM, Cortez, John) Visual inspection and electrical safety passed
1047046 | 200249 | CM | CMPLT | NFF | 2023-12-08 14:56 | 2023-12-08 14:57 | "needs biomed" | (TC 12/08/2023 02:56:30 PM, Campbell, Connor) Perfomed a check in, passed no fault found
1069393 | 200250 | ACC | CMPLT | CMPLT | 2023-12-08 15:00 | 2023-12-13 10:54 | Incoming Inspection | (TC 12/11/2023 10:52:58 AM, Cortez, John) Visual inspection and electrical safety passed
1069394 | 200251 | ACC | CMPLT | CMPLT | 2023-12-08 15:03 | 2023-12-13 10:55 | Incoming Inspection | (TC 12/11/2023 10:55:00 AM, Cortez, John) Visual inspection and electrical safety passed
| 200252 | CM | CMPLT | CMPLT | 2023-12-11 08:09 | 2023-12-11 08:09 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
| 200253 | RDS | CMPLT | CMPLT | 2023-12-11 08:47 | 2023-12-11 08:51 | Weekly Checks For PACU | (TC 12/11/2023 08:50:02 AM, Reid, Emmerson ) Weekly rounds complete.
| 200254 | RDS | CMPLT | CMPLT | 2023-12-11 08:54 | 2023-12-11 08:56 | Weekly Checks For OR | (TC 12/11/2023 08:55:30 AM, Reid, Emmerson ) Rounds complete
1068221 | 200255 | CM | CMPLT | CMPLT | 2023-12-11 08:58 | 2023-12-11 16:33 | MSH 5 US ROOM 6 CANON PRISM ERROR MSG | (TC 12/11/2023 08:59:19 AM, Whittington, Scott) From: Whittington, Scott Sent: December 11, 2023 8:54 AM To: Kumaran.Thayumanavan@medical.canon; Ryan. Francis <Ryan.Francis@medical.canon> Cc: Rice, Murray <Murray.Rice@uhn.ca>; Pejcic, Sonja <Sonja.Pejcic@uhn.ca>; sameh. mekhaeil <sameh.mekhaeil@medical.canon>; Fan, Tak <Tak.Fan@uhn.ca>; Chau, Ken <Ken.Chau@uhn.ca> Subject: MSH CANON PRISM ERROR CODE, no site# Good morning One of our new Canon Prisms at MSH is displaying an error and is down. We don't have a site# yet, so unable to place a service call on 1-800-667-2671 Sn# SHK23X2073 ERROR Hardware data transfer is busy. Cine memory is cleared. Thanks Scott 416-402-2462 (TC 12/11/2023 04:31:55 PM, Whittington, Scott) MSH 5 US ROOM 6 CANON PRISM ERROR MSG SAMEH DOWNLOADED ERROR LOGS, THEN PERFORMED BD TEST RC/RX BOARD FAILED SAMEH HAD A SPARE RC/RX BOARD IN HIS CAR LOAD FIRMWARE ON IT OK | MSH 5 US ROOM 6 CANON PRISM ERROR MSG
1050482 | 200256 | CM | CMPLT | CMPLT | 2023-12-11 08:59 | 2024-01-31 08:43 | ordering parts for Servo-Ns | ordering Philips dongles for Servo-Ns. received Philips dongles for Servo-Ns.
| 200257 | RDS | CMPLT | CMPLT | 2023-12-11 09:09 | 2023-12-11 13:02 | Weekly check for ED
| 200258 | RDS | CMPLT | CMPLT | 2023-12-11 09:10 | 2023-12-12 07:52 | Weekly check for ICU
| 200259 | CM | CMPLT | CMPLT | 2023-12-11 10:23 | 2024-01-11 10:50 | Equipment inventory | (TC 12/11/2023 10:23:39 AM, Baghirzade, Nara) Took inventory at both Ophthalmology clincs of all medical/ophthalmic equipment for inventory purposes and capital. Will be adding the equipment to TMS, creating assets and tagging on another day. (TC 12/11/2023 01:58:28 PM, Baghirzade, Nara) Accessed exam rooms 9 and 20 with Henna. Went over the entire equipment list with Henna and she was able to obtain approximate dates for the equipment that had missing dates. Added lensometers and 1 OCT to the list as well. (TC 12/18/2023 10:08:17 AM, Rhuland, Courtney) Created assets for all untagged equipment (TC 01/04/2024 10:20:41 AM, Rhuland, Courtney) Tagged equipment in clinic (TC 01/11/2024 10:49:22 AM, Rhuland, Courtney) Created assets and tagged all of the slit lamps in clinic Clinic inventory up to date | Updating all equipment that hasn't been tagged in the clinic for inventory and capital purposes.
| 200260 | CM | CMPLT | CMPLT | 2023-12-11 11:11 | 2023-12-11 11:14 | NICU call Zone C | (TC 12/11/2023 11:12:33 AM, Baghirzade, Nara) Call from NICU came in regarding an HL7 error appearing on the overview station in Zone C. Went up to NICU and spoke to a nurse in Zone C; let her know that the HL7 charting issue is pertaining to the research server and does not impact them or patient monitoring in any way. Relayed the same message to a nurse that was at the main central station. | HL7 error on overview
66060 | 200261 | CM | CMPLT | CMPLT | 2023-12-11 11:15 | 2023-12-11 11:19 | Damaged power cord | (TC 12/11/2023 11:15:18 AM, Baghirzade, Nara) Found the power cord very damaged with rips and tears on it. The yellow biomed tag said that an SA got shocked when they touched the power cord. Showed the power cord and tag to Samir. Removed the power cord and tag and left it with Samir. Replaced the power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 70 mohms Leakage current: 97.5 uA
1059550 | 200262 | CM | CMPLT | CMPLT | 2023-12-11 11:30 | 2023-12-11 11:31 | Display shut off | (TC 12/11/2023 11:29:49 AM, Baghirzade, Nara) The display was completely off however the PICIX was on. The UPS was falling off the rack so placed the UPS back on the rack with the PICIX and tightened it. Re-inforced the display cable at the back of the monitor as well as to the PICIX. Display came back on. Overview is fully functional.
| 200263 | PA | CMPLT | CMPLT | 2023-12-11 11:51 | 2023-12-22 13:15 | Formlabs Order - Resin and Tray
1053349 | 200264 | CM | CMPLT | CMPLT | 2023-12-11 12:59 | 2024-02-26 13:23 | MSH BI 1274 H2 mie 1053349 COMPRESSION PADDLE MOTOR FAILED | (TC 12/11/2023 12:58:41 PM, Whittington, Scott) MSH BI 1274 H2 mie 1053349 COMPRESSION PADDLE MOTOR FAILED | 1053349 MSH BI 1274 H2 mie 1053349 COMPRESSION PADDLE MOTOR FAILED
| 200265 | RDS | CMPLT | CMPLT | 2023-12-11 14:02 | 2023-12-11 14:09 | Weekly checks in CCU | (TC 12/11/2023 02:04:04 PM, Scero, Tomlyn) Swaped out a LVP in Echo because the tech said it was dripping fluid even when stopped. (TC 12/11/2023 02:07:16 PM, Scero, Tomlyn) 30 J tests complete.
502044 | 200266 | RQ | CMPLT | CMPLT | 2023-12-11 14:11 | 2023-12-12 07:31 | ERGO brake is tight | (TC 12/12/2023 07:29:03 AM, Green, Paul) Messaged Bernard about the brake and that it should be tight when engaging. We have discussed this with Digirad and we have adjusted the brake accordingly. Informed Bernard of this. | ERGO brake is tight, difficult to lock
1054538 | 200267 | CM | CMPLT | CMPLT | 2023-12-11 15:12 | 2023-12-12 09:26 | Exposed wires | (TC 12/11/2023 03:15:00 PM, Canjimavadivel, Kamallesh ) Multiple damaged leads found. Replaced RL, V1-V3, V4-V5 (TC 12/12/2023 09:21:55 AM, Canjimavadivel, Kamallesh ) Replaced V9 lead. Installed arm mount for PIM. Tested all leads with simulator, working ok.
1046196 | 200268 | CM | CMPLT | CMPLT | 2023-12-11 15:50 | 2023-12-11 15:50 | System not going to standby mode | System not going to standby mode Troubleshooting and parts ordering done by Siemens on Sunday - remotely No communication to the UMAR Further onsite troubleshooting was done the following day UMAR was replaced Tune up and system check up done Systrem returned for use after completing all tests.
| 200269 | OC | CMPLT | CMPLT | 2023-12-11 16:07 | 2023-12-19 14:34 | John Cortez On-Call; Dec 8 - Dec 14, 2023 | (TC 12/08/2023 10:30:00 PM, Cortez, John) Command centre called regarding a fetal monitor issue in L&D triage room G. They mentioned "the monitor cannot associate with the network". Reached out to Kamran to get more info on how the system works. Reached back to command centre and suggested that they can reset the CE box attached to rear of fetal monitor. Command centre support already tried that. I suggested that they can try resetting monitor itself, but the issue seems to be network related. Didn't need to go on-site as it didn't seem like a hardware issue. Told command centre that to reach out to fetal link support team.
1059795 | 200270 | CM | CMPLT | CMPLT | 2023-12-11 19:33 | 2023-12-11 19:32 | External remote monitor output port not working. | Fixed the problem.
1067505 | 200271 | CM | CMPLT | CMPLT | 2023-12-12 07:59 | 2023-12-12 08:00 | Does not turn on | (TC 12/12/2023 07:59:19 AM, Baghirzade, Nara) Replaced batteries but still does not switch on. This unit was originally repurposed and no more spares left in biomed to replace this with. Confirmed with Samir to retire this unit as NICU has submitted a capital request for a purchase of more thermometers.
62200 | 200272 | CM | CMPLT | CMPLT | 2023-12-12 08:02 | 2023-12-12 08:02 | Unit does not turn on | (TC 12/12/2023 08:01:38 AM, Baghirzade, Nara) Replaced batteries but still does not switch on. No repairs/parts available and unit is end of support. Confirmed with Samir to retire this unit as NICU has submitted a capital request for a purchase of more thermometers.
1057611 | 200273 | CM | CMPLT | NFF | 2023-12-12 08:10 | 2023-12-12 08:11 | "temperature module not present" | (TC 12/12/2023 08:09:59 AM, Campbell, Connor) tested functionality of temp, NIBP and SpO2 no fault found returned to service
1057124 | 200274 | CM | CMPLT | CMPLT | 2023-12-12 08:24 | 2023-12-12 08:25 | "bp cuff not inflating" | (TC 12/12/2023 08:23:40 AM, Campbell, Connor) tested temp and SpO2, those worked fine tested NIBP after re steating the connections on multiple modes and with multiple cuffs, no fault found the user likely had the wrong cuff attached for the neonatal setting as it was set to neonate but with an XL adult cuff attached returned to service
66060 | 200275 | CM | CMPLT | CMPLT | 2023-12-12 08:27 | 2023-12-12 08:28 | BME incident report filing | (TC 12/12/2023 08:27:19 AM, Baghirzade, Nara) Filed a BME incident report regarding the damaged power cord and saved to T drive in 2023 folder.
| 200276 | CM | CMPLT | CMPLT | 2023-12-12 08:43 | 2024-01-17 09:28 | Panda warmer parts | (TC 12/12/2023 08:46:46 AM, Rendon Lopez , Elizabeth) L&D panda warmer blender overhaul and spare parts Received
67230 | 200277 | CM | CMPLT | CMPLT | 2023-12-12 09:09 | 2023-12-12 09:10 | Created by Failed PM WO # 199497 : PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/08/2023 09:09:26 AM, Rendon Lopez , Elizabeth) Adjusted blender O2 dial. Tested all functions-OK. PM completed, all tests passed | (TC 12/07/2023 09:08:03 AM, Rendon Lopez , Elizabeth) Found Resus O2 knob loose during PM, removed from room
1062331 | 200278 | CM | CMPLT | CMPLT | 2023-12-12 09:15 | 2023-12-12 09:17 | Low battery and low temp alarm | (TC 12/11/2023 09:15:58 AM, Rendon Lopez , Elizabeth) Replaced batteries and temperature probe. Tested-OK
1059659 | 200279 | CM | CMPLT | CMPLT | 2023-12-12 09:17 | 2023-12-12 09:17 | Button broken | (TC 12/11/2023 09:17:02 AM, Rendon Lopez , Elizabeth) Replaced button, tested-OK
1051290 | 200280 | CM | CMPLT | CMPLT | 2023-12-12 09:18 | 2023-12-12 09:19 | Broken - Physical damage | (TC 12/11/2023 09:17:55 AM, Rendon Lopez , Elizabeth) Replaced device. Updated TMS SN, charged, tested-OK
1069398 | 200281 | ACC | CMPLT | CMPLT | 2023-12-12 09:19 | 2023-12-12 09:19 | Incoming inspection | (TC 12/11/2023 09:19:08 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069399 | 200282 | ACC | CMPLT | CMPLT | 2023-12-12 09:20 | 2023-12-12 09:20 | Incoming inspection | (TC 12/11/2023 09:20:07 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069400 | 200283 | ACC | CMPLT | CMPLT | 2023-12-12 09:21 | 2023-12-12 09:21 | Incoming inspection | (TC 12/11/2023 09:20:56 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069401 | 200284 | ACC | CMPLT | CMPLT | 2023-12-12 09:21 | 2023-12-12 09:22 | Incoming inspection | (TC 12/11/2023 09:21:39 AM, Rendon Lopez , Elizabeth) Incoming inspection
| 200285 | CM | CMPLT | CMPLT | 2023-12-12 09:23 | 2023-12-12 09:23 | OPG otoscope not working | (TC 12/11/2023 09:22:47 AM, Rendon Lopez , Elizabeth) OPG 3rd floor genetics clinic, otoscope not working. Replaced bulb, tested-OK
67233 | 200286 | CM | CMPLT | CMPLT | 2023-12-12 09:25 | 2023-12-12 09:26 | Created by Failed PM WO # 199500 : PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 12/11/2023 09:25:01 AM, Rendon Lopez , Elizabeth) Replaced on/off switch. Replaced blender. PM completed, all tests passed | (TC 12/11/2023 09:23:59 AM, Rendon Lopez , Elizabeth) Resus unit leaking when in off position. Removed warmer from room for repair
67425 | 200287 | PM | CMPLT | CMPLT | 2023-12-12 09:37 | 2023-12-12 09:38 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/11/2023 09:37:29 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67426 | 200288 | PM | CMPLT | CMPLT | 2023-12-12 09:38 | 2023-12-12 09:39 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 12/11/2023 09:38:29 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67914 | 200289 | CM | CMPLT | CMPLT | 2023-12-12 09:40 | 2023-12-12 09:42 | Broken porthole door and SE latch | (TC 12/12/2023 09:40:22 AM, Baghirzade, Nara) Porthole door was broken in half and hinge was also damaged. Replaced the porthole door and hinge. SE latch was found broken. Replaced the SE latch. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.9% RH. Incubator is fully functional.
1059550 | 200290 | CM | CMPLT | CMPLT | 2023-12-12 09:56 | 2023-12-12 10:02 | "monitor powered off" | (TC 12/11/2023 09:56:25 AM, Cortez, John) Nurse mentioned that computer was off, but it was just the monitor that was off. Found power adapter and speaker on the floor so I assumed it was kicked. Reseated cord to adapter, and placed adapter+speaker on top of PC
1061211 | 200291 | CM | CMPLT | CMPLT | 2023-12-12 11:10 | 2023-12-12 11:11 | Stryker tower was not working | (TC 12/12/2023 11:09:57 AM, Reid, Emmerson ) Went up to check a warmer, that I was able to secure onto the stand and repair. Was told about a camera issue in OR10 with the tower. They were using a 1288 instead of a 1588 camera head. I found them a 1588 and it was working.
66463 | 200292 | CM | CMPLT | CMPLT | 2023-12-12 11:34 | 2023-12-12 11:34 | nebulizer not passing during self check | did self check 3 times and run unit overnight. functional. returned to RT and informed Rex.
1046276 | 200293 | CM | CMPLT | CMPLT | 2023-12-12 13:10 | 2024-02-15 15:54 | MSH CFRH 7flr APLIO 300 msh ivf 250 dundas procedure 1 aplio 300 no power | (TC 12/12/2023 01:09:49 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 msh ivf 250 dundas procedure 1 aplio 300 no power | MSH CFRH 7flr APLIO 300 msh ivf 250 dundas procedure 1 aplio 300 no power
1063649 | 200294 | CM | CMPLT | CMPLT | 2023-12-12 13:54 | 2023-12-13 08:36 | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | (TC 12/12/2023 01:55:05 PM, Whittington, Scott) MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT
63963 | 200295 | CM | CMPLT | CMPLT | 2023-12-12 14:34 | 2024-01-30 10:04 | Created by Failed PM WO # 199582 : PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/25/2024 10:02:54 AM, Rendon Lopez , Elizabeth) Replaced blender. Adjusted gas outlet. Left observation light off, replaced bulb. PM completed, all tests passed | (TC 12/12/2023 02:33:20 PM, Rendon Lopez , Elizabeth) Resus gas outlet is loose. Removed device from LDR 18 Waiting for blender overhaul for replacement
| 200296 | CM | CMPLT | CMPLT | 2023-12-12 14:39 | 2023-12-12 11:59 | Ivus set up (volcano intrasight mobil) | (TC 12/05/2023 11:59:00 AM, Smith, Matthew) I worked with Matthew Hobs from Philips to set up and test the IVUS system at Angio/ IVR MSH .
1060149 | 200297 | CM | CMPLT | CMPLT | 2023-12-12 14:53 | 2023-12-12 14:53 | leads connection is faulty/lead pins wiggle on patient interface module | (TC 12/12/2023 02:52:42 PM, Campbell, Connor) cleaned and re-soldered connections on the pins on the patient interace module that were wiggling tested functionality, connection is now good
1054189 | 200298 | CM | CMPLT | CMPLT | 2023-12-12 15:18 | 2023-12-12 15:18 | Sound cutting in and out | (TC 12/12/2023 03:17:53 PM, Rendon Lopez , Elizabeth) Confirmed issue, defective cable. Device repaired several times. Retired
1063183 | 200299 | CM | CMPLT | CMPLT | 2023-12-12 15:19 | 2023-12-12 15:20 | Not working | (TC 12/12/2023 03:19:06 PM, Rendon Lopez , Elizabeth) Tested, NFF
1059746 | 200300 | CM | CMPLT | CMPLT | 2023-12-12 15:26 | 2023-12-12 15:26 | No sound | (TC 12/12/2023 03:25:41 PM, Rendon Lopez , Elizabeth) Confirmed issue with probe. Replaced probe with spare. Sending probe our for warranty investigation
1055928 | 200301 | CM | CMPLT | CMPLT | 2023-12-13 07:24 | 2023-12-13 07:26 | BP not measuring | (TC 12/13/2023 07:23:53 AM, Baghirzade, Nara) Found the BP cuff was ripped near the tubing side. No BP cuffs in biomed so I contacted Sharmaine (PCM) to make sure that they have been ordering BP cuffs for their devices. Used a testing cuff from the biomed shop to test the NIBP. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1069422 | 200302 | ACC | CMPLT | CMPLT | 2023-12-13 07:55 | 2023-12-13 08:54 | Incoming inspection and asset tagging | (TC 12/13/2023 08:53:11 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:53:11 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069423 | 200303 | ACC | CMPLT | CMPLT | 2023-12-13 07:58 | 2023-12-13 08:55 | Incoming inspection and asset tagging | (TC 12/13/2023 08:54:34 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:54:34 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069424 | 200304 | ACC | CMPLT | CMPLT | 2023-12-13 08:00 | 2023-12-13 08:55 | Incoming inspection and asset tagging | (TC 12/13/2023 08:55:05 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:55:05 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069425 | 200305 | ACC | CMPLT | CMPLT | 2023-12-13 08:01 | 2023-12-13 08:56 | Incoming inspection and asset tagging | (TC 12/13/2023 08:55:36 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:55:36 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069426 | 200306 | ACC | CMPLT | CMPLT | 2023-12-13 08:02 | 2023-12-13 08:56 | Incoming inspection and asset tagging | (TC 12/13/2023 08:56:08 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:56:08 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069427 | 200307 | ACC | CMPLT | CMPLT | 2023-12-13 08:03 | 2023-12-13 08:57 | Incoming inspection and asset tagging | (TC 12/13/2023 08:56:43 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:56:43 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069428 | 200308 | ACC | CMPLT | CMPLT | 2023-12-13 08:04 | 2023-12-13 08:57 | Incoming inspection and asset tagging | (TC 12/13/2023 08:57:16 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:57:16 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069429 | 200309 | CM | CMPLT | CMPLT | 2023-12-13 08:06 | 2023-12-13 08:58 | Incoming inspection and asset tagging | (TC 12/13/2023 08:57:57 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:57:57 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069430 | 200310 | ACC | CMPLT | CMPLT | 2023-12-13 08:07 | 2023-12-13 08:59 | Incoming inspection and asset tagging | (TC 12/13/2023 08:58:43 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:58:43 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:58:43 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1058169 | 200311 | CM | CMPLT | CMPLT | 2023-12-13 08:12 | 2023-12-13 08:12 | Library Upgrade | (TC 12/13/2023 08:11:41 AM, Reid, Emmerson ) Pushed APS EPIDURAL 2023 OCT
1058177 | 200312 | CM | CMPLT | CMPLT | 2023-12-13 08:13 | 2023-12-13 08:13 | Library upgrade | (TC 12/13/2023 08:12:47 AM, Reid, Emmerson ) Pushed APS EPIDURAL 2023 OCT
1058174 | 200313 | CM | CMPLT | CMPLT | 2023-12-13 08:14 | 2023-12-13 08:14 | Library upgrade | (TC 12/13/2023 08:13:43 AM, Reid, Emmerson ) Pushed APS EPIDURAL 2023 OCT
1069431 | 200314 | ACC | CMPLT | CMPLT | 2023-12-13 08:16 | 2023-12-13 09:00 | Incoming inspection and asset tagging | (TC 12/13/2023 08:59:26 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 08:59:26 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069432 | 200315 | ACC | CMPLT | CMPLT | 2023-12-13 08:17 | 2023-12-13 09:00 | Incoming inspection and asset tagging | (TC 12/13/2023 09:00:00 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:00:00 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069433 | 200316 | ACC | CMPLT | CMPLT | 2023-12-13 08:18 | 2023-12-13 09:01 | Incoming inspection and asset tagging | (TC 12/13/2023 09:00:33 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:00:33 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069434 | 200317 | ACC | CMPLT | CMPLT | 2023-12-13 08:18 | 2023-12-13 09:01 | Incoming inspection and asset tagging | (TC 12/13/2023 09:01:13 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:01:13 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069435 | 200318 | ACC | CMPLT | CMPLT | 2023-12-13 08:19 | 2023-12-13 09:02 | Incoming inspection and asset tagging | (TC 12/13/2023 09:01:47 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:01:47 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069436 | 200319 | ACC | CMPLT | CMPLT | 2023-12-13 08:20 | 2023-12-13 09:02 | Incoming inspection and asset tagging | (TC 12/13/2023 09:02:19 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:02:19 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069437 | 200320 | ACC | CMPLT | CMPLT | 2023-12-13 08:21 | 2023-12-13 09:03 | Incoming inspection and asset tagging | (TC 12/13/2023 09:02:55 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:02:55 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069438 | 200321 | ACC | CMPLT | CMPLT | 2023-12-13 08:22 | 2023-12-13 09:04 | Incoming inspection and asset tagging | (TC 12/13/2023 09:03:27 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:03:27 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069439 | 200322 | ACC | CMPLT | CMPLT | 2023-12-13 08:23 | 2023-12-13 09:04 | Incoming inspection and asset tagging | (TC 12/13/2023 09:03:58 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:03:58 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069440 | 200323 | ACC | CMPLT | CMPLT | 2023-12-13 08:24 | 2023-12-13 09:05 | Incoming inspection and asset tagging | (TC 12/13/2023 09:05:09 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:05:09 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069441 | 200324 | ACC | CMPLT | CMPLT | 2023-12-13 08:24 | 2023-12-13 09:06 | Incoming inspection and asset tagging | (TC 12/13/2023 09:05:37 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:05:37 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069442 | 200325 | ACC | CMPLT | CMPLT | 2023-12-13 08:26 | 2023-12-13 09:06 | Incoming inspection and asset tagging | (TC 12/13/2023 09:06:13 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:06:13 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069443 | 200326 | ACC | CMPLT | CMPLT | 2023-12-13 08:27 | 2023-12-13 09:07 | Incoming inspection and asset tagging | (TC 12/13/2023 09:07:02 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:07:02 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069444 | 200327 | ACC | CMPLT | CMPLT | 2023-12-13 08:28 | 2023-12-13 09:08 | Incoming inspection and asset tagging | (TC 12/13/2023 09:07:37 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:07:37 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069445 | 200328 | ACC | CMPLT | CMPLT | 2023-12-13 08:29 | 2023-12-13 09:08 | Incoming inspection and asset tagging | (TC 12/13/2023 09:08:10 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:08:10 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069446 | 200329 | ACC | CMPLT | CMPLT | 2023-12-13 08:30 | 2023-12-13 09:09 | Incoming inspection and asset tagging | (TC 12/13/2023 09:08:39 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed. (TC 12/13/2023 09:08:39 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069447 | 200330 | ACC | CMPLT | CMPLT | 2023-12-13 08:31 | 2023-12-13 09:09 | Incoming inspection and asset tagging | (TC 12/13/2023 09:09:10 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069448 | 200331 | ACC | CMPLT | CMPLT | 2023-12-13 08:32 | 2023-12-13 09:10 | Incoming inspection and asset tagging | (TC 12/13/2023 09:09:42 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069449 | 200332 | ACC | CMPLT | CMPLT | 2023-12-13 08:34 | 2023-12-13 09:10 | Incoming inspection and asset tagging | (TC 12/13/2023 09:10:09 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069450 | 200333 | ACC | CMPLT | CMPLT | 2023-12-13 08:35 | 2023-12-13 09:11 | Incoming inspection and asset tagging | (TC 12/13/2023 09:10:42 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069451 | 200334 | ACC | CMPLT | CMPLT | 2023-12-13 08:36 | 2023-12-13 09:11 | Incoming inspection and asset tagging | (TC 12/13/2023 09:11:10 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069452 | 200335 | CM | CMPLT | CMPLT | 2023-12-13 08:37 | 2023-12-13 09:12 | Incoming inspection and asset tagging | (TC 12/13/2023 09:11:40 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1069453 | 200336 | ACC | CMPLT | CMPLT | 2023-12-13 08:38 | 2023-12-13 09:12 | Incoming inspection and asset tagging | (TC 12/13/2023 09:12:14 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
| 200337 | CM | CMPLT | CMPLT | 2023-12-13 09:58 | 2024-02-12 09:59 | ScopeGuides missing covers, base of arm also damaged | (TC 12/08/2023 09:58:29 AM, Cortez, John) Assessed issues on 4 ScopeGuides (TC 12/12/2023 10:00:06 AM, Cortez, John) Reached out to Rob from Olympus regarding ordering parts for ScopeGuide (TC 12/21/2023 08:45:12 AM, Cortez, John) Rob sent ordering info, will order 8 covers each Receiver dish cable connector bent in room 4, will monitor (TC 12/22/2023 10:41:45 AM, Cortez, John) Ordered 4x left, 4x right plastic covers (TC 01/04/2024 03:52:53 PM, Cortez, John) 2x covers arrived (TC 01/11/2024 09:44:40 AM, Cortez, John) 2x more covers arrived. Reached out to Olympus for update/eta on remaining parts. Remaining parts to arrive possibly on Jan 15 (TC 02/09/2024 09:58:14 AM, Cortez, John) Remaining covers arrived
1057054 | 200338 | CM | CMPLT | CMPLT | 2023-12-13 10:29 | 2023-12-13 08:36 | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | (TC 12/13/2023 10:28:49 AM, Whittington, Scott) MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT
1057055 | 200339 | CM | CMPLT | CMPLT | 2023-12-13 10:29 | 2023-12-13 08:37 | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | (TC 12/13/2023 10:29:28 AM, Whittington, Scott) MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT
1057056 | 200340 | CM | CMPLT | CMPLT | 2023-12-13 10:30 | 2023-12-13 08:37 | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | (TC 12/13/2023 10:30:04 AM, Whittington, Scott) MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT | MSH BI RAD REVIEW ROOM SECUREVIEW DIAGNOSTIC WORKSTATION PM BY HOLOGIC WITH UPS BATTERY REPLACEMENT
| 200341 | CM | CMPLT | CMPLT | 2023-12-13 10:58 | 2023-12-12 10:58 | meetings | meetings
| 200342 | CM | CMPLT | CMPLT | 2023-12-13 11:00 | 2023-12-13 11:00 | worked on RT equipment | worked on RT equipment
| 200343 | CM | CMPLT | CMPLT | 2023-12-13 11:01 | 2023-12-13 11:01 | worked on beds | worked on beds
1062247 | 200344 | CM | CMPLT | CMPLT | 2023-12-13 11:22 | 2023-12-13 11:24 | Reported that the bed wasn't working | (TC 12/13/2023 11:22:28 AM, Baghirzade, Nara) Nurse in charge of that room said she reset the bed but none of the motion controls are working. Found that all motion controls were locked. Unlocked the controls and let the nurse know. Tested all motion controls. Bed is fully functional.
1060888 | 200345 | CM | CMPLT | CMPLT | 2023-12-13 11:24 | 2023-12-13 11:25 | Battery replacement | (TC 12/13/2023 11:24:15 AM, Canjimavadivel, Kamallesh ) Replaced batteries and tested ok
| 200346 | CM | CMPLT | CMPLT | 2023-12-13 11:24 | 2023-12-15 14:47 | Placing bed locators on 11 North | (TC 12/13/2023 11:24:13 AM, Baghirzade, Nara) Placed bed locators on the rooms that had brackets installed. Waiting on the rest of the brackets to get installed. Placed locators up to rooms 22. (TC 12/14/2023 07:51:08 AM, Baghirzade, Nara) Placed the rest of the bed locators with exception to: 1116A 1116B Both bed locators labelled as "A"; additionally 1129A is missing a bed locator. Let Naz know and she is to contact Stryker to figure out which locator belongs to which bed/room. (TC 12/14/2023 11:26:34 AM, Baghirzade, Nara) Walkthrough with Naz to check which rooms have the BP and thermometer brackets installed as the sharps container was being placed too closely. (TC 12/15/2023 02:45:44 PM, Baghirzade, Nara) Naz requested to place both locators in room 1116A; she will connect with Stryker to have these bed locators reset as we are not sure which belong to Bay A and Bay B.
1059330 | 200347 | CM | CMPLT | CMPLT | 2023-12-13 11:35 | 2023-12-13 11:35 | PICIX Operating System Patch | (TC 12/12/2023 11:34:49 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059550 | 200348 | CM | CMPLT | CMPLT | 2023-12-13 11:36 | 2023-12-13 11:36 | PICIX Operating System Patch | (TC 12/12/2023 11:35:58 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059546 | 200349 | CM | CMPLT | CMPLT | 2023-12-13 11:37 | 2023-12-13 11:37 | PICIX Operating System Patch | (TC 12/12/2023 11:36:52 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1054599 | 200350 | CM | CMPLT | CMPLT | 2023-12-13 11:37 | 2023-12-13 11:39 | PICIX Operating System Patch | (TC 12/12/2023 11:37:38 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059541 | 200351 | CM | CMPLT | CMPLT | 2023-12-13 11:39 | 2023-12-13 11:41 | PICIX Operating System Patch | (TC 12/12/2023 11:40:14 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059554 | 200352 | CM | CMPLT | CMPLT | 2023-12-13 11:41 | 2023-12-13 11:42 | PICIX Operating System Patch | (TC 12/12/2023 11:41:14 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059556 | 200353 | CM | CMPLT | CMPLT | 2023-12-13 11:42 | 2023-12-13 11:42 | PICIX Operating System Patch | (TC 12/12/2023 11:41:57 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059542 | 200354 | CM | CMPLT | CMPLT | 2023-12-13 11:43 | 2023-12-13 11:43 | PICIX Operating System Patch | (TC 12/12/2023 11:42:48 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059548 | 200355 | CM | CMPLT | CMPLT | 2023-12-13 11:43 | 2023-12-13 11:44 | PICIX Operating System Patch | (TC 12/12/2023 11:43:26 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059557 | 200356 | CM | CMPLT | CMPLT | 2023-12-13 11:44 | 2023-12-13 11:44 | PICIX Operating System Patch | (TC 12/12/2023 11:44:08 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059547 | 200357 | CM | CMPLT | CMPLT | 2023-12-13 11:45 | 2023-12-13 11:45 | PICIX Operating System Patch | (TC 12/12/2023 11:44:47 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059543 | 200358 | CM | CMPLT | CMPLT | 2023-12-13 11:50 | 2023-12-13 11:50 | PICIX Operating System Patch | (TC 12/12/2023 11:49:46 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
1059552 | 200359 | CM | CMPLT | CMPLT | 2023-12-13 11:51 | 2024-02-20 11:45 | PICIX Operating System Patch | (TC 12/12/2023 11:50:41 AM, Baghirzade, Nara) PICIX Operating System Patch failed to install. Waiting on instructions from Kamran on how to proceed next. (TC 02/20/2024 11:44:45 AM, Baghirzade, Nara) Completed the operating system patch/update on 2S central in NICU. Confirmed on update history (local) that 2 updates were successfully completed. Confirmed that all monitors came back up and are displaying in segments.
1059553 | 200360 | CM | CMPLT | CMPLT | 2023-12-13 11:52 | 2023-12-13 11:52 | PICIX Operating System Patch | (TC 12/12/2023 11:51:37 AM, Baghirzade, Nara) PICIX Operating System Patch completed.
| 200361 | CM | CMPLT | CMPLT | 2023-12-13 11:56 | 2023-12-13 11:57 | Meeting with JDMI re: ophthalmic ultrasounds | (TC 12/13/2023 11:56:23 AM, Baghirzade, Nara) Meeting with JDMI re: ophthalmic ultrasounds
| 200362 | CM | CMPLT | CMPLT | 2023-12-13 13:46 | 2023-12-13 13:46 | biomed meeting | Biomed meeting
| 200363 | CM | CMPLT | CMPLT | 2023-12-13 14:06 | 2023-12-13 14:06 | huddle | (TC 12/13/2023 02:05:55 PM, Rhuland, Courtney) Weekly huddle
| 200364 | CM | ACTIV | ISSUE | 2023-12-13 14:17 |  | Ultrasound Devices Management | (TC 12/13/2023 02:16:55 PM, Chukri, Samir) Meeting with Adeel,Naz,Nara and JDMI team.
1058502 | 200365 | CM | CMPLT | NFF | 2023-12-13 14:32 | 2023-12-13 14:32 | No issue stated | (TC 12/13/2023 02:31:40 PM, Rhuland, Courtney) No issue written on tag Functional check - OK No fault found Returned to NICU
1054537 | 200366 | PM | CMPLT | CMPLT | 2023-12-13 15:56 | 2023-12-13 15:57 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/13/2023 03:56:34 PM, Cortez, John) Completed cleaning cycle
1069462 | 200367 | ACC | CMPLT | CMPLT | 2023-12-13 16:09 | 2023-12-13 16:13 | Incoming Inspection | (TC 12/08/2023 04:10:17 PM, Cortez, John) Visual inspection and electrical safety passed Installed in Bronch room
1069463 | 200368 | ACC | CMPLT | CMPLT | 2023-12-13 16:18 | 2023-12-13 16:20 | Incoming Inspection | (TC 12/12/2023 04:18:38 PM, Cortez, John) Visual inspection and electrical safety passed Installed in Endo room 4
1056403 | 200369 | CM | CMPLT | CMPLT | 2023-12-13 16:27 | 2024-01-10 10:41 | To be redeployed to Vaughan clinic | (TC 12/12/2023 04:26:49 PM, Cortez, John) Functionality test ok Changed to Sinai configuration Will reach out to Rita from Vaughan clinic to schedule delivery (TC 01/03/2024 10:20:59 AM, Cortez, John) Coordinated visit to clinic on Jan 9 with Rita (TC 01/09/2024 10:38:22 AM, Cortez, John) Went to Vaughan clinic to deliver monitor, and show nursing staff how to use it Unrelated, but also looked at broken exam light in room 21
1069481 | 200370 | ACC | CMPLT | CMPLT | 2023-12-14 10:27 | 2023-12-27 07:47 | Incoming inspection | (TC 12/15/2023 02:56:33 PM, Canjimavadivel, Kamallesh ) Unboxed and added to database (TC 12/21/2023 07:45:40 AM, Canjimavadivel, Kamallesh ) Dropped off with doctor Dushenski
1069487 | 200371 | ACC | CMPLT | CMPLT | 2023-12-14 10:37 | 2023-12-27 07:47 | Incoming inspection | (TC 12/14/2023 02:55:40 PM, Canjimavadivel, Kamallesh ) Unboxed and added to database (TC 12/21/2023 07:47:05 AM, Canjimavadivel, Kamallesh ) Dropped off with doctor Dushenski
1065160 | 200372 | CM | CMPLT | CMPLT | 2023-12-14 11:00 | 2023-12-14 11:03 | Not scanning barcodes | (TC 12/14/2023 10:59:51 AM, Rhuland, Courtney) Unable to identify clinician - user not found Downloaded configuration from working device Pushed config - issue resolved Functional check - OK
1054771 | 200373 | CM | CMPLT | CMPLT | 2023-12-14 11:09 | 2023-12-14 11:10 | Not identifying clinicians | (TC 12/14/2023 11:08:47 AM, Rhuland, Courtney) Updated location and IP address
1065161 | 200374 | CM | CMPLT | CMPLT | 2023-12-14 11:11 | 2023-12-14 11:11 | Not identifying clinician | (TC 12/14/2023 11:10:45 AM, Rhuland, Courtney) Updated location and IP address
| 200375 | CM | CMPLT | CMPLT | 2023-12-14 11:22 | 2023-12-14 11:26 | EMG meeting with vendor | (TC 12/14/2023 11:24:20 AM, Baghirzade, Nara) Meeting with Vendor, Andrew, and Nima
1069461 | 200376 | CM | CMPLT | CMPLT | 2023-12-14 11:23 | 2023-12-14 11:26 | incoming inspection | (TC 12/14/2023 11:23:21 AM, Rhuland, Courtney) Power cord leakage - 90.7 uA Resistance - 42mOhms Functional - OK (TC 12/15/2023 02:07:25 PM, Rhuland, Courtney) Worked with Rob from Olympus to install device in Room 3
1067341 | 200377 | CM | CMPLT | CMPLT | 2023-12-14 11:51 | 2023-12-14 11:51 | "EDI module error" | contacted Getinge. As per instructed we installed EDI module on a different Servo-N(1067342) and run pre-use check. Error is gone. Confirmed with Getinge and Rex.
| 200378 | CM | CMPLT | CMPLT | 2023-12-14 11:53 | 2023-12-13 11:53 | Mandatory training | Mandatory training
| 200379 | CM | CMPLT | CMPLT | 2023-12-14 12:13 | 2024-02-06 07:50 | Senior Biomedical Tech interview. | (TC 12/14/2023 12:12:53 PM, Chukri, Samir) Senior tech interview for two Tech (TC 12/19/2023 03:04:26 PM, Chukri, Samir) Inerview Rutvik new Sr.technologist with Nima.
| 200380 | CM | CMPLT | CMPLT | 2023-12-14 17:16 | 2023-12-14 17:16 | Site Coverage
65419 | 200381 | CM | CMPLT | CMPLT | 2023-12-15 08:14 | 2023-12-15 08:16 | intermittent pumping, overheating | (TC 12/15/2023 08:13:50 AM, Campbell, Connor) performed a functional check of alarms and pump, intermittent pumping was not an issue pump would run continuously the LCD temp display was erratic for the first 10-15 minutes of operation, but it stabilized after that and was accurate to the temp check adjusted the LCD, overtemp and circulating temp pots let it run for while no issues found returned to service
68034 | 200382 | CM | CMPLT | CMPLT | 2023-12-15 08:21 | 2023-12-15 12:56 | Broken parts | (TC 12/15/2023 12:55:39 PM, Rhuland, Courtney) Replaced broken side door latch Replaced 2 porthole doors Calibrated humidity Functional check - OK
1050068 | 200383 | CM | CMPLT | CMPLT | 2023-12-15 08:45 | 2023-12-15 09:02 | temp probe not working - retry measurement | (TC 12/15/2023 08:46:07 AM, Campbell, Connor) temp probe had wires exposed near the probe, was faulty replaced temp probe and cleaned contacts as it was dusty replaced battery door as it was missing tested functionality of NIBP, SpO2 and temp, all functional now cleaned device, the bottom anti-slip feet seem to be melting and leaving a plastic residue device now fully functional, returned to service
1055799 | 200384 | CM | CMPLT | CMPLT | 2023-12-15 09:04 | 2023-12-15 09:09 | "White vertical lines on ECG waves... 50J charge disappears" | (TC 12/14/2023 09:04:12 AM, Cortez, John) We noticed vertical lines when adjusting ECG leads. Leads were a bit worn out, replaced. Mentioned to nurse that after charging defib, it will timeout after a certain amount of time. Defib tested ok with analyzer.
1063000 | 200385 | PM | CMPLT | CMPLT | 2023-12-15 09:12 | 2023-12-15 09:13 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/14/2023 09:12:53 AM, Cortez, John) Completed cleaning cycle
63475 | 200386 | CM | CMPLT | CMPLT | 2023-12-15 09:32 | 2023-12-15 15:47 | Not Warm Enough | (TC 12/15/2023 03:46:36 PM, Reid, Emmerson ) called tech support who was able to help. Had to unlock display and adjust temperature. went back to check temps went up. woirking now.
| 200387 | CM | CMPLT | CMPLT | 2023-12-15 09:44 | 2023-12-15 09:45 | Equipment inventory - Dental | (TC 12/15/2023 09:44:04 AM, Rhuland, Courtney) Created asset tags for untagged equipment Tagged equipment in clinic
| 200388 | CM | CMPLT | CMPLT | 2023-12-15 09:44 | 2023-12-18 08:22 | Installing sphygs and thermometers on 11N | (TC 12/15/2023 09:44:10 AM, Baghirzade, Nara) Installed all sphygs and thermometers on 11N. Missing thermometer brackets: 1128A thermometer + bracket 1127A thermometer + bracket 1117B thermometer + bracket + mount BP Followed up with Naz regarding this. She is to provide extra thermometers and request the mounting of the brackets. Contacted Michael Badour for the disposal of all the boxes. (TC 12/15/2023 02:31:22 PM, Baghirzade, Nara) Met with Naz and Sarabeth to discuss how to mount the sphygh and thermometer brackets in room 1117 as the sharps container was mounted too closely to the bedside, and there is no space for the sphyghs and thermometer. Sarabeth decided to move the location of the brackets closer to the wall suction. (TC 12/18/2023 08:21:44 AM, Baghirzade, Nara) Completed the installation of sphygs and thermometers in the remaining rooms for 11N transition activities.
1056583 | 200389 | CM | CMPLT | CMPLT | 2023-12-15 10:34 | 2023-12-15 10:35 | Failed to check in for over a week | (TC 12/15/2023 10:33:57 AM, Baghirzade, Nara) Checked network settings and both access points were listed (mshmed, mshdata); did a check-in and it took a while for it to check in, but data transfer was completed successfully. Showed Courtney how to do the check-in.
| 200390 | CM | CMPLT | CMPLT | 2023-12-15 11:03 | 2023-12-15 11:04 | Full team meeting | (TC 12/15/2023 11:03:10 AM, Baghirzade, Nara) Full team meeting
| 200391 | CM | CMPLT | CMPLT | 2023-12-15 11:27 | 2023-12-15 11:27 | worked on patient beds | worked on patient beds.
| 200392 | CM | CMPLT | CMPLT | 2023-12-15 11:28 | 2023-12-15 11:28 | meeting | meeting
| 200393 | CM | CMPLT | CMPLT | 2023-12-15 11:29 | 2023-12-14 11:28 | biomed meeting | biomed meeting
1069503 | 200394 | CM | CMPLT | CMPLT | 2023-12-15 12:31 | 2023-12-15 12:32 | Incoming inspection | (TC 12/15/2023 12:31:40 PM, Rhuland, Courtney) Unboxed and assembled scale and height rod Functional check - OK Electrical safety - OK
| 200395 | RDS | CMPLT | CMPLT | 2023-12-15 14:33 | 2023-12-15 14:35 | Weekly Rounds: 18 South | (TC 12/15/2023 02:33:06 PM, Baghirzade, Nara) Weekly rounds in 18 South completed. See procedure for results.
| 200396 | RDS | CMPLT | CMPLT | 2023-12-15 14:36 | 2023-12-15 14:37 | Weekly Rounds: NICU | (TC 12/15/2023 02:36:08 PM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure for results.
| 200397 | RDS | CMPLT | CMPLT | 2023-12-15 14:38 | 2023-12-15 14:39 | Weekly Rounds: RESUS | (TC 12/15/2023 02:37:36 PM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure for results.
| 200398 | RDS | CMPLT | CMPLT | 2023-12-15 14:39 | 2023-12-15 14:40 | Weekly Rounds: 11 South | (TC 12/15/2023 02:38:59 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 200399 | RDS | CMPLT | CMPLT | 2023-12-15 14:41 | 2023-12-15 14:45 | Weekly Rounds: PACU POD A | (TC 12/15/2023 02:43:13 PM, Baghirzade, Nara) Weekly rounds in PACU POD A completed. See procedure for results.
| 200400 | PA | CMPLT | CMPLT | 2023-12-15 14:52 | 2023-12-18 12:54 | Ordering sterile water for NICU | (TC 12/15/2023 02:54:34 PM, Baghirzade, Nara) SC 1003131060 (TC 12/18/2023 12:53:37 PM, Baghirzade, Nara) Received water. 4150268738
1062999 | 200401 | PM | CMPLT | CMPLT | 2023-12-15 15:46 | 2023-12-15 15:47 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/15/2023 03:46:02 PM, Cortez, John) Completed cleaning cycle
1062002 | 200402 | CM | CMPLT | CMPLT | 2023-12-15 16:00 | 2023-12-15 16:01 | "can't move bed up or down" | (TC 12/15/2023 04:00:34 PM, Cortez, John) Power cycled bed and cleared error log Issue resolved
1067992 | 200403 | ACC | CMPLT | CMPLT | 2023-12-16 13:12 | 2024-02-26 13:23 | MSH 5 US CANON PRISM SN & MAC SWAP | (TC 12/16/2023 01:12:55 PM, Whittington, Scott) MSH 5 US CANON PRISM SN & MAC SWAP From: Whittington, Scott Sent: December 16, 2023 1:10 PM To: helpdesk.MSH@sinaihealth.ca Cc: Riaz, Mirza - Sinai Health System (Mirza.Riaz@sinaihealth.ca) <Mirza.Riaz@sinaihealth.ca> Subject: MSH NETWORK ACCESS REQUEST FOR CANON iSERIES ULTRASOUND - INC29383607 / INC30152619 Hi For the MSH 5th floor Canon Prism Ultrasound MAC A0:36:BC:74:64:CB REPLACES MAC A0:36:BC:74:63:C5 Please Whitelist MAC A0:36:BC:74:64:CB Thanks, Scott 416-402-2462 (TC 12/16/2023 01:12:55 PM, Whittington, Scott) MSH 5 US CANON PRISM SN & MAC SWAP From: Whittington, Scott Sent: December 16, 2023 1:10 PM To: helpdesk.MSH@sinaihealth.ca Cc: Riaz, Mirza - Sinai Health System (Mirza.Riaz@sinaihealth.ca) <Mirza.Riaz@sinaihealth.ca> Subject: MSH NETWORK ACCESS REQUEST FOR CANON iSERIES ULTRASOUND - INC29383607 / INC30152619 Hi For the MSH 5th floor Canon Prism Ultrasound MAC A0:36:BC:74:64:CB REPLACES MAC A0:36:BC:74:63:C5 Please Whitelist MAC A0:36:BC:74:64:CB Thanks, Scott 416-402-2462 | MSH 5 US CANON PRISM SN & MAC SWAP From: Whittington, Scott Sent: December 16, 2023 1:10 PM To: helpdesk.MSH@sinaihealth.ca Cc: Riaz, Mirza - Sinai Health System (Mirza.Riaz@sinaihealth.ca) <Mirza.Riaz@sinaihealth.ca> Subject: MSH NETWORK ACCESS REQUEST FOR CANON iSERIES ULTRASOUND - INC29383607 / INC30152619 Hi For the MSH 5th floor Canon Prism Ultrasound MAC A0:36:BC:74:64:CB REPLACES MAC A0:36:BC:74:63:C5 Please Whitelist MAC A0:36:BC:74:64:CB Thanks, Scott 416-402-2462
1069507 | 200404 | ACC | CMPLT | CMPLT | 2023-12-16 13:51 | 2023-12-28 08:29 | MSH 5 US PROBE PVI-475BX SN# FSC2322329 ACC | (TC 12/27/2023 08:29:34 AM, Whittington, Scott) MSH 5 US PROBE PVI-475BX SN# FSC2322329 ACC | MSH 5 US PROBE PVI-475BX SN# FSC2322329 ACC
1069508 | 200405 | ACC | CMPLT | CMPLT | 2023-12-16 13:54 | 2023-12-28 08:30 | MSH 5 US PROBE PLI-705BX ACC | (TC 12/27/2023 08:30:20 AM, Whittington, Scott) MSH 5 US PROBE PLI-705BX ACC | 5 PROBE PLI-705BX
1069509 | 200406 | ACC | CMPLT | CMPLT | 2023-12-16 13:56 | 2023-12-28 08:31 | MSH 5 US PROBE PLI-1205BX ACC | (TC 12/28/2023 08:30:58 AM, Whittington, Scott) MSH 5 US PROBE PLI-1205BX ACC (TC 12/27/2023 08:30:58 AM, Whittington, Scott) MSH 5 US PROBE PLI-1205BX ACC | MSH 5 US PROBE PLI-1205BX ACC
1069510 | 200407 | ACC | CMPLT | CMPLT | 2023-12-16 13:58 | 2023-12-28 08:32 | MSH 5 US PROBE PVT-781VTE/FS ACC | (TC 12/28/2023 08:31:59 AM, Whittington, Scott) MSH 5 US PROBE PVT-781VTE/FS ACC (TC 12/27/2023 08:31:59 AM, Whittington, Scott) MSH 5 US PROBE PVT-781VTE/FS ACC | MSH 5 US PROBE PVT-781VTE/FS
1058130 | 200408 | CM | CMPLT | CMPLT | 2023-12-18 08:00 | 2023-12-18 08:01 | Repurposed thermometer for 11N transition | (TC 12/18/2023 08:00:32 AM, Baghirzade, Nara) Picked up thermometer from Naz. Repurposed for 11N. Updated details on TMS and printed/placed new asset tag. | Repurposed thermometers from emerg for 11N
1054206 | 200409 | CM | CMPLT | CMPLT | 2023-12-18 08:04 | 2023-12-18 08:05 | Repurposed thermometer for 11N transition | (TC 12/18/2023 08:04:15 AM, Baghirzade, Nara) Picked up thermometer from Naz; repurposed old emerg thermometer for 11N. Updated details on TMS and printed/placed new asset tag. | Repurposed thermometer from emerg for 11N
1069522 | 200410 | CM | CMPLT | CMPLT | 2023-12-18 08:07 | 2023-12-18 08:07 | Incoming inspection and asset tagging | (TC 12/18/2023 08:06:51 AM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1064341 | 200411 | CM | CANCL | CANCL | 2023-12-18 08:46 |  | break stuck, can not move | Or 3 / 2 table ? OR#2 in L&D bed is not working. unable to even release breaks, this is a vital OR because it is used for emergency cases from L&D unit. please sent someone to fix asap. This bed was apparently fixed today because it was broke yesterday. But is now again not working. Nursing as troubleshooting several times in the last couple of hours
1059828 | 200412 | PM | CMPLT | CMPLT | 2023-12-18 08:49 | 2023-12-18 08:50 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/18/2023 08:49:09 AM, Campbell, Connor) Drained, disinfected, rinsed and refilled
69359 | 200413 | CM | CMPLT | CMPLT | 2023-12-18 09:05 | 2023-12-18 09:22 | Reported that the channel is not working when connecting to PCU | (TC 12/18/2023 09:19:46 AM, Baghirzade, Nara) Downloaded error log and noted error "240.4150" referencing high voltage on bottle-side sensor. When opening the pump, noted that the bezel was damaged. Replaced bottle-side sensor and bezel. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
| 200414 | RDS | CMPLT | CMPLT | 2023-12-18 09:08 | 2023-12-18 10:57 | Weekly checks in CCU | (TC 12/18/2023 10:55:46 AM, Scero, Tomlyn) weekly checks complete
| 200415 | RDS | CMPLT | CMPLT | 2023-12-18 09:20 | 2023-12-18 10:03 | Weekly rounds for ED
| 200416 | RDS | CMPLT | CMPLT | 2023-12-18 09:20 | 2023-12-22 07:41 | Weekly rounds for ICU
1050084 | 200417 | CM | CMPLT | CMPLT | 2023-12-18 10:47 | 2023-12-18 10:48 | replacing battery | (TC 12/18/2023 10:46:47 AM, Campbell, Connor) replaced UPS battery
| 200418 | CM | ACTIV | ISSUE | 2023-12-18 10:57 |  | test/not real work order (maisha)
1054540 | 200419 | CM | CMPLT | CMPLT | 2023-12-18 11:14 | 2023-12-18 11:14 | Replaced ECG keyboard cover | Replaced ECG keyboard cover
1064641 | 200420 | CM | CMPLT | CMPLT | 2023-12-18 11:32 | 2023-12-19 11:00 | OR#2 in L&D bed is not working. unable to even release breaks | OR#2 in L&D bed is not working. unable to even release breaks. Check with L&D. They do not know any issue. Arranged a service visit with vendor Medaris.
500198 | 200421 | RQ | CMPLT | CMPLT | 2023-12-18 12:07 | 2024-01-11 14:19 | The Entegra Workstation not working |  | The Entegra workstation is not on. Just see a black screen, Even though the tower is powered on. We did a bone scan on VG and was not able to transfer patient onto Entegra to create report. Both terminals were shut down.
1063746 | 200422 | PM | CMPLT | CMPLT | 2023-12-18 12:20 | 2023-12-18 12:20 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/18/2023 12:20:02 PM, Campbell, Connor) Drained, disinfected, rinsed and refilled
1055029 | 200423 | CM | CMPLT | CMPLT | 2023-12-18 12:39 | 2023-12-18 12:41 | Replaced battery | (TC 12/18/2023 12:38:35 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1055032 | 200424 | CM | CMPLT | CMPLT | 2023-12-18 12:41 | 2023-12-18 12:42 | Replaced battery | (TC 12/18/2023 12:41:30 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1055033 | 200425 | CM | CMPLT | CMPLT | 2023-12-18 12:43 | 2023-12-18 12:43 | Replaced battery | (TC 12/18/2023 12:42:40 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1055973 | 200426 | CM | CMPLT | CMPLT | 2023-12-18 12:43 | 2023-12-18 12:44 | Replaced battery | (TC 12/18/2023 12:43:22 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1055975 | 200427 | CM | CMPLT | CMPLT | 2023-12-18 12:44 | 2023-12-18 12:44 | Replaced battery | (TC 12/18/2023 12:44:07 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1055977 | 200428 | CM | CMPLT | CMPLT | 2023-12-18 12:45 | 2023-12-18 12:45 | Replaced battery | (TC 12/18/2023 12:44:49 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1056493 | 200429 | CM | CMPLT | CMPLT | 2023-12-18 12:45 | 2023-12-18 12:46 | Replaced battery | (TC 12/18/2023 12:45:34 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1056496 | 200430 | CM | CMPLT | CMPLT | 2023-12-18 12:46 | 2023-12-18 12:47 | Replaced battery | (TC 12/18/2023 12:46:15 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1056497 | 200431 | CM | CMPLT | CMPLT | 2023-12-18 12:48 | 2023-12-18 12:48 | Replaced battery | (TC 12/18/2023 12:47:58 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1056498 | 200432 | CM | CMPLT | CMPLT | 2023-12-18 12:49 | 2023-12-18 12:49 | Replaced battery | (TC 12/18/2023 12:48:35 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1056499 | 200433 | CM | CMPLT | CMPLT | 2023-12-18 12:49 | 2023-12-18 12:49 | Replaced battery | (TC 12/18/2023 12:49:08 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
63724 | 200434 | CM | CMPLT | CMPLT | 2023-12-18 12:58 | 2023-12-19 07:43 | Keyboard cover broken | (TC 12/18/2023 01:08:33 PM, Reid, Emmerson ) Replaced the keyboard cover and running a conditioning cycle on the two batteries. (TC 12/19/2023 07:42:02 AM, Reid, Emmerson ) Batteries cycled and are in good condition. Working now.
1055979 | 200435 | CM | CMPLT | CMPLT | 2023-12-18 13:33 | 2023-12-18 13:34 | Replaced battery | (TC 12/18/2023 01:34:11 PM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1066454 | 200436 | FC | CMPLT | CMPLT | 2023-12-18 13:52 | 2023-12-18 13:53 | QC checklist | (TC 12/18/2023 01:52:24 PM, Baghirzade, Nara) QC checklist completed. Dropped off at 11 North
1050702 | 200437 | CM | CMPLT | CMPLT | 2023-12-18 14:14 | 2023-12-18 14:15 | End of support - device has been replaced | (TC 12/18/2023 02:13:50 PM, Baghirzade, Nara) Device is end of support, no parts/repair available. Confirmed with Samir to retire it as 11N has purchased a new fleet of thermometers and has existing Genius 3 thermometers.
62518 | 200438 | CM | CMPLT | CMPLT | 2023-12-18 14:16 | 2023-12-18 14:16 | End of support - retiring device | (TC 12/18/2023 02:15:42 PM, Baghirzade, Nara) Device is end of support, no parts/repair available. Confirmed with Samir to retire it as 11N has purchased a new fleet of thermometers and has existing Genius 3 thermometers.
1062149 | 200439 | CM | CMPLT | CMPLT | 2023-12-18 14:33 | 2023-12-19 11:07 | Not working | (TC 12/18/2023 02:33:51 PM, Canjimavadivel, Kamallesh ) Thermometer intermittently not detecting when probe cover is installed on tip. Opened up device and found evidence of ingress and corrosion on some components on board. Attempted to clean corrosion and test again. Still see intermittent issue. (TC 12/19/2023 11:06:00 AM, Canjimavadivel, Kamallesh ) Tried to clean using contact cleaner and still see issue. Cannot be repaired and warranty has expired. Device to be retired.
1069461 | 200440 | CM | CMPLT | CMPLT | 2023-12-18 14:54 | 2023-12-18 14:55 | Not capturing image on PACSGear | (TC 12/18/2023 02:54:08 PM, Rhuland, Courtney) Worked with Rob from olympus to solve the issue Swapped out connection with 3.5mm audio cable Hand controller on scope now triggering image capture on Hyland PACSGear
1063331 | 200441 | CM | CMPLT | CMPLT | 2023-12-18 16:33 | 2023-12-18 16:33 | MSH OPG 3 SPP RESEARCH VOLUSON 10E NO MWL | (TC 12/18/2023 04:32:49 PM, Whittington, Scott) MSH OPG 3 SPP RESEARCH VOLUSON 10E NO MWL FOUND TIME SOWAS WRONG, ALTHOUG IN MORNING OK RESET OK | MSH OPG 3 SPP RESEARCH VOLUSON 10E NO MWL
501106 | 200442 | CM | CMPLT | CMPLT | 2023-12-18 16:35 | 2024-02-26 13:22 | MSH 5 US IU22 NOT BOOTING PAST PAGE 3 | (TC 12/18/2023 04:34:46 PM, Whittington, Scott) MSH 5 US IU22 NOT BOOTING PAST PAGE 3 | MSH 5 US IU22 NOT BOOTING PAST PAGE 3
68988 | 200443 | CM | CMPLT | CMPLT | 2023-12-19 08:01 | 2023-12-19 09:53 | bent plug | (TC 12/19/2023 08:00:53 AM, Campbell, Connor) used tools to bend the plug back into the proper orientation, it was a recently replaced plug tested electrical safety, passed it with very low numbers, 50 mOhms resistance and 78uA leakage current performed a check in, passed returned to service
| 200444 | CM | CMPLT | CMPLT | 2023-12-19 08:07 | 2023-12-21 08:10 | Zeus SW upgrade to 2.03 | (TC 12/18/2023 08:08:02 AM, Scero, Tomlyn) Helped Draeger locate all Zeus in the hospital and all had SW 2.03 installed. Service reports saved on T drive
68963 | 200445 | CM | CMPLT | CMPLT | 2023-12-19 08:56 | 2023-12-19 09:01 | Damaged power cord | (TC 12/19/2023 08:55:55 AM, Baghirzade, Nara) Power cord was damaged; PCU was also missing the power cord strap/strap holder/screw. Replaced power cord and added the power cord strap/strap holder/screw. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 40 mohms Leakage current: 77.2 uA (TC 12/19/2023 08:55:55 AM, Baghirzade, Nara) Power cord was damaged; PCU was also missing the power cord strap/strap holder/screw. Replaced power cord and added the power cord strap/strap holder/screw. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 40 mohms Leakage current: 77.2 uA
1060149 | 200446 | CM | CMPLT | CMPLT | 2023-12-19 09:11 | 2023-12-19 09:17 | Tech said that the battery wouldn't charge | (TC 12/19/2023 09:15:32 AM, Scero, Tomlyn) Tech brought the unit down and the red bar was present. I plugged it in for the night. When I came in today the red bar was still there but I couldn't press on the battery symbol. I restarted the unit and when it turned back on the battery icon was full and I could see the batteries were fully charged. Possibly a glitch in the SW.
1059555 | 200447 | CM | CMPLT | CMPLT | 2023-12-19 09:57 | 2023-12-19 09:57 | Operating system patch | (TC 12/19/2023 09:57:01 AM, Baghirzade, Nara) Completed the operating system patch/update.
52230 | 200448 | CM | CMPLT | CMPLT | 2023-12-19 09:58 | 2023-12-19 10:00 | Loud fan | (TC 12/19/2023 09:58:32 AM, Baghirzade, Nara) A boot up, the fan was loud and the fan itself was wobbling. Replaced they grey insert that the fan sits on. Powered it on and the fan sits tightly on the grey insert. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.1-75.2% RH. Incubator is fully functional. | End user reported loud fan when switching on the incubator
1059553 | 200449 | CM | CMPLT | CMPLT | 2023-12-19 10:07 | 2023-12-19 10:08 | Display is off | (TC 12/19/2023 10:07:27 AM, Baghirzade, Nara) Display cable was kicked and as a result, the display was down. Reconnected the display cable back to the adapter. 3S is back up and running. Followed up with May. | May reported that the Zone C central was down
| 200450 | CM | CMPLT | CMPLT | 2023-12-19 10:11 | 2023-12-19 10:17 | ECG and RESP not appearing on monitor in RESUS | (TC 12/19/2023 10:11:22 AM, Rhuland, Courtney) Performed functional check on all monitors One monitor (Bay E) had ECG and resp rate turned off Turned these parameters back on All functional checks - OK (TC 12/19/2023 12:55:06 PM, Baghirzade, Nara) Worked with Courtney on the BME incident report. Followed up with May and Janet regarding our findings and recommendations. (TC 12/19/2023 02:50:09 PM, Baghirzade, Nara) Created a troubleshooting/tip sheet for NICU RESUS staff, Janet, and education staff. Tip sheet includes the 2 ways that the staff can revert the MX800 to default settings so that all parameters are being displaced.
| 200451 | RDS | CMPLT | CMPLT | 2023-12-19 10:56 | 2023-12-18 10:56 | weekly check for 14S and 16N | weekly check for 14S and 16N
| 200452 | CM | CMPLT | CMPLT | 2023-12-19 10:59 | 2023-12-19 10:59 | worked with Samir | worked with Samir
1062151 | 200453 | CM | CMPLT | CMPLT | 2023-12-19 11:27 | 2023-12-19 11:28 | Not turning on | (TC 12/19/2023 11:27:08 AM, Canjimavadivel, Kamallesh ) Replaced batteries Verified accuray with tester
1056125 | 200454 | PM | CMPLT | CMPLT | 2023-12-19 11:59 | 2023-12-19 12:00 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 12/19/2023 11:59:29 AM, Campbell, Connor) Drained, disinfected, rinsed and refilled
1057055 | 200455 | CM | CMPLT | CMPLT | 2023-12-19 12:46 | 2024-02-26 13:22 | MSH BI RAD REVIEW WS#2 QWERTY KYBD FAULTY KEYS LOOSE | (TC 12/19/2023 12:46:41 PM, Whittington, Scott) MSH BI RAD REVIEW WS#2 QWERTY KYBD FAULTY KEYS LOOSE (TC 12/29/2023 12:48:08 PM, Whittington, Scott) From: Larocque, Ryan <Ryan.Larocque@hologic.com> Sent: December 29, 2023 12:04 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Chan-Dinevski, Ivy - Sinai Health System <Ivy.Chan-Dinevski@sinaihealth.ca>; Ghai, Sandeep - Women's College Hospital <Sandeep.Ghai@wchospital.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] Keyboard install Secureview @ Mount Sinai It will be shipped to your attention. Thanks, Ryan Get Outlook for iOS ________________________________________ From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Friday, December 29, 2023 12:02:31 PM To: Larocque, Ryan <Ryan.Larocque@hologic.com> Cc: Chan-Dinevski, Ivy - Sinai Health System <Ivy.Chan-Dinevski@sinaihealth.ca>; Ghai, Sandeep - Women's College Hospital <Sandeep.Ghai@wchospital.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] Keyboard install Secureview @ Mount Sinai Hi Ryan Thanks! Is it being shipped to my attention, or directly to BI department? Scott 416-402-2462 Get Outlook for iOS [aka.?ms] From: Larocque, Ryan <Ryan.?Larocque@?hologic.?com> Sent: Friday, December 29, 2023 10:?18 AM To: Whittington, Hi Ryan Thanks! Is it being shipped to my attention, or directly to BI department? Scott 416-402-2462 Get Outlook for iOS [aka.ms] ________________________________________ From: Larocque, Ryan <Ryan.Larocque@hologic.com> Sent: Friday, December 29, 2023 10:18 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] Keyboard install Secureview @ Mount Sinai Hey Scott, I just order a keyboard for the secureview (9780319D0331 SVDX-01200) at Mount Sinai hospital. When it arrives would you be able to install it? Thanks, Ryan (TC 01/09/2024 04:19:48 PM, Whittington, Scott) MSH BI RAD REVIEW WS#2 QWERTY KYBD FAULTY KEYS LOOSE KYBD RECEIVED (TC 01/09/2024 05:19:00 PM, Whittington, Scott) MSH BI RAD REVIEW WS#2 QWERTY KYBD FAULTY KEYS LOOSE KYBD RECEIVED | MSH BI RAD REVIEW WS#2 QWERTY KYBD FAULTY KEYS LOOSE REF# 0142 0340
65901 | 200456 | CM | CMPLT | CMPLT | 2023-12-19 13:00 | 2023-12-20 08:57 | Channel error | (TC 12/19/2023 01:00:34 PM, Rhuland, Courtney) Plunger claws would not close properly Removed claws and cleaned off residue Pump failed plunger head sensor test Reopened device and cleaned plunger head sensor and replaced force sensor seal (TC 12/20/2023 08:56:13 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1061670 | 200457 | CM | CMPLT | NFF | 2023-12-19 13:05 | 2023-12-19 13:30 | "pumps up mattress slower than before" | (TC 12/18/2023 01:05:36 PM, Cortez, John) Anna Kha from wound care had a question about the TruTurn feature on the mattress pumps. Went to room 43B to see how feature works Looks like the user needs to hold down increase/decrease button to adjust turn to desired angle. Mentioned to nurse and reached out to Anna. Will meet with Anna and Vincy tomorrow
1054958 | 200458 | CM | CMPLT | CMPLT | 2023-12-19 13:11 | 2023-12-19 13:24 | Reading low | (TC 12/19/2023 01:23:24 PM, Canjimavadivel, Kamallesh ) Performed calibration with tester, passed after 1 attempt
| 200459 | CM | CMPLT | CMPLT | 2023-12-19 13:12 | 2023-12-19 13:13 | Triage calling that three Fetal monitors not connecting to cerner | (TC 12/19/2023 01:11:49 PM, Chukri, Samir) Went with Olcay to investigate the issue in 15floor, no problem found on our fetal monitors all need IT attantion. Nurse will contact the it to solve this problem.
32510 | 200460 | CM | CMPLT | CMPLT | 2023-12-19 13:15 | 2023-12-19 13:16 | MSH DR1 GUARD PROTECT | (TC 12/19/2023 01:15:16 PM, Whittington, Scott) MSH DR1 GUARD PROTECT (TC 12/19/2023 01:15:53 PM, Dolinar, John) MSH DR1 GUARD PROTECT | MSH DR1 GUARD PROTECT
1066462 | 200461 | FC | CMPLT | CMPLT | 2023-12-19 13:16 | 2023-12-19 13:16 | QC Checklist | (TC 12/19/2023 01:15:52 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South
1062152 | 200462 | CM | CMPLT | CMPLT | 2023-12-19 13:25 | 2023-12-19 14:08 | Temp reading low | (TC 12/19/2023 02:06:01 PM, Canjimavadivel, Kamallesh ) Passed calibration after 2nd attempt.
| 200463 | CM | CMPLT | CMPLT | 2023-12-19 13:32 | 2023-12-19 13:32 | Mandatory training | mandatory training
1054308 | 200464 | CM | CMPLT | CMPLT | 2023-12-19 14:52 | 2023-12-19 14:53 | Not recognizing clinician IDs | (TC 12/19/2023 02:52:19 PM, Rhuland, Courtney) Updated IP address and location Associates clinician
1059545 | 200465 | CM | CMPLT | CMPLT | 2023-12-19 14:54 | 2024-02-20 11:42 | Operating system patch/update | (TC 12/19/2023 02:53:52 PM, Baghirzade, Nara) Operating system patch failed to install. Checked the update history and relayed the information to Kamran. Ran the commands for install and download however it did not work. Rebooted the PIICIX but did not resolve the issue either. Kamran is looking into the solution for this problem. (TC 02/20/2024 11:41:10 AM, Baghirzade, Nara) Update completed as a result of a power cycle when the PICIX PM was performed.
| 200466 | CM | ACTIV | ISSUE | 2023-12-19 15:01 |  | Working in incident reports | (TC 12/19/2023 03:01:19 PM, Chukri, Samir) working with Nara and courtney to complete the incident report # 67574 (TC 12/19/2023 03:02:16 PM, Chukri, Samir) working with Emmerson to complete the incident report # 67315 O.R (TC 12/20/2023 03:10:11 PM, Chukri, Samir) completed the incident report 67315 and attached in SAFER (TC 12/20/2023 03:11:02 PM, Chukri, Samir) Review and complete incident in SAFER. (TC 03/21/2024 02:49:17 PM, Chukri, Samir) Review and completed incident report 69715 O.R Anesthesia ventilator. (TC 04/29/2024 02:37:32 PM, Chukri, Samir) working on incident report and follow up.
| 200467 | CM | CMPLT | CMPLT | 2023-12-20 07:09 | 2023-12-20 07:10 | NICU Huddle
1059828 | 200468 | CM | CMPLT | CMPLT | 2023-12-20 07:55 | 2023-12-20 07:56 | "needs to be cleaned" (they emptied water) | (TC 12/20/2023 07:55:54 AM, Campbell, Connor) refilled and returned to service
1055034 | 200469 | CM | CMPLT | CMPLT | 2023-12-20 08:35 | 2023-12-20 08:35 | Replaced battery | (TC 12/20/2023 08:34:51 AM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1056495 | 200470 | CM | CMPLT | CMPLT | 2023-12-20 08:36 | 2023-12-20 08:36 | Replaced battery | (TC 12/20/2023 08:35:50 AM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1059588 | 200471 | CM | CMPLT | CMPLT | 2023-12-20 08:37 | 2023-12-20 08:37 | battery replacement | (TC 12/20/2023 08:36:41 AM, Campbell, Connor) calibrated (passed) replaced battery
1056492 | 200472 | CM | CMPLT | CMPLT | 2023-12-20 08:37 | 2023-12-20 08:37 | Replaced battery | (TC 12/20/2023 08:36:44 AM, Baghirzade, Nara) Replaced battery. Completed battery upgrade.
1059574 | 200473 | CM | CMPLT | CMPLT | 2023-12-20 08:38 | 2023-12-20 08:38 | battery replacement | (TC 12/20/2023 08:37:41 AM, Campbell, Connor) calibrated (passed) replaced battery
1059586 | 200474 | CM | CMPLT | CMPLT | 2023-12-20 08:39 | 2023-12-22 09:49 | reads temp high- needs calibration | (TC 12/20/2023 08:38:51 AM, Campbell, Connor) calibrated (passed)
1056437 | 200475 | CM | CMPLT | CMPLT | 2023-12-20 09:16 | 2023-12-20 09:18 | Temperature module error | (TC 12/20/2023 09:16:31 AM, Baghirzade, Nara) Re-seated the probe well. Error still persists. Probe well is defective. Replaced the probe well and error was cleared. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 200476 | CM | CMPLT | CMPLT | 2023-12-20 10:18 | 2023-12-20 10:18 | Meeting | (TC 12/20/2023 10:17:56 AM, Rhuland, Courtney) CE meeting in biomed
68684 | 200477 | CM | CMPLT | CMPLT | 2023-12-20 10:38 | 2024-01-04 13:01 | Pushed BH network configuration and data set | (TC 12/20/2023 10:38:24 AM, Baghirzade, Nara) Set BH as the active package on ASM. Pushed network configuration using ASM. Pushed data set "BH DS V.1 01_06_18.gre". Jessica to confirm if dataset is showing up when she is at BH. (TC 01/04/2024 12:59:33 PM, Baghirzade, Nara) Confirmed with Jessica that BH network is associated and the dataset is displayed at the top of the PCU screen. Confirmed that user is able to access the drug library. | Jessica mentioned that at BH the dataset is not showing up. Requested to have the network and dataset pushed and she will check again at BH.
1069549 | 200478 | CM | CMPLT | CMPLT | 2023-12-20 10:56 | 2024-02-26 13:20 | MSH CEOU IU22 C5-1 PROBE ACCEPTANCE C5-1 F09ZWR | (TC 12/20/2023 10:56:15 AM, Whittington, Scott) MSH CEOU IU22 C5-1 PROBE ACCEPTANCE | MSH CEOU IU22 C5-1 PROBE ACCEPTANCE C5-1 F09ZWR
1046203 | 200479 | CM | CMPLT | CMPLT | 2023-12-20 11:53 | 2024-01-02 14:55 | CW artifact/noise | (TC 12/20/2023 11:53:19 AM, Fan, Tak) Confirmed issue. Happens when freezing and unfreezing. Happens on all sector probes. Does not happen on Pedoff probe. Reseated boards, issue persists. Ordered CRX board. (TC 12/21/2023 03:40:36 PM, Fan, Tak) Replaced CRX board but system failed to start up. Swapped back, no problem. Contacted GE. (TC 12/29/2023 02:53:51 PM, Fan, Tak) Installed new CRX board, same issue as before. Contacted GE. Instructions missed that a SW application reload is necessary for replacement of CRX board. Reloaded SW, Issue appears to be gone. | Confirmed issue. Happens when freezing and unfreezing. Happens on all sector probes. Does not happen on Pedoff probe. Reseated boards, issue persists. Ordered CRX board.
1062282 | 200480 | CM | CMPLT | CMPLT | 2023-12-20 13:22 | 2023-12-20 13:25 | Error "obstructed head portion" | (TC 12/20/2023 01:22:07 PM, Baghirzade, Nara) Head portion was not moving down. Noticed that under the "lock" option, the head rest was locked to 30 degrees. Unlocked all motion controls and the headrest portion is working. Checked all the motion controls including patient controls; bed is fully functional. Followed up with Sarabeth regarding this.
68869 | 200481 | CM | CMPLT | CMPLT | 2023-12-20 13:58 | 2024-01-04 13:01 | Pushed BH network configuration and data set | (TC 12/20/2023 01:58:22 PM, Baghirzade, Nara) Set BH package as active on ASM. Transferred network configuration. Transferred data set. Returned to Jessica to take to BH and check if the BH dataset is showing up when she's on site. (TC 01/04/2024 01:00:39 PM, Baghirzade, Nara) Confirmed with Jessica that BH network is associated and the dataset is displayed at the top of the PCU screen. Confirmed that user is able to access the drug library. | Jessica mentioned that at BH the dataset is not showing up. Requested to have the network and dataset pushed and she will check again at BH.
| 200482 | CM | CMPLT | CMPLT | 2023-12-20 14:00 | 2023-12-20 14:00 | Capital/Fleet management meeting | (TC 12/20/2023 01:59:47 PM, Baghirzade, Nara) Capital/Fleet management meeting.
| 200483 | CM | CMPLT | CMPLT | 2023-12-20 14:09 | 2024-01-18 10:11 | PACU Borrowing Simlators | (TC 12/20/2023 02:09:08 PM, Reid, Emmerson ) Gave simulators to Eden for their training tomorrow. Should be done around noon. (TC 01/11/2024 02:34:44 PM, Reid, Emmerson ) Dropped off simulators to PACU
1055603 | 200484 | CM | CMPLT | CMPLT | 2023-12-20 14:10 | 2023-12-20 14:10 | Replaced battery | (TC 12/20/2023 02:09:54 PM, Baghirzade, Nara) Replaced battery. Battery upgrade completed.
1055578 | 200485 | CM | CMPLT | CMPLT | 2023-12-20 14:11 | 2023-12-20 14:16 | Battery depleted | (TC 12/20/2023 02:15:27 PM, Baghirzade, Nara) End user was referring to the radical battery however the root battery was full. Replaced the radical battery. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional. | End user was referring to the radical battery however the root battery was full.
1064753 | 200486 | CM | CMPLT | CMPLT | 2023-12-20 14:30 | 2023-12-20 14:31 | For Safer# 67315 | (TC 12/20/2023 02:30:27 PM, Reid, Emmerson ) Refer to WO 200291
66468 | 200487 | CM | CMPLT | CMPLT | 2023-12-21 07:54 | 2023-12-22 10:42 | Nebulizer test failed on device check | (TC 12/22/2023 10:38:14 AM, Canjimavadivel, Kamallesh ) Logs identified potential issue with pressure regulators. Connected to HIT software and tested pressure sensors, ok. Nebulizer test pass. Checked values of pressure after regulator, slightly higher than accepted range. Adjusted both air and O2 regulators. Values within passing range. If problem returns later, regulators should be replaced.
65262 | 200488 | CM | CMPLT | CMPLT | 2023-12-21 07:54 | 2023-12-21 15:07 | Nebulizer test failed on device check | (TC 12/20/2023 03:30:00 PM, Canjimavadivel, Kamallesh ) Downloaded logs and found reported error. Follow instructions to perform self-tests. Multiple test values not within passing range. (TC 12/21/2023 03:02:08 PM, Canjimavadivel, Kamallesh ) Performed additional self-tests and isolated components. Issue was related to pressure regulator. Verified S2 sensor working correctly. Adjusted air and O2 regulators, decreased pressure in both. Values within range. Closed device and performed device check. Sent copy of logs to draeger for second opinion. (TC 12/21/2023 09:00:00 AM, Canjimavadivel, Kamallesh ) Going through logs and checking pneumatic diagrams
69178 | 200489 | CM | CMPLT | NFF | 2023-12-21 07:55 | 2023-12-21 07:56 | "powers on then shuts off immediately" | (TC 12/21/2023 07:54:56 AM, Campbell, Connor) peformed a check in, passed performed a test infusion no fault found returned to service
69465 | 200490 | CM | CMPLT | CMPLT | 2023-12-21 07:58 | 2023-12-21 07:59 | channel error | (TC 12/21/2023 07:57:51 AM, Campbell, Connor) upper pressure sensor malfunction replaced the upper pressure sensor PM performed immediately after pump is now functional
1053391 | 200491 | CM | CMPLT | CMPLT | 2023-12-21 14:38 | 2023-12-21 14:40 | Check after/before patient use | (TC 12/21/2023 02:39:12 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
1062031 | 200492 | CM | CMPLT | CMPLT | 2023-12-21 15:36 | 2023-12-21 15:39 | "settings not saving" | (TC 12/21/2023 03:35:52 PM, Cortez, John) Unable to replicate issue, tested scale, motion functions, and alarms, Saved weight remains after shutdown Informed ward clerk
1052940 | 200493 | CM | CMPLT | CMPLT | 2023-12-21 15:43 | 2024-01-02 14:54 | "A setup problem has occured. Please reselect probe to continue" | (TC 12/21/2023 03:43:17 PM, Fan, Tak) Confirmed issue. Troubleshot. Other probe OK. Spare probe OK. Ports aren't the problem. Contacted GE for replacement. (TC 12/28/2023 01:11:47 PM, Fan, Tak) No contract apparently. Following up. (TC 01/02/2024 02:52:43 PM, Fan, Tak) Sorted out issue with Tony. 416596TEE under contract but has no probe assigned to it. Ordered under that site ID. (Other side has no contract, will follow up with GE). (TC 01/08/2024 02:26:07 PM, Fan, Tak) 219497 replaced by NO224713 on Jan 08, 2023
67628 | 200494 | CM | CMPLT | CMPLT | 2023-12-22 07:55 | 2023-12-22 08:20 | not working | (TC 12/22/2023 07:55:30 AM, Campbell, Connor) attecmpted to change batteries attempted to calibrate themomter has an error where battery icon flashes and a square box with an X flashes, this is not fixable needs to be retired
1056997 | 200495 | CM | CMPLT | CMPLT | 2023-12-22 07:59 | 2023-12-27 09:58 | SpO2 not working | (TC 12/19/2023 10:30:00 AM, Canjimavadivel, Kamallesh ) Confirmed issue and swapped out X3 with spare. Found evidence of corrosion on MSL connector port. Cleaned with contact cleaner. (TC 12/27/2023 09:43:25 AM, Canjimavadivel, Kamallesh ) Tested with test sensor and cables, SpO2 function working. Added ED config using supporttool
1066679 | 200496 | CM | CMPLT | CMPLT | 2023-12-22 08:21 | 2023-12-22 08:22 | "flow error" | (TC 12/22/2023 08:20:54 AM, Campbell, Connor) primed pump for 15 minutes performed a rate check over 30 min pump rate was within tolerance, delivered 40.804mL in 30 min at 75mL/hour no fault found returned to service
1066891 | 200497 | CM | CMPLT | NFF | 2023-12-22 09:44 | 2023-12-22 09:49 | sent down with no yellow tag | (TC 12/22/2023 09:44:24 AM, Campbell, Connor) primed pump for 15 minutes performed a rate check over 30 minutes, delivered 41.255mL, this is just within acceptable range no fault found returned to service
1064310 | 200498 | CM | CMPLT | CMPLT | 2023-12-22 09:46 | 2023-12-22 09:49 | 6 year PM done | (TC 12/22/2023 09:48:39 AM, Scero, Tomlyn) 6 year PM parts replaced, tests complete and passed. Self test passed. leakage total <100 ml/min in + ex 12 ml/min
1064797 | 200499 | CM | CMPLT | CMPLT | 2023-12-22 09:46 | 2023-12-22 09:48 | 6 year PM done | (TC 12/22/2023 09:47:08 AM, Scero, Tomlyn) 6 year parts replaced, tests complete and passed. Self test passed. leakage total <100 ml/min in + ex 11 ml/min
1054775 | 200500 | CM | CMPLT | CMPLT | 2023-12-22 09:57 | 2023-12-22 09:59 | Temperature not working | (TC 12/22/2023 09:57:57 AM, Ferreira, Qwynn) Tested the probe was not working correctly Had a spare probe here no charge. Working as normal now
1062139 | 200501 | CM | CMPLT | CMPLT | 2023-12-22 10:04 | 2024-01-04 11:29 | Head not lifting | (TC 12/22/2023 10:04:16 AM, Ferreira, Qwynn) Spoke with clerk as no message was left Reset the bed Tested functions - head motor not working Check underside of bed but will need to connect with stryker (TC 01/04/2024 11:27:53 AM, Ferreira, Qwynn) Bed was repaired and notice was left. All functions were tested before hand off
| 200502 | RDS | CMPLT | CMPLT | 2023-12-22 10:11 | 2023-12-22 10:14 | Weekly Rounds - 12S | (TC 12/22/2023 10:12:21 AM, Ferreira, Qwynn) Rounds completed - good unit usually calls for any problems
| 200503 | RDS | CMPLT | CMPLT | 2023-12-22 10:14 | 2023-12-22 10:18 | Weekly Rounds - NICU | (TC 12/22/2023 10:16:50 AM, Ferreira, Qwynn) Performed the check and only notes were the transports were missing some disposable items and one was not plugged in.
| 200504 | RDS | CMPLT | CMPLT | 2023-12-22 10:19 | 2023-12-22 10:24 | Weekly Rounds - 6th Oncology | (TC 12/22/2023 10:19:50 AM, Ferreira, Qwynn) Rounds were good - no issues - all rooms were good as well.
1054599 | 200505 | CM | CMPLT | CMPLT | 2023-12-22 10:35 | 2023-12-22 10:37 | Overview down | (TC 12/22/2023 10:35:05 AM, Baghirzade, Nara) PICIX would not turn on. Noticed that the UPS was operating on battery only. Did not look like the cable was kicked, it was still plugged to AC. Took a multimeter and measured voltage; electrical outlets were functional and had ~115 V. Plugged UPS back on, pressed the "on" button on the UPS however nothing was happening. Asked John to help out. He also reseated the power cable, and held the on button down. UPS turned on and was operating on AC. PICIX booted back up. Overview is functional but will keep monitoring it for any hardware failures.
| 200506 | RDS | CMPLT | CMPLT | 2023-12-22 10:38 | 2023-12-22 10:41 | Weekly Rounds: NICU | (TC 12/22/2023 10:38:55 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure for results.
| 200507 | RDS | CMPLT | CMPLT | 2023-12-22 10:41 | 2023-12-22 10:43 | Weekly Rounds: RESUS | (TC 12/22/2023 10:41:32 AM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure for results.
| 200508 | RDS | CMPLT | CMPLT | 2023-12-22 10:43 | 2023-12-22 10:46 | Weekly Rounds: 11 North | (TC 12/22/2023 10:43:11 AM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 200509 | RDS | CMPLT | CMPLT | 2023-12-22 10:46 | 2023-12-22 10:48 | Weekly Rounds: 11 South | (TC 12/22/2023 10:46:05 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
68788 | 200510 | CM | CMPLT | CMPLT | 2023-12-22 10:49 | 2023-12-22 13:27 | Not populating in iAware (NICU) | (TC 12/22/2023 10:48:36 AM, Baghirzade, Nara) Checked the network status and it was associated. SSID: mshmed Went in to server status and found that the network is configured with a test server instead of production. Followed up with Kamran and John. Ensured that the mount sinai package is active. Went on to ASM and connected the PCU; pushed the network configuration and went back to the server status; found that the server name is "mshalaris01.mountsina.hospital" however when I checked a different PCU, the server name is "alaris.mountsinai.hospital". Waiting for Samir to confirm. (TC 12/22/2023 01:24:54 PM, Baghirzade, Nara) Explained to Samir what happened with the server name and how they differ. He checked through the Alaris Systems Manager by the serial number of the device of its last connection/communication, dataset, and last dataset load. Also checked the second pump with a different server name, both PCUs are connected to the server and are communicating. Samir confirmed that it is functional and can be sent back to NICU after doing a checkin. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 59 mohms Leakage current: 72.2 uA
67628 | 200511 | CM | CMPLT | CMPLT | 2023-12-22 10:53 | 2023-12-22 10:54 | Constant error state | (TC 12/22/2023 10:52:50 AM, Baghirzade, Nara) Thermometer is in a constant error state. Coop student Connor worked on it however was unable to make it work even after changing the batteries. Emailed Sharmaine and she confirmed that the device can be retired and disposed of.
1066454 | 200512 | FC | CMPLT | CMPLT | 2023-12-22 12:18 | 2023-12-22 12:19 | QC Checklist | (TC 12/22/2023 12:18:34 PM, Baghirzade, Nara) QC checklist completed as per Sarabeth's request. Returned to 14 South | As per Sarabeth's request
| 200513 | CM | CMPLT | CMPLT | 2023-12-22 12:20 | 2023-12-22 12:23 | Wiring diagram for INVOS monitor data | (TC 12/22/2023 12:20:19 PM, Baghirzade, Nara) Kamran requested that I create the wiring diagram/connections for the cable that will be used for the INVOS monitors for Dr. Poorva's studies. Traced which wires are connected from the connector to the RJ45 connector that had only wires; checked for continuity for each wire. Then checked and noted each connection from the connector to the RJ45 connector with the cover on. Relayed the diagram to Kamran.
1069562 | 200514 | ACC | CMPLT | CMPLT | 2023-12-22 13:04 | 2024-01-04 14:17 | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE C2-9-D ACCEPTANCE | (TC 01/04/2024 02:17:48 PM, Whittington, Scott) MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE C2-9-D ACCEPTANCE | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE C2-9-D ACCEPTANCE
1069563 | 200515 | ACC | CMPLT | CMPLT | 2023-12-22 13:06 | 2024-01-04 14:14 | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE C1-6-D ACCEPTANCE | (TC 01/04/2024 02:14:15 PM, Whittington, Scott) MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE C1-6-D ACCEPTANCE | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE C1-6-D ACCEPTANCE
1069564 | 200516 | ACC | CMPLT | CMPLT | 2023-12-22 13:08 | 2024-01-04 14:10 | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE RM7C ACCEPTANCE | (TC 01/04/2024 02:10:32 PM, Whittington, Scott) MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE RM7C ACCEPTANCE | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE RM7C ACCEPTANCE
1069565 | 200517 | CM | CMPLT | CMPLT | 2023-12-22 13:09 | 2024-01-04 13:49 | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE eM6C G3 ACCEPTANCE | (TC 01/04/2024 01:48:24 PM, Whittington, Scott) MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE eM6C G3 ACCEPTANCE | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE eM6C G3 ACCEPTANCE
1069566 | 200518 | CM | CMPLT | CMPLT | 2023-12-22 13:11 | 2024-01-04 13:44 | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE RIC5-9-D ACCEPTANCE | (TC 01/04/2024 01:44:28 PM, Whittington, Scott) MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE RIC5-9-D ACCEPTANCE | MSH OPG SPP Ultrasound GE VOLUSON 10 PROBE RIC5-9-D ACCEPTANCE
1069567 | 200519 | CM | CMPLT | CMPLT | 2023-12-22 13:15 | 2023-12-22 13:19 | Not working with Smoke Evac | (TC 12/22/2023 01:17:05 PM, Reid, Emmerson ) Took over from Tommy. Tested the device and changed the foot pedal to the proper one. Also found it was not in TMS and added it. Also re-synced it to the smoke evac. Working now.
1049901 | 200520 | CM | CMPLT | CMPLT | 2023-12-22 13:24 | 2023-12-29 13:18 | MSH R61 COLUMN SENSING SWITCH CABLES TORN LOOSE, AS WELL AS CABLE ANCHOR ON COLUMN, DUE TO LEAD APRONS? | (TC 12/27/2023 08:05:03 AM, Whittington, Scott) DEC 22 MSH R61 COLUMN SENSING SWITCH CABLES TORN LOOSE, AS WELL AS CABLE ANCHOR ON COLUMN, DUE TO LEAD APRONS? TAPED UP LOOSE WIRED, ATTEMPTED TO SECURE MAIN CABLE TO HOLDING HOOL , BUT SYSTEM IN CLINICAL USE (TC 12/28/2023 08:06:28 AM, Dolinar, John) MSH R61 COLUMN SENSING SWITCH CABLES TORN LOOSE, AS WELL AS CABLE ANCHOR ON COLUMN, DUE TO LEAD APRONS? (TC 12/27/2023 08:06:28 AM, Dolinar, John) MSH R61 COLUMN SENSING SWITCH CABLES TORN LOOSE, AS WELL AS CABLE ANCHOR ON COLUMN, DUE TO LEAD APRONS? RECONNECT WIRES, REFASTEN MAIN CABLE ORDER REPLACEMENT ORDER: Part Number Quantity Brief Part Description 354969 1 Sensor, narrow gap SP9J2716 2 Cable, sensor SP9J1992 1 Plug, panel SP8G9958 1 Screws, cover SP9J1994 1 Plug, panel (TC 12/29/2023 01:15:57 PM, Dolinar, John) MSH R61 COLUMN SENSING SWITCH CABLES TORN LOOSE REPLACE: SENSOR PN# 354969 CABLE-COLUMN DOWN PN# SP9J2716 | MSH R61 COLUMN SENSING SWITCH CABLES TORN LOOSE, AS WELL AS CABLE ANCHOR ON COLUMN, DUE TO LEAD APRONS?
65262 | 200521 | CM | CMPLT | CMPLT | 2023-12-22 13:25 | 2024-01-05 08:23 | Failed nebulizer test during device check | (TC 12/22/2023 02:02:04 PM, Canjimavadivel, Kamallesh ) Performed device check and verified issue is reproduced. Downloaded logs. Refer to WO#200779
| 200522 | RDS | CMPLT | CMPLT | 2023-12-27 08:48 | 2023-12-28 07:38 | Weekly rounds for ED
| 200523 | PA | CMPLT | CMPLT | 2023-12-27 08:48 | 2024-01-05 07:53 | Ordering syringe pressure sensors for stock | (TC 12/27/2023 08:50:34 AM, Baghirzade, Nara) SC 1003135959 (TC 01/05/2024 07:52:03 AM, Baghirzade, Nara) PO 7202340674
| 200524 | RDS | CMPLT | CMPLT | 2023-12-27 08:49 | 2024-01-04 14:43 | Weekly rounds for ICU
65416 | 200525 | CM | CMPLT | CMPLT | 2023-12-27 10:00 | 2023-12-27 10:02 | Hood latch broken | (TC 12/27/2023 09:59:57 AM, Baghirzade, Nara) Replaced hood latch. Noted that the tilt was not functioning; tilt brake pad was defective. Replaced tilt brake assembly. Cover for the tilt functions was covered in something sticky; cleaned the cover. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points; calibrated the RH sensor. first measurement was at 71%. After calibration, measured 75.3%. Incubator is fully functional.
| 200526 | PA | CMPLT | CMPLT | 2023-12-27 10:42 | 2024-01-02 14:42 | Cuffs order for stock | (TC 12/27/2023 10:54:25 AM, Canjimavadivel, Kamallesh ) Placed order for cuffs (TC 01/02/2024 02:41:17 PM, Canjimavadivel, Kamallesh ) Received and placed on parts shelf | To be cahrged back as used
63808 | 200527 | CM | CMPLT | CMPLT | 2023-12-27 10:43 | 2024-01-04 09:46 | BP not working | (TC 12/27/2023 11:25:14 AM, Canjimavadivel, Kamallesh ) Device came down without any cables. Tested with spare hose and cuff, BP working normal. Installed BP hose, SpO2 cable and sensor. Needs BP cuff. Replaced stand with extra as it was wobbling. (TC 01/04/2024 09:44:32 AM, Canjimavadivel, Kamallesh ) Installed BP cuff, tested ok.
502044 | 200528 | RQ | CMPLT | CMPLT | 2023-12-27 10:57 | 2024-01-11 14:08 | ERGO brake | (TC 01/11/2024 01:56:04 PM, Green, Paul) Micheal Herz Digirad FSE came onsite. Fine tuned the arm brake as there had been continuous problems reported. Brake adjusted as per Digirad. Micheal showed MRTs how the brake functions and when to notify service for any future adjustment. Brake handle should lock in a 45 degree position and as it gets closer to hardware limit, MRT to report to service for adjustment. | same issue
| 200529 | CM | CMPLT | CMPLT | 2023-12-27 11:45 | 2023-12-27 11:45 | Meeting with 3D Printing Committee | (TC 12/27/2023 11:45:13 AM, Baghirzade, Nara) Meeting with 3D Printing Committee
62966 | 200530 | CM | CMPLT | CMPLT | 2023-12-27 11:56 | 2023-12-27 11:57 | broken door hinge that connects the door to the chassis | (TC 12/27/2023 11:55:41 AM, Campbell, Connor) replaced the plastic pin that holds the hinge onto the door the wrong door was installed, so it was swapped with another one from resus that had the wrong door as well, meaning both units now have the correct door performed a functional check, humidity, temperature and warmer temperature all within range returned to service
1059983 | 200531 | CM | CMPLT | CMPLT | 2023-12-27 15:03 | 2024-03-01 16:29 | Created by Failed PM WO # 199336 : PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 01/03/2024 10:23:36 AM, Cortez, John) Put in service request (TC 01/09/2024 10:08:24 AM, Cortez, John) Bed+Intellidrive batteries were replaced PM aborted, Left rail display flickering (TC 01/10/2024 10:31:03 AM, Cortez, John) Put in service request for defective display (TC 02/23/2024 04:04:09 PM, Cortez, John) Vince completed screen repair and sent me service report Attempted to complete PM, but chair egress test failed. Looks like switch used for footboard detection is faulty? Will reach out to Vince (TC 02/26/2024 10:17:36 AM, Cortez, John) Reached out to Vince and he suggested to place a WO (TC 03/01/2024 04:25:41 PM, Cortez, John) Vince checked bed and replaced switch/cable I tested functionality, tested ok | (TC 12/27/2023 03:01:09 PM, Canjimavadivel, Kamallesh ) Found loose wires on left side near foot board. Swapped out with repaired bed on 18.
1066921 | 200532 | CM | CMPLT | CMPLT | 2023-12-27 15:12 | 2023-12-27 15:16 | Head section not working | (TC 12/27/2023 03:12:06 PM, Canjimavadivel, Kamallesh ) Head board was not set to correct width adn was stuck with head section rail. Freed head board and adjusted to correct width and cleared error.
| 200533 | PA | CMPLT | CMPLT | 2023-12-28 07:42 | 2024-01-02 14:42 | Thermometer probes for stock | (TC 12/28/2023 07:55:14 AM, Canjimavadivel, Kamallesh ) Placed order for temp probes (TC 01/02/2024 02:42:10 PM, Canjimavadivel, Kamallesh ) Received and placed on parts shelf
1055933 | 200534 | CM | CMPLT | CMPLT | 2023-12-28 07:53 | 2023-12-28 07:55 | themometer not working | (TC 12/28/2023 07:53:21 AM, Campbell, Connor) tested NIBP, SpO2 and temp NIBP, SpO2 working fine temp error when measuring, thermometer probe determined to be faulty replaced probe with a spare dug a probe cover that was stuck in the well out thermometer works now returned to service
1051698 | 200535 | CM | CMPLT | CMPLT | 2023-12-28 08:07 | 2023-12-28 08:08 | temperature not working error | (TC 12/28/2023 08:07:30 AM, Campbell, Connor) observed damage to the contacts of the probe replaced probe with a new one, now it functions batteries were low, replaced those as well
| 200536 | RDS | CMPLT | CMPLT | 2023-12-28 08:28 | 2023-12-28 08:32 | Weekly Checks For PACU | (TC 12/28/2023 08:29:29 AM, Reid, Emmerson ) Rounds complete. No issues found.
| 200537 | RDS | CMPLT | CMPLT | 2023-12-28 08:33 | 2023-12-28 08:36 | Weekly Checks For OR | (TC 12/28/2023 08:34:23 AM, Reid, Emmerson ) Also had to check the Defib but it was fine and had been checked by the staff.
63439 | 200538 | CM | CMPLT | CMPLT | 2023-12-28 12:18 | 2023-12-28 12:19 | not charging | (TC 12/28/2023 12:17:51 PM, Campbell, Connor) power cord was faulty, replaced it with a spare from another broken unit that needs to be sent for repairs tested the flow rate it was within range at 41.23mL delivered at 75mL/hr over 30 min returned to service
66463 | 200539 | CM | CMPLT | CMPLT | 2023-12-28 13:45 | 2024-01-05 08:23 | Nebulizer test fail during device check | (TC 12/28/2023 02:55:31 PM, Canjimavadivel, Kamallesh ) Downloaded logs, issue related to regulators. Checked S2 sensor, passed. Adjusted O2 and AIR regulators by decreasing pressure. Self-test pass Refer to WO#200763
64279 | 200540 | CM | CMPLT | CMPLT | 2023-12-28 14:21 |  | missing NICU profile and not conecting to the LAN | (TC 12/28/2023 02:21:27 PM, Chukri, Samir) Me and Kamran switching the monito with one from Biomed , Kamran re-configured and keep it on the biomed for testing .
| 200541 | CM | ACTIV | ISSUE | 2023-12-29 07:37 |  | PFT working calibration and leaking test to Hypair unit. | (TC 12/27/2023 01:00:00 PM, Chukri, Samir) working with Nara and John leaking test and pressure pass. still need more ditals Nara will contact vendor support tech.
1066114 | 200542 | FC | CMPLT | CMPLT | 2023-12-29 09:01 | 2023-12-29 09:45 | QC check | (TC 12/29/2023 09:37:58 AM, Canjimavadivel, Kamallesh ) QC check passed
62966 | 200543 | CM | CMPLT | CMPLT | 2023-12-29 09:13 | 2023-12-29 09:14 | front door broken | (TC 12/29/2023 09:12:54 AM, Campbell, Connor) replaced the front door plastic rivet the holds the door in place performed a functional check humidity, temperature and warmer temperature all passed returned to service
1061965 | 200544 | CM | CMPLT | CMPLT | 2023-12-29 10:21 | 2023-12-29 10:24 | No functions, not weighing | (TC 12/29/2023 10:21:24 AM, Ferreira, Qwynn) Repaired the bed, tested the features. Bed needed to be reset and the patient needed to be removed to zero the scale before weighing.
| 200545 | PA | CMPLT | CMPLT | 2023-12-29 10:30 | 2024-02-13 08:31 | Isopropyl Alcohol 99% | Items have arrived
66468 | 200546 | CM | CMPLT | CMPLT | 2023-12-29 10:55 | 2024-01-19 13:43 | Making a noise from main unit and back of monitor heating up | (TC 12/29/2023 11:00:00 AM, Canjimavadivel, Kamallesh ) Pulled logs and tried to find source of noise (TC 01/02/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Updated Olcay on all the vent issues from previous week. Replaced exp valve heater fan. Replaced both right and left cover of PS500. Tested unit for 2 weeks. Functional.
1058169 | 200547 | CM | CMPLT | NFF | 2023-12-29 10:58 | 2023-12-29 11:13 | Says no cord attached | (TC 12/29/2023 11:12:18 AM, Reid, Emmerson ) Checked and was working fine
1058574 | 200548 | CM | ACTIV | ISSUE | 2023-12-29 11:04 |  | cannot be calibrated | (TC 12/29/2023 11:03:49 AM, Campbell, Connor) attempted to calibrate 5 times on 3 different days calibration failed every time sometimes error 400, sometime 01a needs service
67763 | 200549 | CM | CMPLT | CMPLT | 2023-12-29 11:37 | 2023-12-29 11:40 | Battery lid broken | (TC 12/29/2023 11:39:14 AM, Reid, Emmerson ) Battery lid was broken. I was able to fix it. I also updated the library to APS EPIDURAL 2023 OCT
1057384 | 200550 | CM | CMPLT | CMPLT | 2023-12-29 13:14 | 2024-02-06 08:30 | MSH DR CARESTREAM CR850 MOVE BACK TO IMAGING DR HALLWAY | (TC 12/29/2023 01:14:17 PM, Whittington, Scott) MSH DR CARESTREAM CR850 MOVE BACK TO IMAGING DR HALLWAY ORIGINAL NETWORK JACKS NOT CORRECT ONE IS DISCONNECTED OTHER IS ON 100 FIXED, REQUIRES VALIDATION IT IS ON MI VLAN STILL (TC 12/29/2023 06:14:00 PM, Whittington, Scott) MSH DR CARESTREAM CR850 MOVE BACK TO IMAGING DR HALLWAY ORIGINAL NETWORK JACKS NOT CORRECT ONE IS DISCONNECTED OTHER IS ON 100 FIXED, REQUIRES VALIDATION IT IS ON MI VLAN STILL (TC 12/29/2023 06:14:00 PM, Whittington, Scott) MSH DR CARESTREAM CR850 MOVE BACK TO IMAGING DR HALLWAY ORIGINAL NETWORK JACKS NOT CORRECT ONE IS DISCONNECTED OTHER IS ON 100 FIXED, REQUIRES VALIDATION IT IS ON MI VLAN STILL | MSH DR CARESTREAM CR850 MOVE BACK TO IMAGING DR HALLWAY
28877 | 200551 | PM | CANCL | CANCL | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Scales, Infant
50362 | 200552 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Scales, Infant
50363 | 200553 | PM | CANCL | CANCL | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Scales, Infant
52259 | 200554 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-09-04 11:53 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/04/2024 11:53:06 AM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 79.9 uA Power cord resistance = 180 mohms
53008 | 200555 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
60022 | 200556 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-01-29 11:31 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 01/29/2024 11:29:18 AM, Rhuland, Courtney) PM complete Calibrated REF and main voltage Scale calibrated and verified Replaced 9V rechargeable battery Replaced corner grommet Functional check - OK Set 35C, 75% humidity Read 35C, 74.6% humidity Electrical safety - OK
62124 | 200557 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Scales, Infant
62270 | 200558 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-06-10 13:03 | PM-MAJOR - Annual PM For: Simulators, Multiparameter | (TC 06/10/2024 01:03:07 PM, Chukri, Samir) Performed calibration by Getinge tech, certificate of calibration attached.
62532 | 200559 | PM | CANCL | CANCL | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: LASERS, SURGICAL, CARBON DIOXIDE
62631 | 200560 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-02-16 08:36 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/14/2024 08:35:04 AM, Rendon Lopez , Elizabeth) Warmer in LVL7 OR. Replaced blender. PM completed, all tests passed
62632 | 200561 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-01-31 09:10 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/29/2024 09:09:18 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
62718 | 200562 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
63163 | 200563 | PM | CANCL | CANCL | 2024-01-01 02:17 | 2024-08-02 09:03 | PM-MAJOR - Annual PM For: Aspirators, Emergency
63273 | 200564 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
63274 | 200565 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
63507 | 200566 | PM | CANCL | CANCL | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Holmium Laser System | Laser is out of support by vendor; no corrective and preventive maintenance available. In process of acquiring a new laser as replacement.
63525 | 200567 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Scales, Infant
63526 | 200568 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Scales, Infant
63786 | 200569 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-06-10 13:55 | PM-MAJOR - Annual PM For: Simulators, Multiparameter | (TC 06/10/2024 01:55:12 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached.
63893 | 200570 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Radiofrequency Generators
64964 | 200571 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-06-11 09:21 | PM-MAJOR - Annual PM For: Testers, External Cardiac Pacemaker | (TC 06/11/2024 09:21:24 AM, Chukri, Samir) Performed Calibration by Getinge tech, Calibration Certificate attached.
64996 | 200572 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-06-17 12:28 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/17/2024 12:27:32 PM, DeGuzman, Leila) PM completed. All tests passed.
65194 | 200573 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Microscopes, Light, Examination
65672 | 200574 | PM | CNP | ENL | 2024-01-01 02:17 | 2024-09-05 11:26 | PM-MAJOR - Annual PM For: Aspirators, Emergency
65874 | 200575 | PM | CANCL | CANCL | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM every 3 years - updated PM schedule to reflect
65953 | 200576 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-28 12:59 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 12:58:41 PM, Rhuland, Courtney) PM completed by BD
65958 | 200577 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-05-02 10:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/02/2024 10:44:17 AM, Rhuland, Courtney) Replaced left handle During final force verification, test failed Recalibrated syringe and all tests passed Calibration and check-in complete
65959 | 200578 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-18 10:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/18/2024 10:55:13 AM, Baghirzade, Nara) PM completed by BD.
65998 | 200579 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-08-01 12:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66110 | 200580 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
66112 | 200581 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM - Annual
66439 | 200582 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
66681 | 200583 | PM | CANCL | CANCL | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | PM completed in July 2023 - updated PM schedule to reflect
66704 | 200584 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-08-13 13:10 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/13/2024 01:10:20 PM, DeGuzman, Leila) PM passed and completed. Replaced air probe. Electrical safety passed. Ground resistance: 30 mOhm Leakage current: 11.6 uA
67227 | 200585 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-02-16 08:36 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/15/2024 08:36:05 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67406 | 200586 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-06-11 12:41 | PM-MAJOR - Annual PM For: Testers, ESU | (TC 06/11/2024 12:40:36 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
67417 | 200587 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-01-31 09:05 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 01/25/2024 09:04:26 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67894 | 200588 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-06-20 10:09 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/20/2024 10:09:10 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Upgraded Software and library, replaced batteries
67917 | 200589 | PM | CANCL | CANCL | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | PM completed July 2023 - updated PM schedule to reflect
67922 | 200590 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-05-01 12:28 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/01/2024 12:02:11 PM, Rhuland, Courtney) Calibrated: Line and Ref voltage scale Verified scale Replaced 2 porthole seals Replaced door sticker Functional check - OK Set 35C & 75% humidity Read 35C & 75.4% humidity Electrical safety - OK
67930 | 200591 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-07-11 12:32 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/11/2024 12:32:08 PM, Rhuland, Courtney) PM completed Calibrated: MAINS & LV voltage Scale Replaced 9V battery with rechargeable 9V battery Funtional check - OK Set 35C & 75% humidity Read 35.1C & 75.6% humidity Electrical safety - OK Ground resistance - 67 mOhms Leakage current - 73.1 uA
67962 | 200592 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-08-16 14:02 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/16/2024 02:02:25 PM, Rhuland, Courtney) PM completed Replaced: RH sensor Probe sheild 9V battery Inside door 2 porthole seals Calibrated: RH MAINS & VREF voltage Scale Functional check - OK Set 35C, 75% humidity Read 35C, 75.2% humidity Electrical safety - OK Ground resistance - 58 mOhms Leakage current - 1.1 uA
67988 | 200593 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
68370 | 200594 | PM | ACTIV | ISSUE | 2024-01-01 02:17 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68383 | 200595 | PM | ACTIV | READY | 2024-01-01 02:17 |  | PM Major - Delivery Module SEVO
68384 | 200596 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-09 10:51 | PM Major - Delivery Module SEVO | (TC 04/09/2024 10:50:27 AM, Scero, Tomlyn) PM was completed by Draeger during repair
68391 | 200597 | PM | ACTIV | READY | 2024-01-01 02:17 |  | PM Major - Delivery Module SEVO
68393 | 200598 | PM | ACTIV | READY | 2024-01-01 02:17 |  | PM Major - Delivery Module SEVO
68397 | 200599 | PM | ACTIV | READY | 2024-01-01 02:17 |  | PM Major - Delivery Module SEVO
68541 | 200600 | PM | ACTIV | READY | 2024-01-01 02:17 |  | PM Major - Delivery Module SEVO
68552 | 200601 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-22 12:02 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 02/07/2024 02:32:03 PM, Scero, Tomlyn) PM procedures started after device was repaired. All PM parts replaced. (TC 02/08/2024 02:07:38 PM, Scero, Tomlyn) Replaced batteries and started page 1 of the IPM-L procedures. Emailed Draeger to send me the latest version of IPM-L 27.0 (TC 03/18/2024 02:43:02 PM, Scero, Tomlyn) Continued with PM procedures (TC 03/21/2024 02:20:43 PM, Scero, Tomlyn) PM procedures complete and passed. See attached result sheet for info. replaced Data cable going to EC10.
68605 | 200602 | PM | CMPLT | PMFAIL | 2024-01-01 02:17 | 2024-03-12 08:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 08:31:12 AM, Baghirzade, Nara) Broken front case and rear case, defective right IUI
68606 | 200603 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-07-30 13:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/30/2024 01:21:35 PM, DeGuzman, Leila) PM completed.
68620 | 200604 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-24 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM completed by BD
68627 | 200605 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-15 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68654 | 200606 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-24 09:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68673 | 200607 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-08-16 13:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68704 | 200608 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-02 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68718 | 200609 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-03 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:58:45 PM, Rhuland, Courtney) PM completed by BD
68726 | 200610 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-18 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68737 | 200611 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-18 15:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68744 | 200612 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-20 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68751 | 200613 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-26 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68770 | 200614 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-26 14:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68794 | 200615 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-05 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:43:01 PM, Rhuland, Courtney) PM completed by BD
68799 | 200616 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-26 13:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68806 | 200617 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-18 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68813 | 200618 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-24 14:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68819 | 200619 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-12 14:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68821 | 200620 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-26 14:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68832 | 200621 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-03-27 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68839 | 200622 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-04-24 14:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM completed by BD
68859 | 200623 | PM | CMPLT | CMPLT | 2024-01-01 02:17 | 2024-05-31 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/31/2024 01:11:45 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced by BD. Electrical Safety - Ground resistance: 182 mohms Leakage current: 71.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68879 | 200624 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-12 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68882 | 200625 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-26 12:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
68919 | 200626 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-13 09:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/11/2024 09:09:27 AM, Rendon Lopez , Elizabeth) PM completed by BD
68923 | 200627 | PM | CANCL | CANCL | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68925 | 200628 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-17 08:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 08:52:18 AM, Rhuland, Courtney) PM completed by BD
68939 | 200629 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-18 14:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68945 | 200630 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-26 12:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68959 | 200631 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-12 13:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68964 | 200632 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-08-16 14:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69005 | 200633 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-12 11:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69077 | 200634 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-19 21:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69082 | 200635 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-24 09:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69105 | 200636 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-24 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69129 | 200637 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-23 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69154 | 200638 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-14 14:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:12:38 PM, Huang, Jessica) replaced the male iui connector and completed PM
69165 | 200639 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-14 14:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69210 | 200640 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-14 14:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:28:44 PM, Huang, Jessica) replaced the female iui connector and completed PM
69211 | 200641 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-03 14:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:51:01 PM, Rhuland, Courtney) PM completed by BD
69264 | 200642 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-15 09:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/11/2024 09:07:50 AM, Huang, Jessica) repalced the female IUI connector and completed PM
69272 | 200643 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-04 12:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 12:52:51 PM, Rhuland, Courtney) PM completed by BD
69275 | 200644 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-12 09:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69291 | 200645 | PM | CANCL | CANCL | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69305 | 200646 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-07-19 13:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/19/2024 01:19:12 PM, Baghirzade, Nara) During PM, patient occlusion was failing. Completed analog test and found that both pressure sensors are defective. Replaced bottle-side and patient-side pressure sensors. Replaced latch and spring in the door latch assembly. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69320 | 200647 | PM | CMPLT | PMFAIL | 2024-01-01 02:18 | 2024-03-15 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69325 | 200648 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-03 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:35:50 PM, Rhuland, Courtney) PM completed by BD
69345 | 200649 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-15 14:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69396 | 200650 | PM | CMPLT | PMFAIL | 2024-01-01 02:18 | 2024-03-25 11:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69450 | 200651 | PM | CANCL | CANCL | 2024-01-01 02:18 | 2024-08-16 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69473 | 200652 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-24 09:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69542 | 200653 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-03 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:55:10 PM, Reid, Emmerson ) Ran Calibration and ASM PM, all tests passed. PM Complete.
69554 | 200654 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-03 11:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69576 | 200655 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-19 20:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69579 | 200656 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-25 11:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69603 | 200657 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-24 09:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM completed by BD
69650 | 200658 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-14 07:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69666 | 200659 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-14 15:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69681 | 200660 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-10-28 13:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 10/28/2024 01:51:28 PM, Alvarado, Karla) The device was send to the department because of an error in channel; Turn on the monitor and the PCU screen displayed Error code: 200.5040 According to the manual, the Pump Module software version is not compatible with software flashed into PC Unit. Therefore, reflash the pump module. Performed MSH PM in ASM. All test passed.
67367 | 200661 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
32540 | 200662 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-12 14:39 | PM - Monthly | PM Major done Vitrea Enterprise node setup done IF cal done Backups done to DIRECTORY Slice count - 20008 TOS phantom a(-995.4, 9.9) b(331.7, 14.8) c(132.5, 13.8) d(97.3, 15.0) e(-106.0, 12.7) f(-1.1, 13.3) -System passed all tests
32560 | 200663 | PM | CANCL | CANCL | 2024-01-01 02:18 |  | PM - Monthly
32580 | 200664 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-06 14:47 | PM - Monthly | PM Major done Recon Unit filters and fans cleaned IF cal done Backups completed Slice count - 184741 TOS phantom a(-992.6, 10.3) b(334.6, 14.8) c(134.1, 14.6) d(96.1, 14.1) e(-106.0, 12.3) f(0.9, 13.6) -System passed all tests
501440 | 200665 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-08 20:27 | PM - Semi-Annual | PM
501440 | 200666 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-09 20:26 | HARP - Semi-Annual | HARP
502044 | 200667 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-05-29 12:31 | PM - Semi-Annual | (TC 05/29/2024 12:30:35 PM, Green, Paul) PM as per procedure. Coolant was very low. Topped up. Ordered desscant for next time as we are out. Contaced Digirad and placed the order.
502492 | 200668 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-01-12 13:21 | PM-Major, annual PM | (TC 01/10/2024 03:17:47 PM, Green, Paul) Pulled injector out of room as MRTs informed me contrast side piston does not dock with the syringe plunger. Also, getting a critical error as per John MRT, but don't have further details. When onsite, A syringe was removed, yet display shows that syringe is still installed. (TC 01/12/2024 01:19:54 PM, Green, Paul) Performed full functional tests, ok. From: Green, Paul <Paul.Green@uhn.ca> Sent: January 12, 2024 1:19 PM To: Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca>; Yang, Helen - Sinai Health System <Helen.Yang@sinaihealth.ca>; Heise, Ona - Sinai Health System <Ona.Heise@sinaihealth.ca> Cc: Paz, Ramon <Ramon.Paz@uhn.ca> Subject: RE: [External] RE: Bayer parts needed Hi Jason, The MSH MRXperion on 2 has been repaired. The potentiometers were replaced and position calibrations performed. Also, docking pin ribbon cables were replaced and cam cups cleaned/lubricated. MRX was reinstalled back in 2 this morning, and the spare Spectris Solaris stored again on MSH 5, in 3T equipment room. The injector battery is fully inserted with the power supply connected and plugged in. This will ensure battery is charged for next use. I see there is another spare battery plugged in there, which is great. The spare injector is ready for future use. John will use the MRX and let us know if he has any problems. Regards, Paul
541287 | 200669 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM
1046023 | 200670 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-01-26 08:57 | PM Major - Delivery Module SEVO | (TC 01/25/2024 08:37:26 AM, Scero, Tomlyn) All PM parts replaced and test performed and passed as per IPM-L
1046652 | 200671 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-15 07:29 | PM-MAJOR - Annual PM For: Bladder Scanner | (TC 04/08/2024 07:28:55 AM, Huang, Jessica) PM completed
1047069 | 200672 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-26 12:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1047101 | 200673 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-13 09:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:28:01 AM, Rendon Lopez , Elizabeth) PM completed by BD
1047839 | 200674 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-28 11:26 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 02/28/2024 11:25:26 AM, Baghirzade, Nara) Completed PM after software upgrade to V1.6.5 to ensure full functionality and accuracy of pump. See procedure for results.
1047843 | 200675 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-27 09:09 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 03/27/2024 09:08:44 AM, Rhuland, Courtney) PM completed Refer to procedure for details
1048602 | 200676 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-06-11 09:33 | PM-MAJOR - Annual PM | (TC 06/11/2024 09:32:33 AM, Reid, Emmerson ) PM complete, see proedure for details.
1048627 | 200677 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 09:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048637 | 200678 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 09:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:03:13 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048639 | 200679 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048643 | 200680 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 11:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048651 | 200681 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 11:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:40:52 AM, Huang, Jessica) replaced the female iui connector completed PM
1048654 | 200682 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 11:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:08:44 AM, Huang, Jessica) replaced the female iui connector completed PM
1048655 | 200683 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 11:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:01:51 AM, Huang, Jessica) replaced the female iui connector completed PM
1048663 | 200684 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 11:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:39:54 AM, Huang, Jessica) replaced the female iui connector completed PM
1048666 | 200685 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 11:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:38:04 AM, Huang, Jessica) replaced the female iui connector completed PM
1048676 | 200686 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 13:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048677 | 200687 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 13:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048686 | 200688 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 13:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048688 | 200689 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 13:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:22:13 AM, Huang, Jessica) replaced the female iui connector completed PM
1048691 | 200690 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048694 | 200691 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-19 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 03:44:41 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048697 | 200692 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048700 | 200693 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-22 17:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:16:07 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1049658 | 200694 | PM | CMPLT | PMFAIL | 2024-01-01 02:18 | 2024-01-31 09:15 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 01/29/2024 09:12:59 AM, Rendon Lopez , Elizabeth) Found gas outlet loose during PM, needs blender overhaul. Removed from PACU, waiting for blender overhaul
1050089 | 200695 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
1050476 | 200696 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-01-05 11:18 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | PM parts replaced, all tests passed. Expiratory cassette membrane replaced. Needed 2 new batteries. Rex replaced them.
1050477 | 200697 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-08 09:30 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | Completed annual PM. Replaced expiratory cassette's membrane. Passed device check. Tested. Functional. leakage 20 uA. Resistance 115 mOhm
1050480 | 200698 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-05-03 10:06 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | Annual PM comlpleted. Functional. Replaced air filter holder. It was making noise and stopping fan.
1051379 | 200699 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: Scales, Infant
1051733 | 200700 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-02 08:38 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 03/28/2024 08:35:44 AM, Ferreira, Qwynn) completed PM Changed battery Electrical Safety-OK Cleaned interior of the transport including fan Functional Check - OK
1053033 | 200701 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR-Annual PM For : Aspirator unit battery
1053447 | 200702 | PM | CANCL | CANCL | 2024-01-01 02:18 |  | PM MAJOR-Stretchers, Mobile, Hospital, Patient Transfer | PM completed in June 2023 - updated pm schedule to reflect
1053489 | 200703 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-27 09:54 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult | (TC 02/27/2024 09:49:28 AM, Canjimavadivel, Kamallesh ) Replaced expiratory cassette membrane Replaced filters and nozzles Pre check pass Total op time: 17786 hrs Total vent time: 6188 hrs
1053490 | 200704 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-01-12 14:39 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult | (TC 01/12/2024 11:35:17 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 19855 hrs Total vent time: 10209 hrs
1053491 | 200705 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-27 10:45 | PM-MAJOR - Annual PM For: Ventilators | (TC 02/27/2024 10:36:16 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 19235 hrs Total vent time: 9007 hrs
1053492 | 200706 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-25 10:43 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/25/2024 10:37:08 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 26826 hrs Total vent time: 14458 hrs PM complete
1053493 | 200707 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-09 11:39 | PM-MAJOR - Annual PM For: Ventilators | (TC 02/09/2024 11:34:25 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 15709 hrs Total vent time: 7155 hrs
1053728 | 200708 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-28 13:57 | PM-MAJOR- Annual PM for: Pacemakers, Cardiac, External | (TC 02/28/2024 09:44:22 AM, Scero, Tomlyn) I brought the device down to biomed for PM. When I opened the case I noticed that the patient connection cable was missing. I notified Dianne Locke via email to see if she could locate the cable, or order a new one so the PM could be complete. Waiting for a response. (TC 02/28/2024 01:55:45 PM, Scero, Tomlyn) I used the patient connection cable from the other external pacemaker that had come down for a funcitonal check. This device passed all functional and PM tests
1053995 | 200709 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: Insufflators
1054460 | 200710 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-13 13:02 | Lubrication cycle | (TC 02/13/2024 01:01:34 PM, Canjimavadivel, Kamallesh ) Lubrication cycle complete
1054462 | 200711 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-15 14:46 | Lubrication cycle | (TC 02/15/2024 02:45:43 PM, Canjimavadivel, Kamallesh ) Lubrication cycle complete
1054468 | 200712 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-15 14:46 | Lubrication cycle | (TC 02/15/2024 02:46:09 PM, Canjimavadivel, Kamallesh ) Lubrication cycle complete
1055120 | 200713 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-13 09:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:43:20 AM, Rendon Lopez , Elizabeth) PM completed by BD
1055125 | 200714 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-13 09:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:42:01 AM, Rendon Lopez , Elizabeth) PM completed by BD
1055180 | 200715 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-12 10:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 10:46:42 AM, Rendon Lopez , Elizabeth) PM completed by BD
1055562 | 200716 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-01-09 08:18 | PM-MAJOR - Annual PM For: Ventilators | (TC 01/08/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 14940 hrs Total vent time: 7458 hrs
1058144 | 200717 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-03-12 08:43 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 02/21/2024 01:47:41 PM, Rendon Lopez , Elizabeth) Inspected, cleaned, PM parts replaced (TC 02/22/2024 01:47:41 PM, Rendon Lopez , Elizabeth) Calibration done. Performance verification done. (TC 02/23/2024 01:47:41 PM, Rendon Lopez , Elizabeth) Pinch valve screw broken, ordered replacement (TC 02/27/2024 08:42:12 AM, Rendon Lopez , Elizabeth) Replaced screw, tested-OK. Returned to RT NICU spare
1058145 | 200718 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-23 13:51 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 02/21/2024 01:50:01 PM, Rendon Lopez , Elizabeth) Inspected, cleaned, PM parts replaced (TC 02/22/2024 01:50:01 PM, Rendon Lopez , Elizabeth) Calibration done. Performance verification done. (TC 02/23/2024 01:50:01 PM, Rendon Lopez , Elizabeth) Pinch valve screw broken, ordered replacement
1058146 | 200719 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-23 13:47 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | (TC 02/22/2024 01:45:41 PM, Rendon Lopez , Elizabeth) Inspection, cleaning parts, PM parts replaced. Calibration procedures done. Performance verification done. (TC 02/23/2024 01:45:41 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Sent checklist to Bunnell, returned device to RT NICU
1058257 | 200720 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-08-16 10:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1058542 | 200721 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-28 11:01 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 02/28/2024 11:00:16 AM, Reid, Emmerson ) Battery replaced, PM Complete | Every year needs dust cleaning, every two years the battery is changed.
1058543 | 200722 | PM | CANCL | CANCL | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible |  | Every year needs dust cleaning, every two years the battery is changed.
1058544 | 200723 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-28 11:08 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 02/28/2024 11:07:59 AM, Reid, Emmerson ) Battery Replaced, PM Complete. | Every year needs dust cleaning, every two years the battery is changed.
1058545 | 200724 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-28 10:59 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 02/28/2024 10:58:23 AM, Reid, Emmerson ) Replaced battery, PM Complete. | Every year needs dust cleaning, every two years the battery is changed.
1059180 | 200725 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1061154 | 200726 | PM | CANCL | CANCL | 2024-01-01 02:18 |  | PM, SEMI-ANNUAL, ULTRASOUND GENERAL |  | PM, SEMI-ANNUAL, ULTRASOUND GENERAL CANCELLED NO AGREEMENT IN PLACE REPAIRS ONLY
1061364 | 200727 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-09 19:31 | HARP - Semi Annual | HARP
1061364 | 200728 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-07 19:32 | PM - Semi Annual | PM
1061867 | 200729 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - Annual PM For: Waste-Disposal Units
1063181 | 200730 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-05-22 14:54 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Adult/Pediatric | (TC 05/22/2024 02:54:21 PM, Cortez, John) Visual check ok Replaced PM parts (filters, O2 cells) Completed calibration + general tasks and checks Electrical safety passed: 16uA enclosure leakage Preop check passed
1063182 | 200731 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-05-22 16:18 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Adult/Pediatric | (TC 05/22/2024 04:17:35 PM, Cortez, John) Visual check ok Operating time: 789 hrs Replaced PM parts (filters, O2 cells) Completed calibration + general tasks and checks Electrical safety passed: 9uA enclosure leakage Preop check passed
1063310 | 200732 | PM | ACTIV | ISSUE | 2024-01-01 02:18 |  | PM-MAJOR - ANNUAL PM For: CO2 Laser
1063799 | 200733 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-01-16 11:04 | PM-Major, annual PM | (TC 01/11/2024 01:53:16 PM, Green, Paul) MRTs having problems with contrast A-side. Can't install syringes as piston does not enguage with plungers. Removed mechanical drive and replaced pots. Removed cam cups and cleaned/lubricated. Replaced docking pin ribon cables on both A and B side. Reassemled and performed position calibration, ok. Called away so will return. (TC 01/12/2024 11:02:11 AM, Green, Paul) Fully tested MRXperion and ok. Reinstalled back in MSH-2MRI suite. Stored spare MSH Spectris Solaris back on MSH5- 3T equipment room. Plugged in charger with battery installed on injector. Also, ensure second battery is charging ok from wall. Informed MRTs and Jason of the work done.
1064331 | 200734 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-04-10 13:15 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 04/10/2024 01:14:37 PM, Rhuland, Courtney) Tested with 5kg weight - read 5.000kg Functional check - OK
1066756 | 200735 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-16 13:13 | Lubrication cycle | (TC 02/16/2024 01:12:32 PM, Canjimavadivel, Kamallesh ) Lubrication cycle complete
1066757 | 200736 | PM | CMPLT | CMPLT | 2024-01-01 02:18 | 2024-02-16 13:16 | Lubrication cycle | (TC 02/16/2024 01:12:59 PM, Canjimavadivel, Kamallesh ) Lubrication cycle complete
1067589 | 200737 | PM | CMPLT | NA | 2024-01-01 02:18 | 2024-02-15 08:01 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/15/2024 08:00:42 AM, Baghirzade, Nara) Incorrect PM date. Updated to new PM date.
1067591 | 200738 | PM | CMPLT | NA | 2024-01-01 02:18 | 2024-02-15 08:02 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/15/2024 08:01:38 AM, Baghirzade, Nara) Incorrect PM date. Updated to new PM date.
1046035 | 200739 | CM | CMPLT | CMPLT | 2024-01-01 13:15 | 2024-01-01 13:16 | MSH ER DRX EVO E5-02(ER) WALLSTAND OFF TRACKS | (TC 01/01/2024 01:15:42 PM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) WALLSTAND OFF TRACKS REPAIRED | MSH ER DRX EVO E5-02(ER) WALLSTAND OFF TRACKS
1049362 | 200740 | CM | CANCL | DUP | 2024-01-01 13:18 |  | MSH ER STELLANT INJECTOR FAILED TO LOCK IN SYRINGE |  | MSH ER STELLANT INJECTOR FAILED TO LOCK IN SYRINGE
32520 | 200741 | CM | CMPLT | CMPLT | 2024-01-01 13:21 | 2024-01-12 20:19 | MSH DR2 TUBE LATERAL BRAKE | (TC 01/13/2024 02:20:04 PM, Whittington, Scott) MSH DR2 TUBE LATERAL BRAKE NORTH END REPLACED TUBE BRIDGE HOLDING FIRMLY (TC 01/10/2024 08:20:00 PM, Whittington, Scott) MSH DR2 TUBE LATERAL BRAKE NORTH END REPLACED TUBE BRIDGE HOLDING FIRMLY | MSH DR2 TUBE LATERAL BRAKE
| 200742 | RDS | CMPLT | CMPLT | 2024-01-02 07:47 | 2024-01-02 09:19 | Weekly rounds for ED
66706 | 200743 | CM | CMPLT | CMPLT | 2024-01-02 07:50 | 2024-01-02 10:16 | Membrane switch not working | (TC 01/02/2024 10:15:27 AM, Rhuland, Courtney) Replaced faulty membrane switch Functional check - OK
67948 | 200744 | CM | CMPLT | CMPLT | 2024-01-02 08:31 | 2024-01-02 08:33 | Reported that humidifier water is contaminated | (TC 01/02/2024 08:31:14 AM, Baghirzade, Nara) No cartridge was sent with the incubator. Looking closely near the humidifier, there was brown gunk all over (possibly iodine). Cleaned it with alcohol and then water. Cleaned it again with water. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.1% RH. Incubator is fully functional.
| 200745 | PA | CMPLT | CMPLT | 2024-01-02 08:35 | 2024-01-04 09:36 | Sterile water for NICU altrix | (TC 01/02/2024 08:37:18 AM, Baghirzade, Nara) SC 1003137666 (TC 01/04/2024 09:35:04 AM, Baghirzade, Nara) Order# 4150269409
66710 | 200746 | CM | CMPLT | CMPLT | 2024-01-02 08:43 | 2024-01-02 09:25 | Humidifier leaking | (TC 01/02/2024 09:21:29 AM, Rhuland, Courtney) Humidifier leaking Humidifier gasket dried out and cracked Replaced gasket and reservoir switch button Functional check - OK No leaking
1056125 | 200747 | PM | CMPLT | CMPLT | 2024-01-02 09:13 | 2024-01-04 09:30 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/02/2024 03:01:30 PM, Canjimavadivel, Kamallesh ) Discovered drain valve body was damaged during the start of cleaning cycle. Performed cleaning cycle and discovered a leak from the bottom. Need to replace drain valve body. Can repurpose from Altrix that is to be sent back. (TC 01/03/2024 02:28:16 PM, Canjimavadivel, Kamallesh ) Removed drain valve body from second Altrix and replaced broken part. Assembled all back together and ran it continuously. Still visible drops of water from bottom. (TC 01/04/2024 09:28:18 AM, Canjimavadivel, Kamallesh ) Clsoed up device and returned to unit. Can be used with minor leak for now. Reached out to vendor for quote of parts. Will place order and replace once parts received.
| 200748 | RDS | CMPLT | CMPLT | 2024-01-02 09:34 | 2024-01-02 09:37 | Weekly Checks For PACU | (TC 01/02/2024 09:36:39 AM, Reid, Emmerson ) Round complete, also investigted some nurse call questions for Kamran.
| 200749 | RDS | CMPLT | CMPLT | 2024-01-02 10:12 | 2024-01-02 10:35 | Weekly Checks in CCU | (TC 01/02/2024 10:34:53 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
1053492 | 200750 | CM | CMPLT | CMPLT | 2024-01-02 10:20 | 2024-01-02 10:26 | Exp cassette membrane 7% | (TC 01/02/2024 10:19:40 AM, Canjimavadivel, Kamallesh ) Replaced expiratory cassette membrane and reset counter. Pre-use check pass
65914 | 200751 | CM | CMPLT | CMPLT | 2024-01-02 10:58 | 2024-01-04 12:53 | Not reading syringe size | (TC 01/02/2024 10:57:39 AM, Baghirzade, Nara) Ran a check in test and all tests passed. Ran pump on operating mode and found that the 30 ml and 60 ml BD syringes are not being recognized by the pump. Flashed the syringe pump; confirmed that the version is 9.15.1.2 and confirmed against another syringe pump (same version). Ran pump on operating mode, the syringe is recognizing 30ml BD syringe, but fails to recognize 60ml BD syringe. Next step is to replace the sizer sensor. (TC 01/04/2024 12:52:19 PM, Baghirzade, Nara) Replaced sizer sensor assembly. Barrel clamp calibration completed. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1051156 | 200752 | CM | CMPLT | CMPLT | 2024-01-02 12:15 | 2024-01-02 12:15 | Lens wear and tear | (TC 01/02/2024 12:14:40 PM, Fan, Tak) 306770YP9 replaced by 479877YP3 on Jan 2, 2023.
62985 | 200753 | CM | CMPLT | CMPLT | 2024-01-02 12:45 | 2024-01-02 13:36 | Light appearing orange | (TC 01/02/2024 01:29:34 PM, Rhuland, Courtney) Light is orange when turned on Opened device and replaced intensity switch Issue persisted, resoldered board with help from Samir Issue resolved Functional check - OK
| 200754 | RDS | CMPLT | CMPLT | 2024-01-02 12:52 | 2024-01-02 12:52 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
| 200755 | CM | CMPLT | CMPLT | 2024-01-02 12:55 | 2024-01-02 12:55 | worked on Masimo view stations 16N and 16 Murray | Functional
| 200756 | CM | CMPLT | CMPLT | 2024-01-02 12:56 | 2024-01-02 12:56 | worked on Stryker beds 16 Murray | Checked 2 x beds and requested service from Stryker.
68369 | 200757 | CM | CMPLT | CMPLT | 2024-01-02 14:46 | 2024-01-02 14:50 | O2 measurement failure | (TC 01/02/2024 02:46:37 PM, Scero, Tomlyn) Notified by email that device wasn't working and waiting in the hallway. took the unit into OR 12 for troubleshooting/testing. Confirmed that the issuse is the O2 sensor and the PATO needs to be replaced. We happened to have an extra pato in the shop so I took the device downstairs for repair. Disassembled the unit and replaced the PATO sensor. All device checks passed after reassembly. Self test passed and O2 measurement accurate during ventilation. Returned the AGM to the OR.
1063808 | 200758 | CM | CMPLT | CMPLT | 2024-01-02 16:23 | 2024-01-02 16:24 | MSH 15 SPP PROCEDURE REMOTE VIDEO NOT WORKING | (TC 01/02/2024 04:23:47 PM, Whittington, Scott) MSH 15 SPP PROCEDURE REMOTE VIDEO NOT WORKING REPLACE LONG SVIDEO CABLE FROMLOCAL STOCK PROVIDE STRAIN RELIEF ON US AND TOWER SIDE | MSH 15 SPP PROCEDURE REMOTE VIDEO NOT WORKING
| 200759 | CM | CMPLT | CMPLT | 2024-01-02 16:29 | 2024-01-02 16:30 | MSH CT URGENT REQUEST TO BORROW CONTRAST SOLUTION FROM TWH | (TC 01/02/2024 04:29:17 PM, Whittington, Scott) MSH CT URGENT REQUEST TO BORROW CONTRAST SOLUTION FROM TWH ORGANIZE LOGISTICS DELIVER CONTRAST TO MATT SCHWEITZ | MSH CT URGENT REQUEST TO BORROW CONTRAST SOLUTION FROM TWH
1059638 | 200760 | CM | CMPLT | NFF | 2024-01-03 07:53 | 2024-01-03 07:54 | Error alarm | (TC 01/03/2024 07:52:58 AM, Rhuland, Courtney) Tag stated error alarm Unable to replicate error Functional check - OK No fault found
66453 | 200761 | CM | CMPLT | CMPLT | 2024-01-03 08:04 | 2024-01-09 08:26 | ordering parts for VN500s | pressure regulators for VN500s. received parts. PO#7202342594
1069593 | 200762 | ACC | CMPLT | CMPLT | 2024-01-03 08:20 | 2024-10-15 14:27 | Incoming Inspection
66463 | 200763 | CM | CMPLT | CMPLT | 2024-01-03 10:54 | 2024-01-03 10:54 | unit not passing nebulizer test | replaced both Air and O2 pressure regulators. Calibrated and tested. Functional. Run unit over-night. Passed pre-use check 4 times. Returned back to RT and informed Tom.
| 200764 | CM | CMPLT | CMPLT | 2024-01-03 11:02 | 2024-01-08 14:53 | PATO sensor for stock | (TC 01/03/2024 11:02:37 AM, Scero, Tomlyn) ordering for replacement stock to have in the RT room on hand (TC 01/04/2024 02:56:10 PM, Scero, Tomlyn) price increased. emailing back and forth with draeger to find current price quote. emailed purchasing to approve PO.
| 200765 | CM | CMPLT | CMPLT | 2024-01-03 13:02 | 2024-01-10 09:18 | ordered 2 base storage net | (TC 01/03/2024 01:02:24 PM, Chukri, Samir) emailed Sreyker and got quote for two base storage net and replaced ordered e-req (TC 01/09/2024 09:16:31 AM, Chukri, Samir) Received and installed one on the streacher 17th floor, review with Michelle and Rex .
1055642 | 200766 | CM | CMPLT | CMPLT | 2024-01-03 13:02 | 2024-04-23 09:57 | Touch screens on both side rails don't work | (TC 01/03/2024 01:02:33 PM, Cortez, John) Swapped bed out and put in service request (TC 01/11/2024 03:44:17 PM, Cortez, John) Showed Vince from Hillrom where bed is located
1059828 | 200767 | PM | CMPLT | CMPLT | 2024-01-03 13:41 | 2024-01-05 14:31 | RFU error code 10 | (TC 01/03/2024 02:45:28 PM, Rhuland, Courtney) When cleaning, RFU error code 10 was present Manifld temperature and P1 temperature differentiated from the safety temperature sensors Tried to clear error but still persisted Reaching out to stryker for support (TC 01/05/2024 02:27:49 PM, Baghirzade, Nara) Disassembled the altrix down to the control box. Replaced the temperature probe cable (swapped with a spare unit from emerg). Assembled the altrix back together. Went in to service mode and verified that the temperature on port 1 return is within range of the return manifold and safety sensors. Altrix unit is fully functional. (TC 01/05/2024 02:30:06 PM, Rhuland, Courtney) Disassembled the altrix down to the control box. Replaced the temperature probe cable (swapped with a spare unit from emerg). Assembled the altrix back together. Went in to service mode and verified that the temperature on port 1 return is within range of the return manifold and safety sensors. Altrix unit is fully functional. (TC 01/08/2024 10:22:26 AM, Rhuland, Courtney) Error returned when running - manifold temperature sensor found to be defective as well Replaced with spare from old emerg altrix Verified all temperatures (3 ports, manifold and safety) were within range Ran for an hour with no errors
1049362 | 200768 | CM | CMPLT | CMPLT | 2024-01-03 15:03 | 2024-01-04 14:09 | Contrast spill on side A | (TC 01/03/2024 03:03:23 PM, Debebe, Adam) Removed injector from room Side A piston took apart and cleaned including the housing. (TC 01/04/2024 02:06:38 PM, Debebe, Adam) Injector installed back in room.
1054277 | 200769 | CM | CMPLT | CMPLT | 2024-01-03 15:06 | 2024-01-03 15:07 | Worn lens | (TC 01/03/2024 03:06:37 PM, Fan, Tak) 346321YP3 replaced by 483789YP4
1061205 | 200770 | CM | CMPLT | CMPLT | 2024-01-04 08:38 | 2024-06-17 11:04 | Boom drifting | (TC 01/04/2024 08:38:41 AM, Reid, Emmerson ) I tightened the brakes on the anesthesia boom. The side boom was also drifing, I noticed the handle was activated. The handle had taken damage and was stuck in the depressed position. Note this was the handle for the orange sections. I replaced this handle with one from OR 12. (TC 01/05/2024 03:10:19 PM, Reid, Emmerson ) Asked and recieved a quote for the cover. Ordered the part. (TC 06/17/2024 11:03:32 AM, Reid, Emmerson ) Repaired and installed in OR12
1061211 | 200771 | CM | CMPLT | CMPLT | 2024-01-04 08:49 | 2024-01-04 08:51 | Computer not turning on | (TC 01/04/2024 08:49:47 AM, Reid, Emmerson ) Nurse PC was not turning on. I secured the power cables and reseated it in the PC. After this is came back on.
65926 | 200772 | CM | CMPLT | CMPLT | 2024-01-04 09:21 | 2024-01-04 09:26 | Broken front and rear case; pressure sensor disc error | (TC 01/04/2024 09:21:03 AM, Baghirzade, Nara) Re-seated the pressure sensor disc but could not bypass the error that pops up as soon as the syringe pump is connected to the PCU. Downloaded error log and noted errors pertaining to the pressure sensor disc. Front and rear case were broken. Replaced: -Front case -Rear case -Pressure sensor -Mylar gasket Transferred all parts from old rear case to new case. Transferred display board and pressure sensor holder from old front case to new case. Printed asset sticker and barcode sticker. Completed a full calibration of: -barrel clamp -plunger force -plunger position -pressure sensor disc Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
| 200773 | CM | CMPLT | CMPLT | 2024-01-04 09:58 | 2024-01-04 09:58 | Meeting Biomedical engineering planing and MRI NICU | (TC 01/04/2024 09:58:07 AM, Chukri, Samir) Meeting.
1062999 | 200774 | PM | CMPLT | CMPLT | 2024-01-04 10:19 | 2024-01-04 10:20 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/02/2024 10:19:25 AM, Cortez, John) Completed cleaning cycle
| 200775 | CM | CMPLT | CMPLT | 2024-01-04 10:28 | 2024-02-02 08:45 | misc biomed - January 2024 | (TC 01/02/2024 10:27:44 AM, Cortez, John) Went with Samir to respond to call in 16M about bariatric bed (TC 01/02/2024 03:49:22 PM, Cortez, John) Anita from Endo is asking with the moble vital signs can be swapped with monitors with temp option Informed her we don't have any spares. Mentioned we only have older Welch Allyn models with temp, and handheld thermometers Provided SPO2 sensor ordering info to Anita (TC 01/05/2024 01:13:13 PM, Cortez, John) Endo room 3 monitor not powering on. Checked room and breaker tripped on outlet. Moved power cord back to emergency power. (TC 01/05/2024 01:14:34 PM, Cortez, John) one on one meeting (TC 01/05/2024 04:23:52 PM, Cortez, John) Responded to call from L&D regarding suction tree in OR1, "not enough suction" Looked at it with Gill, seemed to be working normally. No sign of leaks. Nurse unsure if faulty tree was moved. Will mention to Jessica next week. (TC 01/08/2024 10:42:20 AM, Cortez, John) Speech Pathology Lab called about broken microphone. Went with Emmerson to check. Checked continuity with DMM and found out that theres a break also in cable. Mentioned to lab that it's not worth repairing (TC 01/10/2024 01:54:19 PM, Cortez, John) Masimo safetynet cleanup meeting (TC 01/10/2024 09:41:02 AM, Cortez, John) Was told by Iliana from 11N about a non-working, untagged patient lift. Quickly checked lift and confirmed it doesn't work. Forwarded request to Nara (TC 01/16/2024 04:54:12 PM, Cortez, John) systems meeting (TC 01/16/2024 04:56:54 PM, Cortez, John) huddle (TC 01/16/2024 04:57:34 PM, Cortez, John) Masimo safetynet cleanup meeting (TC 01/16/2024 04:57:34 PM, Cortez, John) Masimo safetynet cleanup meeting (TC 01/15/2024 04:58:43 PM, Cortez, John) OR walkthrough with Emmerson (TC 01/15/2024 04:59:15 PM, Cortez, John) Answered Endo call regarding monitor in room1. Not powering on. Power cable was loose. Issue not related to code grey (Intermittent power loss ) (TC 01/15/2024 04:58:43 PM, Cortez, John) OR walkthrough with Emmerson (TC 01/15/2024 09:30:19 AM, Cortez, John) Philips Physio 2 patch/rebuild (TC 01/16/2024 09:31:45 AM, Cortez, John) Philips Physio 2 patch/rebuild con't (TC 01/17/2024 09:32:18 AM, Cortez, John) Rebuild spare Philips server to act as standby (TC 01/18/2024 02:20:19 PM, Cortez, John) Jason from MRI dept requested spare flowmeter. Verified if flowmeter needed to be MRI compatible and Jason said, no. Dropped off flowmeter to MRI reception desk (TC 01/19/2024 02:01:03 PM, Cortez, John) IS meeting (TC 01/19/2024 02:01:03 PM, Cortez, John) IS meeting (TC 01/19/2024 02:03:19 PM, Cortez, John) Replugged and checked in ICU pod c defib (TC 01/23/2024 09:17:27 AM, Cortez, John) Disassembled 4 light sources in Endo 1-4. No signs of fluid in tubing. Tested rooms with scope. (TC 01/25/2024 09:25:24 AM, Cortez, John) Helped Jessica troubleshoot olympus touch monitor issue: "routing not possible because of missing video matrix", in OR 3 (TC 01/26/2024 01:19:55 PM, Cortez, John) Con't OR3 troubleshooting with Jessica and Derrick (Olympus). Coriomatrix firmware installed. OR3 and OR12 operational. Deriick showed process of changing Coriomatrix settings. (TC 01/26/2024 04:12:31 PM, Cortez, John) Answered call from PACU regarding Bay 8 monitor. MSL cable was loose. Reseating resolves issue (TC 01/29/2024 09:45:45 AM, Cortez, John) 10S Masimo viewstation unplugged, plugged it back in (TC 01/29/2024 10:31:11 AM, Cortez, John) Went with Andrew to speak with Anita (Endo), about swapping repaired light sources with loaners (TC 01/30/2024 09:42:30 AM, Cortez, John) Answered OR call regarding lights in OR8, Ipad/controller restarted and needed to be reconnected to network, lights tested ok (TC 01/30/2024 09:44:32 AM, Cortez, John) Updated endo CLV-190 locations spreadsheet and sent to Anita (TC 01/29/2024 11:03:29 AM, Cortez, John) huddle (TC 01/30/2024 01:38:48 PM, Cortez, John) Ivy (interventional radiology) reached out saying they needed a SPO2 trunk cable for the X3 monitor in room 506. Brought up a spare. (TC 01/29/2024 01:46:05 PM, Cortez, John) Installed left and right plastic covers on arms of Endo scope guide receiver dishes in rooms 3 and 4. Waiting on delivery of final 2 covers (TC 01/30/2024 02:11:56 PM, Cortez, John) Left ambulatory light source with Naz (TC 01/30/2024 04:27:34 PM, Cortez, John) OR8 called again about light panel. It disconnected again. Battery at 37%. Replaced charging cable. Lights tested ok. Will check lights again in the morning (TC 01/30/2024 04:33:50 PM, Cortez, John) Stopped by MCR with Kam to help transfer DHCP responsibilities between Physio servers, left early to answer another call (TC 01/31/2024 09:58:52 AM, Cortez, John) Answered 10S call regarding Masimo safetynet alarm. 10-6 Root's Radical was off. Found in room 39 and powered on module, informed nursing station (TC 01/31/2024 09:24:29 AM, Cortez, John) CCU called regarding there overview station. Nurse was unable to access patient history in "review" tab for tele 6. All other telepacks were ok Doesnt seem to be related to Physio being disconnected. Checked physio assignments in system config for that telepack Turns out they were able to view info on surveillance station. Will attempt to restart station later and continue to monitor (TC 01/31/2024 09:28:28 AM, Cortez, John) Tested spare spo2 trunk cables (TC 01/31/2024 09:29:12 AM, Cortez, John) several PIICIXs going to local Turns out DNS settings were incorrect. (final digit of secondary DNS IP set to 8 and not 10) Changed settings in CCU 10th floor and ICU
1061965 | 200776 | CM | CMPLT | CMPLT | 2024-01-04 11:30 | 2024-01-04 11:31 | Could not weigh the patient - not displaying correctly | (TC 01/04/2024 11:29:54 AM, Ferreira, Qwynn) Noticed they did not zero the bed before placing patient causing issues. Patient was lifted, bed was reset and zeroed. Bed was displaying correct weight afterwards.
1056566 | 200777 | CM | CMPLT | CMPLT | 2024-01-04 13:03 | 2024-01-04 13:04 | Failed to report | (TC 01/04/2024 01:02:40 PM, Rhuland, Courtney) Failed to report for >7 days Checked and battery needed replacing Replaced battery and sent status report
| 200778 | CM | CMPLT | CMPLT | 2024-01-04 13:05 | 2024-01-04 13:07 | NICU meeting | (TC 01/04/2024 01:05:54 PM, Rhuland, Courtney) Meeting with Michelle from RT to discuss changing SPO2 histograms on MX800 monitors
65262 | 200779 | CM | CMPLT | CMPLT | 2024-01-04 13:57 | 2024-01-04 13:57 | unit not passing nebulizer test | replaced both Air and O2 pressure regulators. Calibrated and tested. Functional. Run unit over-night. Passed pre-use check 4 times. Returned back to RT and informed Rex.
| 200780 | PA | CMPLT | CMPLT | 2024-01-04 14:12 | 2024-02-09 13:51 | Altrix parts for stock | (TC 01/04/2024 02:27:53 PM, Canjimavadivel, Kamallesh ) Received quote from Stryker for Altrix parts. Placed an order for stock parts (TC 02/09/2024 01:50:39 PM, Canjimavadivel, Kamallesh ) Received all parts, labelled and put on parts shelf
1062393 | 200781 | CM | CMPLT | CMPLT | 2024-01-04 14:22 | 2024-01-04 14:24 | Power cord prongs were bent | (TC 01/04/2024 02:22:26 PM, Baghirzade, Nara) Main power cord prongs were slightly bent; used plyers to bend them straight. Electrical safety results: power cord resistance = 85 mohms leakage current = 115.3 uA Plugged bed to wall outlet; bed battery is charging and bed is fully functional. Followed up with ward clerk who placed the call.
1064418 | 200782 | CM | CMPLT | CMPLT | 2024-01-04 14:22 | 2024-01-04 14:24 | Missing Components | (TC 01/04/2024 02:22:36 PM, Ferreira, Qwynn) Repalced two doors with new black hinges. Repalced door grommets for coaxial cables x2 Removed the old plate secure screw from old MRI and installed into new base plate as they lost the old one (TC 01/04/2024 02:22:36 PM, Ferreira, Qwynn) Repalced two doors with new black hinges. Repalced door grommets for coaxial cables x2 Removed the old plate secure screw from old MRI and installed into new base plate as they lost the old one
63927 | 200783 | CM | CMPLT | CMPLT | 2024-01-04 14:33 | 2024-01-04 14:55 | LL not working | (TC 01/04/2024 02:52:13 PM, Scero, Tomlyn) LL pins on PIM were loose. I tried resoldering them and any other loose pins. Didn't seem to fix the issue. I checked all the leads and they looked ok. Replaced V1-V3 , V2 seemed intermittent. I went to the 6th floor to find a old PIM that I could try. I found one that was working so I tested it on the device. It worked great so I kept it on for now while I look more into the broken PIM.
68780 | 200784 | CM | CMPLT | CMPLT | 2024-01-04 15:56 | 2024-02-14 09:43 | "Doesn't charge" | (TC 01/04/2024 03:55:39 PM, Cortez, John) Seems to charge normally. Will condition battery overnight (TC 02/13/2024 09:14:43 AM, Cortez, John) Replaced battery, Conditioned overnight. Batt capacity: 4169 mAhrs
68044 | 200785 | CM | CMPLT | CMPLT | 2024-01-04 15:59 | 2024-02-14 10:17 | "turns on and off, adjusts setting independently" | (TC 01/03/2024 03:59:00 PM, Cortez, John) Ran test infusion. Encountered system error briefly, but ran without issue for 30 mins otherwise. Didn't notice any setting changes. (TC 02/13/2024 10:16:32 AM, Cortez, John) Ran another test, unable to replicate issue.
1063201 | 200786 | CM | CMPLT | CMPLT | 2024-01-05 08:13 | 2024-01-05 08:13 | battery replacement | (TC 01/05/2024 08:12:56 AM, Rhuland, Courtney) Replaced AAA batteries
69208 | 200787 | CM | CMPLT | CMPLT | 2024-01-05 08:22 | 2024-01-05 08:25 | Detecting air in line | (TC 01/05/2024 08:21:37 AM, Baghirzade, Nara) Ran binary sensors test on ASM multiple times, all binary sensors tests passed. No issue with AIL. Check-in tests passed and completed. IUI connectors and device cleaned. Ran basic infusion at varying rates - no AIL errors. Infusion completed successfully. | L&D staff reported that the pump kept detecting air-in-line despite priming the line
69082 | 200788 | CM | CMPLT | CMPLT | 2024-01-05 08:48 | 2024-01-05 08:50 | Pressure sensor failure | (TC 01/05/2024 08:47:53 AM, Rhuland, Courtney) Replaced bottle side pressure sensor Calibrated pressure and rate Check-in completed All tests passed Refer to ASM for details
67969 | 200789 | CM | CMPLT | CMPLT | 2024-01-05 09:01 | 2024-01-05 09:02 | Broken SW latch | (TC 01/05/2024 09:01:22 AM, Baghirzade, Nara) Replaced SW latch. Checked tilt function - passed. Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74.3%RH Incubator is fully functional.
1050476 | 200790 | CM | CMPLT | CMPLT | 2024-01-05 09:09 | 2024-01-23 10:00 | ordering parts for Servo-Ns | receiveced parts for Servo-Ns
| 200791 | CM | CMPLT | CMPLT | 2024-01-05 11:19 | 2024-01-05 11:19 | worked on TMS | worked on TMS
1063808 | 200792 | CM | CMPLT | CMPLT | 2024-01-05 11:38 | 2024-01-08 14:25 | Probe not being recognized on system | (TC 01/05/2024 11:38:15 AM, Fan, Tak) Probe doesn't recognize in any port. Was just replaced earlier in the week. Some markings on the cord at wheel length. Ordered replacement probe. (TC 01/08/2024 02:24:25 PM, Fan, Tak) Replaced probe.
| 200793 | PA | CMPLT | CMPLT | 2024-01-05 13:54 | 2024-02-06 11:20 | Philips test equipment | (TC 01/05/2024 02:13:46 PM, Canjimavadivel, Kamallesh ) Recevied quote from vendor, placed order for test equipment. (TC 01/26/2024 08:48:32 AM, Canjimavadivel, Kamallesh ) Followed up with Mohawk Medbuy rep for ETA
| 200794 | CM | CMPLT | CMPLT | 2024-01-05 14:27 | 2024-01-05 14:27 | worked on Masimo | worked on Masimo
| 200795 | RDS | CMPLT | CMPLT | 2024-01-05 14:32 | 2024-01-05 14:33 | Weekly rounds: 11 North | (TC 01/05/2024 02:31:58 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 200796 | RDS | CMPLT | CMPLT | 2024-01-05 14:34 | 2024-01-05 14:35 | Weekly Rounds: 11 South | (TC 01/05/2024 02:33:45 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
500857 | 200797 | CM | ACTIV | ISSUE | 2024-01-05 14:44 |  | MSH-5-US RM 3 APLIO XG DEPLOYED AS BACKUP AFTER APLIO 500 POWER FAILURE |  | MSH-5-US RM 3 APLIO XG DEPLOYED AS BACKUP AFTER APLIO 500 POWER FAILURE
531113 | 200798 | CM | CMPLT | CMPLT | 2024-01-05 14:46 | 2024-01-05 14:46 | Not working | (TC 01/05/2024 02:45:35 PM, Rhuland, Courtney) Got a message from Bria stating the device turns on but cannot take measurements Power cycled device and issue resolved Functional check - OK
| 200799 | CM | ACTIV | ISSUE | 2024-01-05 14:47 |  | Training, Learning, and Other Works | (TC 01/08/2024 02:32:43 PM, Ferreira, Qwynn) redesigned the mri cable clamp as there is a better way to position the cable. Samir also requested this to be made Designed from scratch, printed with formlabs as dremel was in use. (TC 01/08/2024 02:33:45 PM, Ferreira, Qwynn) meeting with kam and raayed (TC 01/08/2024 02:51:52 PM, Ferreira, Qwynn) designed a hard model vessel simulation that will allow the IV tube to easily slide over te connectors. Will print tomorrow once the clip is done. (TC 01/09/2024 12:36:55 PM, Ferreira, Qwynn) post processed the aneurysm models printed in ABS. Prepared them for the chemical testing (TC 01/09/2024 12:37:33 PM, Ferreira, Qwynn) Went to canadian tire with emmerson to pick up acetone, xylene, batteries, and lacquer thinner No xylene. (TC 01/09/2024 12:38:12 PM, Ferreira, Qwynn) Printed the hollow aneurysm models in clear and finalized the design with the nozzle attachments Printed halfway before failing. Cleaned printer but not the tray yet. Resent the print with new tray to see. (TC 01/09/2024 12:39:25 PM, Ferreira, Qwynn) Headneck meeting Met a doctor interested in 3d printing a model (TC 01/10/2024 07:30:26 AM, Ferreira, Qwynn) meeting with nicu leadership (TC 01/10/2024 07:30:55 AM, Ferreira, Qwynn) meeting with naz to discuss the mri incbuator issues (TC 01/11/2024 07:31:21 AM, Ferreira, Qwynn) redesigned the MRI clamp as it was a snug fit. Reprinted. Also post processed the blood vessel simulation model and handed over to research team (TC 01/10/2024 07:31:21 AM, Ferreira, Qwynn) redesigned the MRI clamp as it was a snug fit. Reprinted. Also post processed the blood vessel simulation model and handed over to research team (TC 01/10/2024 07:32:45 AM, Ferreira, Qwynn) meeting with adeel and maisha for documentation meeting (TC 01/10/2024 07:33:20 AM, Ferreira, Qwynn) worked on the case for january 16th, doctor from uhn interested in our lab to test out. Set up the files and started segmentation (TC 01/12/2024 02:58:22 PM, Ferreira, Qwynn) cybersecurity meeting (TC 01/12/2024 02:58:43 PM, Ferreira, Qwynn) meeting with nima and kamran on capitol for 3d printing (TC 01/12/2024 02:59:04 PM, Ferreira, Qwynn) segmented the face model and sent for printing. Print is going to take 20hrs. (TC 01/15/2024 09:13:49 AM, Ferreira, Qwynn) biomed meeting meeting with ortho to discuss allograft images and image pickup (TC 01/15/2024 09:14:38 AM, Ferreira, Qwynn) post processed the three models that are for a case on tuesday. 2 SLA 1 FDM (TC 01/15/2024 09:15:45 AM, Ferreira, Qwynn) Modelled more for the aneurysm Started modelling for the case feb 7 for the media (TC 01/15/2024 09:18:27 AM, Ferreira, Qwynn) set up models for Dr John ameida handed them over as well as instructed him on the protocol for sterilization (TC 01/16/2024 03:47:54 PM, Ferreira, Qwynn) had meeting with 12s PCM to discuss issues on the unit relating to biomed or purchasing Reached out to becky as some issues pertained to the mattresses on the bed Needed thermometers (portable) (TC 01/16/2024 03:48:49 PM, Ferreira, Qwynn) worked on ortho case Completing the segmentation Overlaying the models in blender and slicer Found a way to reconstruct the DICOM (TC 01/18/2024 01:32:10 PM, Ferreira, Qwynn) admin, emails, and follow ups for orders (TC 01/18/2024 01:32:36 PM, Ferreira, Qwynn) installed new MRI clips for the coil cables to rest in. These remove some of the stress in the cable from bending. (TC 01/18/2024 01:33:12 PM, Ferreira, Qwynn) Worked on the femur case to practice creating template guides for fitment and process replication Managed to get a good way to match anatomy with a painted guide (TC 01/24/2024 04:21:42 PM, Ferreira, Qwynn) Had meeting with Ortho to discuss the 3d model and make minor adjustments. Finalized one guide for the femur and will print as we will test it in the skills lab. Worked on the mandible case setting up the fibula segments, and orientation as well as worked on the two other cases for segmentations Met with Joel and Karl to review the plan. Made suggestions based on my research and they are moving ahead with plan Showed Kam the basics of 3D slicer (TC 01/25/2024 03:16:19 PM, Ferreira, Qwynn) Modelled the femur guide as revisons were needed to open the face of the guide Ortho team mentioned to have a two peice guide to try and make it fit Worked on Mandible reconstruction for guides Swapped prints and post processed prints for surgical practice in the skills lab (TC 01/25/2024 03:19:06 PM, Ferreira, Qwynn) Meeting with media to capture the 3D printing process and take videos and photos. Worked with the team to provide any supplies or imaging (TC 01/25/2024 03:19:54 PM, Ferreira, Qwynn) Worked on the Femur guide some more as splitting the guide will be challenging. Will continue to work on this. Received the imaging for the case next week thursday and handed to kam and showed him the process for setting up the models in 3D slicer Post processed the mandible for the media as well to aid in filming (TC 01/29/2024 03:40:37 PM, Ferreira, Qwynn) Set up new printer with Kam Segmented models for Case Thursday Placed the general fibula segments Printed the AVM Models for post processing (TC 01/29/2024 03:48:04 PM, Ferreira, Qwynn) Post processed the femur guides Had a meeting with Aaron to discuss the model and next steps Will finish the guides tomorrow and finalize them (TC 01/30/2024 07:39:08 AM, Ferreira, Qwynn) Finished the models for Thursdays case and set up the surgical plan Introduced Kam to Fusion and Meshmixer for modelling - started the heel pricker Post processed some guides and models for Femur case (TC 01/30/2024 07:40:30 AM, Ferreira, Qwynn) received new imaging for next weeks case as the patient will be using their scapula rather than fibula Explained the case to Kam Re printed femur guide as one section broke Infomred ortho group that the guide will need reprinting (TC 01/30/2024 07:41:34 AM, Ferreira, Qwynn) Finalized the Guides for Thursdays case and printed them Showed Kam how to set up the printer and slicing software for basic knowledge Removed femur guides and polished them Will need to write brief report for the case on thursday (TC 01/31/2024 09:26:43 AM, Ferreira, Qwynn) Post processed the guides for Joels case. Had meeting with ortho team for todays case to test fit the guides and go through some brief explanations Worked with Kam on the scapula to show him the basic work flow (TC 01/31/2024 09:29:46 AM, Ferreira, Qwynn) Worked on post processingthe model of the heel pricker with Kam Preped models for hand over to MDRD Gave one model to Jon for his students Recevied parts from amazon for printing (TC 01/31/2024 09:31:28 AM, Ferreira, Qwynn) MRI Incubator meeting with NICU team (TC 02/02/2024 03:28:13 PM, Ferreira, Qwynn) Worked on case for Joel which is on Monday Helped Kam with the heel pricker for NICU Short meeting with NICU staff for the heel pricker study (TC 02/02/2024 03:28:58 PM, Ferreira, Qwynn) Worked on the case for Joel on monday Visted OR 5 - twice First time for confirmation of guides and model in OR as well as test fitting guides and any questions that arise Second was for photos and feedback relating to the guides created and Dr. Kim Tsoi feedback (TC 02/02/2024 03:30:24 PM, Ferreira, Qwynn) Worked on AVM model and sent to printer with two filaments ABS and PVA and had to configure settings to allow this to work. Finished first set of guides for the case monday, will need to print over the weekend. (TC 02/05/2024 10:27:36 AM, Ferreira, Qwynn) Worked on finishing the guides for wednesdays case Post processed the recon mandible for wednesday Provided Karl with both models one for patient and one for plate bending (TC 02/05/2024 10:28:47 AM, Ferreira, Qwynn) Worked on guides for wednesday case - finihsed the chin, fibula, and ramus - need to label them still Swapped vent with Samir on MRI incubator Worked with Kam on the Heel pricker device (TC 02/05/2024 10:29:56 AM, Ferreira, Qwynn) Meeting with Adeel, maisha and kam for Pharmacy (TC 02/06/2024 08:18:14 AM, Ferreira, Qwynn) Post processed guides for Case wednesday Handed guides over to MDRD Made report for wednesdays case Discussion with Nima regarding regulations for guides, software, etc (TC 02/07/2024 08:20:10 AM, Ferreira, Qwynn) resident rounds for head and neck (TC 02/09/2024 08:46:29 AM, Ferreira, Qwynn) Met with joel to work on tuesdays mandible case. made the plan for the case Obtained imaging for an ortho case feb 22 (TC 02/08/2024 08:46:29 AM, Ferreira, Qwynn) Met with joel to work on tuesdays mandible case. made the plan for the case Obtained imaging for an ortho case feb 22 (TC 02/08/2024 09:09:14 AM, Ferreira, Qwynn) Worked on the guides for tuesdays case - Joel agreed on cuts, need to create the guide base for printing (TC 02/08/2024 08:46:29 AM, Ferreira, Qwynn) made the plan for the case, created segmentations and had the models set up for meeting with Joel Obtained imaging for an ortho case feb 22 Worked on AVM model Repaired printer as PVA was causing issues (TC 02/07/2024 10:16:11 AM, Ferreira, Qwynn) Viewed the surgery for the media case that Joel and Karl performed. Brought kam to view the surgery as well (TC 02/07/2024 10:16:59 AM, Ferreira, Qwynn) Met with Aaron and Arieh to discuss a potential next case. Discussed a potenital research paper for the use of guides in Ortho (TC 02/12/2024 03:04:53 PM, Ferreira, Qwynn) General Meeting (TC 02/12/2024 03:05:10 PM, Ferreira, Qwynn) Materialise meeting (TC 02/12/2024 03:05:27 PM, Ferreira, Qwynn) made report for joel Reprinted ramus guide as first one didnt work Post process the guides and handed over to MDRD (TC 02/15/2024 11:49:41 AM, Ferreira, Qwynn) Head and neck meeting (TC 02/13/2024 11:49:41 AM, Ferreira, Qwynn) Head and neck meeting (TC 02/13/2024 11:50:19 AM, Ferreira, Qwynn) Segmented pelvis case for ortho Guided kam in 3d slicer to show him how to do some different tasks. (TC 02/13/2024 11:51:41 AM, Ferreira, Qwynn) Discussion with joel for the case Met with Aaron for the hip case to discuss possible cuts Setup printer for AVM model and sent pelvis print (TC 02/14/2024 11:52:27 AM, Ferreira, Qwynn) Meeting with Aaron and Aieh for the ortho case both pelvis and tibia Discussed the surgical plan (TC 02/14/2024 11:53:17 AM, Ferreira, Qwynn) Segmented tibia case fully Post processed the pelvis model Helped kam with the tumor margins for the left tibia case Sent tibia print Fixed the bambu printer as if had a clog from the PVA - working well now (TC 02/15/2024 11:55:01 AM, Ferreira, Qwynn) Moved desks - to kamrans area Organized working area (TC 02/15/2024 11:55:33 AM, Ferreira, Qwynn) Post processed tibia print Removed FDM pelvis print Helped kam with left tibia margins again and showed him how to export and import into meshmixer (TC 02/16/2024 02:32:12 PM, Ferreira, Qwynn) Meeting with Aaron to go over the pelvis guide and discused the improvements for the model Created new cut margins and made the plan (TC 02/16/2024 02:32:56 PM, Ferreira, Qwynn) post processed the aneurysm models Picked up new gas for the shop Prepped the aneurysm models for pick up (TC 02/16/2024 02:33:38 PM, Ferreira, Qwynn) Sent the pelvis guide to print for the meeting with Aaron Post processed the guide and test fitted as well as measured the guide for accuracy - OK (TC 02/16/2024 02:34:27 PM, Ferreira, Qwynn) Located the SDS for the formalbs resins and other chemicals in the shop Printed and started to attach to the folder (TC 02/16/2024 02:35:05 PM, Ferreira, Qwynn) Cleaned the shop and eye wash areas for easier access and helped kam organize the carts in the shop. (TC 02/20/2024 11:45:35 AM, Ferreira, Qwynn) Finialized the pelvis guide and had ortho review the guide. Post processed previous guide version to see any improvements Showed emmerson the new printer and the general idea to setup the prints Updated the MSDS book with new chemicals and started working on the excel sheet (TC 02/20/2024 11:46:55 AM, Ferreira, Qwynn) 3d printing meeting (TC 02/20/2024 11:47:09 AM, Ferreira, Qwynn) Ortho meeting to review the guide (TC 02/21/2024 02:30:53 PM, Ferreira, Qwynn) Worked on Tibia case to create the cut angles and measurements for the case Worked with Kam and Emmerson to go through the printing steps and setup of te files in slicer Worked with kam and emmerson on the AVM model for setup and the purpose and workflow for the creation of the AVM models (TC 02/22/2024 02:32:04 PM, Ferreira, Qwynn) Met with Aaron for the tibia case for cuts and measurements and decided on general guide format (TC 02/21/2024 02:32:41 PM, Ferreira, Qwynn) Worked with Samir on the transport incident that occured. Met with michelle to get her persepctive on the incident Met with more of the clinical team including Rex to decide on where to position the tanks for transport. (TC 02/22/2024 02:33:46 PM, Ferreira, Qwynn) finialized the guide for the tibia case on monday Sent to printer Post processed the models for the tibia case. Showed emmerson how to setup the bambu printer and send print files from our laptop to the printer via MicroSD (TC 02/21/2024 02:32:04 PM, Ferreira, Qwynn) Met with Aaron for the tibia case for cuts and measurements and decided on general guide format (TC 02/22/2024 04:33:07 PM, Ferreira, Qwynn) Viewed the case for the pelvis, received photos and provided some support with emmerson Finalized report and sent to Dr.Wunder worked with emmerson on the AVM model Post processed one set of guides for mondays case. (TC 02/23/2024 02:02:13 PM, Ferreira, Qwynn) quality meeting with maisha and adeel (TC 02/23/2024 02:02:32 PM, Ferreira, Qwynn) one on one (TC 02/23/2024 02:02:53 PM, Ferreira, Qwynn) Worked with kam to show him the placenta stand and how to navigate within fusion for that project Worked with emmerson on blender for the aneurysm model and AVM model within the shop (TC 02/26/2024 03:36:24 PM, Ferreira, Qwynn) Worked on aneurysm research to find a suitable result for the curing of the elastic 50A. Printed two models over the weekend and washed in IPA for 10 min and then rinsed with warm water. Cured the model in the chamber with water. The mode was submerged and cured for 70deg for 30min. Model came out yellow ish so will try to reduce the curing temp for the next batch. The waxy texture was removed though so new process for rinsing elastic models (TC 02/26/2024 03:40:33 PM, Ferreira, Qwynn) Visited the OR, guides fit well but the tumor margins changed, so some modifications required Will wait to hear from Jay (TC 02/26/2024 03:42:37 PM, Ferreira, Qwynn) Went through the workflow of the mimics software and determined some things great program - some flaws - not completely ironed out. Will attempt a case. (TC 02/26/2024 03:44:40 PM, Ferreira, Qwynn) Obtained more SDS files for the filametns and added chemicals Worked with Emmerson on the AVM stuff to see what we can try next Set up more prints with emmerson for the AVM to cast tomorrow Set up te sonicator for the shop can be used to clean parts, tools, and others (TC 02/27/2024 02:25:42 PM, Ferreira, Qwynn) Worked with emmerson on the printer to repair as it was jammed Showed him the general process for removing the nozzle and cleaning a jam in the print head Added SDS to T drive for printing - more need to be added (TC 02/27/2024 02:27:15 PM, Ferreira, Qwynn) Discussion with Samir and Adeel on the use of chemicals and strong odour that can come from it (TC 02/27/2024 02:27:45 PM, Ferreira, Qwynn) Meeting with Adeel and Maisha on the safety/quality documentation (TC 02/27/2024 02:28:08 PM, Ferreira, Qwynn) Post processed the aneurysm models and soaked in water for 20 min with 60deg C temp Worked well - slight yellowing, maybe less time in the bath/chamber Handed over model to ortho and discussed basic plan - more to discuss as complicated case and started to print new model for cutting/simulation of the case (TC 02/28/2024 10:36:52 AM, Ferreira, Qwynn) Research related to the chemicals produced from 3D printing Discussion with Samir to what we should be doing for safety and what guidelines hed like to have implemented Added SDS to the t drive under 3d printing for all chemicals used in the 3D lab (TC 02/28/2024 10:38:35 AM, Ferreira, Qwynn) Meeting with Zeev for the aneurysm model Tested the aneurym to see if the catheter can be passed through without sticking Met with joel to discuss more cases for april as well as the two cases in march (TC 02/28/2024 10:39:59 AM, Ferreira, Qwynn) Cleaned up the shop with emmerson and labelled all chemicals and made sure the 3D printing equipment is labelled and well placed Organized other sections and moved chemicals to correct location. Emmerson added the printer to TMS (TC 02/28/2024 10:54:21 AM, Ferreira, Qwynn) General Meeting with Nima on Zoom MRI Incubator update meeting (TC 02/29/2024 10:10:50 AM, Ferreira, Qwynn) meeting with ortho for the left tibia case. Discussed the cuts and how they want to include the fibula model as well (TC 02/29/2024 10:11:56 AM, Ferreira, Qwynn) created new model for ortho which included the fibula and pinned the model for printing. Made two versions one for FDM and one for SLA which is hollow Set up both printers (TC 03/01/2024 10:15:50 AM, Ferreira, Qwynn) created a rules/guidelines poster for the 3d printers Looked online for ideas Started to create a safety presentation as well for the hazards of 3dprinting and how we are addressing those concerns (TC 02/29/2024 10:17:15 AM, Ferreira, Qwynn) showed emmerson some tools and tricks within blender Showed kam different workflows for the FDM printer and the slicing software to improve his printing (TC 03/04/2024 03:49:07 PM, Ferreira, Qwynn) Monday huddle (TC 03/04/2024 03:49:24 PM, Ferreira, Qwynn) Worked on the edits for the 3d printing safety meeting (TC 03/04/2024 03:49:43 PM, Ferreira, Qwynn) Dyed model for the left tibia case for dr wunder (TC 03/04/2024 03:50:11 PM, Ferreira, Qwynn) Picked up CD for Joels case, contained CTA (TC 03/04/2024 03:50:40 PM, Ferreira, Qwynn) created rough segmentations for joel case march 14. Created mandible and fibula models Set up for the case for planning (TC 03/04/2024 03:51:27 PM, Ferreira, Qwynn) Ordered parts for the bambu printer Looked into some iso13485 courses (TC 03/04/2024 03:53:45 PM, Ferreira, Qwynn) discussed with emmerson the aneurysm model Discsused him practicing resections for a previous ortho case Gave him 2-3 weeks to create the plan and include his thoughts on it. will work with him so he understands Transferred over data from USB to theNAS (TC 03/05/2024 11:05:24 AM, Ferreira, Qwynn) 3D Printting Meeting (TC 03/05/2024 11:05:42 AM, Ferreira, Qwynn) head/neck meeting (TC 03/05/2024 03:42:33 PM, Ferreira, Qwynn) 3D Printing meeting and OHS review (TC 03/05/2024 03:43:21 PM, Ferreira, Qwynn) Worked on Ortho case for left tibia Added planes and aligned them Met with Aaron regarding the cuts/placement Went through imaging for the case to varfiy resected the model for better visualization (TC 03/05/2024 03:45:25 PM, Ferreira, Qwynn) Set up fusion for Kam on the PC the license was hung so reinstalled and signed in the license Helped emmerson with an old femur case to practice the resection and what are key concepts. For him to research and look into what he thinks is suitable and his idea of guide creation (TC 03/12/2024 08:31:40 AM, Ferreira, Qwynn) WFH Worked on joels case for mandible Predictive plate holes for the mandible Sent print Sent print at home as well to verfiy guide templates (TC 03/11/2024 08:32:57 AM, Ferreira, Qwynn) Worked on joel mandible case MAde a mock resection plan for ortho and separated the parts for Jay Moified the cutting guide based on aarons recommendations - more work needed still (TC 03/14/2024 03:37:58 PM, Ferreira, Qwynn) worked on joels case for segmentations Worked on ortho case for setting up the allograft and making cuts Sent print for the allograft Continued to work on the case (TC 03/14/2024 03:39:11 PM, Ferreira, Qwynn) saw joels case in the OR worked well Met with aaron from ortho about the tibia case Showed him the prints Discussed new plan Meeting with Gerry/Adeel/Nima
| 200800 | CM | CMPLT | CMPLT | 2024-01-05 14:48 | 2024-01-05 14:48 | 1 on 1 meeting | (TC 01/05/2024 02:47:36 PM, Rhuland, Courtney) 1 on 1 meeting with Adeel
| 200801 | RDS | CMPLT | CMPLT | 2024-01-05 15:09 | 2024-01-05 15:10 | Weekly Rounds: NICU | (TC 01/05/2024 03:08:42 PM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure for results.
| 200802 | RDS | CMPLT | CMPLT | 2024-01-05 15:10 | 2024-01-05 15:11 | Weekly Rounds: RESUS | (TC 01/05/2024 03:10:11 PM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure for results.
| 200803 | CM | ACTIV | ISSUE | 2024-01-05 15:21 |  | MSH-Biomed Strategy for new Philips NBP Inlet design in new products | (TC 01/05/2024 03:21:17 PM, Chukri, Samir) Meeting (TC 04/03/2024 09:56:13 AM, Chukri, Samir) Received up grades parts and put them on the shelf and aske Maisha to add the price for the future use.
1069614 | 200804 | ACC | CMPLT | CMPLT | 2024-01-07 14:10 | 2024-01-16 11:21 | PROBE PLI-2004BX i24LX8 FSD23Z4496 ACCEPTANCE | (TC 01/16/2024 11:21:12 AM, Whittington, Scott) PROBE PLI-2004BX i24LX8 FSD23Z4496 ACCEPTANCE | PROBE PLI-2004BX i24LX8 FSD23Z4496 ACCEPTANCE
1069615 | 200805 | ACC | CMPLT | CMPLT | 2024-01-07 14:13 | 2024-01-16 11:25 | PROBE ACC FSD23Y6547 PVT-781VTE 11C3 | (TC 01/16/2024 11:25:41 AM, Whittington, Scott) PROBE ACC FSD23Y6547 PVT-781VTE 11C3 | PROBE FSD23Y6547 PVT-781VTE 11C3
1069616 | 200806 | ACC | CMPLT | CMPLT | 2024-01-07 14:15 | 2024-09-25 10:42 | PROBE FSD99G23Z3774 PVT-712BT 11MMC4 PED HEAD | (TC 01/16/2024 11:46:13 AM, Whittington, Scott) PROBE acceptance FSD99G23Z3774 PVT-712BT 11MMC4 PED HEAD | PROBE FSD99G23Z3774 PVT-712BT 11MMC4 PED HEAD
1069617 | 200807 | ACC | CMPLT | CMPLT | 2024-01-07 14:17 | 2024-01-16 11:49 | PROBE FSC23Z4201 PVI-574BX i10CX1 | (TC 01/16/2024 11:49:38 AM, Whittington, Scott) PROBE ACC FSC23Z4201 PVI-574BX i10CX1 | PROBE ACC FSC23Z4201 PVI-574BX i10CX1
1069618 | 200808 | ACC | CMPLT | CMPLT | 2024-01-07 14:18 | 2024-01-16 11:53 | PROBE FSC2355694 PLT-1202 BT 17LH7 HOCKEY STICK | (TC 01/16/2024 11:53:00 AM, Whittington, Scott) PROBE ACC FSC2355694 PLT-1202 BT 17LH7 HOCKEY STICK | PROBE ACC FSC2355694 PLT-1202 BT 17LH7 HOCKEY STICK
1069619 | 200809 | ACC | CMPLT | CMPLT | 2024-01-07 15:17 | 2024-01-16 11:20 | PROBE 99G23Z3773 PVT-712BT 11MMC4 PED HEAD | (TC 01/17/2024 11:20:01 AM, Whittington, Scott) PROBE ACC 99G23Z3773 PVT-712BT 11MMC4 PED HEAD (TC 01/16/2024 11:20:01 AM, Whittington, Scott) PROBE ACC 99G23Z3773 PVT-712BT 11MMC4 PED HEAD | PROBE ACC 99G23Z3773 PVT-712BT 11MMC4 PED HEAD
1069620 | 200810 | ACC | CMPLT | CMPLT | 2024-01-07 15:18 | 2024-01-16 11:23 | PROBE FSD23Z5774 PLI-1205BX i18LX5 | (TC 01/17/2024 11:22:50 AM, Whittington, Scott) PROBE ACC FSD23Z5774 PLI-1205BX i18LX5 (TC 01/16/2024 11:22:50 AM, Whittington, Scott) PROBE ACC FSD23Z5774 PLI-1205BX i18LX5 | PROBE ACC FSD23Z5774 PLI-1205BX i18LX5
1069621 | 200811 | ACC | CMPLT | CMPLT | 2024-01-07 15:21 | 2024-01-16 11:40 | PROBE 1069621 FSC23Z4199 PVI-574BX i10CX1 | (TC 01/16/2024 11:40:35 AM, Whittington, Scott) PROBE 1069621 ACC FSC23Z4199 PVI-574BX i10CX1 | PROBE 1069621 ACC FSC23Z4199 PVI-574BX i10CX1
1069622 | 200812 | ACC | CMPLT | CMPLT | 2024-01-07 15:26 | 2024-01-17 11:44 | PROBE 1069622 200812 FSD23Z5775 PLI-1205BX i18LX5 | (TC 01/17/2024 11:43:50 AM, Whittington, Scott) PROBE 1069622 200812 ACC FSD23Z5775 PLI-1205BX i18LX5 | PROBE 1069622 200812 ACC FSD23Z5775 PLI-1205BX i18LX5
1069623 | 200813 | ACC | CMPLT | CMPLT | 2024-01-07 15:28 | 2024-01-17 11:46 | PROBE 1069623 FSC23Z4200 PVI-574BX i10CX1 | (TC 01/17/2024 11:46:01 AM, Whittington, Scott) PROBE 1069623 ACC FSC23Z4200 PVI-574BX i10CX1 | PROBE 1069623 ACC FSC23Z4200 PVI-574BX i10CX1
| 200814 | CM | ACTIV | ISSUE | 2024-01-07 16:23 |  | MSH 5 US UHN ASSET 10653076 APLIO 500 BEING USED AT MSH; NO POWER | (TC 01/06/2024 04:23:22 PM, Whittington, Scott) MSH 5 US UHN ASSET 10653076 APLIO 500 BEING USED AT MSH; NO POWER | MSH 5 US UHN ASSET 10653076 APLIO 500 BEING USED AT MSH; NO POWER
65932 | 200815 | CM | CMPLT | CMPLT | 2024-01-08 07:40 | 2024-01-08 07:42 | Reported that "its not connecting" | (TC 01/08/2024 07:40:35 AM, Baghirzade, Nara) Checked IUIs on both sides; no damage or bent pins. Downloaded error and event log; did not note any errors/events pertaining to connectivity or issues. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
| 200816 | RDS | CMPLT | CMPLT | 2024-01-08 08:18 | 2024-01-08 10:41 | Weekly rounds for ED
1055578 | 200817 | CM | CMPLT | CMPLT | 2024-01-08 08:23 | 2024-02-02 13:50 | battery is not charging | received 2 x masimo roots. PO#7202344981
| 200818 | RDS | CMPLT | CMPLT | 2024-01-08 08:24 | 2024-01-12 14:42 | Weekly rounds for ICU
| 200819 | RDS | CMPLT | CMPLT | 2024-01-08 08:33 | 2024-01-08 11:16 | Weekly checks in CCU | (TC 01/08/2024 11:15:19 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
1053378 | 200820 | CM | CMPLT | CMPLT | 2024-01-08 08:42 | 2024-01-11 09:17 | ordering batteries for Servo-Ns | received Servo-N batteries. PO# 7202345116
502089 | 200821 | RQ | CMPLT | CMPLT | 2024-01-08 09:25 | 2024-01-11 14:11 | Symbia 2 light curtain error Det#2 | (TC 01/11/2024 02:09:30 PM, Green, Paul) Parts were received. Replaced LVA on head 2. Light curtain now functions. Closed up detector and tested again, ok. | -unable to contour
| 200822 | RDS | CMPLT | CMPLT | 2024-01-08 09:41 | 2024-01-08 09:41 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
1046035 | 200823 | CM | CMPLT | CMPLT | 2024-01-08 10:13 | 2024-01-08 10:13 | Upright bucky squeaking on lateral movement. |  | Readjusted bottom cover.Tested OK.
| 200824 | CM | CMPLT | CMPLT | 2024-01-08 10:50 | 2024-01-08 10:50 | worked on TMS | worked on TMS
| 200825 | CM | CMPLT | CMPLT | 2024-01-08 10:53 | 2024-01-08 10:53 | worked on Masimo | worked on Masimo
| 200826 | CM | CMPLT | CMPLT | 2024-01-08 10:54 | 2024-01-08 10:54 | worked with Rex from RT | worked with Rex from RT
| 200827 | CM | CMPLT | CMPLT | 2024-01-08 11:32 | 2024-01-08 11:34 | PFT Syringe Calibration | (TC 01/08/2024 11:31:39 AM, Baghirzade, Nara) Gaby performed regular and He-CO cal; all passed. Saved to configuration. Reloaded expair and did SVC cal - also passed. DLCO cal was inconsistent; first test was high (~174), second cal was closer to the desired number (0.6). Switched syringes and third and fourth test was inconsistent. Measured against the bodybox; bodybox passed all cals even with a slight leak. Gaby also showed me how to do the calibration in preparation for the troubleshooting with Medisoft on Jan 11.
| 200828 | CM | CMPLT | CMPLT | 2024-01-08 13:12 | 2024-01-11 13:35 | spare parts for Zeus IE | (TC 01/08/2024 01:14:36 PM, Scero, Tomlyn) Ordering parts for ZEUS IE. test and AGS valves (TC 01/11/2024 01:34:58 PM, Scero, Tomlyn) Order received.
| 200829 | CM | CMPLT | CMPLT | 2024-01-08 13:33 | 2024-01-08 13:37 | Doppler check | (TC 01/08/2024 01:34:50 PM, Rhuland, Courtney) Replaced 2 cables on dopplers Replaced battery on one Functional check - OK
1059828 | 200830 | PM | CMPLT | CMPLT | 2024-01-08 13:38 | 2024-01-08 13:44 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/08/2024 01:37:39 PM, Rhuland, Courtney) Disinfected, rinsed and refilled with sterile water Returned to service
1050088 | 200831 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
1058056 | 200832 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
1058057 | 200833 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
1058058 | 200834 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
1058059 | 200835 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68367 | 200836 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68368 | 200837 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68369 | 200838 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68370 | 200839 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68371 | 200840 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68372 | 200841 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68373 | 200842 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68374 | 200843 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68375 | 200844 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68376 | 200845 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68377 | 200846 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68378 | 200847 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68537 | 200848 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68542 | 200849 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68543 | 200850 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68544 | 200851 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68545 | 200852 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68546 | 200853 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68547 | 200854 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68552 | 200855 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68553 | 200856 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68554 | 200857 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
68555 | 200858 | CM | CMPLT | CMPLT | 2024-01-08 13:44 | 2024-01-08 13:44 | SW Upgraded to 2.03 TSB #19 |  | SW Upgrade 2.03 done by Salur from Draeger.
66706 | 200859 | CM | CMPLT | CMPLT | 2024-01-08 13:47 | 2024-01-08 13:54 | Broken door hinge and missing the smaller inner door | (TC 01/08/2024 01:47:33 PM, Baghirzade, Nara) Replaced door hinge. Took the spare inner door from biomed. When testing the tilt function, the bed was not locking and was moving freely. Replaced the brake pad - bed was still not locking in position. Replaced the tilt rod/ball - bed was still not locking in position. Re-seated the entire tilt mechanism. Replaced the white plastic where the ball sits into. Tilt is now fully functional and the bed locks and doesn't move at any angle. During functional check, the RH sensor was calibrated twice as it was out of range (more than 4%). Functional tests passed and completed. Scale is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74.8 %RH Incubator is fully functional. (TC 01/08/2024 01:53:26 PM, Rhuland, Courtney) Helped troubleshoot the tilt mechanism issue.
1067169 | 200860 | CM | CMPLT | CMPLT | 2024-01-08 14:01 | 2024-01-08 14:06 | Microphone broken | (TC 01/08/2024 02:02:33 PM, Reid, Emmerson ) Went with John to check the issue. The microphone wire had ripped out. The unit did have a spare microphone they installed. We tried to repair with sodder but were not able to because of another break somewhere along the cable. We let the unit know we could not repair.
| 200861 | RDS | CMPLT | CMPLT | 2024-01-08 14:07 | 2024-01-08 14:11 | Weekly Checks For PACU | (TC 01/08/2024 02:09:37 PM, Reid, Emmerson ) Rounds complete. Also checked some nurse call settings between OR and PACU with Kamran
| 200862 | CM | CMPLT | CMPLT | 2024-01-08 14:11 | 2024-01-08 14:11 | training | E-learning
1053391 | 200863 | CM | CMPLT | CMPLT | 2024-01-08 14:23 | 2024-01-08 14:24 | QC before/after patient use | (TC 01/08/2024 02:23:05 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 124 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S (Sri - CNS)
| 200864 | PA | CMPLT | CMPLT | 2024-01-08 14:25 | 2024-01-10 12:29 | Ordering sterile water for NICU altrix | (TC 01/08/2024 02:33:10 PM, Baghirzade, Nara) SC 1003142098 (TC 01/10/2024 12:28:52 PM, Baghirzade, Nara) Order # 4150269849
1069625 | 200865 | CM | ACTIV | ISSUE | 2024-01-08 14:42 |  | MSH20 DOVIGI SPORTS PLI-705 BX i11LX3 FSD23Z6280 |  | MSH20 DOVIGI SPORTS PLI-705 BX i11LX3 FSD23Z6280
1069626 | 200866 | CM | ACTIV | ISSUE | 2024-01-08 14:45 |  | MSH 20 DOVIG SPORTS PLI-1205BX i18LX5 FSD23Z5776 |  | MSH 20 DOVIG SPORTS PLI-1205BX i18LX5 FSD23Z5776
1069627 | 200867 | CM | ACTIV | ISSUE | 2024-01-08 14:46 |  | MSH20 DOVIGI PVI-475 BX i8CX1 FSD2383632 |  | MSH20 DOVIGI PVI-475 BX i8CX1 FSD2383632
65878 | 200868 | CM | CMPLT | CMPLT | 2024-01-09 07:58 | 2024-01-09 07:58 | Fluids not associating with iAware | (TC 01/09/2024 07:57:44 AM, Baghirzade, Nara) Downloaded error log; did not note any errors on the pump. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. | Came down with PCU 66019
66019 | 200869 | CM | CMPLT | CMPLT | 2024-01-09 08:00 | 2024-01-09 08:10 | Fluids not associating on iAware | (TC 01/09/2024 07:59:56 AM, Baghirzade, Nara) Checked the network status, its associated and connected to mshmed. Checked the server status and confirmed that it is connected to the correct server. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 50 mohms Leakage current: 91.0 uA | Came down with syringe pump 65878
| 200870 | CM | CMPLT | CMPLT | 2024-01-09 09:49 | 2024-03-13 13:22 | Altrix temperature probe cable order | (TC 01/09/2024 09:53:50 AM, Rhuland, Courtney) Ordered atrix temperature probe cables for stock
1063746 | 200871 | PM | CMPLT | CMPLT | 2024-01-09 10:07 | 2024-01-09 10:08 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/09/2024 10:06:45 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1057962 | 200872 | CM | CMPLT | CMPLT | 2024-01-09 10:36 | 2024-01-09 10:37 | Update membrane change interval | (TC 01/09/2024 10:36:29 AM, Rhuland, Courtney) Updated membrane change date to 31 days
64276 | 200873 | CM | CMPLT | CMPLT | 2024-01-09 12:19 | 2024-01-09 12:20 | Histogram Update Policy | (TC 01/09/2024 12:19:35 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64296 | 200874 | CM | CMPLT | CMPLT | 2024-01-09 12:20 | 2024-01-09 12:20 | Histogram Update Policy | (TC 01/09/2024 12:20:15 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64203 | 200875 | CM | CMPLT | CMPLT | 2024-01-09 12:21 | 2024-01-09 12:21 | Histogram Update Policy | (TC 01/09/2024 12:20:46 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64295 | 200876 | CM | CMPLT | CMPLT | 2024-01-09 12:21 | 2024-01-09 12:21 | Histogram Update Policy | (TC 01/09/2024 12:21:17 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64292 | 200877 | CM | CMPLT | CMPLT | 2024-01-09 12:22 | 2024-01-09 12:22 | Histogram Update Policy | (TC 01/09/2024 12:21:51 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64301 | 200878 | CM | CMPLT | CMPLT | 2024-01-09 12:22 | 2024-01-09 12:23 | Histogram Update Policy | (TC 01/09/2024 12:22:26 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64206 | 200879 | CM | CMPLT | CMPLT | 2024-01-09 12:23 | 2024-01-09 12:23 | Histogram Update Policy | (TC 01/09/2024 12:22:59 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64204 | 200880 | CM | CMPLT | CMPLT | 2024-01-09 12:23 | 2024-01-09 12:24 | Histogram Update Policy | (TC 01/09/2024 12:23:32 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64263 | 200881 | CM | CMPLT | CMPLT | 2024-01-09 12:24 | 2024-01-09 12:24 | Histogram Update Policy | (TC 01/09/2024 12:24:10 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64282 | 200882 | CM | CMPLT | CMPLT | 2024-01-09 12:25 | 2024-01-09 12:25 | Histogram Update Policy | (TC 01/09/2024 12:25:01 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
1053347 | 200883 | CM | CMPLT | CMPLT | 2024-01-09 12:26 | 2024-01-30 09:11 | MSH BI 1270 H1 CHAO REPORTED SCREW FELL OUT, MAYBE FROM SYSTEM? | (TC 01/09/2024 12:25:44 PM, Whittington, Scott) MSH BI 1270 H1 CHAO REPORTED SCREW FELL OUT, MAYBE FROM SYSTEM? ALSO P/U 2 BOXES PARTS FOR RETURN TO NOOR, LEFT IN RAD REVIEW. P/U QWERTY KEYBOARD, FOR RWS 2 (TC 01/12/2024 10:41:46 AM, Whittington, Scott) MSH BI 1270 H1 CHAO REPORTED SCREW FELL OUT, MAYBE FROM SYSTEM? (TC 01/12/2024 10:41:58 AM, Chau, Ken) MSH BI 1270 H1 CHAO REPORTED SCREW FELL OUT, MAYBE FROM SYSTEM? (TC 01/12/2024 10:42:20 AM, Fan, Tak) MSH BI 1270 H1 CHAO REPORTED SCREW FELL OUT, MAYBE FROM SYSTEM? (TC 01/12/2024 10:41:46 AM, Whittington, Scott) MSH BI 1270 H1 CHAO REPORTED SCREW FELL OUT, MAYBE FROM SYSTEM? SCREW WAS FROM COMPRESSION PADDLE HOLDER REPLACEMENT COMPRESSION PADDLE HOLDER ORDERED | MSH BI 1270 H1 CHAO REPORTED SCREW FELL OUT, MAYBE FROM SYSTEM? FROM COMPRESSION PADDLE HOLDER
64250 | 200884 | CM | CMPLT | CMPLT | 2024-01-09 12:26 | 2024-01-09 12:26 | Histogram Update Policy | (TC 01/09/2024 12:25:48 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
1049532 | 200885 | CM | CMPLT | CMPLT | 2024-01-09 12:26 | 2024-01-09 12:27 | Histogram Update Policy | (TC 01/09/2024 12:26:37 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64299 | 200886 | CM | CMPLT | CMPLT | 2024-01-09 12:27 | 2024-01-09 12:28 | Histogram Update Policy | (TC 01/09/2024 12:27:20 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64255 | 200887 | CM | CMPLT | CMPLT | 2024-01-09 12:28 | 2024-01-09 12:29 | Histogram Update Policy | (TC 01/09/2024 12:28:28 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64265 | 200888 | CM | CMPLT | CMPLT | 2024-01-09 12:29 | 2024-01-09 12:29 | Histogram Update Policy | (TC 01/09/2024 12:29:03 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64256 | 200889 | CM | CMPLT | CMPLT | 2024-01-09 12:30 | 2024-01-09 12:30 | Histogram Update Policy | (TC 01/09/2024 12:29:40 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64207 | 200890 | CM | CMPLT | CMPLT | 2024-01-09 12:30 | 2024-01-09 12:30 | Histogram Update Policy | (TC 01/09/2024 12:30:17 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
67214 | 200891 | CM | CMPLT | CMPLT | 2024-01-09 12:31 | 2024-01-09 12:31 | Histogram Update Policy | (TC 01/09/2024 12:30:58 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64254 | 200892 | CM | CMPLT | CMPLT | 2024-01-09 12:31 | 2024-01-09 12:32 | Histogram Update Policy | (TC 01/09/2024 12:31:31 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64262 | 200893 | CM | CMPLT | CMPLT | 2024-01-09 12:32 | 2024-01-09 12:32 | Histogram Update Policy | (TC 01/09/2024 12:32:09 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64289 | 200894 | CM | CMPLT | CMPLT | 2024-01-09 12:33 | 2024-01-09 12:33 | Histogram Update Policy | (TC 01/09/2024 12:32:42 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64281 | 200895 | CM | CMPLT | CMPLT | 2024-01-09 12:33 | 2024-01-09 12:34 | Histogram Update Policy | (TC 01/09/2024 12:33:21 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64275 | 200896 | CM | CMPLT | CMPLT | 2024-01-09 12:34 | 2024-01-09 12:34 | Histogram Update Policy | (TC 01/09/2024 12:34:14 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64269 | 200897 | CM | CMPLT | CMPLT | 2024-01-09 12:35 | 2024-01-09 12:35 | Histogram Update Policy | (TC 01/09/2024 12:34:54 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
68392 | 200898 | CM | CMPLT | CMPLT | 2024-01-09 13:47 | 2024-01-09 14:00 | RT/AA said the sevo module wasn't passing self test | (TC 01/09/2024 01:58:01 PM, Scero, Tomlyn) Self test was done and passed. Functional tests completed with no issues. DUMA flow sensor working correctly and SEVO delivery as expected. No error messages occured.
| 200899 | CM | CMPLT | CMPLT | 2024-01-09 13:49 | 2024-01-09 13:49 | worked on TMS | cleaning and updating TMS
| 200900 | CM | CMPLT | CMPLT | 2024-01-09 13:50 | 2024-01-09 13:50 | worked with RT | worked with RT staff
| 200901 | CM | CMPLT | CMPLT | 2024-01-09 13:51 | 2024-01-09 13:51 | training | e-learning
64274 | 200902 | CM | CMPLT | CMPLT | 2024-01-09 14:20 | 2024-01-09 14:20 | Histogram Update Policy | (TC 01/09/2024 02:19:48 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64267 | 200903 | CM | CMPLT | CMPLT | 2024-01-09 14:22 | 2024-01-09 14:25 | Histogram Update Policy | (TC 01/09/2024 02:25:03 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64205 | 200904 | CM | CMPLT | CMPLT | 2024-01-09 14:26 | 2024-01-09 14:26 | Histogram Update Policy | (TC 01/09/2024 02:25:42 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64264 | 200905 | CM | CMPLT | CMPLT | 2024-01-09 14:27 | 2024-01-09 14:27 | Histogram Update Policy | (TC 01/09/2024 02:26:42 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
1066942 | 200906 | CM | CMPLT | CMPLT | 2024-01-09 14:27 | 2024-01-09 14:31 | Not connecting to screen when plugged in and turned on | (TC 01/09/2024 02:27:13 PM, Canjimavadivel, Kamallesh ) Screen input was not correct, changed input to SDI and camera showed up on screen. Called RT and informed them
67211 | 200907 | CM | CMPLT | CMPLT | 2024-01-09 14:28 | 2024-01-09 14:28 | Histogram Update Policy | (TC 01/09/2024 02:27:36 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64268 | 200908 | CM | CMPLT | CMPLT | 2024-01-09 14:28 | 2024-01-09 14:29 | Histogram Update Policy | (TC 01/09/2024 02:28:02 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64278 | 200909 | CM | CMPLT | CMPLT | 2024-01-09 14:29 | 2024-01-09 14:29 | Histogram Update Policy | (TC 01/09/2024 02:28:59 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64290 | 200910 | CM | CMPLT | CMPLT | 2024-01-09 14:30 | 2024-01-09 14:30 | Histogram Update Policy | (TC 01/09/2024 02:29:43 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
1067933 | 200911 | CM | CMPLT | CMPLT | 2024-01-09 14:33 | 2024-01-09 14:34 | V lead off meassge from monitor and no RESP | (TC 01/09/2024 02:32:37 PM, Canjimavadivel, Kamallesh ) Leads needed to be connected to vitals bridge box and not SimMan directly. Used zip ties to tidy up excess cords
64294 | 200912 | CM | CANCL | DUP | 2024-01-09 14:46 |  | Histogram update
64291 | 200913 | CM | CMPLT | CMPLT | 2024-01-09 14:49 | 2024-01-09 14:49 | Histogram Update Policy | (TC 01/09/2024 02:49:03 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64302 | 200914 | CM | CMPLT | CMPLT | 2024-01-09 14:53 | 2024-01-09 14:53 | Histogram Update Policy | (TC 01/09/2024 02:52:40 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64247 | 200915 | CM | CMPLT | CMPLT | 2024-01-09 14:53 | 2024-01-09 14:54 | Histogram Update Policy | (TC 01/09/2024 02:53:21 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64298 | 200916 | CM | CMPLT | CMPLT | 2024-01-09 14:54 | 2024-01-09 14:54 | Histogram Update Policy | (TC 01/09/2024 02:54:03 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64272 | 200917 | CM | CMPLT | CMPLT | 2024-01-09 14:55 | 2024-01-09 14:55 | Histogram Update Policy | (TC 01/09/2024 02:54:42 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64277 | 200918 | CM | CMPLT | CMPLT | 2024-01-09 14:55 | 2024-01-09 14:55 | Histogram Update Policy | (TC 01/09/2024 02:55:12 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64277 | 200919 | CM | CANCL | DUP | 2024-01-09 14:56 | 2024-01-09 14:56 | Histogram Update Policy
64259 | 200920 | CM | CMPLT | CMPLT | 2024-01-09 14:56 | 2024-01-09 14:57 | Histogram Update Policy | (TC 01/09/2024 02:56:30 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64207 | 200921 | CM | CANCL | DUP | 2024-01-09 14:57 | 2024-01-09 14:57 | Histogram Update Policy
64302 | 200922 | CM | CMPLT | CMPLT | 2024-01-09 14:59 | 2024-01-09 14:59 | Histogram Update Policy | (TC 01/09/2024 02:58:50 PM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
| 200923 | PA | CMPLT | CMPLT | 2024-01-09 14:59 | 2024-01-11 08:12 | Tools order | (TC 01/09/2024 03:04:41 PM, Canjimavadivel, Kamallesh ) Placed order for tools (TC 01/10/2024 03:00:00 PM, Canjimavadivel, Kamallesh ) Received tools
67209 | 200924 | CM | CMPLT | CMPLT | 2024-01-09 14:59 | 2024-01-09 14:59 | Histogram Update Policy | (TC 01/09/2024 02:59:12 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
| 200925 | CM | CMPLT | CMPLT | 2024-01-09 15:45 | 2024-01-09 15:45 | NICU Huddle | (TC 01/09/2024 03:45:09 PM, Baghirzade, Nara) NICU huddle
69541 | 200926 | CM | CMPLT | CMPLT | 2024-01-09 15:46 | 2024-01-16 09:19 | Reported that it disconnects and shuts off mid-infusion | (TC 01/09/2024 03:46:22 PM, Baghirzade, Nara) Ran the LVP for more than 2 hours on left IUI; pump did not disconnect or encounter any issues. Changes rates mid-infusion but it did not cause any issues. Ran the LVP for more than 2 hours on the right IUI; same outcome as the left IUI. No issues encountered. PCU was also functional. (TC 01/16/2024 09:19:11 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. | LVP was sent down with PCU (1050252)
1050252 | 200927 | CM | CMPLT | CMPLT | 2024-01-09 15:48 | 2024-01-16 09:20 | Pump disconnected and shut off mid-infusion | (TC 01/09/2024 03:48:24 PM, Baghirzade, Nara) Downloaded error and event logs; did not note anything unusual. Ranthe LVP for more than 2 hours on left IUI; pump did not disconnect or encounter any issues. Changes rates mid-infusion but it did not cause any issues. Ran the LVP for more than 2 hours on the right IUI; same outcome as the left IUI. No issues encountered. (TC 01/16/2024 09:19:48 AM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 84 mohms Leakage current: 61.1 uA | Came down with LVP (69541)
67212 | 200928 | CM | CMPLT | CMPLT | 2024-01-10 07:10 | 2024-01-09 14:11 | Histogram Update Policy | (TC 01/09/2024 02:09:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
67213 | 200929 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:19 | Histogram Update Policy | (TC 01/09/2024 02:19:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
67215 | 200930 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:21 | Histogram Update Policy | (TC 01/10/2024 02:20:38 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
67210 | 200931 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:24 | Histogram Update Policy | (TC 01/09/2024 02:24:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64253 | 200932 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:25 | Histogram Update Policy | (TC 01/09/2024 02:25:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64252 | 200933 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:26 | Histogram Update Policy | (TC 01/09/2024 02:26:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64248 | 200934 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:27 | Histogram Update Policy | (TC 01/09/2024 02:27:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64297 | 200935 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:28 | Histogram Update Policy | (TC 01/09/2024 02:28:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64270 | 200936 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:30 | Histogram Update Policy | (TC 01/09/2024 02:29:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64271 | 200937 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:31 | Histogram Update Policy | (TC 01/09/2024 02:30:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64288 | 200938 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:32 | Histogram Update Policy | (TC 01/09/2024 02:31:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64293 | 200939 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:33 | Histogram Update Policy | (TC 01/09/2024 02:32:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64287 | 200940 | CM | CMPLT | CMPLT | 2024-01-10 07:19 | 2024-01-09 14:34 | Histogram Update Policy | (TC 01/09/2024 02:33:00 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
1061203 | 200941 | CM | CMPLT | CMPLT | 2024-01-10 07:53 | 2024-01-10 08:01 | DVI port on multi AUX not working | (TC 01/09/2024 07:53:36 AM, Reid, Emmerson ) DVI port on the foot boom was damaged. I replaced it with a new one and tested the scope. Working now.
67216 | 200942 | CM | CMPLT | CMPLT | 2024-01-10 08:00 | 2024-01-10 08:01 | Histogram Update Policy | (TC 01/10/2024 08:00:27 AM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
67217 | 200943 | CM | CMPLT | CMPLT | 2024-01-10 08:01 | 2024-01-10 08:01 | Histogram Update Policy | (TC 01/10/2024 08:01:01 AM, Baghirzade, Nara) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64284 | 200944 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 08:47 | Histogram Update Policy | (TC 01/10/2024 08:46:57 AM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64251 | 200945 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 08:48 | Histogram Update Policy | (TC 01/10/2024 08:47:32 AM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64280 | 200946 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 08:48 | Histogram Update Policy | (TC 01/10/2024 08:48:03 AM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64300 | 200947 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 08:49 | Histogram Update Policy | (TC 01/10/2024 08:48:45 AM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64257 | 200948 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 08:49 | Histogram Update Policy | (TC 01/10/2024 08:49:14 AM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64260 | 200949 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 08:51 | Histogram Update Policy | (TC 01/10/2024 08:49:50 AM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64283 | 200950 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 12:06 | Histogram Update Policy | (TC 01/10/2024 12:05:52 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64202 | 200951 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 12:07 | Histogram Update Policy | (TC 01/10/2024 12:06:23 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64285 | 200952 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 13:52 | Histogram Update Policy | (TC 01/10/2024 01:51:42 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64273 | 200953 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 13:53 | Histogram Update Policy | (TC 01/10/2024 01:52:38 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64286 | 200954 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 13:53 | Histogram Update Policy | (TC 01/10/2024 01:53:11 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64249 | 200955 | CM | CMPLT | CMPLT | 2024-01-10 08:44 | 2024-01-10 14:47 | Histogram Update Policy | (TC 01/10/2024 02:45:55 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
67000 | 200956 | CM | CMPLT | CMPLT | 2024-01-10 09:42 | 2024-01-10 09:55 | Battery replacement | (TC 01/10/2024 09:41:57 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66910 | 200957 | CM | CMPLT | CMPLT | 2024-01-10 09:42 | 2024-01-10 09:56 | Battery replacement | (TC 01/10/2024 09:55:30 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66880 | 200958 | CM | CMPLT | CMPLT | 2024-01-10 09:42 | 2024-01-10 09:56 | Battery replacement | (TC 01/10/2024 09:56:01 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66896 | 200959 | CM | CMPLT | CMPLT | 2024-01-10 09:42 | 2024-01-10 09:57 | Battery replacement | (TC 01/10/2024 09:56:31 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66944 | 200960 | CM | CMPLT | CMPLT | 2024-01-10 09:42 | 2024-01-10 09:58 | Battery replacement | (TC 01/10/2024 09:58:00 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66908 | 200961 | CM | CMPLT | CMPLT | 2024-01-10 09:42 | 2024-01-10 09:59 | Battery replacement | (TC 01/10/2024 09:58:33 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66912 | 200962 | CM | CMPLT | CMPLT | 2024-01-10 10:01 | 2024-01-10 10:02 | Battery replacement | (TC 01/10/2024 10:01:13 AM, Baghirzade, Nara) Updated account/location for asset. Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
502089 | 200963 | RQ | CMPLT | CMPLT | 2024-01-10 10:10 | 2024-01-11 14:12 | Symbia 2 gantry frozen, service message | (TC 01/10/2024 02:10:52 PM, Green, Paul) Performed a complete system shutdown. Problem resolved. | Symbia 2 gantry frozen, service message
67006 | 200964 | CM | CMPLT | CMPLT | 2024-01-10 10:12 | 2024-01-10 10:12 | Battery replacement | (TC 01/10/2024 10:11:47 AM, Baghirzade, Nara) Updated account/location for asset. Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1067334 | 200965 | CM | CMPLT | CMPLT | 2024-01-10 10:12 | 2024-01-17 15:11 | Broken power adapter | (TC 01/10/2024 11:21:47 AM, Canjimavadivel, Kamallesh ) Prong broken from adapter, reached out to vendor for quote. Also installed backup batteries so scale can still be used (TC 01/10/2024 02:01:52 PM, Canjimavadivel, Kamallesh ) Received quote, placed order for adapter (TC 01/12/2024 10:25:58 AM, Canjimavadivel, Kamallesh ) Placed order again as there was error in request on SAP (TC 01/17/2024 03:10:20 PM, Canjimavadivel, Kamallesh ) Received adapter, installed on scale
66900 | 200966 | CM | CMPLT | CMPLT | 2024-01-10 10:14 | 2024-01-10 10:14 | Battery replacement | (TC 01/10/2024 10:13:35 AM, Baghirzade, Nara) Updated account/location for asset. Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1046303 | 200967 | CM | CMPLT | CMPLT | 2024-01-10 10:14 | 2024-01-10 10:15 | Battery replacement | (TC 01/10/2024 10:14:22 AM, Baghirzade, Nara) Updated account/location for asset. Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
| 200968 | CM | CMPLT | CMPLT | 2024-01-10 10:15 | 2024-01-10 10:16 | Copied most updated configuration for NICU MX800 and X2 | (TC 01/10/2024 10:15:20 AM, Baghirzade, Nara) Copied the most updated MX800 and X2 configuration to biomed laptop using the Philips Support Tool. | As per Kamran's request
1049023 | 200969 | CM | CMPLT | CMPLT | 2024-01-10 10:37 | 2024-04-02 13:50 | request a replacement unit | (TC 02/07/2024 01:51:48 PM, Huang, Jessica) SN 190267716 is replaced by SN 230452744 in Feb 2024
| 200970 | CM | CMPLT | CMPLT | 2024-01-10 11:53 | 2024-01-10 11:54 | Tested all EEG leads | (TC 01/10/2024 11:53:31 AM, Baghirzade, Nara) Tested 36 EEG leads for continuity. Found 2 defective leads. Followed up with Vinod and dropped off EEG amplifier along with all the leads.
64279 | 200971 | CM | CMPLT | CMPLT | 2024-01-10 12:29 | 2024-01-10 12:30 | Histogram Update Policy | (TC 01/10/2024 12:29:31 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64294 | 200972 | CM | CMPLT | CMPLT | 2024-01-10 12:35 | 2024-01-10 12:37 | Histogram Update Policy | (TC 01/10/2024 12:36:27 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
1062259 | 200973 | CM | CMPLT | CMPLT | 2024-01-10 13:16 | 2024-05-28 13:30 | Bed frame and mattress burned | (TC 01/10/2024 01:16:11 PM, Baghirzade, Nara) Patient lit a lighter and caused the mattress to catch fire. The rail also sustained damage. Moved the bed to 12th fl bed shop. Created service request with Stryker. (TC 01/10/2024 02:46:45 PM, Baghirzade, Nara) Met with Mark inspect the bed. He will be ordering a new rail as well as a fowler motor for the head portion of the bed (very loud when moving it up/down). Mark to follow up with Melissa and then let me know when he will be coming back on site. (TC 01/12/2024 09:57:06 AM, Baghirzade, Nara) Attached SAFER report submitted by Security. (TC 05/28/2024 01:30:10 PM, Baghirzade, Nara) Bed rail was replaced. Attached service report. | Bed was involved in a fire incident
| 200974 | CM | CMPLT | CMPLT | 2024-01-10 13:33 | 2024-04-12 09:53 | order three ECG simulator PR-SL-8 | (TC 01/10/2024 01:39:51 PM, Chukri, Samir) emailed and got a quotation replaced order. (TC 04/12/2024 09:52:20 AM, Chukri, Samir) ordered and Received three ECG
| 200975 | CM | CMPLT | CMPLT | 2024-01-10 13:39 | 2024-01-10 13:43 | Suction regulator INT not working correctly | (TC 01/10/2024 01:39:24 PM, Canjimavadivel, Kamallesh ) Verified regulator working, and determined issue with seal between canister and liner. Informed nurse to swap liner and try again | ED Annex
1054741 | 200976 | CM | CMPLT | CMPLT | 2024-01-10 13:47 | 2024-01-10 13:47 | Histogram Update Policy | (TC 01/10/2024 01:46:49 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
64258 | 200977 | CM | CMPLT | CMPLT | 2024-01-10 13:48 | 2024-01-10 13:49 | Histogram Update Policy | (TC 01/10/2024 01:48:19 PM, Rhuland, Courtney) MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy. Checked the histogram cursor and confirmed that the lower limit is 90% and higher limit is 95%.
66768 | 200978 | CM | CMPLT | CMPLT | 2024-01-10 14:02 | 2024-01-10 14:04 | BP not working | (TC 01/10/2024 02:02:23 PM, Scero, Tomlyn) Replaced BP pump and front cover assembly from an X2 that was removed from service. BP leak test ok. Passed functional tests. Put on spare cart for now.
| 200979 | CM | CMPLT | CMPLT | 2024-01-10 14:03 | 2024-01-10 14:03 | worked with getinge tech | worked with Getinge tech
| 200980 | CM | CMPLT | CMPLT | 2024-01-10 14:04 | 2024-01-10 14:04 | training | E-learning
| 200981 | CM | CMPLT | CMPLT | 2024-01-10 14:49 | 2024-01-11 10:13 | Nitrogen tank changes on cryo machine | (TC 01/10/2024 02:48:45 PM, Baghirzade, Nara) Elena requested the 2 tanks (nitrogen) on the cryo machine. Replaced only one tank as the second one wouldnt fit, it was the wrong size tank. (TC 01/10/2024 02:51:14 PM, Rhuland, Courtney) Elena requested the 2 tanks (nitrogen) on the cryo machine. Replaced only one tank as the second one wouldnt fit, it was the wrong size tank. (TC 01/11/2024 10:12:07 AM, Baghirzade, Nara) Correct tank was brought in by Elena. Replaced tank and turn the tanks on. Tested that it works. (TC 01/11/2024 10:12:33 AM, Rhuland, Courtney) Correct tank was brought in by Elena. Replaced tank and turn the tanks on. Tested that it works.
66914 | 200982 | CM | CMPLT | CMPLT | 2024-01-10 14:53 | 2024-01-16 07:47 | Fails battery upgrade | (TC 01/10/2024 02:53:10 PM, Baghirzade, Nara) Unable to perform battery upgrade on the Radical 7. When putting in a new battery along with upgrade SD card, the unit does not turn on. Tried with 3 different batteries and SD cards, none work. Will send to Masimo for replacement/repair. (TC 01/16/2024 07:45:25 AM, Baghirzade, Nara) Samuel from Masimo advised to do the upgrade on an old battery first and then on the new battery. Followed his instructions and successfully completed the battery upgrade. Radical now has a new battery and it has been upgraded. Followed up with Samuel regarding the results.
| 200983 | CM | CMPLT | CMPLT | 2024-01-10 14:54 | 2024-01-10 14:56 | Otoscopes not working | (TC 01/10/2024 02:54:29 PM, Rhuland, Courtney) 2 otoscopes from ENT saying the lights were not working Replaced battery on one as it was >2 years old Second one battery was new so brought it down to biomed and charged battery Confirmed with DMM that battery was still good
| 200984 | CON | CMPLT | CMPLT | 2024-01-10 15:56 | 2024-01-10 15:56 | MSH OPG SPP MOVE ULTRASOUND SYSTEMS TO REASSIGNED ROOMS 3rd Floor | (TC 01/10/2024 03:55:47 PM, Whittington, Scott) MSH OPG SPP MOVE ULTRASOUND SYSTEMS TO REASSIGNED ROOMS 3rd Floor (TC 01/10/2024 03:55:57 PM, Dolinar, John) MSH OPG SPP MOVE ULTRASOUND SYSTEMS TO REASSIGNED ROOMS 3rd Floor | MSH OPG SPP MOVE ULTRASOUND SYSTEMS TO REASSIGNED ROOMS 3rd Floor Hi Scott, We need to transfer two ultrasound machines next week please. A Phillips machine needs to be moved from E hallway (3-520) to C hallway (3-535). The GE from room 3-535 needs to be moved from C hallway to D hallway (3-525). Let me know if this is possible next week. Thank you! Andrea
| 200985 | CM | CMPLT | CMPLT | 2024-01-10 16:01 | 2024-09-25 09:39 | MSH OPG SPP DATA CONNECTION TO HSC | (TC 01/10/2024 04:01:32 PM, Dolinar, John) MSH OPG SPP DATA CONNECTION TO HSC (TC 06/14/2024 11:48:39 AM, Whittington, Scott) MSH OPG SPP DATA CONNECTION TO HSC (TC 07/30/2024 11:38:01 AM, Whittington, Scott) MSH OPG SPP DATA CONNECTION TO HSC ATTEMPT TO OBTAIN VOLUSON FOR TESTING UNABLE TO, DUE TO CLINICAL LOAD SETUP TEAMS MEET FOR MSH & HSC (TC 06/14/2024 11:48:00 PM, Whittington, Scott) MSH OPG SPP DATA CONNECTION TO HSC | MSH OPG SPP DATA CONNECTION TO HSC
1046035 | 200986 | CM | CMPLT | CMPLT | 2024-01-10 16:09 | 2024-01-10 16:09 | Water spilt on clerical computer keyboard. |  | Replace the keyboard.
1053348 | 200987 | CM | CMPLT | CMPLT | 2024-01-10 16:14 | 2024-01-11 14:51 | MSH BI H3 BIOPSY DECK LOCK STIFF | (TC 01/10/2024 04:14:34 PM, Whittington, Scott) MSH BI H3 BIOPSY DECK LOCK STIFF | MSH BI H3 BIOPSY DECK LOCK STIFF
| 200988 | CM | CMPLT | CMPLT | 2024-01-11 07:58 | 2024-01-11 07:59 | Histogram policy update on all NICU MX800 Monitors | (TC 01/11/2024 07:57:49 AM, Baghirzade, Nara) Completed the MX800 histogram lower changed from 91% to 90%. Reflects the most up-to-date SpO2 policy.
| 200989 | CM | CMPLT | CMPLT | 2024-01-11 08:33 | 2024-01-16 11:25 | Ording Neptune Filters | (TC 01/11/2024 08:34:58 AM, Reid, Emmerson ) Placed an order for more Neptune Filters (TC 01/16/2024 11:24:25 AM, Reid, Emmerson ) Recieved
1066475 | 200990 | CM | CMPLT | CMPLT | 2024-01-11 08:36 | 2024-01-11 08:38 | Filter needed to be changed | (TC 01/11/2024 08:36:40 AM, Reid, Emmerson ) Went and checked all of the Neptune Filters. This one needed to be replaced. Working now.
62971 | 200991 | CM | CMPLT | CMPLT | 2024-01-11 09:36 | 2024-01-11 09:38 | Bar won't move over incubator | (TC 01/11/2024 09:36:07 AM, Rhuland, Courtney) Loosened bolt to allow bar to move Orange light was present when turned on and bumped Opened device and replaced high/low switch Issue resolved and light intensity - OK
1056583 | 200992 | CM | CMPLT | CMPLT | 2024-01-11 10:47 | 2024-01-11 10:49 | Failed to check in | (TC 01/11/2024 10:47:10 AM, Rhuland, Courtney) Device had not checked in for over a week Tried to send data over wifi but could device could not find access point Both access points were listed Turned device off and on again and was able to send report.
1060462 | 200993 | CM | CMPLT | CMPLT | 2024-01-11 10:48 | 2024-01-11 10:53 | Arm loose, can't be tightened | (TC 01/10/2024 10:48:13 AM, Cortez, John) Set screw stuck, can't tighten. Lubricating threads fixes issue. Locking pin broken on receiver dish arm. Unlikely, but need to look into if it can be ordered. Otherwise, may need to order entire arm. In the meantime, tightened arm and used tape to prevent pin from falling out.
| 200994 | CM | CMPLT | CMPLT | 2024-01-11 12:14 | 2024-01-26 12:46 | Ordering incubator parts | (TC 01/26/2024 12:45:57 PM, Rhuland, Courtney) Unboxed and put away all parts received
| 200995 | CM | CMPLT | CMPLT | 2024-01-11 14:05 | 2024-01-11 14:05 | worked with Getinge tech | worked with Getinge tech
| 200996 | CM | CMPLT | CMPLT | 2024-01-11 14:06 | 2024-01-11 14:06 | Worked on Masimo | Worked on Masimo
| 200997 | CM | CMPLT | CMPLT | 2024-01-11 14:07 | 2024-01-11 14:07 | training | E-learning
1053348 | 200998 | CM | CMPLT | CMPLT | 2024-01-11 14:13 | 2024-02-26 13:19 | MSH BI 1272 H3 BIOPSY DECK LEFT HAND LOCK HANDLE VERY DIIFICULT TO RELEASE | (TC 01/12/2024 11:02:19 AM, Whittington, Scott) MSH BI 1272 H3 BIOPSY DECK LEFT HAND LOCK HANDLE VERY DIIFICULT TO RELEASE NOOR & RYAN ON SITE (TC 01/12/2024 11:02:50 AM, Chau, Ken) MSH BI 1272 H3 BIOPSY DECK LEFT HAND LOCK HANDLE VERY DIIFICULT TO RELEASE NOOR & RYAN ON SITE (TC 01/12/2024 11:03:20 AM, Fan, Tak) MSH BI 1272 H3 BIOPSY DECK LEFT HAND LOCK HANDLE VERY DIIFICULT TO RELEASE NOOR & RYAN ON SITE (TC 02/23/2024 08:16:49 PM, Whittington, Scott) MSH BI 1272 H3 BIOPSY DECK LEFT HAND LOCK HANDLE VERY DIIFICULT TO RELEASE REPLACEMENT BIOPSY DECK RECEIVED | MSH BI 1272 H3 BIOPSY DECK LEFT HAND LOCK HANDLE VERY DIIFICULT TO RELEASE BIOPSY DECK: STLC-00004 SN# 82001193037 SYSTEM SN# 3DM160700414 REF# 01432903 SPOKE TO LARRY IN TECH SUPPORT LCUGINI@HOLOGIC.COM RECORDS INDICATE THER IS NO SERVICE CONRACT ON THE AFFIRM BIOPSY DECK
1057506 | 200999 | CM | CMPLT | CMPLT | 2024-01-11 14:52 | 2024-01-11 15:08 | ECG not working | (TC 01/11/2024 02:51:53 PM, Canjimavadivel, Kamallesh ) Went to dental room and tested with simulator, ECG working normal. Emailed unit
| 201000 | OC | CMPLT | CMPLT | 2024-01-11 17:58 | 2024-01-12 07:52 | On-call Jan 5-11, 2024 | (TC 01/11/2024 05:58:43 PM, Canjimavadivel, Kamallesh ) Received call from 15 Antenatal at 5:40 pm. Reported Masimo computer alarming with message client down 12S. Explained to them alarm is related to 12S computer being down. Requested them to call 12S and ask them to turn it back on. Called locating to connect with 12S as well, line was busy. Logged in remotely after some time to verify 12S station back online.
1060032 | 201001 | CM | CMPLT | CMPLT | 2024-01-12 10:17 | 2024-01-17 14:16 | Ordering patient valve | (TC 01/12/2024 02:36:50 PM, Baghirzade, Nara) SC 1003146083 (TC 01/17/2024 02:15:22 PM, Baghirzade, Nara) PO 7202347771
32520 | 201002 | CM | CMPLT | CMPLT | 2024-01-12 11:37 | 2023-12-13 09:01 | MSH DR2 COLLIMATOR REPLACED DUE TO CAN NETWORK COMMUNICATION FAULTS | (TC 12/11/2023 10:00:00 PM, Whittington, Scott) MSH DR2 COLLIMATOR REPLACED DUE TO CAN NETWORK COMMUNICATION FAULTS | MSH DR2 COLLIMATOR REPLACED DUE TO CAN NETWORK COMMUNICATION FAULTS
32520 | 201003 | CM | CMPLT | CMPLT | 2024-01-12 11:43 | 2023-12-19 09:03 | MSH DR2 TUBE HEAD CONTROLS DEFECTIVE | (TC 12/18/2023 09:02:00 PM, Whittington, Scott) MSH DR2 TUBE HEAD CONTROLS DEFECTIVE REPLACE | MSH DR2 TUBE HEAD CONTROLS DEFECTIVE
32510 | 201004 | CM | CMPLT | CMPLT | 2024-01-12 11:50 | 2023-12-21 09:47 | MSH DR1 GUARD CARD REPLACE/CAL | (TC 12/20/2023 08:40:00 PM, Whittington, Scott) MSH DR1 TUBE GUARD CARD REPLACE/CAL | MSH DR1 GUARD CARD REPLACE/CAL
32520 | 201005 | CM | CMPLT | CMPLT | 2024-01-12 12:13 | 2024-01-08 09:30 | MSH DR2 DETECTOR & HOUSING | (TC 01/02/2024 11:28:00 PM, Whittington, Scott) MSH DR2 DETECTOR & HOUSING | MSH DR2 DETECTOR & HOUSING
32510 | 201006 | CM | CMPLT | CMPLT | 2024-01-12 12:14 | 2024-01-04 09:37 | MSH DR1 DETECTOR | (TC 01/03/2024 08:36:00 PM, Whittington, Scott) MSH DR1 DETECTOR & CAL | MSH DR1 DETECTOR INCLUDING CALIBRATION
26791 | 201007 | PM | CMPLT | CMPLT | 2024-01-12 12:36 | 2024-01-12 12:37 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 01/12/2024 12:36:16 PM, Rhuland, Courtney) PM complete See attached checklist for details
| 201008 | CM | CMPLT | CMPLT | 2024-01-12 13:10 | 2024-01-12 13:10 | training | E-learning
| 201009 | CM | CMPLT | CMPLT | 2024-01-12 13:11 | 2024-01-12 13:11 | worked with RT | worked with RT
1055949 | 201010 | CM | CMPLT | CMPLT | 2024-01-12 14:30 | 2024-01-12 14:30 | power is not working | tested unit for 3 days. Power, SpO2, BP and Temp are functional. Replaced Temp probe.
1066454 | 201011 | FC | CMPLT | CMPLT | 2024-01-12 14:49 | 2024-01-12 14:49 | QC Checklist | (TC 01/12/2024 02:49:08 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
1056440 | 201012 | CM | CMPLT | NFF | 2024-01-12 14:50 | 2024-01-12 14:57 | Reported that NIBP is not working | (TC 01/12/2024 02:50:17 PM, Baghirzade, Nara) Battery was very depleted - placed Root to charge fully. Charge completed. Tested NIBP several times did not find fault. SpO2, NIBP, and Temp passed functional checks.
502044 | 201013 | CM | CMPLT | CMPLT | 2024-01-12 15:39 | 2024-01-18 11:06 | UPS defective | (TC 01/12/2024 03:39:07 PM, Green, Paul) UPS defective. Travelled to site and took out UPS. Connected directly to power. Looking into UPS replacement with Digirad. Filled water as it was half full. (TC 01/16/2024 10:58:30 AM, Green, Paul) Travelled to site. UPS was plugged in overnight, however not powering up. Same problem. UPS was opened up and batteries tested. 12V batteries reading 7V. Notified Digirad and they will send new batteries. (TC 01/18/2024 11:04:40 AM, Green, Paul) Installed batteries in replacement UPS. Powered up overnight, ok. Installed back in Ergo system, tested ok.
32520 | 201014 | CM | CMPLT | CMPLT | 2024-01-13 14:27 | 2024-01-15 09:56 | MSH DR2 TUBE LATERAL DRIVE CAUSING OGP MOVEMENT INTERUPPTIONS | (TC 01/13/2024 02:26:59 PM, Whittington, Scott) MSH DR2 TUBE LATERAL DRIVE CAUSING OGP MOVEMENT INTERUPPTIONS FOUND LATERAL TRAVEL CONSTANTLY INTERUPPTED CHECKED POSITIONING DEVICE READOUTS, APPERED TO ALWYS RETURN TO SAME POSITION & ENCODER VALUES IN CHEST PA FOUND LATERAL DRIVE WITH DRIVE WHEEL WEAR USED AN UNUSED NEW DRIVE FROM LOCAL STOCK THAT HAD DIFFICULTY HOMING REST/DRIVE POSITION REPAIRED BY REALIGNING CUTOFF SWITCVHH ACTUATORS INSTALLED, BUT FOUND THAT DRIVE TO EAST (DOWNHILL ) OK BUT TO SEST (UPHILL) CONSTANT INTERUPPTIONS, EVEN IN SERVICE MODE MOVEMENT TEST NORMAL SPEED IN SLOW SPEED OK WILL CONSIDER SWAPPING PRESENT DRIVE WHEEL ONTO FORMER DRIVE (TC 01/13/2024 02:26:59 PM, Whittington, Scott) MSH DR2 TUBE LATERAL DRIVE CAUSING OGP MOVEMENT INTERUPPTIONS FOUND LATERAL TRAVEL CONSTANTLY INTERUPPTED CHECKED POSITIONING DEVICE READOUTS, APPERED TO ALWYS RETURN TO SAME POSITION & ENCODER VALUES IN CHEST PA FOUND LATERAL DRIVE WITH DRIVE WHEEL WEAR USED AN UNUSED NEW DRIVE FROM LOCAL STOCK THAT HAD DIFFICULTY HOMING REST/DRIVE POSITION REPAIRED BY REALIGNING CUTOFF SWITCVHH ACTUATORS INSTALLED, BUT FOUND THAT DRIVE TO EAST (DOWNHILL ) OK BUT TO SEST (UPHILL) CONSTANT INTERUPPTIONS, EVEN IN SERVICE MODE MOVEMENT TEST NORMAL SPEED IN SLOW SPEED OK WILL CONSIDER SWAPPING PRESENT DRIVE WHEEL ONTO FORMER DRIVE (TC 01/13/2024 09:54:00 PM, Whittington, Scott) MSH DR2 TUBE LATERAL DRIVE CAUSING OGP MOVEMENT INTERUPPTIONS FOUND LATERAL TRAVEL CONSTANTLY INTERUPPTED CHECKED POSITIONING DEVICE READOUTS, APPERED TO ALWYS RETURN TO SAME POSITION & ENCODER VALUES IN CHEST PA FOUND LATERAL DRIVE WITH DRIVE WHEEL WEAR USED AN UNUSED NEW DRIVE FROM LOCAL STOCK THAT HAD DIFFICULTY HOMING REST/DRIVE POSITION REPAIRED BY REALIGNING CUTOFF SWITCVHH ACTUATORS INSTALLED, BUT FOUND THAT DRIVE TO EAST (DOWNHILL ) OK BUT TO SEST (UPHILL) CONSTANT INTERUPPTIONS, EVEN IN SERVICE MODE MOVEMENT TEST NORMAL SPEED IN SLOW SPEED OK WILL CONSIDER SWAPPING PRESENT DRIVE WHEEL ONTO FORMER DRIVE REBUILD DRIVE ABLE TO RETURN TO CLINICAL USE | MSH DR2 TUBE LATERAL DRIVE CAUSING OGP MOVEMENT INTERUPPTIONS
65915 | 201015 | CM | CMPLT | CMPLT | 2024-01-15 07:38 | 2024-01-15 13:27 | Returned from Sick Kids | (TC 01/15/2024 01:26:45 PM, Rhuland, Courtney) Check-in completed All tests passed Refer to ASM for details
65917 | 201016 | CM | CMPLT | CMPLT | 2024-01-15 07:38 | 2024-01-15 13:26 | Returned from Sick Kids | (TC 01/15/2024 01:26:01 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
65895 | 201017 | CM | CMPLT | CMPLT | 2024-01-15 07:38 | 2024-01-15 13:25 | Returned from Sick Kids | (TC 01/15/2024 01:23:18 PM, Rhuland, Courtney) Device failed check-in at pressure accuracy test Failed calibration after flashing device Replaced pressure sensor Replaced rear case Replaced right handle Calibrated syringe Replaced latch assembly Check-in completed See ASM for details
65898 | 201018 | CM | CMPLT | CMPLT | 2024-01-15 07:38 | 2024-01-15 13:22 | Returned from Sick Kids | (TC 01/15/2024 01:21:51 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
65911 | 201019 | CM | CMPLT | CMPLT | 2024-01-15 07:38 | 2024-01-15 13:20 | Returned from Sick Kids | (TC 01/15/2024 01:13:15 PM, Rhuland, Courtney) Device failed check-in at pressure sensor accuracy test Failed calibration after flashing syringe Broken front case replaced Pressure sensor replaced Mylar gasket replaced Flag disk holder replaced Syringe calibrated All tests passed Refer to ASM for details
1048526 | 201020 | CM | CMPLT | CMPLT | 2024-01-15 07:38 | 2024-01-15 13:12 | Returned from Sick Kids | (TC 01/15/2024 01:06:36 PM, Rhuland, Courtney) Check-in failed pressure sensor accuracy test Replaced pressure sensor with spare Flashed syringe Calibrated syringe Check-in completed See ASM for details
| 201021 | RDS | CMPLT | CMPLT | 2024-01-15 13:25 | 2024-01-18 12:25 | Weekly checks in CCU | (TC 01/15/2024 12:23:38 PM, Scero, Tomlyn) Weekly check in complete. 30J test done. 1 transport defib missing, found on a crash cart in ICU. Notified Diane to have it returned to CCU
1062147 | 201022 | CM | CMPLT | CMPLT | 2024-01-15 13:26 | 2024-01-18 12:47 | No power when pluggd in | (TC 01/15/2024 12:41:28 PM, Scero, Tomlyn) Bed was left in the hallway and I was told that there was no power to the bed. Plugged it in and all buttons and functions were working properly. Put the bed back into room 10.
67948 | 201023 | CM | CMPLT | CMPLT | 2024-01-15 13:37 | 2024-01-15 13:38 | Broken side wall latch | (TC 01/15/2024 01:37:38 PM, Rhuland, Courtney) Replaced side wall latch Wrong right side door was sent down with device, replaced door with spare Functional check - OK
67959 | 201024 | CM | CMPLT | CMPLT | 2024-01-15 13:40 | 2024-01-15 13:42 | Broken porthole hinge | (TC 01/15/2024 01:40:31 PM, Rhuland, Courtney) Replaced broken porthole middle hinge Repaired porthole hinge Functional check - OK
502453 | 201025 | CM | CMPLT | CMPLT | 2024-01-15 13:44 | 2024-02-26 13:16 | Temporary move | (TC 01/15/2024 01:43:42 PM, Fan, Tak) Swapped from 3-526 with the Philips in 3-539. (TC 01/22/2024 11:14:40 AM, Fan, Tak) Systems moved out for room maintenance. Moved system back to correct rooms. Took inventory of missing HSC volusons. 1057302 3-525 (D7), 3N-125D, c400ad541fd7, 44a56e49fc78 1069363 3-529 (D4), 3N-126D, 0013955af331, n/a 1063331 3-537 (C3), 3N-D071A, c400ad9ee7e4, c89e43a1a0ce 502453 3-539 (C2), 3N-D069B, c400ad4b123d, 60a44cec43e5
| 201026 | CM | CMPLT | CMPLT | 2024-01-15 15:57 | 2024-01-15 15:57 | mmeting | Biomed meeting
| 201027 | RDS | CMPLT | CMPLT | 2024-01-15 15:58 | 2024-01-15 15:58 | weekly check for 14S and 16N | weekly check for 14S and 16N
| 201028 | CM | CMPLT | CMPLT | 2024-01-15 16:01 | 2024-01-15 16:01 | worked at L&D with IT | worked at L&D with IT
1053349 | 201029 | CM | CMPLT | CMPLT | 2024-01-15 17:03 | 2024-02-26 13:16 | H2 Artifact in image and Error GCN AEC calculated mAs too high |  | Filter wheel artifact. Reset unit system tested OK. Jon says to call it in if it happens again.
1063808 | 201030 | CM | CMPLT | CMPLT | 2024-01-15 17:13 | 2024-01-15 17:13 | Keyboard keys sticking. |  | Air dusted keyboard. Tested Ok
| 201031 | CM | CMPLT | CMPLT | 2024-01-15 17:15 | 2024-01-15 17:15 | Site coverage
1059778 | 201032 | FC | CMPLT | CMPLT | 2024-01-16 06:58 | 2024-01-16 06:59 | QC checklist | (TC 01/16/2024 06:58:02 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1054978 | 201033 | CM | CMPLT | CMPLT | 2024-01-16 07:17 | 2024-02-21 10:19 | Defective battery | (TC 01/16/2024 07:16:44 AM, Baghirzade, Nara) Left Root to charge for a day however the battery capacity does not exceed 2500 mAH. Contact Masimo for support/replacement. (TC 01/17/2024 07:55:25 AM, Baghirzade, Nara) Initiated service request - will be sending the root for a replacement. Submitted PO request. SC 1003148447 (TC 01/17/2024 02:12:39 PM, Baghirzade, Nara) Shipped device for replacement. PO : 7500200697 Waybill #: 412483 (TC 02/21/2024 09:49:21 AM, Baghirzade, Nara) Received. PO 7500200697 (TC 02/21/2024 10:17:21 AM, Baghirzade, Nara) Installed new Root on the mount/stand. Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number.
| 201034 | MEP | ACTIV | ISSUE | 2024-01-16 08:00 |  | Systems committee | (TC 01/15/2024 08:00:03 AM, Canjimavadivel, Kamallesh ) Started patching Physio2 server and saw server was crashing. Rebuilt using usb (TC 01/16/2024 01:51:13 PM, Canjimavadivel, Kamallesh ) Monthly update (TC 01/16/2024 01:51:26 PM, Canjimavadivel, Kamallesh ) Restarted Standy physio server to complete patch install (TC 01/16/2024 03:27:44 PM, Canjimavadivel, Kamallesh ) Switching standy server back to physio 2 and completing setup (TC 01/18/2024 08:24:21 AM, Canjimavadivel, Kamallesh ) Installed VNC and verified patch complete on standby server (TC 01/22/2024 03:09:03 PM, Canjimavadivel, Kamallesh ) Adding new Philips network equipment into database and labelling before Brandon's arrival (TC 01/23/2024 11:16:37 AM, Canjimavadivel, Kamallesh ) Added all network equipment to database and asset tagged. Closed all WOs (TC 01/26/2024 02:29:02 PM, Canjimavadivel, Kamallesh ) Investigated notification feature on Zoll dashboard (TC 01/30/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Worked with Kamran to validate configurations on Physio 1 server (TC 01/31/2024 03:19:18 PM, Canjimavadivel, Kamallesh ) Validating DNS settings on PICiX in ICU and ER and some centrals' settigns were changed. (TC 02/01/2024 01:12:15 PM, Canjimavadivel, Kamallesh ) Checking PICiX in PACU to validate DNS settings. Assisted Nara with PICiX in NICU. Started rebuilding PACU3S PICiX (TC 02/01/2024 03:15:24 PM, Canjimavadivel, Kamallesh ) Started system setup for spare PICiX (TC 02/02/2024 03:29:48 PM, Canjimavadivel, Kamallesh ) Completed setup of spare PICiX (TC 02/05/2024 10:29:00 AM, Canjimavadivel, Kamallesh ) Rebuilt physio1 server and got it ready for patching (TC 02/07/2024 09:56:45 AM, Canjimavadivel, Kamallesh ) Configured physio 1 server back to original settings after patch (TC 02/07/2024 11:28:05 AM, Canjimavadivel, Kamallesh ) Testing server connections with Kamran (TC 02/07/2024 07:00:00 PM, Canjimavadivel, Kamallesh ) Working on DHCP issue with monitors not connecting to PICiX. Worked with Kamran and John. (TC 02/08/2024 02:30:00 PM, Canjimavadivel, Kamallesh ) Resolved DHCP issue with ICU and CCU monitors. Worked with John and reset network for all monitors. (TC 02/08/2024 02:30:00 PM, Canjimavadivel, Kamallesh ) Resolved DHCP issue with ICU and CCU monitors. Worked with John and reset network for all monitors. (TC 02/13/2024 03:27:12 PM, Canjimavadivel, Kamallesh ) Debrief on DHCP issue from last week (TC 02/21/2024 10:53:48 AM, Canjimavadivel, Kamallesh ) Bringing standby server back online (TC 03/07/2024 02:45:23 PM, Canjimavadivel, Kamallesh ) Updates (TC 04/25/2024 01:28:33 PM, Canjimavadivel, Kamallesh ) Going through task list and updating (TC 06/26/2024 03:31:11 PM, Canjimavadivel, Kamallesh ) Monthly meeting (TC 07/05/2024 02:48:40 PM, Canjimavadivel, Kamallesh ) Focal point alerts review (TC 07/09/2024 10:34:53 AM, Canjimavadivel, Kamallesh ) Checking the WiFi setting on TC70s (TC 07/18/2024 02:41:09 PM, Canjimavadivel, Kamallesh ) Monthly meeting (TC 07/22/2024 02:40:55 PM, Canjimavadivel, Kamallesh ) Scheduling PICiX patch updates and communicationg with units (TC 07/23/2024 11:01:09 AM, Canjimavadivel, Kamallesh ) Updating excel spreadsheet for patch update (TC 07/25/2024 02:20:00 PM, Canjimavadivel, Kamallesh ) Completed patch udate on all PICiX in biomed shop and remaining PACU monitors. Added entries in excel spreadsheet (TC 08/02/2024 02:16:58 PM, Canjimavadivel, Kamallesh ) Scheduling patches and updating spreadsheet (TC 08/07/2024 02:11:19 PM, Canjimavadivel, Kamallesh ) Going over PICiX upgrade plan with Kamran (TC 08/08/2024 02:36:59 PM, Canjimavadivel, Kamallesh ) Working on PICiX patch installs (TC 08/09/2024 01:47:58 PM, Canjimavadivel, Kamallesh ) Working on trivia deck with John (TC 08/09/2024 02:00:22 PM, Canjimavadivel, Kamallesh ) Scheduling patch updates and updating spreadsheet (TC 08/14/2024 01:50:23 PM, Canjimavadivel, Kamallesh ) Working on quiz deck (TC 08/13/2024 01:50:23 PM, Canjimavadivel, Kamallesh ) Working on quiz deck (TC 08/20/2024 02:59:28 PM, Canjimavadivel, Kamallesh ) Finalizing quiz deck edits (TC 08/21/2024 02:42:00 PM, Canjimavadivel, Kamallesh ) Moving PICiX hardware to storage (TC 08/22/2024 11:17:27 AM, Canjimavadivel, Kamallesh ) Final edits for quiz and full run through (TC 09/06/2024 10:41:55 AM, Canjimavadivel, Kamallesh ) Testing PICiX alarms with Kamran (TC 09/09/2024 12:00:00 PM, Canjimavadivel, Kamallesh ) PICiX Rev 4 alarm testing in ICU (TC 09/11/2024 12:31:15 PM, Canjimavadivel, Kamallesh ) Compiled stats from Quiz (TC 09/26/2024 12:49:29 PM, Canjimavadivel, Kamallesh ) Rev 4 plan kick-off meeting (TC 10/01/2024 01:05:32 PM, Canjimavadivel, Kamallesh ) Rev 4 secondary alerting testing with Kamran (TC 10/10/2024 09:56:19 AM, Canjimavadivel, Kamallesh ) PICiX rev 4 alarm testing in NICU (TC 10/24/2024 02:04:32 PM, Canjimavadivel, Kamallesh ) Secondary alert testing for Rev 4 in NICU (TC 10/30/2024 02:23:03 PM, Canjimavadivel, Kamallesh ) Secondary alerting testing teams call with Philips
| 201035 | PA | ACTIV | ISSUE | 2024-01-16 08:06 |  | Sterile water for ED | (TC 01/16/2024 08:12:29 AM, Canjimavadivel, Kamallesh ) Placed order for sterile water (TC 02/26/2024 12:47:52 PM, Canjimavadivel, Kamallesh ) Placed order for water (TC 03/25/2024 08:10:43 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 04/24/2024 10:00:41 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 07/02/2024 08:13:40 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 07/15/2024 02:59:01 PM, Canjimavadivel, Kamallesh ) Placed order for water (TC 07/29/2024 08:28:28 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 08/12/2024 09:05:44 AM, Canjimavadivel, Kamallesh ) Placed order for water (TC 09/11/2024 08:08:22 AM, Canjimavadivel, Kamallesh ) Placed order for sterile water
| 201036 | RDS | CMPLT | CMPLT | 2024-01-16 08:17 | 2024-01-18 10:27 | Weekly rounds for ED
| 201037 | RDS | CMPLT | CMPLT | 2024-01-16 08:18 | 2024-01-17 09:48 | Weekly rounds for ICU
| 201038 | CM | CMPLT | CMPLT | 2024-01-16 08:20 | 2024-01-23 10:31 | Warranty replacement for regulators | (TC 01/16/2024 08:27:14 AM, Canjimavadivel, Kamallesh ) 2 regulators covered under warranty, shipping not covered. Requested PO for freight charges (TC 01/16/2024 01:54:31 PM, Canjimavadivel, Kamallesh ) Sent PO to vendor, waiting for next steps (TC 01/23/2024 10:25:07 AM, Canjimavadivel, Kamallesh ) Received replacement regulators. Packed up defective ones and used shipping label provided by vendor to ship. Dropped off box to receiving docks
65935 | 201039 | CM | CMPLT | CMPLT | 2024-01-16 08:51 | 2024-01-16 08:52 | Not showing amount infused properly | (TC 01/16/2024 08:51:08 AM, Rhuland, Courtney) Failed pressure sensor accuracy test Reseated pressure sensor Replaced rear case Calibrated and performed check-in All tests passed Display working fine Refer to ASM for details
1062022 | 201040 | CM | CMPLT | CMPLT | 2024-01-16 08:55 | 2024-01-16 08:58 | Reported that the head portion makes a loud noise when moving up/down | (TC 01/16/2024 08:54:52 AM, Baghirzade, Nara) First time the head portion was raised, it made a loud noise when close to the 36-38 degree angle. After moving the head up and down, there were no more noises. Checked all motion controls including the patient motion controls. Raised the bed to the maximum height and raised head up and down, no noise. Lowered the bed to the minimum height and raised head up and down, no noise. All functions of the bed are working. There was noise coming from the contact between the mattress and the frame, so I put a cloth on each corner and when moving the head up/down, it did not make any noise. Removed the cloth and put in the soiled bin. Let the clerk know at the 11South nurisng station that the bed is functional and that once they put bed sheet covers, the mattress wont make any noise either.
| 201041 | CM | CMPLT | CMPLT | 2024-01-16 08:56 | 2024-01-15 14:56 | Capital monthly meeting
1065669 | 201042 | CM | CMPLT | CMPLT | 2024-01-16 09:04 | 2024-01-18 14:34 | Error code E5.0 | (TC 01/16/2024 09:53:20 AM, Canjimavadivel, Kamallesh ) Confirmed error code in logs. Tested with tester to verify error, did not reappear. (TC 01/18/2024 11:18:31 AM, Canjimavadivel, Kamallesh ) Checked again and found error code. Replaced probe
| 201043 | CM | CMPLT | CMPLT | 2024-01-16 10:44 | 2024-01-18 07:33 | Huddle
| 201044 | CM | CMPLT | CMPLT | 2024-01-16 10:47 | 2024-01-30 08:10 | NICU rounds | (TC 01/16/2024 10:48:17 AM, Rhuland, Courtney) Rounds completed for NICU (TC 01/16/2024 10:48:30 AM, Baghirzade, Nara) Rounds completed for NICU (TC 01/24/2024 08:31:14 AM, Rhuland, Courtney) Rounds completed for NICU
| 201045 | CM | CMPLT | CMPLT | 2024-01-16 10:49 | 2024-01-30 08:10 | RESUS rounds | (TC 01/16/2024 10:50:27 AM, Rhuland, Courtney) Rounds completed for RESUS (TC 01/16/2024 10:50:40 AM, Baghirzade, Nara) Rounds completed for RESUS
| 201046 | CM | CMPLT | CMPLT | 2024-01-16 10:53 | 2024-01-16 10:53 | PM huddle/update | (TC 01/16/2024 10:52:50 AM, Baghirzade, Nara) PM huddle/update in biomed.
1061212 | 201047 | CM | CMPLT | CMPLT | 2024-01-16 11:12 | 2024-01-16 11:13 | VMC not connecting to coriomatrix | (TC 01/15/2024 11:11:50 AM, Reid, Emmerson ) Need to call olympus and troubleshoot this issue. Ended up needed to restart the VMC and CorioMatrix before they connected. Likely due to the code grey power.
1061217 | 201048 | CM | CMPLT | CMPLT | 2024-01-16 11:13 | 2024-01-16 11:14 | VMC not connecting to CorioMatrix | (TC 01/15/2024 11:13:20 AM, Reid, Emmerson ) Need to call olympus and troubleshoot this issue. Ended up needed to restart the VMC and CorioMatrix before they connected. Likely due to the code grey power.
1069703 | 201049 | ACC | CMPLT | CMPLT | 2024-01-16 11:14 | 2024-01-16 13:39 | Incoming Inspection | (TC 01/16/2024 01:39:02 PM, Scero, Tomlyn) Incoming inspection complete. Info entered into TMS for demo
1061207 | 201050 | CM | CMPLT | CMPLT | 2024-01-16 11:14 | 2024-01-16 11:15 | VMC not connecting to CORIOmatrix | (TC 01/15/2024 11:14:25 AM, Reid, Emmerson ) Need to call olympus and troubleshoot this issue. Ended up needed to restart the VMC and CorioMatrix before they connected. Likely due to the code grey power.
1061207 | 201051 | CM | CMPLT | CMPLT | 2024-01-16 11:20 | 2024-01-16 11:21 | Music/speakers not working | (TC 01/16/2024 11:19:43 AM, Reid, Emmerson ) Tried connecting my phone and playing music. It connected but no music played. I restarted the VMC and tried again and it worked.
1061209 | 201052 | CM | CMPLT | CMPLT | 2024-01-16 11:21 | 2024-01-16 11:22 | Lights iPad in weird mode | (TC 01/16/2024 11:21:16 AM, Reid, Emmerson ) The iPad was on the app screen, I just repoened app and it was working
| 201053 | CM | CMPLT | CMPLT | 2024-01-16 11:26 | 2024-03-13 13:20 | Philips MX Series training | (TC 01/16/2024 11:25:49 AM, Rhuland, Courtney) Requested and recieved quote from Philips for all-inclusive package (travel, hotel, flights) PO 7500200642 (TC 03/08/2024 01:16:38 PM, Rhuland, Courtney) MX series training March 6-8 (TC 03/08/2024 01:19:08 PM, Rhuland, Courtney) Travel time to/from training in San Francisco
1061333 | 201054 | FC | CMPLT | CMPLT | 2024-01-16 12:55 | 2024-01-16 12:56 | QC Checklist | (TC 01/16/2024 12:55:16 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
| 201055 | CM | CMPLT | CMPLT | 2024-01-16 13:54 | 2024-01-16 13:54 | meeting | meeting
| 201056 | CM | CMPLT | CMPLT | 2024-01-16 13:56 | 2024-01-16 13:56 | worked at L&D | worked at L&D
65916 | 201057 | CM | CMPLT | CMPLT | 2024-01-16 15:01 | 2024-01-16 15:22 | Broken front and rear case, left handle, and latch | (TC 01/16/2024 03:01:00 PM, Baghirzade, Nara) During rounds, noticed that the pump had a broken rear case, front case, cracked left handle, and the latch was broken. It was also missing the SN sticker on the rear case. Replaced: -rear case -front case -left handle -latch -mylar gasket Printed new asset tag and bar code sticker. Placed new SN sticker on the rear case. (TC 01/16/2024 03:21:18 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65417 | 201058 | CM | CMPLT | CMPLT | 2024-01-17 08:12 | 2024-01-17 08:15 | Wrong door placement | (TC 01/17/2024 08:12:10 AM, Baghirzade, Nara) Doors were placed in reversed order. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74.3 %RH Incubator is fully functional.
| 201059 | CM | CMPLT | CMPLT | 2024-01-17 10:35 | 2024-01-17 10:36 | OR Support Training | (TC 01/17/2024 10:35:25 AM, Baghirzade, Nara) OR Support Training
| 201060 | CM | ACTIV | ISSUE | 2024-01-17 12:21 |  | Philips TEST PICiX Rev. C to Rev. 4 upgrade
| 201061 | CM | CMPLT | CMPLT | 2024-01-17 12:31 | 2024-09-25 09:35 | MSH OPG 3 SPP SURVEY 7 IDENTIFY VOLUSON SYSTEMS THAT ARE NO LONGER CONNECTED TO HSC , AND REQUEST RECONNECTION | (TC 01/17/2024 12:31:37 PM, Whittington, Scott) MSH OPG 3 SPP SURVEY 7 IDENTIFY VOLUSON SYSTEMS THAT ARE NO LONGER CONNECTED TO HSC , AND REQUEST RECONNECTION (TC 04/11/2024 09:56:32 AM, Whittington, Scott) MSH OPG 3 SPP SURVEY 7 IDENTIFY VOLUSON SYSTEMS THAT ARE NO LONGER CONNECTED TO HSC , AND REQUEST RECONNECTION CONTINUE RECONNECTION EFFORTS (TC 04/11/2024 09:56:32 AM, Whittington, Scott) MSH OPG 3 SPP SURVEY 7 IDENTIFY VOLUSON SYSTEMS THAT ARE NO LONGER CONNECTED TO HSC , AND REQUEST RECONNECTION CONTINUE RECONNECTION EFFORTS (TC 05/17/2024 10:23:19 AM, Whittington, Scott) MSH SPP>HSC NETWORK MEETING | MSH OPG 3 SPP SURVEY 7 IDENTIFY VOLUSON SYSTEMS THAT ARE NO LONGER CONNECTED TO HSC , AND REQUEST RECONNECTION
67958 | 201062 | CM | CMPLT | CMPLT | 2024-01-17 13:47 | 2024-01-17 15:06 | Broken hood and latch | (TC 01/17/2024 01:47:15 PM, Baghirzade, Nara) Removed all components that were on the old hood. Disposed of the old hood. (TC 01/17/2024 01:47:32 PM, Baghirzade, Nara) Placed new hood on the incubator. Attached porthole door, humdity and air sensors and the bracket that hold the humidity/air cord to the hood. Placed all the spare parts on the incubator parts shelf. (TC 01/17/2024 03:05:13 PM, Baghirzade, Nara) Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
1061447 | 201063 | CM | CMPLT | CMPLT | 2024-01-17 13:51 | 2024-02-01 10:22 | Right side panel motion controls not working | (TC 01/17/2024 01:51:07 PM, Baghirzade, Nara) Tested all motion controls - right side panel is not working. Reset the bed but it did not mitigate the issue. Power cycled the bed and also did not fix the issue. Will contact Stryker for service. (TC 01/31/2024 03:06:36 PM, Baghirzade, Nara) Checked all the bed motion controls; all functional. Waiting on service report from Stryker. (TC 02/01/2024 10:21:00 AM, Baghirzade, Nara) Attached service report
1067933 | 201064 | CM | CMPLT | CMPLT | 2024-01-17 14:10 | 2024-01-17 14:12 | Vitals not showing up from software to monitor | (TC 01/17/2024 02:09:47 PM, Canjimavadivel, Kamallesh ) Monitor settings were changed and needed to be changed back. ECG leads also had intermittent issue, swapped with spare and working ok
| 201065 | CM | CMPLT | CMPLT | 2024-01-17 14:21 | 2024-01-17 14:21 | training | OR training
| 201066 | CM | CMPLT | CMPLT | 2024-01-17 14:31 | 2024-07-22 14:40 | Training/meetings | (TC 01/17/2024 02:31:10 PM, Rhuland, Courtney) OR overview training (TC 01/19/2024 11:01:51 AM, Rhuland, Courtney) Full team monthly meeting (TC 01/22/2024 02:15:43 PM, Rhuland, Courtney) Monday huddle (TC 01/29/2024 02:05:39 PM, Rhuland, Courtney) Monday huddle (TC 02/05/2024 01:45:52 PM, Rhuland, Courtney) Monday huddle (TC 02/07/2024 01:33:54 PM, Rhuland, Courtney) Bassinet deployment meeting (TC 02/12/2024 02:08:49 PM, Rhuland, Courtney) Monday huddle (TC 02/16/2024 11:04:07 AM, Rhuland, Courtney) Full department meeting (TC 02/21/2024 02:48:47 PM, Rhuland, Courtney) Oncology pump remediaton meeting to plan schedule of taking/replacing pumps during upgrade. (TC 02/26/2024 11:10:22 AM, Rhuland, Courtney) BD pump remediation plan meeting (TC 02/26/2024 02:03:07 PM, Rhuland, Courtney) Monday huddle (TC 02/27/2024 02:38:22 PM, Rhuland, Courtney) PM meeting (TC 02/28/2024 10:30:30 AM, Rhuland, Courtney) Medfusion 4000 firmware webinar and test of software update (TC 02/28/2024 11:37:22 AM, Rhuland, Courtney) WatchBP demo for heart function clinic (TC 02/28/2024 01:50:03 PM, Rhuland, Courtney) Capital monthly meeting (TC 02/29/2024 12:50:44 PM, Rhuland, Courtney) Incoming incubator plan meeting (TC 02/29/2024 02:07:03 PM, Rhuland, Courtney) NICU pump remediation meeting (TC 03/04/2024 01:37:09 PM, Rhuland, Courtney) Monday huddle (TC 03/15/2024 10:45:20 AM, Rhuland, Courtney) 1 on 1 w Adeel (TC 03/21/2024 03:45:00 PM, Rhuland, Courtney) NICU huddle (TC 03/27/2024 01:23:41 PM, Rhuland, Courtney) PM meeting (TC 03/28/2024 11:34:59 AM, Rhuland, Courtney) Meeting with Draegar and NICU to discuss boom PM/repairs (TC 04/08/2024 12:58:01 PM, Rhuland, Courtney) Meeting with Draegar rep to review issues with the booms (TC 04/10/2024 02:26:31 PM, Rhuland, Courtney) Disposal plan for incubators meeting (TC 04/16/2024 09:55:53 AM, Rhuland, Courtney) Huddle (TC 04/17/2024 09:42:04 AM, Rhuland, Courtney) Capital meeting (TC 04/19/2024 02:05:48 PM, Rhuland, Courtney) Full team meeting (TC 04/22/2024 02:36:42 PM, Rhuland, Courtney) Monday huddle (TC 04/24/2024 01:54:59 PM, Rhuland, Courtney) PM update meeting (TC 04/29/2024 02:15:12 PM, Rhuland, Courtney) monday huddle (TC 05/07/2024 01:30:22 PM, Rhuland, Courtney) NICU walkthrough with John and Naz to discuss room changes (TC 05/27/2024 01:50:19 PM, Rhuland, Courtney) Monday huddle (TC 06/03/2024 02:51:22 PM, Rhuland, Courtney) Monday huddle (TC 06/05/2024 02:30:40 PM, Rhuland, Courtney) 1 on 1 (TC 06/19/2024 01:54:35 PM, Rhuland, Courtney) Huddle (TC 06/20/2024 02:50:39 PM, Rhuland, Courtney) Meeting with NICU staff to discuss configuration of oxygen tank, phototherapy light and exam light on omnibed (TC 06/21/2024 10:45:39 AM, Rhuland, Courtney) Full team meeting (TC 06/25/2024 10:48:43 AM, Rhuland, Courtney) NICU meeting with John & Naz (TC 06/26/2024 02:39:58 PM, Rhuland, Courtney) PM meeting (TC 07/04/2024 10:30:05 AM, Rhuland, Courtney) Meeting with building services to discuss nitrogen leaks in NICU booms (TC 07/09/2024 02:17:10 PM, Rhuland, Courtney) Monday huddle (TC 07/22/2024 02:40:10 PM, Rhuland, Courtney) Monday huddle
1051571 | 201067 | CM | CMPLT | CMPLT | 2024-01-17 15:13 | 2024-01-17 15:15 | Not working, batteries changed still not working | (TC 01/17/2024 03:13:40 PM, Rendon Lopez , Elizabeth) Confirmed issue, unable to turn on device. Device is unrepairable, retired from service, will request quote for replacement
| 201068 | CM | CMPLT | CMPLT | 2024-01-17 17:20 | 2024-01-17 17:20 | MRI Wheel Chair brake not working. |  | Adjusted and repaired wheelchair brake.
1060518 | 201069 | CM | CMPLT | CMPLT | 2024-01-17 18:24 | 2024-01-17 07:22 | No table height motion. | Replaced Motor assembly, replaced AMC motor control and replaced the potentiometer, re-installed the AMC SOFTWARE, recalibrated height motion didn't fix the problem. Ordered new potentiometer, and the Force sensor, will replace it once recieved.
1056125 | 201070 | PM | CMPLT | CMPLT | 2024-01-18 08:47 | 2024-01-18 09:54 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/18/2024 09:53:16 AM, Canjimavadivel, Kamallesh ) Cleaned hose ends and valves Cleaning cycle complete
1056209 | 201071 | CM | CMPLT | NFF | 2024-01-18 09:29 | 2024-01-18 09:29 | "not working well" | (TC 01/17/2024 09:28:44 AM, Cortez, John) No issue found, All parameters tested ok
| 201072 | RDS | CMPLT | CMPLT | 2024-01-18 10:11 | 2024-01-18 10:13 | Weekly Checks For PACU | (TC 01/17/2024 10:12:22 AM, Reid, Emmerson ) Rounds complete
62504 | 201073 | CM | CMPLT | CMPLT | 2024-01-18 10:12 | 2024-01-18 10:19 | BP cuff leaking | (TC 01/18/2024 10:17:45 AM, Scero, Tomlyn) Tested unit, confirmed BP cuff was leaking. Replaced cuff and retested. No leaks. Returned device to 1609 clinic room.
1056421 | 201074 | CM | CMPLT | CMPLT | 2024-01-18 10:31 | 2024-01-18 10:55 | "needs calibration" | (TC 01/18/2024 10:30:52 AM, Cortez, John) Someone attached Philips cuff to Masimo hose. Swapped with proper cuff/hose All parameters tested ok
67951 | 201075 | CM | CMPLT | CMPLT | 2024-01-18 11:12 | 2024-01-18 11:12 | Broken latch and air probe cover | (TC 01/18/2024 11:11:47 AM, Baghirzade, Nara) Replaced latch. Replaced air probe cover. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74.6 %RH Incubator is fully functional.
67921 | 201076 | CM | CMPLT | CMPLT | 2024-01-18 11:13 | 2024-01-18 11:13 | Broken latch | (TC 01/18/2024 11:12:42 AM, Baghirzade, Nara) Replaced broken latch. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
66828 | 201077 | CM | CMPLT | CMPLT | 2024-01-18 11:14 | 2024-01-18 11:20 | Broken screen | (TC 01/18/2024 11:19:30 AM, Rhuland, Courtney) Replaced display screen Functional check - OK
1059548 | 201078 | CM | CMPLT | CMPLT | 2024-01-18 11:16 | 2024-01-18 11:17 | Received alert that it had an unexpected shutdown due to hardware failure | (TC 01/18/2024 11:15:39 AM, Baghirzade, Nara) Checked all the connections from the PIICIX - all good. Checked UPS - functional. Nurses did not notice that it had been shut down. Overview is functional.
| 201079 | CM | CMPLT | CMPLT | 2024-01-18 12:48 | 2024-03-20 10:35 | Zeus IE monitor configuration change | (TC 01/24/2024 10:24:10 AM, Scero, Tomlyn) Quote received from BOMIMed for 22 rails. PO ordered from MedBuy. Quote attached to WO. (TC 03/20/2024 10:33:27 AM, Scero, Tomlyn) Received Rails and moved them to the RT room for storage until use.
62966 | 201080 | CM | CMPLT | CMPLT | 2024-01-18 12:58 | 2024-01-18 13:01 | Reported that the canopy does not go up or down | (TC 01/18/2024 12:58:08 PM, Baghirzade, Nara) Canopy covers were not going up when using the right-hand side switch. Replaced right-hand side membrane switch. Replaced tank holder as it was broken. Replaced SW latch. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74.1-74.5%RH. Incubator is fully functional.
64867 | 201081 | CM | CMPLT | CMPLT | 2024-01-18 13:10 | 2024-02-28 10:10 | Leaking water | (TC 01/18/2024 03:22:31 PM, Rendon Lopez , Elizabeth) Confirmed issue, cracked tank cover. inquired pricing, ordered part (TC 02/27/2024 10:09:11 AM, Rendon Lopez , Elizabeth) Replaced tank cover and gasket. Refill with concentration fluid. Tested all functions, OK
1066843 | 201082 | CM | CMPLT | CMPLT | 2024-01-18 14:17 | 2024-02-07 11:23 | Water ejected from scope connector port | (TC 01/18/2024 02:18:13 PM, Cortez, John) Requested RMA, shipped to Olympus (TC 01/26/2024 03:35:09 PM, Cortez, John) Received repair estimate. Reached out to Anita for approval. Requested PO. (TC 02/02/2024 02:37:53 PM, Cortez, John) Received back from repair unboxed, visual inspection passed. Opened up casing to check scope connector tubing, brand new. Electrical safety passed- 60 mohms, 74 uAs Planning to install in Rm1 on Monday, pending approval from Anita (TC 02/05/2024 04:49:06 PM, Cortez, John) Installed in room 1. Tested ok with a scope. Removed loaner, which will be shipped back to Olympus
1061204 | 201083 | CM | CMPLT | CMPLT | 2024-01-18 14:20 | 2024-01-18 14:21 | CORIOmatrix not connecting | (TC 01/18/2024 02:20:31 PM, Reid, Emmerson ) Resterted CORIOmatrix and VMC, working now.
1053391 | 201084 | CM | CMPLT | CMPLT | 2024-01-19 07:19 | 2024-01-19 07:24 | QC and filter change before/after patient use | (TC 01/19/2024 07:23:23 AM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 51 mmHg Set 125 mmHg - Measured 124 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Checked service timer; set to off. Device is functional. Returned to 11S.
1067437 | 201085 | FC | CMPLT | CMPLT | 2024-01-19 08:55 | 2024-01-19 09:12 | QC Checklist | (TC 01/19/2024 09:11:52 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
67889 | 201086 | CM | CMPLT | CMPLT | 2024-01-19 09:22 | 2024-02-27 13:30 | Downstream occlusion alert | (TC 01/17/2024 09:24:31 AM, Rendon Lopez , Elizabeth) Dimmed LCD screen, to be sent out for repair, requested RMA (TC 01/19/2024 09:24:31 AM, Rendon Lopez , Elizabeth) Received RMA:318174. Sent out (TC 02/09/2024 08:45:45 AM, Rendon Lopez , Elizabeth) Received repaire estimate, pump over 7 years and repair is $1000. Approved, updated PO. (TC 02/26/2024 01:29:48 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK. See attached report. PM completed, all tests passed
63360 | 201087 | CM | CMPLT | CMPLT | 2024-01-19 09:27 | 2024-01-19 09:28 | missing foot | (TC 01/18/2024 09:26:47 AM, Cortez, John) borrowed one foot from spare, will place order for more feet/spacers
67881 | 201088 | CM | CMPLT | CMPLT | 2024-01-19 09:29 | 2024-02-27 13:29 | Created by Failed PM WO # 198754 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 01/19/2024 09:29:16 AM, Rendon Lopez , Elizabeth) RMA:318174 Sent out (TC 02/26/2024 01:28:36 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK. See attached report. PM completed, all tests passed | (TC 01/18/2024 09:27:26 AM, Rendon Lopez , Elizabeth) Testing device, ran rate accuracy several times with different cassettes and device over delivering. requested RMA to send out for repair
64997 | 201089 | CM | CMPLT | NFF | 2024-01-19 09:32 | 2024-01-19 09:33 | Not working | (TC 01/17/2024 09:31:55 AM, Rendon Lopez , Elizabeth) Tested on user mode, NFF. All checks passed
| 201090 | CM | CMPLT | CMPLT | 2024-01-19 09:34 | 2024-04-23 09:46 | Order for feet/spacers for SR555 scale, biomed stock | (TC 01/19/2024 09:40:06 AM, Cortez, John) Placed order for 4x feet, 4x spacers (TC 01/31/2024 08:42:48 AM, Cortez, John) Reached out for ETA. OOS and no ETA (TC 02/09/2024 04:31:41 PM, Cortez, John) spacers arrived, screws BO
1056444 | 201091 | CM | CMPLT | CMPLT | 2024-01-19 10:14 | 2024-01-19 10:14 | Reported that temperature probe is not working | (TC 01/19/2024 10:13:58 AM, Baghirzade, Nara) Probe well was missing - replaced the probe well. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1059828 | 201092 | CM | CMPLT | CMPLT | 2024-01-19 10:36 | 2024-01-19 10:36 | Cleaning | (TC 01/19/2024 10:35:59 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1061333 | 201093 | FC | CMPLT | CMPLT | 2024-01-19 10:44 | 2024-01-19 10:47 | QC Checklist | (TC 01/19/2024 10:46:13 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1055575 | 201094 | CM | CMPLT | CMPLT | 2024-01-19 10:52 | 2024-02-05 12:34 | 2 x root unit needs battery replacement | 2 x root unit needs battery replacement. we received 2 x roots. PO#7202350410
| 201095 | CM | CMPLT | CMPLT | 2024-01-19 11:06 | 2024-01-19 11:06 | meeting | meeting
| 201096 | CM | CMPLT | CMPLT | 2024-01-19 11:09 | 2024-01-19 11:09 | Full team meeting | (TC 01/19/2024 11:08:46 AM, Baghirzade, Nara) Full team meeting
1058505 | 201097 | CM | CMPLT | CMPLT | 2024-01-19 11:10 | 2024-01-19 11:12 | Reported that the enter button is not working | (TC 01/19/2024 11:10:25 AM, Baghirzade, Nara) Found no issue with the enter button. Tested all the buttons on the front panel; ran the warmer for a few minutes - no issues or errors. Turned the unit on and off multiple times, testing all buttons again. Found no issues with it. Returned to NICU.
| 201098 | CM | ACTIV | ISSUE | 2024-01-19 11:47 |  | Meetings and biomed projects | (TC 01/25/2024 10:01:46 AM, Rendon Lopez , Elizabeth) Fetal monitor inventory L&D and antenatal (TC 01/30/2024 10:00:33 AM, Rendon Lopez , Elizabeth) e-learning (TC 01/30/2024 10:00:33 AM, Rendon Lopez , Elizabeth) WOs input (TC 01/31/2024 03:48:27 PM, Rendon Lopez , Elizabeth) Fetal monitor inventory (TC 02/01/2024 12:54:28 PM, Rendon Lopez , Elizabeth) Fetal monitor inventory (TC 02/09/2024 03:54:29 PM, Rendon Lopez , Elizabeth) WOs and fetal monitor inventory location. (TC 02/26/2024 01:31:26 PM, Rendon Lopez , Elizabeth) Pump software upgrade config (TC 03/07/2024 02:11:06 PM, Rendon Lopez , Elizabeth) Fetalink replacement project (TC 03/15/2024 03:38:44 PM, Rendon Lopez , Elizabeth) WOs input (TC 03/14/2024 03:38:44 PM, Rendon Lopez , Elizabeth) WOs input (TC 07/26/2024 03:49:48 PM, Rendon Lopez , Elizabeth) Avalon CL and wireless transducers SW upgrade testing with Kamran.
| 201099 | CM | ACTIV | ISSUE | 2024-01-19 13:07 |  | working with Naz to order transport incubator. | (TC 01/19/2024 01:07:19 PM, Chukri, Samir) Meeting with Vayago transport to check their streachers compatable with our IAD (TC 01/24/2024 11:21:24 AM, Chukri, Samir) Disscuse with Sandy (R.T) and explain about the IAD for the new transport incubators to be compatabled with their streachers.
1051351 | 201100 | CM | CMPLT | CMPLT | 2024-01-19 13:36 | 2024-01-19 13:37 | Power cord broken | (TC 01/19/2024 01:35:52 PM, Reid, Emmerson ) Power cord was bent badly, likely from a table coming down on it. I disposed of the old power cord and replace with a new one. Working now.
| 201101 | RDS | CMPLT | CMPLT | 2024-01-19 13:42 | 2024-01-19 13:44 | Weekly Rounds: 11 South | (TC 01/19/2024 01:42:27 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 201102 | RDS | CMPLT | CMPLT | 2024-01-19 13:44 | 2024-01-19 13:45 | Weekly Rounds: 11 North | (TC 01/19/2024 01:44:18 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 201103 | CM | CMPLT | CMPLT | 2024-01-19 13:45 | 2024-01-19 13:45 | worked on Masimo | worked on Masimo
1064421 | 201104 | CM | CMPLT | CMPLT | 2024-01-19 13:56 | 2024-01-19 14:00 | Keyboard cover ripped | (TC 01/19/2024 01:56:15 PM, Cortez, John) Replaced keyboard cover
1069679 | 201105 | PM | CMPLT | CMPLT | 2024-01-19 14:01 | 2024-01-19 14:23 | PM-MAJOR - Annual PM For: High Frequency Desiccator with Cut | (TC 01/19/2024 02:00:36 PM, Rhuland, Courtney) Device was found in urology untagged Brought over and performed PM All tests passed Refer to attatched checklist for details Electrical safety: leakage current: 9.9uA Resistance: 226mOhms
1063746 | 201106 | CM | CMPLT | CMPLT | 2024-01-19 14:11 | 2024-01-19 14:11 | Cleaning | (TC 01/19/2024 02:11:12 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
69292 | 201107 | CM | CMPLT | CMPLT | 2024-01-19 14:12 | 2024-01-19 14:12 | Functional check/ check in | (TC 01/19/2024 02:11:47 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
| 201108 | RDS | CMPLT | CMPLT | 2024-01-19 14:24 | 2024-01-19 14:26 | Weekly Rounds: NICU and RESUS | (TC 01/19/2024 02:24:32 PM, Baghirzade, Nara) Weekly rounds in NICU and RESUS completed. See procedure for results.
| 201109 | MEP | CMPLT | CMPLT | 2024-01-19 15:13 | 2024-09-25 09:36 | MSH US DATABASE REVIEW OF SYSTEM STATUS | (TC 01/25/2024 02:37:15 PM, Whittington, Scott) MSH OPG SPP ZONARE LOCATED IN OPG 3-804 B1 MATERNAL/FETAL MEDICINE MSH OPG SPP ZONARE 502237 201219 | MSH US DATABASE REVIEW OF SYSTEM STATUS
1049901 | 201110 | CM | CMPLT | CMPLT | 2024-01-20 12:56 | 2024-01-20 12:57 | MSH R61 NICU Revo - Reported Not Working (Top Monitor) | (TC 01/19/2024 12:56:09 PM, Chau, Ken) User Reported top monitor not working -went to inspect unit and fully functional -restarted machine and discuss with tech -tech reported later that it was his badge not working, tested on other machine and his badge is working except that one machine. Super user to add user into R61
65417 | 201111 | CM | CMPLT | CMPLT | 2024-01-22 08:13 | 2024-01-22 08:15 | Broken door hinge | (TC 01/22/2024 08:14:04 AM, Rhuland, Courtney) Replaced broken side door hinge piece Tightened drawer screws Functional check - OK
| 201112 | OC | CMPLT | CMPLT | 2024-01-22 08:25 | 2024-02-26 08:54 | Emmerson Reid On-Call Jan 19 - Jan 25 2024 | (TC 01/20/2024 08:25:01 AM, Reid, Emmerson ) Recieved a call from 10 M&B Their Masimo monitor (Asset# 1057283) had a redical that would not stop beeping. I instructed them to turn off the redical by holding the power button for 5-10 seconds and then turning it back on. It stopped beeping after that and was working. I had asked them if there was any error messages but they did not see any.
| 201113 | RDS | CMPLT | CMPLT | 2024-01-22 08:27 | 2024-01-22 08:27 | weekly check for 14S and 16N | weekly check for 14S and 16N
66680 | 201114 | CM | CMPLT | CMPLT | 2024-01-22 09:02 | 2024-01-22 09:06 | Broken latch and drawer | (TC 01/22/2024 09:02:24 AM, Baghirzade, Nara) Wrong door was placed resulting in the latch to break. Replaced with the correct side door. Repaired the latch as well. Drawer had all screws missing on the right rail; replaced screws and secured drawer to the rail. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
| 201115 | RDS | CMPLT | CMPLT | 2024-01-22 09:11 | 2024-01-22 09:13 | Weekly Checks For PACU | (TC 01/22/2024 09:12:31 AM, Reid, Emmerson ) Rounds complete
69274 | 201116 | CM | CMPLT | CMPLT | 2024-01-22 09:26 | 2024-01-22 09:29 | Reported "channel error" | (TC 01/22/2024 09:26:14 AM, Baghirzade, Nara) Downloaded error log - found errors pertaining to pressure sensor failure. Performed analog sensors test; bottle side sensor fails as it holds 4.9V and does not decrease. Replaced bottle side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
1055103 | 201117 | CM | CMPLT | CMPLT | 2024-01-22 09:47 | 2024-01-22 09:48 | Check in/Functional test | (TC 01/22/2024 09:47:16 AM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 98 mohms Leakage current: 67.5 uA
66054 | 201118 | CM | CMPLT | CMPLT | 2024-01-22 09:48 | 2024-01-22 09:49 | Check in/Functional test | (TC 01/22/2024 09:48:21 AM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 56 mohms Leakage current: 95.7 uA
1069765 | 201119 | ACC | CMPLT | CMPLT | 2024-01-22 09:57 | 2024-01-23 10:49 | Incoming inspection | (TC 01/23/2024 10:46:38 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1048340 | 201120 | CM | CMPLT | CMPLT | 2024-01-22 10:03 | 2024-01-22 10:06 | Functional check - repurposed or urology | (TC 01/22/2024 10:03:30 AM, Rhuland, Courtney) Functional check - OK Relocated unit to 60 Murray - Urology Replaced SPO2 finger sensor Showed PCM how to use device
1059778 | 201121 | FC | CMPLT | CMPLT | 2024-01-22 10:03 | 2024-01-22 10:04 | QC checklist | (TC 01/22/2024 10:03:08 AM, Baghirzade, Nara) QC Checklist completed. Returned to 11 South
| 201122 | CM | ACTIV | ISSUE | 2024-01-22 10:07 |  | Philips NIBP upgrade with new connectors and Fetal monitor upgrade | (TC 01/22/2024 10:07:34 AM, Chukri, Samir) Review the list of the unit, and the quotation from Philips. (TC 02/06/2024 11:01:03 AM, Chukri, Samir) meeting with Maish, Kamran, tomlyn and Becky. (TC 02/07/2024 02:09:23 PM, Chukri, Samir) Discuss with Maisha of the quantity of unit need nibp upgrade (TC 02/27/2024 01:53:13 PM, Chukri, Samir) NIBP update meeting approved by Adeel (TC 04/03/2024 12:41:37 PM, Chukri, Samir) meeting with Adeel, Kamran and Elizabeth for Philips Fetal moitor upgrade
1069766 | 201123 | ACC | CMPLT | CMPLT | 2024-01-22 10:12 | 2024-01-23 10:50 | Incoming incoming | (TC 01/23/2024 10:49:23 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
68922 | 201124 | CM | CMPLT | CMPLT | 2024-01-22 10:13 | 2024-01-22 10:16 | Broken power cord | (TC 01/22/2024 10:15:55 AM, Baghirzade, Nara) Replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 48 mohms Leakage current: 74.4 uA
1069767 | 201125 | ACC | CMPLT | CMPLT | 2024-01-22 10:25 | 2024-01-23 10:50 | Incoming inspection | (TC 01/23/2024 10:49:43 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1054522 | 201126 | CM | CMPLT | CMPLT | 2024-01-22 10:36 | 2024-01-22 15:08 | No reading on leads | (TC 01/22/2024 11:20:33 AM, Canjimavadivel, Kamallesh ) Tested with simulator, all leads working fine. No visible breaks seen on leads
68291 | 201127 | CM | CMPLT | CMPLT | 2024-01-22 10:50 | 2024-01-22 10:54 | Not populating to iAware | (TC 01/22/2024 10:50:45 AM, Baghirzade, Nara) Checked the network settings; unit is associated and has the correct server. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 78 mohms Leakage current: 74.9 uA
| 201128 | CM | CMPLT | CMPLT | 2024-01-22 10:56 | 2024-01-22 10:56 | meeting | meeting
69086 | 201129 | CM | CMPLT | CMPLT | 2024-01-22 12:52 | 2024-01-22 12:52 | Check-in/Functional check | (TC 01/22/2024 12:52:03 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. | Failed rate accuracy during check in
1050274 | 201130 | CM | CMPLT | CMPLT | 2024-01-22 12:53 | 2024-01-22 12:53 | Check-in/functional check | (TC 01/22/2024 12:53:09 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. | Failed rate accuracy during check-in
| 201131 | CM | CMPLT | CMPLT | 2024-01-22 14:03 | 2024-01-22 14:03 | PM Huddle/Update | (TC 01/22/2024 02:02:58 PM, Baghirzade, Nara) PM Huddle/Update
1057119 | 201132 | CM | CMPLT | CMPLT | 2024-01-22 14:11 | 2024-01-22 14:13 | Reported that the thermometer is not working | (TC 01/22/2024 02:11:26 PM, Baghirzade, Nara) No issue found with the temperature probe. Checked the temperature on ORAL, ADULT AUX, and PED AUX - all modes are functional . Functional testing passed; NIBP and SpO2 modules are functional.
| 201133 | CM | CMPLT | CMPLT | 2024-01-22 14:15 | 2024-01-22 14:15 | meeting | meeting
| 201134 | CM | CMPLT | CMPLT | 2024-01-22 14:16 | 2024-01-24 07:40 | Ordering sterile water for NICU | (TC 01/22/2024 02:17:25 PM, Baghirzade, Nara) SC 1003152242 (TC 01/24/2024 07:39:22 AM, Baghirzade, Nara) Order# 4150270666
| 201135 | CM | CMPLT | CMPLT | 2024-01-22 14:17 | 2024-01-22 14:17 | worked with RT | worked with RT
1060030 | 201136 | CM | CMPLT | CMPLT | 2024-01-22 14:21 | 2024-02-14 13:01 | Ordering PM Kits | (TC 01/22/2024 02:32:55 PM, Baghirzade, Nara) SC 1003152246 (TC 01/26/2024 07:20:58 AM, Baghirzade, Nara) Received all parts except for the Hypair/Body Box 02 cell UFO-130. PO 7202351463 (TC 02/14/2024 01:00:25 PM, Baghirzade, Nara) Received complete order. PO 7202351463
1046060 | 201137 | CM | ACTIV | ISSUE | 2024-01-22 14:27 |  | Lamp Burnt Out | (TC 01/22/2024 02:27:22 PM, Reid, Emmerson ) Investigated and found that the lamp is burnt out. A new one needs to be ordered. Adding Jessica to take over as I will be away. (TC 04/12/2024 11:09:22 AM, Reid, Emmerson ) Reached out to valleymed the vendor for a quote
66509 | 201138 | CM | CMPLT | CMPLT | 2024-01-22 15:56 | 2024-01-29 08:53 | SpO2 not reading | (TC 01/22/2024 04:18:12 PM, Rendon Lopez , Elizabeth) Testing monitor in room, alarm unknown SpO2 sensor. Removed from room Board was previously replaced. ordering replacement boards for Troubleshooting (TC 01/25/2024 08:51:36 AM, Rendon Lopez , Elizabeth) Noticied fluid on boards. Cleaned Spo2 and API board, tested, working properly. All functions working. **kept new boards as spare**
1049903 | 201139 | CM | CMPLT | CMPLT | 2024-01-22 17:33 | 2024-01-24 17:21 | MSH R63 Short life on system batteries. | (TC 01/22/2024 05:36:39 PM, Dolinar, John) MSH R63 Short life on system batteries. Ordered new batteries Batteries replaced. (TC 01/24/2024 05:24:22 PM, Dolinar, John) MSH R63 Short life on system batteries. Ordered new batteries Batteries replaced. (TC 01/24/2024 05:24:30 PM, Fan, Tak) MSH R63 Short life on system batteries. Ordered new batteries Batteries replaced. | MSH R63 Short life on system batteries. Ordered new batteries Batteries replaced.
| 201140 | CM | CMPLT | CMPLT | 2024-01-22 17:35 | 2024-01-22 17:35 | Site coverage
1060518 | 201141 | CM | CMPLT | CMPLT | 2024-01-22 19:03 | 2024-01-22 19:02 | Table height motion problem. | Work in progress, ordered more parts.
67958 | 201142 | CM | CMPLT | CMPLT | 2024-01-23 07:52 | 2024-01-23 10:05 | Broken latch | (TC 01/23/2024 10:04:20 AM, Rhuland, Courtney) Replaced broken latch Humdifier test failed Calibrated humidity Functional check - OK
| 201143 | CM | CMPLT | CMPLT | 2024-01-23 10:01 | 2024-01-23 10:07 | CO2 laser servicing | (TC 01/23/2024 10:01:01 AM, Baghirzade, Nara) Tyler from NOI was on site to service/replace the CO2 arm on the laser. Fired the laser at a thermal paper - no issue with beam dispersion. Energy tests at low and high end were good. Waited for Dr. Witterick as he was the one who reported a double beam issue. When coupled with the microscope, there are 2 beams that appear on the subject; the beam also appears to be dispersing. When looking at the micromanipulator, the lens looks scratch. Tyler tried to clean it multiple times however it did not help mitigate the issue. There are deep scratches on the micromanipulator lens (glass). Let Matt and Dr. Witterick know about the findings. Dr did a couple of fire tests and was comfortable to use it for the upcoming cases. Updated Jessica regarding the findings and pictures of the scratches, the beam dispersion, and part number for the micromanipulator. Jessica will contact Wally from Boston Scientific to purchase another micromanipulator. Will update Tyler and NOI regarding the double beam and beam dispersion once part arrives.
| 201144 | CM | CMPLT | CMPLT | 2024-01-23 10:02 | 2024-01-23 10:02 | worked on Masimo | worked on Masimo
1056895 | 201145 | CM | CMPLT | CMPLT | 2024-01-23 10:03 | 2024-01-23 10:10 | Needed to be balanced | (TC 01/23/2024 10:09:10 AM, Reid, Emmerson ) went up and balanced the microscope, working now.
1065046 | 201146 | CM | CMPLT | CMPLT | 2024-01-23 10:07 | 2024-01-23 10:10 | Not turning on | (TC 01/23/2024 10:07:16 AM, Rhuland, Courtney) Device would not turn on - no AC/battery light Tried replacing battery - no change Fuses were OK on board and plug Tried replacing power control board - no change Front case found to be faulty Replaced old power control board and new front case Device turned on Functional check - OK
| 201147 | RDS | CMPLT | CMPLT | 2024-01-23 10:08 | 2024-01-23 11:55 | Weekly Checks in CCU | (TC 01/23/2024 11:54:36 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
1061217 | 201148 | CM | CMPLT | CMPLT | 2024-01-23 10:11 | 2024-01-23 10:11 | Music cable broke inside PC | (TC 01/23/2024 10:10:45 AM, Reid, Emmerson ) Tried to remove with tweesers but was not able to. Called hepdesk to repair as they look after the PC
1061209 | 201149 | CM | CMPLT | CMPLT | 2024-01-23 10:12 | 2024-01-23 10:13 | Light drifting up | (TC 01/23/2024 10:11:53 AM, Reid, Emmerson ) Light needed to be balanced. I adjusted the arm, working now. (TC 01/23/2024 10:11:53 AM, Reid, Emmerson ) Light needed to be balanced. I adjusted the arm, working now.
1061333 | 201150 | CM | CMPLT | CMPLT | 2024-01-23 10:13 | 2024-03-21 09:37 | Shuts off on its own in middle of therapy | (TC 01/23/2024 10:12:52 AM, Baghirzade, Nara) Sarabeth requested to meet to discuss the issue with the VAC; it shut off in the middle of therapy on its own. The screen went blank and the unit powered down. Confirmed that it was plugged to AC the entire time. They switched it back on and it's currently on a patient. E-mailed Ifrah from 3M to help connect with the technical team to troubleshoot. This VAC was previously sent to 3M/KCI for the same issue but they were unable to find the problem or replicate the issue. (TC 01/30/2024 12:31:06 PM, Baghirzade, Nara) Met with Sarabeth to see if the issue occurred again; it did not happen again the entire time of the therapy. Requested that Sarabeth takes a video or picture of the VAC when it shuts off to provide it to KCI. DId not hear back from KCI regarding the issue; will send a follow up. (TC 02/13/2024 02:02:10 PM, Baghirzade, Nara) Troubleshooting with Keith from KCI/3M; looked over the previous service report and noted the tests that were done. I let him know that last time the entire therapy was interrupted while it was on a patient. He suggested that main board will have to be tested and most likely replaced. Spoke about the downtime and they will treat it as high priority. No cost associated with the repair yet as they need to troubleshoot to determine if the board will have to be replaced. Contacted Sarabeth for approval to send this device out and let her know all the details. Will contact Ifrah to schedule a pickup. (TC 02/13/2024 02:09:07 PM, Baghirzade, Nara) Contacted Ifrah to request pickup. Waiting on the pickup confirmation and reference #. (TC 02/14/2024 07:48:29 AM, Baghirzade, Nara) RMA# 13093142 (TC 02/14/2024 01:58:53 PM, Baghirzade, Nara) Pickup receipt attached. (TC 02/27/2024 11:57:55 AM, Baghirzade, Nara) Followed up with KCI/3M team, their response: "We did find that the therapy turned off in our testing. This will be discussed further with our team at our team at head office. Rizwan will return tomorrow, so we can discuss this further." Updated Sarabeth and Katherine. (TC 02/29/2024 03:04:05 PM, Baghirzade, Nara) Followed up with the vendor; they mentioned that they ran the unit again and encountered the issue of the display switching off in the middle of therapy. Rizwan (service tech) suspects that the display board may be faulty; they sent the most recent logs over to the main office for inspection and will follow up with me when they have an answer. (TC 03/01/2024 03:13:22 PM, Baghirzade, Nara) KCI sent a repair quote and details; the main board was replaced as it was shutting off during therapy. Forwarded the quote to Katie and Sarabeth for approval. (TC 03/07/2024 10:28:46 AM, Baghirzade, Nara) SC 1003184857 (TC 03/11/2024 07:30:54 AM, Baghirzade, Nara) Sent signed repair estimate form and PO to KCI for repair approval. Attached the signed repair estimate to this WO. (TC 03/21/2024 09:33:49 AM, Baghirzade, Nara) Received unit from KCI. Main board was replaced. Attached QC checklist from KCI, drop-off receipt, and estimate containing main board replacement/service report details. Sent VAC to 11N.
1061211 | 201151 | CM | CMPLT | CMPLT | 2024-01-23 10:13 | 2024-01-23 10:14 | VMC acting abnormally | (TC 01/23/2024 10:13:21 AM, Reid, Emmerson ) VMC needed to be restarted. Working now.
| 201152 | CM | CMPLT | CMPLT | 2024-01-23 10:16 | 2024-01-23 10:18 | Meeting with Sri regarding Stryker in-servicing | (TC 01/23/2024 10:16:00 AM, Baghirzade, Nara) Sri requested to meet regarding the Stryker beds in-servicing. Shes concerned that the clinical staff don't know how to use the beds properly as there has been a high staff turnover. There is confusing regarding alarms, i-bed, the different color of lights and their meaning. Contacted Becky for her to request the in-servicing.
52117 | 201153 | CM | CMPLT | CMPLT | 2024-01-23 10:20 | 2024-02-09 08:29 | Ordering parts for repair | (TC 01/23/2024 10:27:25 AM, Baghirzade, Nara) SC 1003152733 (TC 01/25/2024 09:33:24 AM, Baghirzade, Nara) Received east lift rail and arrow fastener. Backorder: nyl patch PO 7202351747 (TC 02/09/2024 08:27:59 AM, Baghirzade, Nara) Received full order. PO 7202351747 | Broken belt; unable to move canopy.
65149 | 201154 | CM | CMPLT | CMPLT | 2024-01-23 10:36 | 2024-01-23 10:37 | Screen wont come on | (TC 01/23/2024 10:36:41 AM, Reid, Emmerson ) There were no batteries. I added batteries and it turned on. Working now.
1069751 | 201155 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 10:51 | Incoming inspection | (TC 01/23/2024 10:50:37 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069752 | 201156 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 10:52 | Incoming inspection | (TC 01/23/2024 10:51:55 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069753 | 201157 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:09 | Incoming inspection | (TC 01/23/2024 11:08:55 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069754 | 201158 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:10 | Incoming inspection | (TC 01/23/2024 11:09:22 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069755 | 201159 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:10 | Incoming inspection | (TC 01/23/2024 11:09:43 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069756 | 201160 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:10 | Incoming inspection | (TC 01/23/2024 11:10:08 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069757 | 201161 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:11 | Incoming inspection | (TC 01/23/2024 11:10:30 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069758 | 201162 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:12 | Incoming inspection | (TC 01/23/2024 11:10:48 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069759 | 201163 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:12 | Incoming inspection | (TC 01/23/2024 11:12:16 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069760 | 201164 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:13 | Incoming inspection | (TC 01/23/2024 11:12:35 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069761 | 201165 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:13 | Incoming inspection | (TC 01/23/2024 11:12:57 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069762 | 201166 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:14 | Incoming inspection | (TC 01/23/2024 11:13:19 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069763 | 201167 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:14 | Incoming inspection | (TC 01/23/2024 11:13:53 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069764 | 201168 | ACC | CMPLT | CMPLT | 2024-01-23 10:39 | 2024-01-23 11:15 | Incoming inspection | (TC 01/23/2024 11:14:24 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1069765 | 201169 | ACC | CANCL | DUP | 2024-01-23 10:39 |  | Incoming inspection
1069766 | 201170 | ACC | CANCL | DUP | 2024-01-23 10:39 |  | Incoming inspection
1069767 | 201171 | ACC | CANCL | DUP | 2024-01-23 10:39 |  | Incoming inspection
67160 | 201172 | CM | CMPLT | CMPLT | 2024-01-23 11:04 | 2024-01-23 11:06 | Not reading BP | (TC 01/23/2024 11:04:07 AM, Rhuland, Courtney) Brought extra cuff over No cuffs were attatched to unit Reseated BP tube and replaced with spare cuff Functional check - OK
1056570 | 201173 | CM | CMPLT | CMPLT | 2024-01-23 11:20 | 2024-01-23 11:22 | Relocate from cabinet to cart | (TC 01/23/2024 11:20:31 AM, Canjimavadivel, Kamallesh ) Request from RT to move AED from cabinet to cart in the same area. Cabinet no longer in use
1055928 | 201174 | CM | CMPLT | CMPLT | 2024-01-23 11:57 | 2024-01-23 11:58 | Reported that the BP cuff does not inflate | (TC 01/23/2024 11:56:38 AM, Baghirzade, Nara) Checked BP multiple times with different cuffs - no issue with BP inflation. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1069776 | 201175 | CM | CMPLT | CMPLT | 2024-01-23 13:58 | 2024-01-23 13:59 | Reading higher than expected | (TC 01/23/2024 01:57:59 PM, Canjimavadivel, Kamallesh ) SN was not found in TMS. Added asset and tagged. Tested with calibrator, passed after 1 attempt.
1058274 | 201176 | CM | CMPLT | CMPLT | 2024-01-23 14:08 | 2024-01-23 14:09 | Does not turn on | (TC 01/23/2024 02:08:03 PM, Baghirzade, Nara) Replaced battery. Device switches on. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 78 mohms Leakage current: 65.3 uA
1058282 | 201177 | CM | CMPLT | CMPLT | 2024-01-23 14:10 | 2024-01-23 14:10 | Reported that it beeps and is stuck on the message to change the battery | (TC 01/23/2024 02:09:41 PM, Baghirzade, Nara) Replaced battery. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 50 mohms Leakage current: 67.7 uA
1063802 | 201178 | PM | CMPLT | CMPLT | 2024-01-23 14:10 | 2024-01-23 14:14 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 01/23/2024 02:12:30 PM, Rhuland, Courtney) PM complete RF leakage: 6.7uA See attached document for details
68703 | 201179 | CM | CMPLT | CMPLT | 2024-01-23 14:12 | 2024-01-23 14:13 | Repeated unknown errors (user reported) | (TC 01/23/2024 02:12:11 PM, Baghirzade, Nara) Power cord was damaged; replaced power cord. Downloaded error log; only noted errors that are log only and transmit error log only. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 38 mohms Leakage current: 74.2 uA
1062022 | 201180 | CM | CMPLT | CMPLT | 2024-01-23 14:41 | 2024-02-01 10:20 | Loud noise when raising the head portion | (TC 01/23/2024 02:40:58 PM, Baghirzade, Nara) When raising the head portion of the bed, its very loud. When lowering it, there is no noise. Created a service request with Stryker (TC 01/31/2024 02:50:00 PM, Baghirzade, Nara) Checked all the bed motion controls; all functional. Waiting on service report from Stryker. (TC 02/01/2024 10:19:38 AM, Baghirzade, Nara) Attached service report
| 201181 | CM | CMPLT | CMPLT | 2024-01-23 15:27 | 2024-01-23 15:28 | NICU Huddle | (TC 01/23/2024 03:27:27 PM, Baghirzade, Nara) NICU huddle
| 201182 | PA | CMPLT | CMPLT | 2024-01-24 06:24 | 2024-01-30 09:38 | Ordering parts for stock | (TC 01/24/2024 06:28:58 AM, Baghirzade, Nara) SC 1003153552 (TC 01/30/2024 09:37:54 AM, Baghirzade, Nara) PO 7202352416
69083 | 201183 | CM | CMPLT | CMPLT | 2024-01-24 06:52 | 2024-01-24 06:53 | Functional check / check-in | (TC 01/24/2024 06:52:31 AM, Baghirzade, Nara) Check-in tests completed. IUI connectors and device cleaned.
69361 | 201184 | CM | CMPLT | CMPLT | 2024-01-24 07:05 | 2024-01-24 07:07 | Functional test/check in | (TC 01/24/2024 07:05:26 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
| 201185 | CM | CMPLT | CMPLT | 2024-01-24 07:51 | 2024-01-24 07:55 | Check light source for leaks | (TC 01/23/2024 07:51:10 AM, Canjimavadivel, Kamallesh ) Worked with John and opened up light sources in room 1,2,3 to check to liquid residue in tubes
| 201186 | RDS | CMPLT | CMPLT | 2024-01-24 08:09 | 2024-01-25 08:52 | Weekly check for ED
| 201187 | RDS | CMPLT | CMPLT | 2024-01-24 08:11 | 2024-01-25 14:07 | Weekly check for ICU
1061316 | 201188 | CM | CMPLT | CMPLT | 2024-01-24 08:17 | 2024-01-30 15:17 | Can't burn pt CD from VP | (TC 01/23/2024 04:17:00 PM, Fan, Tak) Met with Dianne to advise about how to store images externally/anonymously. The researchers request anonymizing by changing the patient name then storing, but the DICOM tags still have patient names attached to image properties, which is a privacy concern they probably don't know about. Unable to burn CD, encrypted USB does not work either. Contacted Sinai IT team. #30733283 (TC 01/24/2024 01:56:16 PM, Fan, Tak) Worked with Ihor from IT. Encrypted USB is no longer compatible with security software. Disc permissions were gone due to reimaging at one point. Tested Echo's encrypted usb with new SW, works, Ihor upgraded the security software. Burned patient discs on vivid but it failed 2x on finalizing. Copied via usb and burned. (TC 01/25/2024 03:12:10 PM, Fan, Tak) Computer being replaced by IT. (TC 01/30/2024 03:16:11 PM, Fan, Tak) Computer replaced. DVD permissions added. Tested, OK.
1064418 | 201189 | CM | CMPLT | CMPLT | 2024-01-24 08:50 | 2024-02-01 09:38 | Head Coil not working | (TC 01/24/2024 08:50:14 AM, Ferreira, Qwynn) MRI Coil was not working Followed troubleshooting tips from LMT and found there was one pin that was open Notified LMT and shipped the coil out (TC 01/24/2024 08:51:05 AM, Ferreira, Qwynn) Michael dropped the coil off at imaging to Jason - waiting on report (TC 02/01/2024 09:37:56 AM, Ferreira, Qwynn) Report was sent over and attached. Coil was tested and working Back in use now
| 201190 | PA | CMPLT | CMPLT | 2024-01-24 09:05 | 2024-02-13 07:45 | Ordering PM Kits for T1 ventilators | (TC 01/24/2024 09:22:52 AM, Canjimavadivel, Kamallesh ) Placed order for PM Kits (TC 02/12/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Received PM Kits
1047913 | 201191 | CM | CMPLT | CMPLT | 2024-01-24 09:23 | 2024-04-23 09:48 | "water pressure low" | (TC 01/23/2024 09:23:23 AM, Cortez, John) Was able to repeat error. Error pops up after 200-300 shocks with F-SW handpiece. Renewed water, and it temporarily resolves issue Reached out to Remington Medical tech support (TC 01/24/2024 09:51:27 AM, Cortez, John) Tom from Remington Medical reached back and requested pictures of hand piece and F-SW unit. He suggested it could be the hand piece or water pump. Included Gabby (Dovigi clinic manager) in email thread to verify clinic scheduling and proper next steps. Tom suggested to send out loaner hand piece, to swap in when users come across error again. Will request n/c PO (TC 01/25/2024 10:43:01 AM, Cortez, John) Requesting no-charge PO from plexxus (TC 01/25/2024 01:15:36 PM, Cortez, John) Sent PO 7500201181 to Tom (TC 01/30/2024 10:24:10 AM, Cortez, John) Received loaner hand piece from Tom (TC 01/30/2024 04:26:39 PM, Cortez, John) Dropped off loaner HP in room, informed Gabby (TC 02/20/2024 02:11:37 PM, Cortez, John) Melanie (physiotherapist, Dovigi clinic called saying they tried swapping F-SW hand piece, but control unit did not recognize it. Also said there was a leak underneath unit Reached out saying I would check it the next morning (TC 02/21/2024 02:15:05 PM, Cortez, John) Checked unit and confirmed leak underneath F-SW unit. Was able to swap handpiece with no issue, but control unit needed to update software for compatibility. Also noticed the R-SW was making repetitive noise (pump?) Called Tom (service tech, Remington Medical) for assistance. He suggested leak could be from tubing, reservoir or pump. Will need to send it in for repair. He suggested sound from R-SW unit might be related to faulty F-SW module. Informed nurse desk that I would take unit down to be sent out. (TC 02/21/2024 02:21:12 PM, Cortez, John) Disassembled cart and shipped FSW module to Remington Medical Reassembled rest of cart and sent back to procedure room Informed Melanie, and Gabby (TC 02/23/2024 03:21:13 PM, Cortez, John) shipped back loaner hand piece (TC 02/28/2024 01:49:16 PM, Cortez, John) Requested PO for repair, updated Gabby (PCM) (TC 03/05/2024 04:24:16 PM, Cortez, John) Repaired unit returned. Reassembled F-SW to cart Refilled water circuit Tested handpiece up to 500 shocks, no issues Returned unit to Dovigi clinic, informed Gabby (TC 03/04/2024 04:24:16 PM, Cortez, John) Repaired unit returned. Reassembled F-SW to cart Refilled water circuit Tested handpiece up to 500 shocks, no issues Returned unit to Dovigi clinic, informed Gabby
1055048 | 201192 | CM | CMPLT | CMPLT | 2024-01-24 10:16 | 2024-01-24 10:26 | Battery replacement | (TC 01/24/2024 10:24:43 AM, Rhuland, Courtney) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055049 | 201193 | CM | CMPLT | CMPLT | 2024-01-24 10:27 | 2024-01-24 10:27 | Battery replacement | (TC 01/24/2024 10:27:06 AM, Rhuland, Courtney) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
68384 | 201194 | CM | CMPLT | CMPLT | 2024-01-24 10:40 | 2024-04-03 14:18 | Not passing Self test. | (TC 01/24/2024 11:00:18 AM, Scero, Tomlyn) Confirmed with Diva tests that the Refill valve does not open. Valve and leakage test failed. Will request a PO from Medbuy to send the device to Draeger for repair in Germany. Waiting on a quote from Draeger. (TC 01/24/2024 01:48:06 PM, Scero, Tomlyn) Recieved RMA 309000871 and shipped box with KJV (TC 03/27/2024 11:10:27 AM, Scero, Tomlyn) NOtified by Draeger that repair was complete and unit is being shipped back to me tdoay. Service report attached. (TC 04/03/2024 02:11:42 PM, Scero, Tomlyn) device received and unit put back into service.
1054224 | 201195 | CM | CMPLT | CMPLT | 2024-01-24 11:01 | 2024-01-24 11:25 | Not reading properly | (TC 01/24/2024 11:24:02 AM, Canjimavadivel, Kamallesh ) Cleaned residue from lense tip. Tested with calibrator, passed after 1 attempt.
66055 | 201196 | CM | CMPLT | CMPLT | 2024-01-24 11:13 | 2024-01-26 12:14 | Broken front case | (TC 01/24/2024 01:38:57 PM, Rhuland, Courtney) Replaced cracked front case Replaced battery - conditioning battery (TC 01/26/2024 12:13:26 PM, Rhuland, Courtney) Check in complete All tests passed Refer to ASM for details
| 201197 | CM | ACTIV | ISSUE | 2024-01-24 11:34 |  | Supervisor Miscellaneous | (TC 01/24/2024 11:34:01 AM, Chukri, Samir) Helpping Elizabeth and ordered tube connector for her test equipment (TC 01/30/2024 10:58:27 AM, Chukri, Samir) Helpping Qwyn and ordered iteams from amazon for 3D lab and search for Masimo nurse call for Nara. (TC 02/01/2024 09:32:33 AM, Chukri, Samir) Read andd replied my email . (TC 02/05/2024 08:59:36 AM, Chukri, Samir) Reply to my emails (TC 02/07/2024 12:54:49 PM, Chukri, Samir) Repaly for phone call from Naz regarding baby basent disposable and Courteny refer ECG check up in 60 Murray st clinic (TC 02/08/2024 01:01:49 PM, Chukri, Samir) explain and traning about the SAP and ordering parts to Rutvik (TC 02/12/2024 10:03:49 AM, Chukri, Samir) review and reply to my emails- follow up with Rutvik regarding the brother printer and test software download into his new computer (TC 02/12/2024 12:38:37 PM, Chukri, Samir) BD meeting with Adeel and Maisha, went to the 18 floor to check empty room for BD Tech . (TC 02/13/2024 02:38:40 PM, Chukri, Samir) Reply to my emails, and help Maisha for the BD steps , review and approved shopping carts. (TC 02/14/2024 11:13:08 AM, Chukri, Samir) Reply to my emails and review and approved shopping carts (TC 02/16/2024 11:16:06 AM, Chukri, Samir) Review the Medline Pump Loaner Agreement and emailed to Becky and Adeel. working on the JHSC inspection reports, printing a monthly and annual log sheet for the Ey/Face wash. reply to Cole regarding the esoph/rectal probe. (TC 02/20/2024 08:30:04 AM, Chukri, Samir) requested ceiling tiles repair from building service. (TC 02/20/2024 09:57:40 AM, Chukri, Samir) Received Nara new workstation and help Nara to moved all the boxes to biomed main door (TC 02/20/2024 02:37:23 PM, Chukri, Samir) Completed the second Eye/wash station inspection in workshop room. (TC 02/21/2024 09:35:08 AM, Chukri, Samir) Working with Rutvik to solve the Demo devices arrived to Endo unit misssing the CSA and missing the incoming inspection and safety. (TC 02/23/2024 09:41:52 AM, Chukri, Samir) working on Co-Op summer student for interview schedule (TC 02/23/2024 02:29:03 PM, Chukri, Samir) contact service support for ceiling tile repair and clean a corner beside the sink for mouse dirt (TC 02/26/2024 11:16:48 AM, Chukri, Samir) BD upgrade work (TC 02/28/2024 09:22:33 AM, Chukri, Samir) ordered fan for Qwynn. Amazon. (TC 02/29/2024 01:48:03 PM, Chukri, Samir) meeting with Jason from Gatinge (TC 03/04/2024 02:59:50 PM, Chukri, Samir) Airseal insufflator meeting (TC 03/06/2024 08:44:30 AM, Chukri, Samir) Talking and helping Rutvik for mangeing a power bar 20Am for Endo. (TC 03/13/2024 11:19:00 AM, Chukri, Samir) TMS and OR asset meeting (TC 03/14/2024 09:23:05 AM, Chukri, Samir) Helping Courtney and import 20 assets for Symphone Breasr pump (TC 03/18/2024 10:21:00 AM, Chukri, Samir) meeting with BD tech regarding alaris pump with IUI connection issue (TC 03/18/2024 11:22:00 AM, Chukri, Samir) working and investigating for Lifenet webside to correct Kam roll. (TC 03/18/2024 02:24:52 PM, Chukri, Samir) emiled Alexandra from Med-Set for X3 Batteries and working on the PFT lab patient block to replaced valve disck. (TC 03/19/2024 09:55:52 AM, Chukri, Samir) checked the two Eye-wash Stations (TC 03/19/2024 02:11:36 PM, Chukri, Samir) helping Rutvik for CO@ cylinder regulator leaking issue in Endo . (TC 03/20/2024 11:05:08 AM, Chukri, Samir) O.R equipment repair discussion meeting (TC 03/20/2024 02:14:59 PM, Chukri, Samir) explained to Courtney the Micron Filter position with Lipids isue. (TC 03/21/2024 02:47:28 PM, Chukri, Samir) Meeting with Adeel for Emergency Power interruption on L&D (TC 03/25/2024 02:29:15 PM, Chukri, Samir) Replaced order and follow up on Amazon for DC adpeter to Milk bank and emailed to them, also ordered filter for Qwynn . (TC 03/26/2024 02:48:45 PM, Chukri, Samir) ordered and received mail box and installed. (TC 03/27/2024 01:16:39 PM, Chukri, Samir) Keep searching for the BD pump with the same serial number work with Kamran and courtney try to find the dublicated serial number pump (TC 04/02/2024 09:21:03 AM, Chukri, Samir) Helping Rutvik to complete his Petty cash form . (TC 04/03/2024 09:54:23 AM, Chukri, Samir) working with Courtney and emailed to Mirza to get the Pump with duplicated serial number location . (TC 04/05/2024 10:56:05 AM, Chukri, Samir) Meeting with Will from Airseal inseflator (TC 04/09/2024 10:05:34 AM, Chukri, Samir) Meeting with Nima, Andrew and Emmerson regarding Neptune Touch Base and Action Plan. (TC 04/18/2024 01:42:09 PM, Chukri, Samir) contact to Kamran and John regarding the normal power off on Apr 18 2024 (TC 04/29/2024 02:35:34 PM, Chukri, Samir) meet with Kevin regarding the AP in 14 N (TC 07/24/2024 01:11:50 PM, Chukri, Samir) Meeting with Milk Bank department for HPP new unit.
| 201198 | CM | CMPLT | CMPLT | 2024-01-24 12:14 | 2024-01-24 12:14 | meeting | meeting
| 201199 | CM | CMPLT | CMPLT | 2024-01-24 12:15 | 2024-01-24 12:15 | worked at 16N | worked at 16N
1063310 | 201200 | CM | CMPLT | CMPLT | 2024-01-24 12:19 | 2024-05-02 08:06 | User reported that Laser does not work at 5W | (TC 01/24/2024 12:18:47 PM, Baghirzade, Nara) Contacted Tyler from NOI and he requested that we check if the Laser lases above 5W. Jessica and I set up the CO2 laser to test at CW (continous wave) on continous mode, and output checks at: 1-4 W (no error messages, power divergance passes) however at 5W, 5.5W, 6W, 6.5W, and 7W the power divergance fails and laser defaults to standby mode. From 7.5W to 10W the power divergance passes and were able to lase. Contacted Tyler back and he suggested to contact NOI directly for service. Turned off laser, stowed the arm and foot pedal safely, and put the laser back in its place. (TC 02/09/2024 02:34:20 PM, Baghirzade, Nara) Rehan from NOI was on site. He did 2 calibrations but both calibrations failed and were not capturing in the Laser. NOI has yet to diagnose the root of the problem. Rehan gave us 2 options, either to wait to be on site next time with another tool for calibration and parts, second option to send it to the NOI repair facility to get it repaired there. Presented the options to Tania (PCM) and Maureen from OR, final decision to be made by Dr. Witterick. Will try to find a loaner by contacting Wally from Boston Scientific. (TC 02/12/2024 02:54:25 PM, Huang, Jessica) informed OR team and coodinated with NOI for picking up the defective laser on Tuesday Feb 13th (TC 02/28/2024 03:25:27 PM, Huang, Jessica) the laser was sent back from NOI and met Dillon at Murray entrance. took the laser back to the OR noticed the small arm was glued and was gone during transport. infromed Melissa and asked to get the part Dillon tested the laser in OR 15 and confrimed it was working fine Informed the OR team (TC 03/01/2024 03:28:16 PM, Huang, Jessica) got the call from Irina and was informed the laser was not working properly. went to OR 3 to discuss the issues with Dr. Mentiro and Irina. noticed the laser was giving double beams and contacted Rehan. Rehan mentioned the device needed to be removed from service right away. coordinated with Melissa to have Rehan onsite to inspect the device (TC 03/05/2024 03:31:30 PM, Huang, Jessica) Rehan was onsite to check the laser. Contacted Irina to get the instruments for testing. Rehan calibrated the device and ensured the laser beam was properly alighted; however, it would be the temperate soluction. the main arm needed to be replaced. infromed Adeel and got approved contacted Melissa for the quote and coodinated Rehan to be onsite to replace the arm (TC 03/19/2024 03:36:08 PM, Huang, Jessica) received the email from Tania regarding the repair status on the CO2 laser beacause there would be a case on March 25th and Dr. Mentiro would like to have Rehan to be onsite on March 25th. coordinated with Rehan, Melissa and the OR team (TC 03/25/2024 03:38:57 PM, Huang, Jessica) stayed in OR 3 for Dr. Mentiro's laser case to ensure the laser and micromanipulator worked properly (TC 05/02/2024 08:05:10 AM, Baghirzade, Nara) Laser was serviced. Micromanipulator is working properly. No issues reported.
69557 | 201201 | CM | CMPLT | CMPLT | 2024-01-24 12:24 | 2024-01-24 12:26 | Doesn't latch properly to PCU | (TC 01/24/2024 12:24:30 PM, Baghirzade, Nara) Latch and left IUI were heavily coated in some kind of sticky sugar feed or medication. Took the latch off and doused it in alcohol to clean off the gunk. Placed the latch back on the LVP. Cleaned the left IUI as well. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
| 201202 | CM | CMPLT | CMPLT | 2024-01-24 12:27 | 2024-05-02 10:30 | CMAC blade; damaged lens and no visible image | (TC 01/24/2024 12:27:06 PM, Baghirzade, Nara) Met with Michelle to get the blade and information on the issue. Connected the blade to the CMAC in NICU, no image is visible, very blurry, and the blade intermittently disconnects. Contacted Kim from Karl Storz to get a list of the blades that are under warranty for MSH. If not under warranty, will e-mail Steve for approval on sending in the blades for repair. (TC 01/24/2024 01:32:05 PM, Baghirzade, Nara) Requested a warranty list from Kim so we can check for our own warranty. Kim was nearby so she came by for a quick meeting; informed me that the warranty has changed from 2 years to 1 year and only covers manufacturing defects. Kim will follow up with MSH MDRD to confirm that cleaning practices have not changed that could be affecting the blades. Kim will send over an updated list of blades under warranty as well as contract information for MSH pricing and promos. Informed her that we have an additional 3 blades that if they fall under the promo, we would like to trade in for newer model. I tested blade SN: NQ32210 Courtney tested 3 blades: SN: SS8251 (no light) SN: TS9169 (blurry image) SN: TS9161 (blurry image) (TC 01/24/2024 01:35:46 PM, Rhuland, Courtney) Courtney tested 3 blades: SN: SS8251 (no light) SN: TS9169 (blurry image) SN: TS9161 (blurry image) (TC 05/02/2024 10:29:23 AM, Baghirzade, Nara) Followed up with Steve Arai regarding the repair and exchange costs for the blades. He does not want to proceed with it as new blades were already ordered and more will be request through capital.
1065736 | 201203 | CM | ACTIV | PV | 2024-01-24 12:41 |  | PM Quote | (TC 01/24/2024 12:41:32 PM, Baghirzade, Nara) Contact customer service at Clarion for the annual PM quote. Henna mentioned that one of the doctors is having an issue with the joystick as it doesn't move as smoothly. (TC 01/26/2024 07:19:28 AM, Baghirzade, Nara) Received quote from Clarion. Forwarded it to Henna to have Dr. Yan sign and then proceed with scheduling the PM. (TC 02/09/2024 08:27:01 AM, Baghirzade, Nara) PO request. Attached signed copy of quote. SC 1003165915 (TC 02/29/2024 01:21:21 PM, Baghirzade, Nara) Sent a follow up e-mail to Garima from customer service as the service tech has not contacted to schedule for annual PM.
68679 | 201204 | CM | CMPLT | CMPLT | 2024-01-24 12:54 | 2024-01-24 12:56 | User reported that the battery won't hold charge even after charging all day | (TC 01/24/2024 12:54:26 PM, Baghirzade, Nara) Replaced battery. Check-in tests passed. Electrical Safety - Ground resistance: 140 mohms Leakage current: 73.7 uA IUIs and device cleaned.
68898 | 201205 | CM | CMPLT | CMPLT | 2024-01-24 13:13 | 2024-01-24 13:14 | Doesnt hold battery charge | (TC 01/24/2024 01:13:42 PM, Baghirzade, Nara) Battery was replaced. Electrical Safety - Ground resistance: 121 mohms Leakage current: 72.3 uA Error and event logs downloaded and cleared. IUIs and device cleaned. (TC 01/24/2024 01:13:42 PM, Baghirzade, Nara) Battery was replaced. Check-in tests passed. Electrical Safety - Ground resistance: 121 mohms Leakage current: 72.3 uA IUIs and device cleaned.
67008 | 201206 | CM | CMPLT | CMPLT | 2024-01-24 13:29 | 2024-01-24 13:31 | Battery replacement | (TC 01/24/2024 01:30:50 PM, Rhuland, Courtney) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056205 | 201207 | CM | CMPLT | CMPLT | 2024-01-24 15:18 | 2024-01-26 11:32 | Reinstall and patching | (TC 01/24/2024 03:18:30 PM, Canjimavadivel, Kamallesh ) Removed internal usb and reimaged server. Once completed, ran setup and changed server to standby. Ready for patching. (TC 01/24/2024 09:46:25 AM, Cortez, John) Shutdown and restarted server after we completed reimaging. Focal point was sending out communication errors messages. (TC 01/26/2024 11:30:29 AM, Canjimavadivel, Kamallesh ) Changed server back to Physio3 and updated config
| 201208 | CM | CMPLT | CMPLT | 2024-01-24 17:22 | 2024-01-24 17:22 | Site coverage.
1069815 | 201209 | CM | CMPLT | CMPLT | 2024-01-25 07:34 | 2024-01-25 08:14 | Incoming inspection | (TC 01/25/2024 08:14:00 AM, Rhuland, Courtney) Tagged and tested all settings on device
1069816 | 201210 | CM | CMPLT | CMPLT | 2024-01-25 07:45 | 2024-01-25 08:14 | Incoming inspection | (TC 01/25/2024 08:13:42 AM, Rhuland, Courtney) Tagged and tested all settings on device
1069817 | 201211 | CM | CMPLT | CMPLT | 2024-01-25 07:47 | 2024-01-25 08:13 | Incoming inspection | (TC 01/25/2024 08:12:56 AM, Rhuland, Courtney) Tagged and tested all settings on device
| 201212 | CM | CMPLT | CMPLT | 2024-01-25 08:26 | 2024-01-24 08:26 | worked with RT | worked with RT
1067868 | 201213 | CM | CMPLT | CMPLT | 2024-01-25 09:15 | 2024-01-25 13:36 | Otoscope not working | (TC 01/25/2024 09:31:49 AM, Canjimavadivel, Kamallesh ) Bulb was inserted upside down. Reinsterted bulb correctly and tested, working ok.
67932 | 201214 | CM | CMPLT | CMPLT | 2024-01-25 09:27 | 2024-01-25 09:29 | Broken latch | (TC 01/25/2024 09:27:37 AM, Rhuland, Courtney) Found in 18 storage with broken door latch Brought down to biomed and repaired latch Humidity test failed functional and calibration Replaced and calibrated humidity sensor Functional check - OK
1062145 | 201215 | CM | CMPLT | NFF | 2024-01-25 09:50 | 2024-01-25 09:51 | "controls not working" | (TC 01/23/2024 09:49:47 AM, Cortez, John) Unable to repeat error. Issue possibly resolved after power cycle. Informed ward clerk.
| 201216 | CM | CMPLT | CMPLT | 2024-01-25 09:53 | 2024-02-16 14:52 | Transport Battery
502047 | 201217 | CM | CMPLT | CMPLT | 2024-01-25 14:22 | 2024-02-26 13:15 | MSH OPG SPP C1 IU22 61855330 PROBE C5-1 B1MKVM | (TC 01/25/2024 02:28:53 PM, Whittington, Scott) MSH OPG SPP C1 IU22 61855330 PROBE C5-1 B1MKVM REF# 0123013116 RHONDA CORRECT SITE 52668736 PROBE ORDERED FOR DELIVERY TO MSH 600 UNIVERSITY RECEIVING (TC 02/05/2024 04:26:39 PM, Whittington, Scott) MSH OPG SPP C1 IU22 61855330 PROBE C5-1 B1MKVM REF# 0123013116 RHONDA CORRECT SITE 52668736 REPLACEMENT PROBE SN# F0D2T4 (TC 02/05/2024 04:27:55 PM, Dolinar, John) MSH OPG SPP C1 IU22 61855330 PROBE C5-1 B1MKVM REF# 0123013116 RHONDA CORRECT SITE 52668736 REPLACEMENT PROBE SN# F0D2T4 | MSH OPG SPP C1 IU22 61855330 PROBE C5-1 B1MKVM REF# 0123013116 RHONDA CORRECT SITE 52668736
52117 | 201218 | CM | CMPLT | CMPLT | 2024-01-25 14:30 | 2024-01-26 12:03 | Unable to move the canopy up or down | (TC 01/25/2024 02:30:35 PM, Baghirzade, Nara) Initial observation of the omnibed with Samir; took off the canopy covers and the covers for the left rail and found the belt is broken. Belt is from an older generation and is plastic (new one is metal). (TC 01/25/2024 02:31:28 PM, Chukri, Samir) Initial observation of the omnibed; took off the canopy covers and the covers for the left rail and found the belt is broken. Belt is from an older generation and is plastic (new one is metal). (TC 01/25/2024 02:33:03 PM, Baghirzade, Nara) Completely took the canopy off. Disassembled the left rail and took it off the incubator along with the wiring for the left rail. Took off the back panel and cover for the display. Assembled the new rail, wired it in to the main board, put the transformer and motor back in place. Assembled the canopy partially and put the display panel and covers back on. Put on any remaining covers for the sides of the rails. (TC 01/25/2024 02:34:55 PM, Chukri, Samir) Completely took the canopy off. Disassembled the left rail and took it off the incubator along with the wiring for the left rail. Took off the back panel and cover for the display. Assembled the new rail, wired it in to the main board, put the transformer and motor back in place. Assembled the canopy partially and put the display panel and covers back on. Put on any remaining covers for the sides of the rails. (TC 01/25/2024 02:35:06 PM, Rhuland, Courtney) Completely took the canopy off. Disassembled the left rail and took it off the incubator along with the wiring for the left rail. Took off the back panel and cover for the display. Assembled the new rail, wired it in to the main board, put the transformer and motor back in place. Assembled the canopy partially and put the display panel and covers back on. Put on any remaining covers for the sides of the rails. (TC 01/26/2024 09:05:45 AM, Baghirzade, Nara) Placed canopy covers back on the omnibed. (TC 01/26/2024 09:06:03 AM, Rhuland, Courtney) Placed canopy covers back on the omnibed (TC 01/26/2024 12:01:38 PM, Rhuland, Courtney) PM complete LV and main voltage calibrated Scale calibrated RH calibrated 9V rechargeable battery replaced Functional check - OK Electrical safety - OK
502237 | 201219 | CM | CMPLT | CMPLT | 2024-01-25 14:32 | 2024-09-25 09:36 | MSH OPG SPP ZONARE LOCATED IN OPG 3-804 B1 MATERNAL/FETAL MEDICINE | (TC 01/25/2024 02:32:11 PM, Whittington, Scott) MSH OPG SPP ZONARE LOCATED IN OPG 3-804 B1 MATERNAL/FETAL MEDICINE | MSH OPG SPP ZONARE LOCATED IN OPG 3-804 B1 MATERNAL/FETAL MEDICINE
66375 | 201220 | CM | CMPLT | CMPLT | 2024-01-25 14:40 | 2024-01-25 14:41 | Broken BP hose/cuff | (TC 01/25/2024 02:40:16 PM, Rhuland, Courtney) Checked device and replaced broken cuff connector and leaky adult large cuff with spares Functional check - OK
68392 | 201221 | PM | CMPLT | CMPLT | 2024-01-25 15:10 | 2024-01-25 15:12 | PM-MAJOR - Annual PM For: Delivery Module | (TC 01/25/2024 03:10:42 PM, Scero, Tomlyn) PM parts replaced and tests from IPM-L complete and passed.
| 201222 | PA | CMPLT | CMPLT | 2024-01-26 08:28 | 2024-02-06 11:20 | Otoscope/ophthalmoscope parts for ED | (TC 01/26/2024 08:47:41 AM, Canjimavadivel, Kamallesh ) Parts needed for few rooms in ED. Emailed vendor for quote (TC 01/29/2024 08:53:20 AM, Canjimavadivel, Kamallesh ) Placed order for parts
1061031 | 201223 | CM | CMPLT | CMPLT | 2024-01-26 09:16 | 2024-01-26 09:17 | Patient Safety Net: Client Down Error | (TC 01/26/2024 09:16:12 AM, Baghirzade, Nara) PSN monitor was unplugged from the wall. Plugged it back in and booted it up. Boot-up sequence good. All sectors showing on the display. No longer alarming that it's down.
| 201224 | CM | CMPLT | CMPLT | 2024-01-26 09:18 | 2024-01-26 09:51 | Patient Safety Net: 15 Murray Client Down | (TC 01/26/2024 09:17:37 AM, Baghirzade, Nara) Monitor was unplugged from the wall outlet. Plugged it back in and booted it up; boot up sequence was ok, no errors. All sectors are displaying on the monitor and no longer alarming on other PSNs.
1062936 | 201225 | CM | CMPLT | CMPLT | 2024-01-26 09:53 | 2024-01-26 09:54 | Reported that it does not switch on | (TC 01/26/2024 09:53:16 AM, Baghirzade, Nara) No power cord was on the Masimo, Root battery was fully depleted. Left it to charge for a couple of hours. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056447 | 201226 | CM | CMPLT | CMPLT | 2024-01-26 09:55 | 2024-01-26 09:55 | Reported "not turning on, blinks and shuts off" | (TC 01/26/2024 09:54:35 AM, Baghirzade, Nara) Root battery was fully depleted. Left it to charge for a couple of hours. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056505 | 201227 | CM | CMPLT | CMPLT | 2024-01-26 10:11 | 2024-01-26 10:11 | Battery replacement | (TC 01/26/2024 10:10:49 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055617 | 201228 | CM | CMPLT | CMPLT | 2024-01-26 10:11 | 2024-01-26 10:12 | Battery replacement | (TC 01/26/2024 10:11:25 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1054074 | 201229 | CM | CMPLT | CMPLT | 2024-01-26 10:37 | 2024-01-26 10:39 | Functional check - move to ambulatory pod | (TC 01/26/2024 10:37:16 AM, Rhuland, Courtney) Functional check and relocated for ambulatory pods
| 201230 | CM | CMPLT | CMPLT | 2024-01-26 10:40 | 2024-01-26 10:42 | Shared documents for tracking (NICU) | (TC 01/26/2024 10:39:48 AM, Baghirzade, Nara) Created a shared folder in staff files for myself and Courtney to keep track of shared work such as PMs, repairs, replacements, etc. for NICU and RESUS equipment. Created documents for: - Root NIBP tracking - Radical 7 battery replacements - Stand-alone Radical 7 battery replacements - NeoPuff PMs Updated Courtney.
1063193 | 201231 | CM | CMPLT | CMPLT | 2024-01-26 12:45 | 2024-01-30 09:09 | MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ | (TC 01/26/2024 12:59:35 PM, Whittington, Scott) MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ QUYEN EXPLAINED THAT IT WAS LINEAR PROBE MIE 1051739 PLT 705AT HAD A GRAINY IMAGE CONVEX PROBE PVT 375 HAS VERY LOW VIDEO O/P. NOT VISIBLE WHEN PROBE ACTIVATED, 2D HAS TO BE TURNED TO MAX TO SEE IMAGE | MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ QUYEN EXPLAINED THAT IT WAS LINEAR PROBE MIE 1051739 PLT 705AT HAD A GRAINY IMAGE CONVEX PROBE PVT 375 HAS VERY LOW VIDEO O/P. NOT VISIBLE WHEN PROBE ACTIVATED, 2D HAS TO BE TURNED TO MAX TO SEE IMAGE
1051739 | 201232 | CM | CMPLT | CMPLT | 2024-01-26 13:01 | 2024-01-30 09:10 | MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ | (TC 01/26/2024 01:01:27 PM, Whittington, Scott) MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ QUYEN EXPLAINED THAT IT WAS LINEAR PROBE MIE 1051739 PLT 705AT HAD A GRAINY IMAGE CONVEX PROBE PVT 375 HAS VERY LOW VIDEO O/P. NOT VISIBLE WHEN PROBE ACTIVATED, 2D HAS TO BE TURNED TO MAX TO SEE IMAGE REPLACED WITH MIE 500882 SN# 99A1533158 | MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ QUYEN EXPLAINED THAT IT WAS LINEAR PROBE MIE 1051739 PLT 705AT HAD A GRAINY IMAGE CONVEX PROBE PVT 375 HAS VERY LOW VIDEO O/P. NOT VISIBLE WHEN PROBE ACTIVATED, 2D HAS TO BE TURNED TO MAX TO SEE IMAGE
500882 | 201233 | CM | CMPLT | CMPLT | 2024-01-26 13:06 | 2024-01-29 14:51 | MSH IRUS ANGIO 1 APLIO XG PLT 705 PROBE SWAPPED IN | (TC 01/29/2024 02:51:28 PM, Whittington, Scott) (TC 01/26/2024 01:01:27 PM, Whittington, Scott) MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ QUYEN EXPLAINED THAT IT WAS LINEAR PROBE MIE 1051739 PLT 705AT HAD A GRAINY IMAGE CONVEX PROBE PVT 375 HAS VERY LOW VIDEO O/P. NOT VISIBLE WHEN PROBE ACTIVATED, 2D HAS TO BE TURNED TO MAX TO SEE IMAGE REPLACED WITH MIE 500882 SN# 99A1533158 | (TC 01/26/2024 01:01:27 PM, Whittington, Scott) MSH IRUS APLIO XG ANGIO 1 OMED, FELLOW, COMPLAINED OF INADEQUATE IQ QUYEN EXPLAINED THAT IT WAS LINEAR PROBE MIE 1051739 PLT 705AT HAD A GRAINY IMAGE CONVEX PROBE PVT 375 HAS VERY LOW VIDEO O/P. NOT VISIBLE WHEN PROBE ACTIVATED, 2D HAS TO BE TURNED TO MAX TO SEE IMAGE REPLACED WITH MIE 500882 SN# 99A1533158
1060032 | 201234 | CM | CMPLT | CMPLT | 2024-01-26 14:10 | 2024-02-02 12:08 | Ordering pneumotachograph heaters and o-rings | (TC 01/26/2024 02:10:13 PM, Baghirzade, Nara) Requested a quote from Chantal for heaters and o-rings for the Hypair and other PFT systems. (TC 01/30/2024 11:04:09 AM, Baghirzade, Nara) SC 1003157883 (TC 02/02/2024 12:07:01 PM, Baghirzade, Nara) PO 7202355053 (TC 02/13/2024 09:14:35 AM, Baghirzade, Nara) Did not receive O-rings. Followed up with Chantal. (TC 02/23/2024 10:36:03 AM, Baghirzade, Nara) Received O-rings. PO7202355053
1054986 | 201235 | CM | CMPLT | CMPLT | 2024-01-26 14:11 | 2024-01-26 14:19 | Respiratory accessories | (TC 01/26/2024 02:16:07 PM, Baghirzade, Nara) Met with Sarabeth regarding the respiratory monitoring accessories as they were short in supply. Gave her: 1 box of RAS-45 acoustic sensors 1 box of RAS-126C acoustic sensors 1 box of LNCS SpO2 sensors 1 box of SpO2 replacement tapes
1056442 | 201236 | CM | CMPLT | CMPLT | 2024-01-26 14:19 | 2024-01-26 14:21 | Respiratory Accessories | (TC 01/26/2024 02:19:36 PM, Baghirzade, Nara) Met with Sri regarding the respiratory monitoring accessories as they were short in supply. Gave her: 1 box of RAS-45 acoustic sensors 1 box of RAS-126C acoustic sensors 1 box of LNCS SpO2 sensors 1 box of SpO2 replacement tapes
| 201237 | CM | CMPLT | CMPLT | 2024-01-26 14:25 | 2024-01-26 14:28 | Mount stuck to the bracket | (TC 01/26/2024 02:26:18 PM, Baghirzade, Nara) During rounds, noticed that there is an old blender from 2018. When trying to take it off the bracket, it was stuck to the mount as some of the metal was warped/bent. Used a mallet to get the bracket off and put in a new blender. (TC 01/26/2024 02:27:59 PM, Rhuland, Courtney) During rounds, noticed that there is an old blender from 2018. When trying to take it off the bracket, it was stuck to the mount as some of the metal was warped/bent. Used a mallet to get the bracket off and put in a new blender.
| 201238 | CM | CMPLT | CMPLT | 2024-01-26 14:26 | 2024-02-06 09:16 | Panda warmer blender overhaul | (TC 01/26/2024 02:27:48 PM, Rendon Lopez , Elizabeth) Sending out two panda warmer blender for overhaul Received
| 201239 | RDS | CMPLT | CMPLT | 2024-01-26 14:29 | 2024-01-26 14:30 | Weekly Rounds: 11 South | (TC 01/26/2024 02:29:09 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 201240 | RDS | CMPLT | CMPLT | 2024-01-26 14:31 | 2024-01-26 14:32 | Weekly Rounds: 11 North | (TC 01/26/2024 02:30:36 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 201241 | RDS | CMPLT | CMPLT | 2024-01-26 14:32 | 2024-01-26 14:33 | Weekly Rounds: RESUS and NICU | (TC 01/26/2024 02:31:56 PM, Baghirzade, Nara) Weekly rounds in RESUS and NICU completed. See procedure for results.
1057546 | 201242 | CM | CMPLT | CMPLT | 2024-01-26 14:51 | 2024-01-29 14:52 | SPO2 not reading, possibly blocked port | (TC 01/29/2024 02:51:27 PM, Scero, Tomlyn) tested SPO2 with tester cables that are known to be working. No issues reading spo2 on the monitor. possibly a bad cable.
1054968 | 201243 | CM | CMPLT | CMPLT | 2024-01-26 15:46 | 2024-01-29 09:45 | Missing temp probe well | (TC 01/26/2024 03:46:01 PM, Cortez, John) Not medical grade power cord, replaced Cleaned filter Added missing screw to base of roll stand Replaced temp probe well (TC 01/29/2024 09:44:04 AM, Cortez, John) Charged overnight, All parameters tested ok
1054537 | 201244 | PM | CMPLT | CMPLT | 2024-01-26 16:15 | 2024-01-26 16:16 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/26/2024 04:15:40 PM, Cortez, John) Cleaning cycle completed
1060518 | 201245 | CM | CMPLT | CMPLT | 2024-01-26 18:28 | 2024-01-26 18:27 | Replaced Motor height assembly. | Repair is not complete, need to replace the motor assembly on Monday.
66713 | 201246 | CM | CMPLT | CMPLT | 2024-01-29 07:34 | 2024-01-29 08:30 | Broken porthole door | (TC 01/29/2024 08:29:37 AM, Rhuland, Courtney) Replaced broken porthole door & springs on 2 porthole doors Calibrated RH Functional check - OK
1050672 | 201247 | CM | CMPLT | CMPLT | 2024-01-29 07:53 | 2024-01-29 08:05 | Exposed inner wires | (TC 01/29/2024 08:04:40 AM, Canjimavadivel, Kamallesh ) Replaced damaged wires, tested ok with simulator
1050342 | 201248 | CM | CMPLT | CMPLT | 2024-01-29 08:57 | 2024-02-09 11:28 | Physical damage | (TC 01/23/2024 08:57:21 AM, Rendon Lopez , Elizabeth) Confirmed issue, to be set out for repair (TC 02/09/2024 11:27:38 AM, Rendon Lopez , Elizabeth) Incoming inspection-oK
1055658 | 201249 | CM | CMPLT | CMPLT | 2024-01-29 09:00 | 2024-02-09 11:29 | Physical damage | (TC 01/29/2024 08:59:45 AM, Rendon Lopez , Elizabeth) Confirmed issue, to be set out for repair (TC 02/09/2024 11:28:50 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1054117 | 201250 | CM | CMPLT | CMPLT | 2024-01-29 09:01 | 2024-02-09 10:09 | Not detected by monitor | (TC 01/23/2024 09:01:26 AM, Rendon Lopez , Elizabeth) Confirmed issue, to be set out for repair (TC 02/09/2024 10:08:34 AM, Rendon Lopez , Elizabeth) Deemed defective, unable to repair. Retired
1067322 | 201251 | CM | CMPLT | CMPLT | 2024-01-29 09:03 | 2024-02-09 10:24 | Loose piece inside, no sound | (TC 01/23/2024 09:04:26 AM, Rendon Lopez , Elizabeth) Confirmed issue, to be set out for repair under warranty (TC 02/09/2024 10:18:25 AM, Rendon Lopez , Elizabeth) Received device back, tested-OK
1069827 | 201252 | CM | CMPLT | CMPLT | 2024-01-29 09:09 | 2024-02-13 11:13 | Adding to TMS -Sending out for PM | (TC 02/13/2024 11:12:09 AM, Rhuland, Courtney) (TC 02/13/2024 11:10:38 AM, Baghirzade, Nara) Functional testing completed. 21% = 20.9% 40% = 39% 60% = 59.8% 80% = 80.1% 100% = 99.9% When disconnecting air OR oxygen, blender alarm is functional. Attached service report.
1055792 | 201253 | CM | CMPLT | CMPLT | 2024-01-29 09:13 | 2024-01-29 09:56 | check in failed | (TC 01/29/2024 09:55:25 AM, Scero, Tomlyn) Moved crash cart back to alcove and performed 30J test after plugging in unit. Device passed and check mark displayed. | OR Suite - OR Desk – Connect cable to electrodes – Plug in AC
1069828 | 201254 | CM | CMPLT | CMPLT | 2024-01-29 09:16 | 2024-02-13 11:24 | Adding to TMS - Sending out for PM | (TC 02/13/2024 11:21:02 AM, Rhuland, Courtney) Device determined to be beyond repair Returned from Spectramedx Retiring from service
1069829 | 201255 | CM | CMPLT | CMPLT | 2024-01-29 09:22 | 2024-02-13 11:26 | Adding to TMS - Sending out for PM | (TC 01/30/2024 02:45:17 PM, Rhuland, Courtney) Bubble wrapped and boxed up all blenders to be sent out (TC 02/13/2024 11:25:17 AM, Rhuland, Courtney) Device determined to be beyond repair Returned from Spectramedx Retiring from service
| 201256 | CM | CMPLT | CMPLT | 2024-01-29 09:45 | 2024-01-26 09:45 | Mandatory annual e-learning | (TC 01/26/2024 09:45:06 AM, Chukri, Samir) Completed all the Mandatory annul E-learning .
| 201257 | RDS | CMPLT | CMPLT | 2024-01-29 10:15 | 2024-01-29 10:46 | Weekly checks in CCU | (TC 01/29/2024 10:45:11 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
501071 | 201258 | CM | CMPLT | CMPLT | 2024-01-29 10:48 | 2024-01-30 15:22 | System crashed, won't boot up. Stays on black screen | (TC 01/29/2024 10:47:52 AM, Fan, Tak) Checked system. Was stuck on a white screen. Shutdown system, was never able to boot it up. Stays on black screen. Placed call. (TC 01/30/2024 03:19:08 PM, Fan, Tak) HD replacement had issues. Cloned drive instead. System started up. Tested OK. Jillian did the PM.
1062322 | 201259 | CM | CMPLT | NFF | 2024-01-29 14:02 | 2024-01-29 14:04 | USB connected cord shorted | (TC 01/29/2024 02:02:29 PM, Rhuland, Courtney) 11N called saying patient used cord to charge device and cord got very hot and melted Cord used was defective - no issue with bed itself Checked bed with Samir - no fault found with bed
1054779 | 201260 | CM | CMPLT | CMPLT | 2024-01-29 14:33 | 2024-01-29 14:42 | Rubber feet melting | (TC 01/29/2024 02:33:11 PM, Rhuland, Courtney) Rubber feet under device were melting and dripping onto device/accessories Removed all 4 feet as device is on rolling stand Cleaned device and accessories of gunk Functional check - OK
1057384 | 201261 | CM | CMPLT | CMPLT | 2024-01-29 14:59 | 2024-02-07 13:02 | MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 | (TC 01/29/2024 03:08:08 PM, Dolinar, John) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 REF# Ticket tracking #919185 Case Number: 919185 K#: 41873354 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: Power Issue not turning on Opening Date: 2024-01-29T14:01:06 EST Name of Agent: UPS INSTALL TEMP BATTERIES CLEAN SYSTEM (TC 01/29/2024 03:08:52 PM, Fan, Tak) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 REF# Ticket tracking #919185 Case Number: 919185 K#: 41873354 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: Power Issue not turning on Opening Date: 2024-01-29T14:01:06 EST Name of Agent: ATTEMPT TO POWER UP UPS INSTALL TEMP BATTERIES CLEAN SYSTEM (TC 01/30/2024 10:22:37 AM, Dolinar, John) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 REF# Ticket tracking #919185 Case Number: 919185 K#: 41873354 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: Power Issue not turning on Opening Date: 2024-01-29T14:01:06 EST Name of Agent: (TC 01/30/2024 10:22:51 AM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 REF# Ticket tracking #919185 Case Number: 919185 K#: 41873354 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: Power Issue not turning on Opening Date: 2024-01-29T14:01:06 EST Name of Agent: (TC 01/30/2024 10:08:00 AM, Fan, Tak) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 REF# Ticket tracking #919185 Case Number: 919185 K#: 41873354 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: Power Issue not turning on Opening Date: 2024-01-29T14:01:06 EST Name of Agent: ATTEMPT TO POWER UP UPS INSTALL TEMP BATTERIES CLEAN SYSTEM (TC 01/31/2024 07:51:19 AM, Fan, Tak) Front door open message. Not able to clear. Took out UPS and placed in other CR reader. (TC 02/01/2024 10:47:10 AM, Fan, Tak) Cassette not loading. Loaded fine when checked. They will call again if anything. (TC 02/01/2024 03:09:32 PM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 REF# Ticket tracking #919185 Case Number: 919185 K#: 41873354 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: Power Issue not turning on Opening Date: 2024-01-29T14:01:06 EST Name of Agent: 2 SETS OF UPS BATTERIES DELIVERED PN# 3F4029 INSTALLED IN THIS SYSTEM, AS UPS WERE SWAPPED THIS MORNING BY TAK UPS OK WITH NEW BATTERIES, BUT FRONT DOOR OPEN ERROR PERSISTENT TRIED JUMPERING SWITCH, NO RESULT (TC 02/07/2024 10:52:16 AM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 FRONT DOOR OPEN ERROR SEN & 2 COLLEAGUES ON SITE (TC 02/07/2024 10:52:16 AM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 FRONT DOOR OPEN ERROR SEAN & 2 COLLEAGUES AAREN & ERIC ON SITE PN#1F6473 MSC BOARD REPLACED FAULTY KEPT ON SITE AS TEST/SPARE FOR REBUILD SEAN CLEANED AND CHECKED OPERATION (TC 02/20/2024 03:25:19 PM, Dolinar, John) MSH CR MOVE FROM CR ROOM TO ALCOVE (TC 02/21/2024 01:05:28 PM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 MOVE SYSTEM INTO ALCOVE | MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3354 REF# Ticket tracking #919185 Case Number: 919185 K#: 41873354 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: Power Issue not turning on Opening Date: 2024-01-29T14:01:06 EST Name of Agent:
1057383 | 201262 | CM | CMPLT | CMPLT | 2024-01-29 15:05 | 2024-03-04 11:20 | MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3265 | (TC 01/30/2024 10:21:29 AM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3265 REF# Ticket tracking #919468 Case Number: 919468 K#: 41873265 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: No Power Opening Date: 2024-01-29T13:58:47 EST Name of Agent: (TC 01/30/2024 10:21:41 AM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3265 REF# Ticket tracking #919468 Case Number: 919468 K#: 41873265 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: No Power Opening Date: 2024-01-29T13:58:47 EST Name of Agent: (TC 01/30/2024 10:21:41 AM, Dolinar, John) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3265 REF# Ticket tracking #919468 Case Number: 919468 K#: 41873265 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: No Power Opening Date: 2024-01-29T13:58:47 EST Name of Agent: (TC 01/31/2024 07:50:08 AM, Fan, Tak) Moved UPS from other CR reader due inability to clear front door open message. Tested, OK. (TC 02/07/2024 01:03:45 PM, Whittington, Scott) MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3265 UPS BATTERIES REQUIRE REPLACEMENT WITH BATTERIES RECEIVED FROM CARESTREAM (TC 02/08/2024 03:37:44 PM, Whittington, Scott) MSH CR READER NETWORK JACK ACTIVATION (TC 02/13/2024 09:05:40 AM, Whittington, Scott) From: HELP Desk (MSH) <helpdesk.MSH@sinaihealth.ca> Sent: February 12, 2024 6:11 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] INC30965425 -RE: Reconnect and activate on Medical Imaging VLAN MSH 5 DR STRETCHER ALCOVE Hello , A ticket has been logged for this email. Your ticket number is INC30965425, and was sent to the onsite team. If you have any questions or concerns, please contact us and reference the above ticket number. Pascal Jaeger MSH Service Desk Toll Free: 1-866-244-9909 Phone: 416-586-4800 Ext: 4356 Email: helpdesk.msh@sinaihealth.ca From: Whittington, Scott [mailto:Scott.Whittington@uhn.ca] Sent: Monday, February 12, 2024 4:50 PM To: HELP Desk (MSH) <helpdesk.MSH@sinaihealth.ca> Cc: Riaz, Mirza <Mirza.Riaz@sinaihealth.ca>; Schwietz, Matthew <Matthew.Schwietz@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Reconnect and activate on Medical Imaging VLAN MSH 5 DR STRETCHER ALCOVE CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Please Reconnect and activate on Medical Imaging VLAN 5NNW-84D Jack. This was previously operational prior to construction activity. Please duplicate settings currently active on 5NNW-85D Thanks Scott 416-402-2462 (TC 02/20/2024 03:27:56 PM, Fan, Tak) MOVE CR FROM ROOM TO ALCOVE (TC 03/04/2024 11:21:53 AM, Dolinar, John) Replaced batteries. | MSH CR 850 WILL NOT BOOT, UPS BATTERIES DEAD K4187-3265 REF# Ticket tracking #919468 BATTERY Scanner keeps reinitializing. Close front door. Ref# 923116 Replaced batteries Please Reconnect and activate on Medical Imaging VLAN 5NNW-84D Jack. This was previously operational prior to construction activity. Please duplicate settings currently active on 5NNW-85D INC30965425 Case Number: 919468 K#: 41873265 Equipment Description: DIRECTVIEW 850 SYSTEM Problem Description: No Power Opening Date: 2024-01-29T13:58:47 EST Name of Agent:
1062026 | 201263 | CM | CMPLT | CMPLT | 2024-01-29 15:50 | 2024-01-29 15:51 | No controls working | (TC 01/29/2024 03:49:59 PM, Ferreira, Qwynn) Reset the bed and bed controls were working Discussed with the team lead for an in service to show the staff how to use the beds Explained some basic steps to take before calling as it could save time
| 201264 | TEDU | CMPLT | CMPLT | 2024-01-30 07:27 | 2024-01-30 07:28 | Draeger V500/VN500 E-learning modules and training | (TC 01/29/2024 03:00:00 PM, Baghirzade, Nara) Draeger V500/VN500 E-learning modules and training
| 201265 | MEP | CMPLT | CMPLT | 2024-01-30 07:55 | 2024-05-16 15:08 | 3D Printing projects | (TC 01/29/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Discuss ongoing projects and scope of work (TC 01/29/2024 11:30:00 AM, Canjimavadivel, Kamallesh ) Set up new printer in shop (TC 01/29/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Segmenting for case (TC 01/30/2024 02:45:00 PM, Canjimavadivel, Kamallesh ) Worked on heel pricker case. Modified design and ready for print. (TC 01/31/2024 03:18:08 PM, Canjimavadivel, Kamallesh ) Post-print processing of heel pricker project prints. Segmentation of scapula for mandible case. (TC 02/01/2024 01:16:03 PM, Canjimavadivel, Kamallesh ) Finishing up scapula segmentation (TC 02/01/2024 03:13:27 PM, Canjimavadivel, Kamallesh ) Stepping in on cases in the OR (TC 02/02/2024 03:26:46 PM, Canjimavadivel, Kamallesh ) Finishing up model for heel pricker case (TC 02/02/2024 03:27:15 PM, Canjimavadivel, Kamallesh ) Went to OR to observe case (TC 02/05/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Pharmacy bag solution (TC 02/05/2024 10:27:04 AM, Canjimavadivel, Kamallesh ) Updated heel pricker model and printed latetst version (TC 02/05/2024 03:00:00 PM, Canjimavadivel, Kamallesh ) Observed case in OR (TC 02/09/2024 10:34:10 AM, Canjimavadivel, Kamallesh ) Prepping new tray for Qwynn to send print. Start post print process for guides (TC 02/13/2024 03:25:39 PM, Canjimavadivel, Kamallesh ) Segmentation (TC 02/13/2024 03:25:53 PM, Canjimavadivel, Kamallesh ) Follow up with heel-pricker case with the NICU (TC 02/13/2024 03:26:21 PM, Canjimavadivel, Kamallesh ) committee meeting (TC 02/14/2024 03:15:00 PM, Canjimavadivel, Kamallesh ) Segmentation for tibia case (TC 02/15/2024 01:48:44 PM, Canjimavadivel, Kamallesh ) Finish up tibia case, ready for print (TC 02/20/2024 03:25:58 PM, Canjimavadivel, Kamallesh ) Updates meeting (TC 02/20/2024 03:26:22 PM, Canjimavadivel, Kamallesh ) Casting avm model with urethane (TC 02/21/2024 10:54:24 AM, Canjimavadivel, Kamallesh ) Checking up on urethane casting (TC 02/21/2024 02:43:06 PM, Canjimavadivel, Kamallesh ) Working on casting avm with resin (TC 02/23/2024 03:18:47 PM, Canjimavadivel, Kamallesh ) Working on placenta stand model (TC 02/26/2024 11:11:37 AM, Canjimavadivel, Kamallesh ) Consulatation for placenta stand model improvements (TC 02/29/2024 03:10:47 PM, Canjimavadivel, Kamallesh ) Designing and printing tray for placenta stand modification (TC 03/05/2024 03:18:48 PM, Canjimavadivel, Kamallesh ) Weekly update (TC 03/05/2024 03:19:02 PM, Canjimavadivel, Kamallesh ) Placenta stand design (TC 03/07/2024 02:47:25 PM, Canjimavadivel, Kamallesh ) Finishing design for placenta stand and sent for print (TC 03/08/2024 01:56:07 PM, Canjimavadivel, Kamallesh ) Assembling placenta stand tray with super glue. Used epoxy to seal edges. Went to Canadian Tire to grab some supplies. Clear coated tray (TC 03/11/2024 03:14:46 PM, Canjimavadivel, Kamallesh ) Testing placenta tray flow rate and added another cleat coat layer after sanding (TC 03/19/2024 02:52:23 PM, Canjimavadivel, Kamallesh ) Updates (TC 03/20/2024 01:30:00 PM, Canjimavadivel, Kamallesh ) Finishing off placenta stand testing and hand off to Emily (TC 03/21/2024 11:35:02 AM, Canjimavadivel, Kamallesh ) OHS meeting (TC 03/21/2024 03:01:41 PM, Canjimavadivel, Kamallesh ) Got new project for refrigerator lock in ED med room. Went to see space and equipment and took measurements. Started process to renew fusion license (TC 03/22/2024 03:13:23 PM, Canjimavadivel, Kamallesh ) Reinstalled fusion 360 and tried to set up account license (TC 03/25/2024 02:52:59 PM, Canjimavadivel, Kamallesh ) Introduction to rhinoplasty simulation project (TC 03/26/2024 07:31:47 AM, Canjimavadivel, Kamallesh ) Meeting with Balram from bone bank regarding new project (TC 03/26/2024 07:32:26 AM, Canjimavadivel, Kamallesh ) Started design for fridge lock project (TC 04/02/2024 03:00:00 PM, Canjimavadivel, Kamallesh ) weekly update (TC 04/03/2024 09:49:13 AM, Canjimavadivel, Kamallesh ) Designed bracket for pharmacy fridge lock and started print (TC 04/04/2024 10:22:48 AM, Canjimavadivel, Kamallesh ) Post processing of the fridge lock print (TC 04/05/2024 10:23:39 AM, Canjimavadivel, Kamallesh ) Tested fridge lock mount on the fridge (TC 04/05/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Meet with Sarah to demo fridge lock mount prototype. Set up 3 mounts for print and started (TC 04/08/2024 01:44:13 PM, Canjimavadivel, Kamallesh ) Post print processing for the 3 fridge lock mounts. Sanded surface and washed. Emailed Sarah to inform they are ready for install. (TC 04/10/2024 03:00:38 PM, Canjimavadivel, Kamallesh ) Mocked up prototype for bone bank tray project and printed. Showed prototype to team and received feedback (TC 04/16/2024 01:34:24 PM, Canjimavadivel, Kamallesh ) Went to install fridge lock bracket. Bracket needed to be modified to better accomodate installation. Modified design with changes. Sent for print. (TC 04/23/2024 03:04:25 PM, Canjimavadivel, Kamallesh ) Installed one lock in ED East med room. Prepped brackets to print. Prepped remaining locks for install (TC 04/24/2024 09:32:57 AM, Canjimavadivel, Kamallesh ) Installation of fridge lock bracket with Pharmacy team (TC 04/25/2024 01:27:51 PM, Canjimavadivel, Kamallesh ) Following up on todoist tasks (TC 04/26/2024 01:24:26 PM, Canjimavadivel, Kamallesh ) Following SOP for anonymization procedure (TC 04/29/2024 02:59:04 PM, Canjimavadivel, Kamallesh ) Working on design for bone bank project (TC 05/01/2024 09:37:12 AM, Canjimavadivel, Kamallesh ) Privacy assessment discussion (TC 05/01/2024 03:00:12 PM, Canjimavadivel, Kamallesh ) Spoke with Balram to provide updates on the project (TC 05/01/2024 03:00:54 PM, Canjimavadivel, Kamallesh ) Installing fan filter box into Bambu printer (TC 05/02/2024 02:35:05 PM, Canjimavadivel, Kamallesh ) Set up bento filter box with switch. Soldered parts together and tested functionality. (TC 05/02/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Set up formlabs printer with biomed durable resin and started print (TC 05/03/2024 02:45:34 PM, Canjimavadivel, Kamallesh ) Cleared error in printer and started print. Designed box for Bambu printer accessory. Set up print and started. Troubleshooted issue with printer. Resent print. Started post processing for resin print after completion. (TC 05/06/2024 11:09:10 AM, Canjimavadivel, Kamallesh ) Post processing of coupon for bone bank project. Dropped off with bone bank to test freezing conditions. (TC 05/07/2024 01:13:27 PM, Canjimavadivel, Kamallesh ) Went to bone bank to check on test coupon that was left in freezer. Took base measuremenets. Cut polycarbonate acrylic into sections and gave back. (TC 05/08/2024 03:23:13 PM, Canjimavadivel, Kamallesh ) Post processing removing supports and sanding. Put print back into cure second time. (TC 05/09/2024 09:06:53 AM, Canjimavadivel, Kamallesh ) Dropped off finished tray to bone bank. (TC 05/09/2024 10:05:23 AM, Canjimavadivel, Kamallesh ) Mocked up prototype for new fridge lock adapter. Set up print and started overnight (TC 05/10/2024 02:44:10 PM, Canjimavadivel, Kamallesh ) Mixed silicone and poured into rhinoplasty mold (TC 05/13/2024 10:20:05 AM, Canjimavadivel, Kamallesh ) Met with Brandon from Pharmacy to validate fridge lock prototype. Need to make modifcations and validate again next week. (TC 05/13/2024 03:51:47 PM, Canjimavadivel, Kamallesh ) Worked on rhinoplasty model. Tested out combining 2 layers of silicone. Prepped second mold with 2 layers and new version of mold. Left to cure overnight (TC 05/14/2024 03:58:21 PM, Canjimavadivel, Kamallesh ) Meeting update (TC 05/14/2024 03:58:21 PM, Canjimavadivel, Kamallesh ) Worked on rhinoplasty mold (TC 05/14/2024 03:58:21 PM, Canjimavadivel, Kamallesh ) Worked on rhinoplasty mold (TC 05/15/2024 10:26:10 AM, Canjimavadivel, Kamallesh ) Set up fusion account and transfer projects to team hub (TC 05/15/2024 03:14:44 PM, Canjimavadivel, Kamallesh ) Worked on project requested my Steve for a tool to reduce wrist turn when turning cylinder knobs on ventilators. Designed prototype knob kit (TC 05/16/2024 01:37:54 PM, Canjimavadivel, Kamallesh ) Pouring nose mold with latest design (TC 05/16/2024 03:07:59 PM, Canjimavadivel, Kamallesh ) Post processing of nose mold
| 201266 | CM | CMPLT | CMPLT | 2024-01-30 08:11 | 2024-01-30 08:12 | NICU rounds | (TC 01/30/2024 08:10:53 AM, Rhuland, Courtney) NICU rounds complete
| 201267 | CM | CMPLT | CMPLT | 2024-01-30 08:13 | 2024-01-30 08:14 | RESUS rounds | (TC 01/30/2024 08:12:43 AM, Rhuland, Courtney) RESUS rounds complete Nurse alerted me to a bracket that had fallen off a suction canister Will bring screws up to repair
1069750 | 201268 | ACC | CMPLT | CMPLT | 2024-01-30 08:18 | 2024-01-30 08:22 | Incoming Inspection | (TC 01/22/2024 08:19:41 AM, Rendon Lopez , Elizabeth) Assembled device, asset tagged, incoming inspection tests-OK. Delivered to L&D OR
66555 | 201269 | CM | CMPLT | CMPLT | 2024-01-30 08:25 | 2024-01-30 08:29 | BP not inflating pass 80 | (TC 01/22/2024 08:25:36 AM, Rendon Lopez , Elizabeth) Tested, NFF. Checked BP calibration-OK.
1057048 | 201270 | CM | CMPLT | CMPLT | 2024-01-30 08:35 | 2024-01-30 08:35 | Battery replacement | (TC 01/30/2024 08:34:55 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055602 | 201271 | CM | CMPLT | CMPLT | 2024-01-30 08:35 | 2024-01-30 08:36 | Battery replacement | (TC 01/30/2024 08:35:33 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66565 | 201272 | CM | CMPLT | CMPLT | 2024-01-30 08:36 | 2024-01-30 08:38 | Load L&D Config - Return to service for L&D | (TC 01/23/2024 08:36:18 AM, Rendon Lopez , Elizabeth) L&D needed extra X2 monitor. Loaded L&D config to Biomed spare monitor, Replaced battery, tested all functions-OK
67817 | 201273 | CM | CMPLT | CMPLT | 2024-01-30 08:38 | 2024-01-30 08:38 | L&D needed spare monitor | (TC 01/23/2024 08:38:00 AM, Rendon Lopez , Elizabeth) L&D needed extra X2 monitor. Loaded L&D config to Biomed spare monitor. Replaced battery, tested all functions-OK
1057577 | 201274 | CM | CMPLT | NFF | 2024-01-30 08:42 | 2024-01-30 08:43 | Not working | (TC 01/23/2024 08:42:33 AM, Rendon Lopez , Elizabeth) Tested functions, NFF
1057453 | 201275 | CM | CMPLT | CMPLT | 2024-01-30 08:43 | 2024-01-30 08:45 | Knob broken | (TC 01/23/2024 08:43:41 AM, Rendon Lopez , Elizabeth) Replaced knob from spare, tested functions-OK
1054192 | 201276 | CM | CMPLT | CMPLT | 2024-01-30 08:45 | 2024-01-30 08:45 | Knob broken | (TC 01/23/2024 08:45:06 AM, Rendon Lopez , Elizabeth) Replaced knob from spare, tested functions-OK
1050348 | 201277 | CM | CMPLT | CMPLT | 2024-01-30 08:46 | 2024-01-30 08:47 | Knob broken | (TC 01/23/2024 08:46:11 AM, Rendon Lopez , Elizabeth) Replaced knob from spare, tested functions-OK
1069775 | 201278 | CM | CMPLT | NFF | 2024-01-30 08:47 | 2024-01-30 08:48 | Not working | (TC 01/23/2024 08:47:00 AM, Rendon Lopez , Elizabeth) Tested, NFF. Added device to TMS, returned to 15M
1069777 | 201279 | ACC | CMPLT | CMPLT | 2024-01-30 09:25 | 2024-01-30 09:26 | Incoming inspection | (TC 01/23/2024 09:26:01 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069778 | 201280 | ACC | CMPLT | CMPLT | 2024-01-30 09:27 | 2024-01-30 09:27 | Incoming inspection | (TC 01/23/2024 09:26:45 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069779 | 201281 | ACC | CMPLT | CMPLT | 2024-01-30 09:27 | 2024-01-30 09:28 | Incoming inspection | (TC 01/23/2024 09:27:28 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069780 | 201282 | ACC | CMPLT | CMPLT | 2024-01-30 09:28 | 2024-01-30 09:28 | Incoming inspection | (TC 01/23/2024 09:28:08 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069781 | 201283 | ACC | CMPLT | CMPLT | 2024-01-30 09:29 | 2024-01-30 09:29 | Incoming inspection | (TC 01/23/2024 09:28:53 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069782 | 201284 | ACC | CMPLT | CMPLT | 2024-01-30 09:30 | 2024-01-30 09:30 | Incoming inspection | (TC 01/23/2024 09:29:35 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069783 | 201285 | ACC | CMPLT | CMPLT | 2024-01-30 09:30 | 2024-01-30 09:33 | Incoming inspection | (TC 01/23/2024 09:32:42 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069784 | 201286 | ACC | CMPLT | CMPLT | 2024-01-30 09:33 | 2024-01-30 09:34 | Incoming inspection | (TC 01/23/2024 09:33:24 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069785 | 201287 | ACC | CMPLT | CMPLT | 2024-01-30 09:34 | 2024-01-30 09:35 | Incoming inspection | (TC 01/23/2024 09:34:40 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069786 | 201288 | ACC | CMPLT | CMPLT | 2024-01-30 09:35 | 2024-01-30 09:39 | Incoming inspection | (TC 01/23/2024 09:35:18 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069787 | 201289 | ACC | CMPLT | CMPLT | 2024-01-30 09:40 | 2024-01-30 09:41 | Incoming inspection | (TC 01/24/2024 09:40:36 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069788 | 201290 | ACC | CMPLT | CMPLT | 2024-01-30 09:46 | 2024-01-30 09:47 | Incoming inspection | (TC 01/24/2024 09:46:02 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069789 | 201291 | ACC | CMPLT | CMPLT | 2024-01-30 09:47 | 2024-01-30 09:47 | Incoming inspection | (TC 01/23/2024 09:46:59 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069790 | 201292 | ACC | CMPLT | CMPLT | 2024-01-30 09:48 | 2024-01-30 09:48 | Incoming inspection | (TC 01/24/2024 09:47:38 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069791 | 201293 | ACC | CMPLT | CMPLT | 2024-01-30 09:48 | 2024-01-30 09:49 | Incoming inspection | (TC 01/24/2024 09:48:18 AM, Rendon Lopez , Elizabeth) Incoming inspection
1047101 | 201294 | CM | CMPLT | CMPLT | 2024-01-30 09:49 | 2024-01-30 09:50 | Battery does not charge, unable to switch it on | (TC 01/30/2024 09:49:08 AM, Baghirzade, Nara) Battery replaced. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 133 mohms Leakage current: 63.3 uA
1069792 | 201295 | ACC | CMPLT | CMPLT | 2024-01-30 09:49 | 2024-01-30 09:49 | Incoming inspection | (TC 01/24/2024 09:49:07 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069793 | 201296 | ACC | CMPLT | CMPLT | 2024-01-30 09:50 | 2024-01-30 09:50 | Incoming inspection | (TC 01/24/2024 09:50:14 AM, Rendon Lopez , Elizabeth) Incoming inspection
68862 | 201297 | CM | CMPLT | CMPLT | 2024-01-30 09:50 | 2024-01-30 09:50 | Battery does not charge or hold charge, unable to switch it on | (TC 01/30/2024 09:50:04 AM, Baghirzade, Nara) Battery replaced. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 145 mohms Leakage current: 71.1 uA
1069794 | 201298 | ACC | CMPLT | CMPLT | 2024-01-30 09:51 | 2024-01-30 09:51 | Incoming inspection | (TC 01/24/2024 09:50:48 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069795 | 201299 | ACC | CMPLT | CMPLT | 2024-01-30 09:51 | 2024-01-30 09:52 | Incoming inspection | (TC 01/24/2024 09:51:26 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069796 | 201300 | ACC | CMPLT | CMPLT | 2024-01-30 09:52 | 2024-01-30 09:52 | Incoming inspection | (TC 01/24/2024 09:51:59 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069797 | 201301 | ACC | CMPLT | CMPLT | 2024-01-30 09:54 | 2024-01-30 09:54 | Incoming inspection | (TC 01/24/2024 09:53:40 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069798 | 201302 | ACC | CMPLT | CMPLT | 2024-01-30 09:54 | 2024-01-30 09:55 | Incoming inspection | (TC 01/24/2024 09:54:35 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069799 | 201303 | ACC | CMPLT | CMPLT | 2024-01-30 09:55 | 2024-01-30 09:55 | Incoming inspection | (TC 01/24/2024 09:55:14 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069780 | 201304 | ACC | CMPLT | CMPLT | 2024-01-30 09:56 | 2024-01-30 09:56 | Incoming inspection | (TC 01/24/2024 09:55:49 AM, Rendon Lopez , Elizabeth) Incoming inspection
1059828 | 201305 | PM | CMPLT | CMPLT | 2024-01-30 09:56 | 2024-01-30 09:57 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/30/2024 09:56:30 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1069801 | 201306 | ACC | CMPLT | CMPLT | 2024-01-30 09:57 | 2024-01-30 09:57 | Incoming inspection | (TC 01/24/2024 09:56:36 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069802 | 201307 | ACC | CMPLT | CMPLT | 2024-01-30 09:57 | 2024-01-30 09:57 | Incoming inspection | (TC 01/24/2024 09:57:10 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069803 | 201308 | ACC | CMPLT | CMPLT | 2024-01-30 09:58 | 2024-01-30 09:59 | Incoming inspection | (TC 01/24/2024 09:58:30 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069804 | 201309 | ACC | CMPLT | CMPLT | 2024-01-30 09:59 | 2024-01-30 09:59 | Incoming inspection | (TC 01/24/2024 09:59:10 AM, Rendon Lopez , Elizabeth) Incoming inspection
1069800 | 201310 | ACC | CMPLT | CMPLT | 2024-01-30 10:00 | 2024-01-30 10:01 | Incoming inspection | (TC 01/24/2024 10:00:35 AM, Rendon Lopez , Elizabeth) Incoming inspection
1057047 | 201311 | CM | CMPLT | CMPLT | 2024-01-30 10:52 | 2024-01-30 10:56 | Reported that the root is not turning on | (TC 01/30/2024 10:52:34 AM, Baghirzade, Nara) Root was left to charge overnight by Courtney. Booted the Root with no issues; battery fully charged and operating. Checked battery capacity; capacity within nominal range. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055577 | 201312 | CM | CMPLT | CMPLT | 2024-01-30 10:58 | 2024-01-30 10:58 | Reported that the root is not switching on | (TC 01/30/2024 10:57:53 AM, Baghirzade, Nara) Root was left to charge overnight by Courtney. Booted the Root with no issues; battery fully charged and operating. Checked battery capacity; capacity within nominal range. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1069463 | 201313 | CM | CMPLT | CMPLT | 2024-01-30 11:00 | 2024-01-31 10:46 | Loaner removed from service, repaired unit swapped in | (TC 01/30/2024 11:00:26 AM, Cortez, John) Opened up unit to inspect tubing. Seems to be an accumulation of dust in some areas of the tube. No signs of water. Unit to be returned to Olympus (TC 01/31/2024 10:53:49 AM, Cortez, John) Packaged and shipped loaner to olympus
| 201314 | CM | CMPLT | CMPLT | 2024-01-30 11:06 | 2024-01-30 11:07 | Bringing down the radical 7 batteries from 6th fl. storage | (TC 01/30/2024 11:06:28 AM, Rhuland, Courtney) Brought down 2 boxes of radical 7 batteries from 6th fl. storage. (TC 01/30/2024 11:06:07 AM, Baghirzade, Nara) Brought down 2 boxes of radical 7 batteries from 6th fl. storage.
1063746 | 201315 | PM | CMPLT | CMPLT | 2024-01-30 11:21 | 2024-01-30 11:22 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 01/30/2024 11:21:47 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1069386 | 201316 | CM | CMPLT | CMPLT | 2024-01-30 11:49 | 2024-01-30 11:52 | Removed from service, borrowed by endo | (TC 01/30/2024 11:49:23 AM, Cortez, John) Removed from Bronch room, Opened case to inspect tubing. Tubing was good. Device will be held by clinical engineers (Naz?)
| 201317 | ACC | ACTIV | ISSUE | 2024-01-30 12:04 |  | MSH 12 BI Trident HD - Specimen Imaging System DEMO | (TC 01/30/2024 12:04:03 PM, Whittington, Scott) MSH 12 BI Trident HD - Specimen Imaging System DEMO (TC 02/07/2024 10:53:54 AM, Whittington, Scott) MSH 12 BI Trident HD - Specimen Imaging System DEMO (TC 02/07/2024 10:53:54 AM, Whittington, Scott) MSH 12 BI Trident HD - Specimen Imaging System DEMO (TC 02/09/2024 08:34:28 AM, Whittington, Scott) MSH 12 BI Trident HD - Specimen Imaging System DEMO MEETING CHECK ELECTRICAL (TC 02/12/2024 01:31:51 PM, Whittington, Scott) MSH 12 BI Trident HD - Specimen Imaging System DEMO CHECK ELECTRICAL CAPACITYOF INTENDED INSTALL LOCATION WITH IVY RECOMMEND SPLIT RECEPTACLE FOR USE BY TRIDENT AND CONTRAST WARMER IVY FORWARDED REQ TO FACILITIES | MSH 12 BI Trident HD - Specimen Imaging System DEMO MSH_BISPC_DEMO CR1837
| 201318 | CM | CMPLT | CMPLT | 2024-01-30 12:14 | 2024-03-12 11:38 | order Tempcheck DSTA-40 | (TC 01/30/2024 01:40:34 PM, Chukri, Samir) Replaced PO for new tester. attached the PO. (TC 03/12/2024 11:37:02 AM, Chukri, Samir) Received it and give it to Elizabeth
| 201319 | PA | CMPLT | CMPLT | 2024-01-30 12:37 | 2024-02-13 07:58 | Ordering respiratory monitoring accessories for stock | (TC 01/30/2024 12:42:54 PM, Baghirzade, Nara) SC 1003158078 (TC 02/09/2024 02:41:38 PM, Baghirzade, Nara) Received 3/4 items ordered. PO 7500201769 Did not receive item: RAS-45 acoustic sensors. (TC 02/13/2024 07:57:26 AM, Baghirzade, Nara) Received full order. PO 7500201769
1061333 | 201320 | FC | CMPLT | CMPLT | 2024-01-30 12:45 | 2024-01-30 12:45 | QC Checklist | (TC 01/30/2024 12:44:51 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
| 201321 | PA | CMPLT | CMPLT | 2024-01-30 12:49 | 2024-02-09 08:35 | Ordering monophasic neutral plate lead for ESU tester | (TC 01/30/2024 01:12:28 PM, Baghirzade, Nara) SC 1003158107 (TC 02/09/2024 08:30:05 AM, Baghirzade, Nara) PO 7202356274
65976 | 201322 | CM | CMPLT | CMPLT | 2024-01-30 13:20 | 2024-01-30 13:21 | Check in/functional check | (TC 01/30/2024 01:20:35 PM, Baghirzade, Nara) Picked up the syringe pump from SickKids. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. | Picked up the syringe pump from SickKids
1059938 | 201323 | CM | CMPLT | CMPLT | 2024-01-30 13:25 | 2024-01-30 13:26 | Lead 4 not working | (TC 01/30/2024 01:25:28 PM, Canjimavadivel, Kamallesh ) Lead 4 has exposed wires. Replaced and tested ok
66001 | 201324 | CM | CMPLT | CMPLT | 2024-01-30 13:40 | 2024-02-09 14:03 | Returned from SickKids | (TC 01/30/2024 01:40:12 PM, Rhuland, Courtney) During inspection, when device was turned off, watchdog alarm would sound Opened device and found logic board to be faulty Ordering logic board (TC 02/07/2024 10:42:04 AM, Rhuland, Courtney) Replaced logic board Replaced cracked front and rear cases Device is throwing error 100.5010 - Boot block software version not compatible with software flashed into PCU Unable to put device into service mode (TC 02/09/2024 02:01:53 PM, Rhuland, Courtney) Ran external flash kit to update boot block Transferred network settings using ASM Errors cleared Check-in complete All tests passed
1055745 | 201325 | CM | CMPLT | CMPLT | 2024-01-30 14:18 | 2024-01-30 14:19 | Functional check / check-in | (TC 01/30/2024 02:18:53 PM, Baghirzade, Nara) Picked up pump from SickKids. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65897 | 201326 | CM | CMPLT | CMPLT | 2024-01-30 14:21 | 2024-01-30 14:21 | Functional check / check-in | (TC 01/30/2024 02:20:45 PM, Baghirzade, Nara) Picked up pump from SickKids. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65882 | 201327 | CM | CMPLT | CMPLT | 2024-01-30 14:28 | 2024-01-30 14:29 | Functional check / check-in | (TC 01/30/2024 02:29:02 PM, Baghirzade, Nara) Picked up pump from SickKids. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1065889 | 201328 | CM | CMPLT | CMPLT | 2024-01-30 14:44 | 2024-01-30 14:45 | Functional check / check-in | (TC 01/30/2024 02:44:12 PM, Baghirzade, Nara) Downloaded error log. No errors noted. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. | Was sent with no yellow tag/no issues indicated
| 201329 | CON | CMPLT | CMPLT | 2024-01-30 14:45 | 2024-02-02 12:04 | Stryker beds in-service training for 11 South | (TC 01/30/2024 02:45:39 PM, Baghirzade, Nara) Met with Sri; she requested in-service training for 11 South staff as there has been a high turnover of staff in the department. Reached out to Becky to organize 2 in-service sessions; Melissa from Stryker got back with 3 possible dates, Sri confirmed, and waiting on confirmation from Stryker and Melissa. (TC 02/02/2024 12:03:16 PM, Baghirzade, Nara) Melissa, Melissa, and Aleks all came to do the in-service with 11S regarding bed alarms, bed lights, lock/unlock motion controls features, patient helper, and general functionality of bed. Another in-service session will be held on Feb 6.
501224 | 201330 | CM | CMPLT | CMPLT | 2024-01-30 15:35 | 2024-02-22 14:16 | Soon to be replaced by 502075 (UHN asset)
1049658 | 201331 | CM | CMPLT | CMPLT | 2024-01-31 09:15 | 2024-02-16 08:35 | Created by Failed PM WO # 200694 : PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 02/14/2024 08:33:54 AM, Rendon Lopez , Elizabeth) Replaced blender, adjusted gas outlet. PM completed, all tests passed | (TC 01/29/2024 09:12:59 AM, Rendon Lopez , Elizabeth) Found gas outlet loose during PM, needs blender overhaul. Removed from PACU, waiting for blender overhaul
65369 | 201332 | CM | CMPLT | CMPLT | 2024-01-31 09:27 | 2024-02-13 08:39 | Device Labelled Broken | (TC 01/31/2024 09:38:00 AM, Ferreira, Qwynn) Replaced the BP hose and cuff Will need new sp02 cable Needs power cord
63343 | 201333 | CM | CMPLT | NFF | 2024-01-31 09:39 | 2024-01-31 09:41 | Leads are not working | (TC 01/31/2024 09:39:27 AM, Ferreira, Qwynn) Tested the ECG with simulator All waveforms present No issues reading waveforms Tested moving the cables for any noise in signal - no noise present Sent back to unit No issues found
1056583 | 201334 | CM | CMPLT | CMPLT | 2024-01-31 09:45 | 2024-02-02 08:43 | Failed to check in for over 2 weeks | (TC 01/31/2024 09:45:31 AM, Rhuland, Courtney) Device did not check in for 20 days Swapped AED with biomed spare to check connectivity Waiting to see if urology AED checks in from biomed and if biomed AED checks in from urology (TC 02/02/2024 08:42:17 AM, Rhuland, Courtney) Both AED's checked in as expected Swapped devices back and will continue to monitor
66065 | 201335 | CM | CMPLT | CMPLT | 2024-01-31 09:49 | 2024-01-31 13:26 | Broken front/rear cases | (TC 01/31/2024 09:48:56 AM, Rhuland, Courtney) Replaced front and rear cases Replaced bent power cord (TC 01/31/2024 01:24:55 PM, Rhuland, Courtney) Check in complete All tests passed Refer to ASM for details
66060 | 201336 | CM | CMPLT | CMPLT | 2024-01-31 10:48 | 2024-01-31 10:51 | Broken rear case | (TC 01/31/2024 10:48:09 AM, Baghirzade, Nara) Replaced rear case; took longer than usual as the parts inside were covered in something sticky. Cleaned as many parts as possible including screws and IUIs. After replacing the rear case, printed new asset tag, placed new sticker for the SN, and moved the latch from the old case to the new rear case. After assembling the PCU together, plugged it to the AC and noted that the front case needed to be replaced under recall as the AC light did not come on when plugged in. Replaced the front case. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 128 mohms Leakage current: 87.7
1046023 | 201337 | CM | CMPLT | CMPLT | 2024-01-31 10:58 | 2024-02-08 13:57 | Failing self test 4054 - Safety Function Test | (TC 02/08/2024 01:57:00 PM, Scero, Tomlyn) As of SW 2.02 and higher, a pass and fail test result is both permitted and considered a pass. | Safety cut-out too late
| 201338 | CM | CMPLT | CMPLT | 2024-01-31 10:59 | 2024-01-31 11:01 | Sent email to Sharmaine regarding dirty wound VACs | (TC 01/31/2024 10:59:03 AM, Baghirzade, Nara) Noted that regularly when wound VACs are sent down from 11S, the biomed tag notes "to be cleaned and checked". Sent email to Sharmaine to ask the staff to clean the VACs before sending it down to biomed for a QC check as it can pose an infection issue for us.
68392 | 201339 | CM | CMPLT | CMPLT | 2024-01-31 10:59 | 2024-02-08 13:57 | Failing Self test Safety Function Test | (TC 02/08/2024 01:55:53 PM, Scero, Tomlyn) As of SW 2.02 and higher, a pass and fail test result is both permitted and considered a pass. | Safety cut-out too late
1059778 | 201340 | FC | CMPLT | CMPLT | 2024-01-31 11:02 | 2024-01-31 11:05 | QC Checklist | (TC 01/31/2024 11:04:24 AM, Baghirzade, Nara) QC Checklist completed. Returned to 11 South
| 201341 | CM | CMPLT | CMPLT | 2024-01-31 11:16 | 2024-01-31 11:18 | Doppler screen not working | (TC 01/31/2024 11:16:28 AM, Rhuland, Courtney) Went over to OPG to look at their doppler Reporter mentioned the device had been dropped or knocked Gently pressed on the screen and all numbers appeared Functional check - OK
66029 | 201342 | CM | CMPLT | CMPLT | 2024-01-31 11:41 | 2024-01-31 11:41 | Functional check | (TC 01/31/2024 11:40:49 AM, Baghirzade, Nara) Picked up from SickKids. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 184 mohms Leakage current: 92.3 uA
1054599 | 201343 | CM | CMPLT | CMPLT | 2024-01-31 13:16 | 2024-01-31 13:20 | Overview was operating on local mode only | (TC 01/31/2024 01:15:53 PM, Baghirzade, Nara) Received email from focal point stating that the overview is unable to connect to the server and is operating on local mode only. When trying to enter the system configuration, the login window would not pop up. Did a reboot and it took an unusually long time to reboot. The reboot did not mitigate the issue with the local mode connection and it was taking a very long time for any of the main setup buttons/options to load. Connected with Kamran regarding this issue. He requested that courtney and i run the PICIX setup; ran the setup however during the server connection part, it would not connect to MSHPHIIXPRI server. Kamran remoted in, deleted the certificates, but it still would not connect to the server. Disabled the firewall. Ran a command to ping the server, the overview was able to ping the server so the network connection was up and running. Kamran checked the network settings of the adapter and found that the DNS server was changed somehow (unknown cause of the change). Changed the DNS server, ran the PICIX setup and it was successful. Rebooted the overview and was able to connect to patient monitoring successfully. Checked on main central, checked on overview itself, overview is connected to the server.
| 201344 | CM | ACTIV | ISSUE | 2024-01-31 13:23 |  | defective blade inspection | (TC 01/31/2024 01:46:54 PM, Huang, Jessica) got the call from OR desk and picked up the blade contacted Kral Storz and was informed the blade would need to be sent for repair (TC 02/22/2024 01:48:18 PM, Huang, Jessica) received the repair quote from the vendor and forwarded it to Steve for approval.
65968 | 201345 | CM | CMPLT | CMPLT | 2024-01-31 14:00 | 2024-01-31 14:02 | Channel error | (TC 01/31/2024 01:59:45 PM, Rhuland, Courtney) Channel error 354.6770 apparent Opened device and reseated pressure sensor Error cleared During check-in, pressure failed test Calibrated and verified pressure Check-in complete All tests passed post calibration Refer to ASM for details
| 201346 | CM | CMPLT | CMPLT | 2024-01-31 14:03 | 2024-01-31 14:03 | Laser Meeting | (TC 01/31/2024 02:02:57 PM, Baghirzade, Nara) Meeting with Adeel and Jessica re: lasers at MSH (TC 01/31/2024 02:02:57 PM, Baghirzade, Nara) Meeting with Adeel and Jessica re: lasers at MSH
| 201347 | CM | CMPLT | CMPLT | 2024-01-31 14:31 | 2024-01-31 14:46 | BP Cuffs for 11N P3A | (TC 01/31/2024 02:33:05 PM, Baghirzade, Nara) Met with Sarabeth to drop off a box of 32 normal cuffs, 4 large, and 2 small adult. Tried to find a space on 11N to store them however we were not able to find the space. Sarabeth took all the large, small, and a couple of the normal adult cuffs and i agreed to store the rest of them in my area.
68547 | 201348 | CM | CMPLT | CMPLT | 2024-01-31 14:50 | 2024-02-14 14:33 | Left side DIVA port failure | (TC 01/31/2024 02:50:59 PM, Scero, Tomlyn) Brought device downstairs to workshop for troubleshooting. Confirmed that left side DIVA 1 port failing selftest. Downloaded logs for review. DSR showing "Left NGF valve doesn't close" with 3 possibly valves not working. I went through the tests that it says to perform to see which actual valve is faulty. Based on my tests, I think it's the 2.59 saturated vapor valve that isn't fully closing. (TC 02/01/2024 02:30:27 PM, Scero, Tomlyn) Completely disassembled the AGM to remove the DIVA supply. Disassembly of the diva supply also necessary to locate valve 2.59 (TC 02/13/2024 03:08:23 PM, Scero, Tomlyn) replaced 2.59 saturated vapor valve Reassembled most of unit to test that its functional. (TC 02/14/2024 02:26:05 PM, Scero, Tomlyn) Completely assembled the unit and tested the function of the gas valves using the test procedures in the IPM-L. All tests passed. Self test was done again and passed. Taken back to the OR. (TC 02/15/2024 09:48:52 AM, Scero, Tomlyn) Filled out BME Incident Report follow up form for incident 69025 | Incident report 69025
68373 | 201349 | CM | CMPLT | CMPLT | 2024-01-31 14:51 | 2024-02-05 10:20 | Pressure sensor failure | (TC 02/05/2024 10:12:31 AM, Scero, Tomlyn) Confirmed pressure sensor failure during self test. Downloaded DSR to save and look over. Bag sensor faulty. Replaced bag sensor and confirmed with IPM-L tests that it's working correctly. Reassembled device and completed self test with no errors. Volume control simulation complete, no errors. RTs notified that AGM ready for use.
1062551 | 201350 | CM | CMPLT | CMPLT | 2024-01-31 15:36 | 2024-01-31 15:37 | bed alarming | (TC 01/31/2024 03:36:21 PM, Rendon Lopez , Elizabeth) Reconnected bed after nurse unplugged it. bed width on head wider, adjusted rails, working-OK
32520 | 201351 | CM | CMPLT | CMPLT | 2024-01-31 16:00 | 2024-02-01 08:35 | MSH DR2 DETECTOR Z NOT MOVING | (TC 01/31/2024 04:00:58 PM, Whittington, Scott) MSH DR2 DETECTOR Z NOT MOVING 15A FUSE BLOWN REPLACEMENT FUSE BLOWN MOTOR CONTROLLSIDE OF 20A CONTROLLER FAULTY (TC 01/31/2024 04:14:40 PM, Whittington, Scott) MSH DR2 DETECTOR Z NOT MOVING (TC 02/01/2024 08:14:00 AM, Whittington, Scott) MSH DR2 DETECTOR Z NOT MOVING | MSH DR2 DETECTOR Z NOT MOVING
52170 | 201352 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | Incubator to be retired
53076 | 201353 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Transport
62500 | 201354 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-07-12 13:46 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 07/12/2024 01:45:42 PM, Ferreira, Qwynn) PM Completed Battery Replaced Cleaned under the mattress Functional Check - OK Tested 35C - 35.1C No loose parts, cracks, or worns parts noticed. Condition - Good (TC 07/12/2024 01:45:42 PM, Ferreira, Qwynn) PM Completed Battery Replaced Cleaned under the mattress Functional Check - OK Sunction - Good Tested 35C - 35.1C No loose parts, cracks, or worns parts noticed. Condition - Good
62501 | 201355 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
62728 | 201356 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-15 09:10 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | (TC 03/27/2024 01:38:36 PM, Chukri, Samir) Called rex ( R.T tech) to check with Michelle and Asteve if they are still use this ventilator, no parts and maintenance support by company should retir this unit. Retired
62729 | 201357 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | (TC 04/15/2024 09:10:44 AM, Chukri, Samir) Retired this vent requested by Steve
62942 | 201358 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: High Flow Insufflator
63523 | 201359 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-02-08 10:00 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 12/18/2023 09:59:16 AM, Huang, Jessica) PM completed
63937 | 201360 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Pulse-Free Flow Peristaltic Pump
64240 | 201361 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM Major Annual: Microstream CO2 extension
64241 | 201362 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM Major Annual: Microstream CO2 extension
65290 | 201363 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
65889 | 201364 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-25 13:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:21:50 PM, Rhuland, Courtney) Calibration and PM completed All tests passed Refer to ASM for details
65904 | 201365 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-26 14:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/26/2024 02:23:24 PM, Rhuland, Courtney) PM completed by BD
65908 | 201366 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-04 13:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/18/2024 01:44:37 PM, Huang, Jessica) replaced the rear case and completed PM
65919 | 201367 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65927 | 201368 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65936 | 201369 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-15 11:56 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/15/2024 11:55:21 AM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65939 | 201370 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65948 | 201371 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-17 12:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/17/2024 12:23:52 PM, Baghirzade, Nara) Replaced: -front case -rear case -mylar gasket Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65950 | 201372 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65978 | 201373 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-04 13:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/18/2024 01:40:48 PM, Huang, Jessica) replaced the rear case and completed PM
65994 | 201374 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-01 12:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
66000 | 201375 | PM | CMPLT | PMFAIL | 2024-02-01 02:13 | 2024-03-25 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
66012 | 201376 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-26 12:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
66036 | 201377 | PM | CMPLT | PMFAIL | 2024-02-01 02:13 | 2024-03-13 10:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66045 | 201378 | PM | CMPLT | PMFAIL | 2024-02-01 02:13 | 2024-03-15 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
66068 | 201379 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-11-06 13:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/06/2024 01:36:59 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 62 mohms Leakage current: 57.0 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66462 | 201380 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-09 11:02 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced V-battery and air filter. Run battery check. 220 minutes, 43 Ah. Tested. Functional. SW2.51.01 Leakage: 91 uA and 90 ohm.
66472 | 201381 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-06 07:34 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Battery date: 11.2023. Run battey check. 230 min(31Ah) and 25 min(2Ah).Performed 4th year PM after 3nd year PM last year. New PS500 with Li-Ion battery. It has GS500. Tested. Functional. SW2.51 Leakage: 77uA and 88 ohm.
66804 | 201382 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66806 | 201383 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-20 09:08 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/20/2024 09:08:20 AM, DeGuzman, Leila) PM passed and completed. See procedure for details.
67354 | 201384 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
67879 | 201385 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-15 14:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 03/15/2024 02:57:43 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67880 | 201386 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-01 14:28 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 03/01/2024 02:27:24 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
67943 | 201387 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-02-27 12:23 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 02/27/2024 12:04:08 PM, Rhuland, Courtney) Repaired broken door latch Calibrated: Line voltage Main voltage RH Scale Replaced 9V rechargeable battery Functional check - OK Set 35C & 75% humidity Read 35C & 75.7% humidity Electrical safety - OK 95 mOhms 80.7 uA
68022 | 201388 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-08 15:29 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Transport vent. Replaced PS500 batteries and V-battery.Performed 4th year PM after 3nd year PM last year. It has GS500. Tested. Functional. SW2.51 Leakage: 79uA and 91 ohm.
68023 | 201389 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-13 15:39 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Transport vent. Replaced PS500 batteries and V-battery.Performed 4th year PM after 3nd year PM last year. It has GS500. Tested. Functional. SW2.51 Leakage: 81uA and 91 ohm.
68110 | 201390 | PM | CNP | ENL | 2024-02-01 02:13 | 2024-09-05 11:28 | PM-MAJOR - Annual PM For: Aspirators, Emergency
68544 | 201391 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-05-15 10:39 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 04/29/2024 10:36:32 AM, Rendon Lopez , Elizabeth) Removed AGM from OR 2, brought down for PM (TC 04/30/2024 10:36:32 AM, Rendon Lopez , Elizabeth) 1 Year PM parts replaced. PM procedures started (TC 05/02/2024 10:36:32 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. See attached checklist
68609 | 201392 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-16 12:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD.
68623 | 201393 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-18 14:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68628 | 201394 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-23 10:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/23/2024 10:06:40 AM, Rhuland, Courtney) PM completed by BD
68644 | 201395 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-04 12:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 12:50:44 PM, Rhuland, Courtney) PM completed by BD
68730 | 201396 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-25 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68742 | 201397 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-11 15:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68743 | 201398 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-15 11:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68753 | 201399 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-12 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68755 | 201400 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-17 09:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:40:59 AM, Rhuland, Courtney) PM completed by BD
68756 | 201401 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-18 12:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68818 | 201402 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-16 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68835 | 201403 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-25 10:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68848 | 201404 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-21 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68889 | 201405 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68918 | 201406 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-03 13:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:16:53 PM, Rhuland, Courtney) PM completed by BD
68966 | 201407 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-27 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68970 | 201408 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68995 | 201409 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-26 14:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69036 | 201410 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-21 20:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69075 | 201411 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-17 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69076 | 201412 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-12 09:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69088 | 201413 | PM | CMPLT | PMFAIL | 2024-02-01 02:13 | 2024-03-14 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69122 | 201414 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69184 | 201415 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-28 10:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69204 | 201416 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-05-16 09:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/16/2024 09:03:52 AM, DeGuzman, Leila) PM completed - all tests passed. cleaned.
69229 | 201417 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-09 11:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 11:40:54 AM, Patel, Rutvik) by BD
69230 | 201418 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-28 10:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69233 | 201419 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-12 09:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69242 | 201420 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-20 10:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69244 | 201421 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-24 14:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69304 | 201422 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-28 09:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69312 | 201423 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-26 11:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69367 | 201424 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-14 13:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 01:24:23 PM, Huang, Jessica) replaced the female iui connector and completed PM
69380 | 201425 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-27 10:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69414 | 201426 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-11 16:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69429 | 201427 | PM | CMPLT | PMFAIL | 2024-02-01 02:13 | 2024-03-15 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69459 | 201428 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-26 09:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69467 | 201429 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-21 10:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/21/2024 10:05:36 AM, Reid, Emmerson ) Ran rate calibration, ASM PM, all tests passed. PM Complete.
69487 | 201430 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69488 | 201431 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-23 10:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69512 | 201432 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-18 13:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69556 | 201433 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-27 10:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69573 | 201434 | PM | CMPLT | PMFAIL | 2024-02-01 02:13 | 2024-03-14 15:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69630 | 201435 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-15 11:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69656 | 201436 | PM | CMPLT | PMFAIL | 2024-02-01 02:13 | 2024-03-14 13:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69686 | 201437 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-05-22 14:15 | PM-MAJOR - Annual PM For: Portable Patient Monitor | (TC 05/03/2024 11:20:44 AM, Scero, Tomlyn) Ordered a new battery for PM (TC 05/22/2024 02:14:45 PM, Scero, Tomlyn) All PM tests complete and passed. Battery replaced. LC 69 uA GWR .07 ohm
65316 | 201438 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-20 15:45 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 03/20/2024 03:04:49 PM, Cortez, John) Physical inspection passed, cleaned unit, Functional test passed, Mains leakage (uA): NC, SFC - HI 7.3, 15.1 - LO 6.2, 12.8 - P/P 10, 20.6 - BI 1 1.8 - handswitch 0.8, 1.2 all within safety limits RF output passed, see attached document RF leakage test passed -35HI 53 mA -35BI 26 mA
32540 | 201439 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-08 15:34 | PM - Monthly | PM completed
32560 | 201440 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-05 15:09 | PM - Monthly | PM Major done Recon Unit cleaned VIMS Enterprise node setup done AQNET send error troubleshooting done - PACs team is working on it on their end. Slice count - 1378371 TOS phantom a(-990.2, 9.4) b(332.5, 14.5) c(132.1, 12.8) d(96.3, 13.3) e(-106.0, 12.3) f(-1.3, 14.2) -System passed all tests
32580 | 201441 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-06 14:49 | PM - Monthly | PM Major done DICOM node configuration done for VIMS Clinical/Enterprise NRA/AIR cal done for M FOV Slice count - 201513 TOS phantom a(-992.8, 10.6) b(334.3, 14.1) c(133.9, 144.1) d(97.3, 13.8) e(-106.2, 12.2) f(0.2, 14.4) -System passed all tests
541124 | 201442 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541125 | 201443 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541127 | 201444 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541128 | 201445 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541205 | 201446 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541206 | 201447 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541207 | 201448 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541208 | 201449 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1046273 | 201450 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-04 16:03 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 03/01/2024 10:28:34 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/07/2024 07:51:05 AM, Whittington, Scott) 201450 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL ABDOMINAL PROBES LESS THA IDEAL IMAGING
1046274 | 201451 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-04 16:04 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 03/01/2024 10:29:01 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/07/2024 07:50:29 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL ABDOMINAL PROBES LESS THA IDEAL IMAGING | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL ABDOMINAL PROBES LESS THA IDEAL IMAGING
1046275 | 201452 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-04 16:05 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 03/01/2024 10:29:25 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/07/2024 07:51:50 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL ABDOMINAL PROBES LESS THA IDEAL IMAGING | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL ABDOMINAL PROBES LESS THA IDEAL IMAGING
1046276 | 201453 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-04 16:04 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 03/01/2024 10:29:49 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 03/07/2024 07:52:24 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL ABDOMINAL PROBES LESS THA IDEAL IMAGING | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL ABDOMINAL PROBES LESS THA IDEAL IMAGING
1046294 | 201454 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-26 10:24 | PM-MAJOR - Annual PM for: Electrosurgical Units | (TC 04/22/2024 10:23:43 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached results
1046407 | 201455 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1046408 | 201456 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047039 | 201457 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-14 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 01:57:27 PM, Huang, Jessica) replaced the female iui connector and completed PM
1048472 | 201458 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-01 12:20 | HARP - Semi Annual |  | HARP completed
1048472 | 201459 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-01 12:22 | PM |  | PM performed.
1048524 | 201460 | PM | CANCL | DUP | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1048537 | 201461 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-22 14:40 | HARP - Semi-Annual | HARP
1048537 | 201462 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-21 16:41 | PM - Semi-Annual | PM
1048629 | 201463 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 09:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048632 | 201464 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-22 08:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048636 | 201465 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048641 | 201466 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 09:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048647 | 201467 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:14:35 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048648 | 201468 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:04:24 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048649 | 201469 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 05:19:40 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048650 | 201470 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:08:27 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048652 | 201471 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:06:38 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048656 | 201472 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:41:58 AM, Huang, Jessica) replaced the female iui connector completed PM
1048659 | 201473 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048667 | 201474 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:31:56 AM, Huang, Jessica) replaced the female iui connector completed PM
1048669 | 201475 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048671 | 201476 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 11:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048679 | 201477 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 13:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048682 | 201478 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 13:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:21:04 AM, Huang, Jessica) replaced the female iui connector completed PM
1048687 | 201479 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 13:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:38:59 AM, Huang, Jessica) replaced the female iui connector completed PM
1048690 | 201480 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 13:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 05:23:22 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048692 | 201481 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-19 13:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:06:58 AM, Huang, Jessica) replaced the female iui connector completed PM
1048698 | 201482 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-22 17:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:10:32 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048699 | 201483 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049416 | 201484 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-29 09:18 | PM-MAJOR - Annual PM For: Scale
1051372 | 201485 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-22 13:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2024 01:34:23 PM, Reid, Emmerson ) ASM PM Complete.
1051373 | 201486 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-28 11:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1053034 | 201487 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-23 08:35 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 04/23/2024 08:32:02 AM, Ferreira, Qwynn) PM Passed Functional Check - OK Replaced battery Ordered new hood - repaired old hood until new one comes in Replaced power cord Replaced port grommet Replaced Iris sleeve (TC 04/22/2024 08:32:02 AM, Ferreira, Qwynn) PM Passed Functional Check - OK Replaced battery Ordered new hood - repaired old hood until new one comes in Replaced power cord Replaced port grommet Replaced Iris sleeve
1053348 | 201488 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-08 17:24 | MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG | (TC 03/08/2024 05:23:05 PM, Dolinar, John) MSH BI 1272 H3 HARP (TC 03/08/2024 05:54:00 PM, Whittington, Scott) MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG (TC 03/08/2024 05:54:00 PM, Whittington, Scott) MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG | MSH BI 1272 H3 HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG
1053378 | 201489 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-06-04 13:28 | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric | (TC 06/04/2024 01:27:50 PM, Rendon Lopez , Elizabeth) PM parts replaced, all tests passed. Replaced exp cassette membrane (SN:103075). Replaced 1x filter.
1053380 | 201490 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-05-10 10:07 | PM-MAJOR - Annual PM For: Ventilators | performed annual PM. passed pre-use check. functional. replaced 2 x expiratory cassette membrane. | 2 years warranty covers PMs, and repairs
1053381 | 201491 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-16 12:08 | PM-MAJOR - Annual PM For: Ventilators | performed annual PM. replaced expiratory casette membrane too. Passed pre-use check. Functional. | 2 years warranty covers PMs and repairs
1054037 | 201492 | PM | CANCL | CANCL | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM | PM not required
1054424 | 201493 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-07-26 08:09 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/25/2024 08:08:38 AM, Baghirzade, Nara) PM completed. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
1054427 | 201494 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
1054428 | 201495 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
1054434 | 201496 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
1054441 | 201497 | PM | ACTIV | ISSUE | 2024-02-01 02:13 |  | PM-MAJOR - Annual PM
1055121 | 201498 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-04-03 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:20:08 PM, Rhuland, Courtney) PM completed by BD
1055128 | 201499 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-08-16 10:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055750 | 201500 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-21 08:28 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 08:27:23 AM, Rhuland, Courtney) PM completed by BD
1056936 | 201501 | PM | CMPLT | CMPLT | 2024-02-01 02:13 | 2024-03-01 14:34 | PM Major Annual: Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 03/01/2024 02:33:49 PM, Canjimavadivel, Kamallesh ) PM complete, see procedure
1057601 | 201502 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-24 10:37 | PM-MAJOR ANNUAL PM for: Ventilators, Intensive Care, Adult | (TC 04/23/2024 10:34:36 AM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 13315hrs Total vent time: 8295 Membrane capacity: 100%
1057603 | 201503 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-07-24 16:55 | PM-MAJOR ANNUAL PM for: Ventilators, Intensive Care, Adult | (TC 07/24/2024 04:54:47 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 8344 hrs Total vent time: 2560 hrs Membrane capacity: 80%
1057680 | 201504 | PM | ACTIV | READY | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: TRANSCUTANEOUS MONITORS, CARBON DIO
1057704 | 201505 | PM | CNP | ENL | 2024-02-01 02:14 | 2024-07-19 07:57 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1057706 | 201506 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-04 09:49 | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide | (TC 04/04/2024 09:47:51 AM, Canjimavadivel, Kamallesh ) Calibration complete, PM pass
1057892 | 201507 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-09-10 14:29 | PM MAJOR - ANNUAL PM for: High Flow, Warmer | (TC 09/10/2024 01:34:49 PM, Reid, Emmerson ) General Alarm Test - OK Low Recirculating Fluid Alarm Test - OK Over Temp Alarm Test - OK Disposable Alarm Test - OK Fluid Warmer Alarm Test - OK Drained old fluid and added new recirculating fluid Performed Functional Test - OK Temperature 41.6 C Replaced O-rings Cleaned fan filter Refilled with new recirculating fluid Electrical Safety - Current Leakage - 73 uA - Ground Resistance - 50 mohms PM Comlplete (TC 09/10/2024 02:29:09 PM, Reid, Emmerson ) Checked the internal tubing for cracking due to recall announcement. Everything looked to be in very good condition.
1058056 | 201508 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: Anesthesia Units
1058058 | 201509 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-11-22 14:21 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 10/29/2024 12:22:39 PM, Scero, Tomlyn) 3 year PM parts replaced. (TC 11/21/2024 02:18:21 PM, Scero, Tomlyn) Continued with PM tests as per IPM-L (TC 11/22/2024 02:20:50 PM, Scero, Tomlyn) All PM tests complete and passed. Results sheet attached to asset for review.
1058059 | 201510 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: Anesthesia Units
1058060 | 201511 | PM | ACTIV | READY | 2024-02-01 02:14 |  | PM Major - Delivery Module SEVO
1058061 | 201512 | PM | ACTIV | READY | 2024-02-01 02:14 |  | PM Major - Delivery Module SEVO
1058062 | 201513 | PM | ACTIV | READY | 2024-02-01 02:14 |  | PM Major - Delivery Module SEVO
1058063 | 201514 | PM | ACTIV | READY | 2024-02-01 02:14 |  | PM Major - Delivery Module SEVO
1058255 | 201515 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-19 14:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058289 | 201516 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-13 10:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 10:00:23 AM, Rendon Lopez , Elizabeth) PM completed by BD
1059202 | 201517 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-05-14 11:14 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/14/2024 11:11:04 AM, Scero, Tomlyn) PM checks complete and passed.
1059204 | 201518 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-06-05 14:29 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 06/05/2024 02:29:14 PM, Scero, Tomlyn) PM tests complete and passed. Battery replaced.
1059206 | 201519 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059207 | 201520 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-11 10:53 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/11/2024 10:09:28 AM, Scero, Tomlyn) PM tests complete and passed. BP leak <5mmhg Replaced battery. (TC 03/11/2024 10:09:28 AM, Scero, Tomlyn) PM tests complete and passed. BP leak <5mmhg Replaced battery.
1059208 | 201521 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1059209 | 201522 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-05-27 10:07 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/27/2024 10:07:01 AM, Scero, Tomlyn) PM checks complete and passed.
1059223 | 201523 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059655 | 201524 | PM | CANCL | DUP | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | PM done 2023, rescheduled every two years
1059795 | 201525 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-25 23:45 | HARP | completed.
1059795 | 201526 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-22 23:53 | PM - Semi-Annual | Completed.
1059796 | 201527 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-15 23:32 | HARP | Completed.
1059796 | 201528 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-14 23:31 | PM - Semi-Annual | Completed.
1059810 | 201529 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-08 23:47 | HARP | Completed.
1059810 | 201530 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-07 23:48 | PM - Semi-Annual | Completed.
1059828 | 201531 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-10-03 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 10/04/2024 07:44:13 AM, Alvarado, Karla) Disassembled to replace battery, filters and tubing. Replaced air vent from a spare altrix. Reassembled and performed PM. Performed functional test: LCD functionality, visual, and audible alarms; and high temp cutout. Performed probe resistance verification, where: - Port 1= device temperature at 33.4 C, measured 34 C; device flow at 1.4 l/m, measured 1.5 l/m. - Port 2= device temperature at 17.8 C, measured 17.8 C; device flow at 1.9 l/m, measured 2.0 l/m. - Port 3= device temperature at 13.9 C, measured 13.9 C; device flow at 1.4 l/m, measured 1.6 l/m. Electrical safety results: - Leakage current = 182 uA - GND Resistance = 168 mohms All test passed. Cleaning cycle performed (TC 10/03/2024 07:47:04 AM, Rhuland, Courtney) supervised
1060030 | 201532 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-11 12:04 | PM | (TC 03/11/2024 12:01:11 PM, Baghirzade, Nara) Annual PM completed. Replaced valve o-rings, pneumotachograph mesh, and flow/pressure regulator. Cleaned pistons and box seal. Completed tests/checks as per manual. Gaby completed the DLCO and regular cal successfully.
1060031 | 201533 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-11 14:24 | PM | (TC 03/11/2024 02:22:35 PM, Baghirzade, Nara) Annual PM completed. Replaced mesh filter and O-rings. Completed linearisation calibration. QC & daily cal completed successfully.
1060032 | 201534 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-15 13:27 | PM | (TC 03/15/2024 01:24:05 PM, Baghirzade, Nara) PM passed and completed. Replaced valve o-rings, O2 fuel cell, and pressure regulator. All checks passed. Calibration and biological test passed. One issue encountered was when the Expair folder was copied and pasted back in to C:/ drive, the calibration page had one parameter changed from He-Co to CH4-C2H2-CO which is a test not used. When making the original folder active (not the copied one), then the parameter is back to He-CO. Will follow up with Novus regarding this.
1062779 | 201535 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-14 13:20 | PM | (TC 03/14/2024 01:17:13 PM, Baghirzade, Nara) Annual PM completed. Replaced umbilical, sample o ring and filter, flow regulator and O2 electrochemical cell. During calibration, ran into issue with zeroing the O2 cell. Contacted Antony and Chris; Chris remoted in and guided John and I to O2 cell calibration and troubleshooting. John adjusted the potentiometer on O2 cell board, we adjusted the offset values on ExpairTec and all O2 calibration tests passed.
1062999 | 201536 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-19 13:36 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 02/13/2024 10:04:19 AM, Cortez, John) Ordered pm kit (TC 03/18/2024 01:33:52 PM, Cortez, John) Replaced PM parts, 70 mohms, 20uA Completed cleaning cycle
1063401 | 201537 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-11 10:36 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/11/2024 10:27:31 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 15746 hrs Total vent time: 6519 hrs PM complete
1063402 | 201538 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-16 16:28 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/16/2024 04:26:04 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 8674 hrs Total vent time: 707 hrs Membrane capacity: 95%
1063403 | 201539 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-22 16:18 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/22/2024 04:15:20 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 17039hrs Total vent time: 5984hrs Membrane capacity: 46%
1063404 | 201540 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-19 12:58 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/18/2024 12:55:32 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 15445hrs Total vent time: 6584hrs Membrane capacity: 42%
1063405 | 201541 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-22 16:25 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/22/2024 04:22:51 PM, Cortez, John) PM parts replaced, expiratory cassette membrane replaced Pre-op check passed Total operating time: 16326hrs Total vent time: 7718hrs Membrane capacity: 21%, replaced
1063406 | 201542 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-16 16:21 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/16/2024 04:18:49 PM, Cortez, John) Replaced PM parts Pre-use check passed Total operating time: 16046 hrs Total vent time: 7205 hrs Membrane capacity: 30%
1063407 | 201543 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-22 09:48 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/19/2024 09:45:29 AM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 16814hr Total vent time: 6579hr Membrane capacity: 24%
1063408 | 201544 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-22 16:21 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/22/2024 04:20:00 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 8702hrs Total vent time: 3013hrs Membrane capacity: 72%
1063409 | 201545 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-22 09:53 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/19/2024 09:51:53 AM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 10195hr Total vent time: 3899hr Membrane capacity: 62%
1063746 | 201546 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-06-25 07:19 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 06/24/2024 07:18:41 AM, Rhuland, Courtney) Replaced the back filters Replaced Battery Replaced air filter Replaced port/connector covers PM completed All tests passed Unit was disinfected, rinsed, and refilld with sterile water and sent back to unit.
1064201 | 201547 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-19 13:07 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/18/2024 01:05:05 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 14496hrs Total vent time: 4358hrs Membrane capacity: 56%
1064202 | 201548 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-19 12:20 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/17/2024 12:12:23 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 16186 hrs Total vent time: 4712 hrs Membrane capacity: 51%
1064203 | 201549 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-19 13:11 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/18/2024 01:08:27 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 15751hrs Total vent time: 4987hrs Membrane capacity: 58%
1064204 | 201550 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: Ventilators
1064205 | 201551 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-05-10 16:03 | PM-MAJOR - Annual PM For: Ventilators | (TC 05/10/2024 04:00:40 PM, Cortez, John) PM parts replaced Pre-op check passed Upgraded humidifier software to 6.1 Total operating time: 17458 hrs Total vent time: 7821 hrs Membrane capacity: 10%, membrane replaced
1064206 | 201552 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-19 13:03 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/18/2024 12:59:30 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 8181hrs Total vent time: 2265hrs Membrane capacity: 78%
1064207 | 201553 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-16 13:08 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/16/2024 01:02:24 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 8082 hrs Total vent time: 2337 hrs Membrane capacity: 82%
1064208 | 201554 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-05-09 16:41 | PM-MAJOR - Annual PM For: Ventilators | (TC 05/09/2024 04:38:17 PM, Cortez, John) PM parts replaced Pre-op check passed Upgraded humidifier software to 6.1 Total operating time: 16664 hrs Total vent time: 5594 hrs Membrane capacity: 46%
1064209 | 201555 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-03-25 09:52 | PM-MAJOR - Annual PM For: Ventilators | (TC 03/25/2024 09:39:55 AM, Canjimavadivel, Kamallesh ) Replaced filters and nozzles Pre check pass Total op time: 15980 hrs Total vent time: 6441 hrs PM complete
1064210 | 201556 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-04-19 12:38 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/17/2024 12:22:11 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 12035hrs Total vent time: 3478hrs Membrane capacity: 62%
1064211 | 201557 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-05-09 16:44 | PM-MAJOR - Annual PM For: Ventilators | (TC 05/09/2024 04:42:19 PM, Cortez, John) PM parts replaced Pre-op check passed Upgraded humidifier software to 6.1 Total operating time: 14068 hrs Total vent time: 4535 hrs Membrane capacity: 57%
1064687 | 201558 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
1066598 | 201559 | PM | ACTIV | ISSUE | 2024-02-01 02:14 |  | PM MAJOR-Annual PM For: ELECTRICAL SAFETY ANALYZER
1066599 | 201560 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-06-11 12:23 | Annual PM | (TC 06/11/2024 12:23:08 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1066600 | 201561 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-06-11 12:25 | Annual PM | (TC 06/11/2024 12:24:55 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1066604 | 201562 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-06-11 12:27 | Annual PM | (TC 06/11/2024 12:26:51 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1066605 | 201563 | PM | CMPLT | CMPLT | 2024-02-01 02:14 | 2024-06-11 12:29 | Annual PM | (TC 06/11/2024 12:28:52 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
| 201564 | OC | CMPLT | CMPLT | 2024-02-01 06:26 | 2024-02-02 07:19 | Nara Baghirzade On-Call; January 25 - Feb 1 2024 | (TC 01/31/2024 09:30:00 PM, Baghirzade, Nara) Received an on-call from PACU POD C regarding their PICIX central. It was completely shut off and they were unable to switch it on. Did troubleshooting over the phone by asking them to switch off the UPS, switch it back on, and try to turn on the central separately. Did not work. Asked them to force shutdown the PICIX and boot it back up, upon boot up the central would freeze and not move past the "HP" logo. Came on-site and tried to bypass UPS entirely; PICIX still would not boot up properly. Swapped the PICIX with the spare biomed PICIX, had to delete the MSHPACU3S off the philips network first and then ran a full setup to configure the host name and IP address; installed database and did a complete rebuild of the central with the sectors. Set it up and had the staff admit patients and associate the patients on cerner; all was successful. While I was troubleshooting the PACU PICIX I noticed that the CCUTELE3O in L&D had switched over to local mode. After PACU, I went to L&D, logged out of patient monitoring for CCUTELE3O, logged in as a supportuser and updated the network settings DNS primary and secondary to the correct address. Rebooted CCUTELE3O and confirmed that it was operating properly. (TC 02/01/2024 10:30:00 PM, Baghirzade, Nara) Received a call from CCU, their central was down. Asked the nurse to make sure UPS was on, he turned it on and confirmed that 3 green LEDs were lit. After the central came back up, nurse informed me that they were still no getting any data on the central but were able to see the sectors only. Tried to remote in but all CCU centrals and overviews were not accessible remotely. There was a power outage in downtown Toronto. Came on site, and when passing by the IT closet, I could hear a loud beeping. First went to ccu to check on them, then came back down to biomed to the get IT closet key. Went back up to the IT closet and found that the UPS did not boot up properly when power came back on, which caused the switch to be off. After power cycling the UPS, the battery symbol was amber and it was last replaced in May 2022. Replaced the UPS battery that is connected to switch 10U_DISTA. Checked on CCU again and they were getting all their data from patient rooms again. Central and overviews are functional. Placed the master key for the IT closet back in the biomed key box.
| 201565 | PA | CMPLT | CMPLT | 2024-02-01 06:56 | 2024-02-02 12:09 | Ordering sterile water for altrix | (TC 02/01/2024 07:01:43 AM, Baghirzade, Nara) SC 1003159597 (TC 02/02/2024 12:08:56 PM, Baghirzade, Nara) Order # 4150271225
1052940 | 201566 | CM | CMPLT | CMPLT | 2024-02-01 07:54 | 2024-02-01 07:58 | Meeting with GE for contract and TEE discussion | (TC 01/31/2024 12:30:00 PM, Fan, Tak) Met with Babak. Confirmed contract and roll over options. made sure we distinguish between echo and other departments for the contract and coverage.
| 201567 | RDS | CMPLT | CMPLT | 2024-02-01 08:46 | 2024-01-31 08:45 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
| 201568 | CM | CMPLT | CMPLT | 2024-02-01 08:54 | 2024-01-31 08:53 | worked with RT | worked with RT
1055575 | 201569 | CM | CMPLT | CMPLT | 2024-02-01 09:06 | 2024-01-31 09:06 | Root battery is not holding charge | Root battery is not holding charge. nit shiped to Masimo. Masimo replaced the unit with sn.2000165454. Programmed new unit and radical 7. Updated config. tested bp, temp and spo2. functional.
1055581 | 201570 | CM | CMPLT | CMPLT | 2024-02-01 10:59 | 2024-02-01 10:59 | Root battery is not holding charge | Root battery is not holding charge. unit shiped to Masimo. Masimo replaced the unit with sn.2000186850. Programmed new unit and radical 7. Updated config. tested bp, temp and spo2. functional.
1062237 | 201571 | CM | CMPLT | CMPLT | 2024-02-01 12:23 | 2024-02-01 12:25 | Reported that bed functions were not working | (TC 02/01/2024 12:23:22 PM, Baghirzade, Nara) Tested bed functions and found that one of the nursing staff locked the motion controls. Unlocked all motions and tested all motion controls (side panels, patient side panels, and main display). Bed is fully functional. Let the staff know and showed them how to unlock/lock the controls.
1062389 | 201572 | CM | CMPLT | CMPLT | 2024-02-01 12:51 | 2024-02-01 13:04 | Functional check | (TC 02/01/2024 12:50:47 PM, Baghirzade, Nara) Staff reported bed not working however could not find an issue with the bed. Checked all motion controls including side panels and patient panels. Bed is functional. | Staff reported bed not working
| 201573 | CM | CMPLT | CMPLT | 2024-02-01 13:57 | 2024-02-01 13:57 | worked on TMS | worked on TMS
1058237 | 201574 | CM | CMPLT | CMPLT | 2024-02-02 07:55 | 2024-03-19 09:22 | Not working | (TC 02/01/2024 02:54:00 PM, Rhuland, Courtney) Pressure calibration failed Changed pressure sensor - still failed Tried to change logic board - still failed (TC 02/02/2024 07:56:05 AM, Rhuland, Courtney) Replaced pressure sensor with one from working LVP Pressure sensor good Tried pressure sensor from different order and failed Pressure sensors may be defective (TC 03/19/2024 09:21:47 AM, Reid, Emmerson ) flashed pump to update SN
| 201575 | CM | CMPLT | CMPLT | 2024-02-02 08:45 | 2024-03-01 15:34 | misc biomed - february 2024 | (TC 02/01/2024 08:45:32 AM, Cortez, John) travel to Vaughan clinic (TC 02/01/2024 08:46:10 AM, Cortez, John) Dropped off 3x pediatric SPO2 sensors, and answered nurse's questions about Masimo root monitor (TC 02/01/2024 08:57:09 AM, Cortez, John) Samir and I looked into CLV190 (sn7773111) invoice. Vendor matched invoice with wrong PO. (TC 02/01/2024 08:59:22 AM, Cortez, John) Cleaned out area (TC 02/02/2024 02:04:51 PM, Cortez, John) Found spare working Radical for Jessica (TC 02/06/2024 04:50:31 PM, Cortez, John) Boston scientific tubing trial, for endo (TC 02/05/2024 04:51:13 PM, Cortez, John) monday huddle (TC 02/06/2024 04:53:59 PM, Cortez, John) Went with Kam and Rutvik to respond to Endo Rm4 scope guide issue. Scope not appearing on display as expected. Appears that calibration is off. After looking into how UPD operates and calling Rob from Olympus, we determined how to recalibrate scope position. Also recalibrated scope position in other 3 rooms. (TC 02/05/2024 04:57:27 PM, Cortez, John) monday huddle (TC 02/06/2024 04:58:19 PM, Cortez, John) Showed Rutvik around ICU, for on-call (TC 02/05/2024 04:59:22 PM, Cortez, John) Looked for spare switches and syncs in 6th floor storage, with Kam (TC 02/05/2024 04:57:27 PM, Cortez, John) monday huddle (TC 02/07/2024 03:33:46 PM, Cortez, John) Noticed ICU room 29, 33, and 35 monitors were not connecting to central. Kamran, Kam and I looked into issue. Eventually we found out there were duplicate IPs distributed to monitors around ICU, CCU, Emerg and NICU. Worked with Kamran and Philips support to resolve monitors with no connection to central in CCU and ICU. Needed to manually assign IPs, then return setting back to to DHCP. (TC 02/08/2024 03:55:26 PM, Cortez, John) Continued to resolve IP conflicts in ICU and CCU with Kam (TC 02/09/2024 04:14:01 PM, Cortez, John) Helped Kamran run fibers in MCR for 12th floor (TC 02/08/2024 04:26:49 PM, Cortez, John) Tried to figure out why microstream wasn't completing calibration. Tried PM with different setups. Eventually was able to complete a cal successfully after connecting microstream to MX800 monitor setup. Attempted PM again with microstream connected to charging dock, and cal completed this time. Still unsure why this is the case. (TC 02/12/2024 04:32:01 PM, Cortez, John) huddle (TC 02/13/2024 09:11:25 AM, Cortez, John) Helped Jessica install video cable in OR1 (TC 02/14/2024 10:12:09 AM, Cortez, John) Philips DHCP incident debrief meeting (TC 02/13/2024 10:12:09 AM, Cortez, John) Philips DHCP incident debrief meeting (TC 02/16/2024 03:27:17 PM, Cortez, John) IS meeting (TC 02/16/2024 03:27:45 PM, Cortez, John) Created asset tags for Masimo viewstations on 15M and 16M. Confirmed SN for 12S viewstation, for Samuel (Masimo) (TC 02/16/2024 03:29:05 PM, Cortez, John) Moved Hillrom bed from 6th floor hallway to storage room (TC 02/16/2024 03:50:51 PM, Cortez, John) On-call daily check (TC 02/14/2024 03:53:03 PM, Cortez, John) ICU sent down an IV pole, saying it was loose. Its missing the screw washer securing the rod and bottom of pole. Will hold onto IV pole as my spare (TC 02/22/2024 02:08:37 PM, Cortez, John) eleaning modules (TC 02/22/2024 02:09:09 PM, Cortez, John) on-call daily check (TC 02/21/2024 02:09:39 PM, Cortez, John) on-call daily check (TC 02/20/2024 02:10:02 PM, Cortez, John) on-call daily check (TC 02/22/2024 02:38:08 PM, Cortez, John) TMS work orders (TC 02/23/2024 04:07:45 PM, Cortez, John) Helped troubleshoot LVP that was failing pressure calibration (TC 02/26/2024 11:07:50 AM, Cortez, John) BD remediation meeting (TC 02/26/2024 11:09:18 AM, Cortez, John) 10N pump remediation meeting (TC 02/26/2024 11:09:53 AM, Cortez, John) huddle (TC 02/27/2024 11:10:15 AM, Cortez, John) pm meeting (TC 02/26/2024 11:11:26 AM, Cortez, John) Ivy from IR is asking for extension cords, provided spare power bar as a short term solution. Looked into ordering a proper power bard for IR rooms (TC 02/27/2024 11:13:26 AM, Cortez, John) Delivered power bar to 506, Noticed multiple non-medical grade power bars being used in room. Will reach out to Ivy to recommend switching those out (TC 02/28/2024 11:19:19 AM, Cortez, John) Delivered training aed to 12S (TC 02/28/2024 11:19:48 AM, Cortez, John) went over medfusion 4000 software upgrades (TC 02/28/2024 01:56:16 PM, Cortez, John) ce meeting (TC 02/29/2024 04:05:19 PM, Cortez, John) Prep+work to resolve issue where 2 APs were "unregistered" on 12th floor. Just needed to reseat ethernet cable (TC 02/29/2024 04:07:31 PM, Cortez, John) Lunch and learn regarding Servo-U software upgrade with Getinge reps (TC 02/29/2024 04:07:31 PM, Cortez, John) Lunch and learn regarding Servo-U software upgrade with Getinge reps (TC 02/29/2024 04:37:09 PM, Cortez, John) Reached out to Ivy(IR) regarding power bar use in procedure rooms
1056214 | 201576 | CM | CMPLT | CMPLT | 2024-02-02 08:51 | 2024-02-02 08:56 | "missing BP cable and oxygen probe" | (TC 02/01/2024 08:50:42 AM, Cortez, John) Borrowed hose/cuff from spare Cleaned filter Installed Radical battery Added SPO2 probe Calibrated BP (off by 4-5 mmhg) All parameters tested ok
| 201577 | CM | CMPLT | CMPLT | 2024-02-02 09:03 | 2024-02-12 09:57 | Masimo accessories order, for biomed stock | (TC 02/02/2024 09:28:44 AM, Cortez, John) Placed order for 6x bp hose 5x SPO2 sensor 1x box adult cuffs (TC 02/05/2024 02:43:24 PM, Cortez, John) spo2 sensors came in (TC 02/09/2024 04:32:23 PM, Cortez, John) Hoses and cuffs arrived
1054613 | 201578 | CM | CMPLT | CMPLT | 2024-02-02 09:39 | 2024-02-02 09:42 | Checked on battery | (TC 02/02/2024 09:39:36 AM, Scero, Tomlyn) The battery was replaced last night by Nara after it failure during a power outtage. Samir asked me to check to confirm that the battery and the UPS were working fine. I checked to make sure the UPS was connected to Emergency power and notified CCU that I would be working in the closet. I unplugged the UPS and the battery was fine and the switch stayed on. I confirmed with CCU that they didn't experience any problems. Battery and UPS are working properly.
1054613 | 201579 | CM | CMPLT | CMPLT | 2024-02-02 11:24 | 2024-02-02 11:40 | Battery replacement | (TC 02/02/2024 11:24:04 AM, Baghirzade, Nara) Battery on UPS was defective; did not operate properly when power outage occurred causing the switch to go down as well. Replaced battery | Battery in UPS was faulty
1055578 | 201580 | CM | CMPLT | CMPLT | 2024-02-02 11:35 | 2024-02-02 11:35 | Batery is not holding the charge | unit shipped to Masimo. Sn.2000060564 has been replaced to Sn.2000186663 on Feb.02.2024. programmed. configured. tested bp, spo2 and temp. replaced temp probe. functional.
1055712 | 201581 | CM | CMPLT | CMPLT | 2024-02-02 11:42 | 2024-02-02 11:43 | UPS was defective, shuts off when not plugged in to AC | (TC 02/02/2024 11:41:35 AM, Baghirzade, Nara) Tested UPS; does not work when unplugged from AC and does not boot up on its own. Caused overview to remain off; no monitoring. Removed UPS and replaced with a spare one from biomed. Overview 4O (zone A) (TC 02/02/2024 11:43:03 AM, Rhuland, Courtney) Tested UPS; does not work when unplugged from AC and does not boot up on its own. Caused overview to remain off; no monitoring. Removed UPS and replaced with a spare one from biomed. Overview 4O (zone A)
64069 | 201582 | CM | CMPLT | CMPLT | 2024-02-02 11:54 | 2024-02-02 11:57 | Spare biomed UPS for NICU Overview | (TC 02/02/2024 11:54:12 AM, Baghirzade, Nara) The original UPS in NICU overview 4 failed, replaced with this UPS which is a spare from biomed. Replaced battery. Tested that it works when unplugged from AC. Tested that it comes back up on its own when plugged back in to AC. (TC 02/02/2024 11:55:55 AM, Rhuland, Courtney) The original UPS in NICU overview 4 failed, replaced with this UPS which is a spare from biomed. Replaced battery. Tested that it works when unplugged from AC. Tested that it comes back up on its own when plugged back in to AC.
1059546 | 201583 | CM | CMPLT | CMPLT | 2024-02-02 11:59 | 2024-02-02 12:00 | Exposed wires on network cable | (TC 02/02/2024 11:59:09 AM, Baghirzade, Nara) Replaced network cable as the original one had exposed wires. Checked that overview network status was connected. (TC 02/02/2024 11:59:41 AM, Rhuland, Courtney) Replaced network cable as the original one had exposed wires. Checked that overview network status was connected.
1059546 | 201584 | CM | CMPLT | CMPLT | 2024-02-02 12:00 | 2024-02-02 12:02 | Philips focal point email stating that overview is operating in disconnected mode | (TC 02/02/2024 12:00:32 PM, Baghirzade, Nara) Overview was on local mode despite having correct network configuration. Rebooted the PICIX. Network status was first pending reboot. Later it changed to connected mode. Will keep monitoring this overview. (TC 02/02/2024 12:01:26 PM, Rhuland, Courtney) Overview was on local mode despite having correct network configuration. Rebooted the PICIX. Network status was first pending reboot. Later it changed to connected mode. Will keep monitoring this overview.
1049217 | 201585 | CM | CMPLT | CMPLT | 2024-02-02 13:14 | 2024-02-02 13:21 | "recorder not working" | (TC 02/02/2024 01:13:46 PM, Cortez, John) Confirmed issue. Paper installed correctly. Restarting doesn't resolve issue. Grabbed spare printer and usb cable+power supply, but that printer also doesnt work after restart. Eventually noticed a strip of paper lodged inside. Removing seems to resolve problem. Coincidentally, original printer also had a loose strip of paper inside. Left original printer.
| 201586 | CM | CMPLT | CMPLT | 2024-02-02 13:46 | 2024-02-02 13:46 | worked with rt | worked with RT
| 201587 | CM | CMPLT | CMPLT | 2024-02-02 13:47 | 2024-02-02 13:47 | daily checks | daily checks
500230 | 201588 | RQ | CMPLT | CMPLT | 2024-02-02 14:00 | 2024-02-12 11:43 | THe cobra welll counter is not powering up | (TC 02/02/2024 02:38:31 PM, Green, Paul) Power problems in Toronto last night. Suspect UPS went on battery power and did not come back up. Went to site and inspected. UPS batteries are defective and notified Barry to order 3 replacement batteries. | The well counter in hot lab is not powering up. Came in this morning and saw it was off. Tried turning the on and off switch with no luck. The battery pack next to it seems to be off. Not sure how to turn on,
| 201589 | CM | CMPLT | CMPLT | 2024-02-02 14:22 | 2024-02-13 14:17 | Inventory ENT | (TC 02/02/2024 02:21:55 PM, Rhuland, Courtney) Took inventory of all untagged equipment in both clinc A and B (TC 02/13/2024 02:17:06 PM, Rhuland, Courtney) Created assets for untagged devices
68745 | 201590 | CM | CMPLT | NFF | 2024-02-02 14:45 | 2024-02-02 14:46 | Not charging | (TC 02/02/2024 02:44:50 PM, Rhuland, Courtney) Charged device overnight Battery fully charged and shows 10 hours of battery life Fucntional check - OK No fault found
| 201591 | CM | CMPLT | CMPLT | 2024-02-02 15:05 | 2024-04-23 09:49 | Welch Allyn 6000 battery order, for biomed stock | (TC 02/02/2024 03:43:51 PM, Cortez, John) Placed order for 5 batteries (TC 02/12/2024 04:31:13 PM, Cortez, John) Reached out to Fan for update, only 1 item in stock. BBM will ship one for now. (TC 02/14/2024 11:12:17 AM, Cortez, John) 1 battery arrived, the rest BO
66531 | 201592 | CM | CMPLT | CMPLT | 2024-02-02 15:32 | 2024-02-02 15:35 | Email from Kim, fetal monitor not connecting to Fetalink | (TC 02/02/2024 03:33:13 PM, Rendon Lopez , Elizabeth) Email from Kim to look into Room3 since IT wasn't able to resolve connection issue. Checked room, unplugged connections, restarted CE, connection was sucessful. | Issue placed with IT, reassigned to Biomed multiple times, coordinator emailed.
65431 | 201593 | CM | CMPLT | CMPLT | 2024-02-02 15:47 | 2024-02-05 10:15 | "missing cord, battery not charging" | (TC 02/02/2024 03:47:13 PM, Cortez, John) Confirmed issue, Replaced battery, Removed rubber feet, Added power cord, Charging overnight (TC 02/05/2024 10:14:05 AM, Cortez, John) Secured power cord Added Spo2 sensor All parameters tested ok (TC 02/05/2024 10:15:07 AM, Cortez, John) Changed time, Tightened screws on rollstand
67933 | 201594 | CM | CMPLT | CMPLT | 2024-02-05 07:18 | 2024-02-05 08:11 | Not turning on | (TC 02/05/2024 08:10:06 AM, Rhuland, Courtney) Main power switch on the back was turned off Calibrated humidity Functional check - OK
| 201595 | RDS | CMPLT | CMPLT | 2024-02-05 08:04 | 2024-02-05 08:04 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
| 201596 | RDS | CMPLT | CMPLT | 2024-02-05 08:27 | 2024-02-12 10:11 | Weekly checks in CCU | (TC 02/05/2024 10:09:51 AM, Scero, Tomlyn) weekly checks complete. 30J tests passed. (TC 02/12/2024 10:10:26 AM, Scero, Tomlyn) weekly checks complete. 30J tests passed.
67921 | 201597 | CM | CMPLT | CMPLT | 2024-02-05 09:13 | 2024-02-05 09:14 | Broken latch | (TC 02/05/2024 09:13:54 AM, Rhuland, Courtney) Replaced broken latch Functional check - OK
| 201598 | CM | ACTIV | ISSUE | 2024-02-05 09:22 |  | Sterile water orders (for ICU Altrix) | (TC 02/05/2024 09:24:30 AM, Cortez, John) Ordered 36x bottles (TC 02/20/2024 10:19:17 AM, Cortez, John) ordered 36x bottles (TC 04/22/2024 10:49:13 AM, Cortez, John) ordered 36x bottles (TC 05/21/2024 03:20:54 PM, Cortez, John) ordered 36x bottles (TC 06/25/2024 10:29:06 AM, Cortez, John) ordered x36 bottles (TC 08/29/2024 09:39:50 AM, Cortez, John) ordered x36 (TC 09/10/2024 01:21:17 PM, Cortez, John) previous order cancelled, ordered x3 cases
68552 | 201599 | CM | CMPLT | CMPLT | 2024-02-05 10:12 | 2024-02-06 14:46 | O2 sensor error | (TC 02/06/2024 02:42:29 PM, Scero, Tomlyn) Confirmed O2 sensor not working. Downloaded DSR. Replaced PATO sensor and ordered a new one to have on hand. Self test complete and passed. Simulation compelted as well with no issues. O2 sensor working fine now.
1064169 | 201600 | CM | CMPLT | CMPLT | 2024-02-05 10:42 | 2024-02-05 10:42 | new incoming inspection
1061900 | 201601 | CM | CMPLT | CMPLT | 2024-02-05 10:47 | 2024-02-05 10:47 | not working | reset. tested. functional. fixed aux. power plug
| 201602 | CM | CMPLT | CMPLT | 2024-02-05 10:48 | 2024-02-05 10:48 | daily check | daily check
1048551 | 201603 | CM | CMPLT | CMPLT | 2024-02-05 11:03 | 2024-02-05 11:03 | not working | (TC 11/01/2023 11:03:37 AM, Huang, Jessica) got the call from OR and picked the device ordered the new cuff and replaced it tested it. working fine sent it back to OR
65698 | 201604 | CM | CMPLT | CMPLT | 2024-02-05 11:05 | 2024-02-05 11:05 | broken hose | (TC 11/01/2023 11:05:28 AM, Huang, Jessica) contacted the vendor to get the quote and ordered the part replaced the hose and tested it sent it back to OR
66008 | 201605 | CM | CMPLT | CMPLT | 2024-02-05 12:38 | 2024-02-06 13:22 | Returned from SickKids | (TC 02/05/2024 12:37:56 PM, Rhuland, Courtney) Replaced broken front and rear cases Replaced & conditioned battery (TC 02/06/2024 01:14:36 PM, Rhuland, Courtney) Check-in complete All tests passed Printed new asset tag
66050 | 201606 | CM | CMPLT | CMPLT | 2024-02-05 12:55 | 2024-02-06 13:10 | Returned from SickKids | (TC 02/05/2024 12:55:27 PM, Rhuland, Courtney) Replaced broken rear case Replaced power cord retainer Replaced and conditioned battery (TC 02/06/2024 01:09:17 PM, Rhuland, Courtney) check in complete All tests passed Printed new asset tag
| 201607 | CM | CMPLT | CMPLT | 2024-02-05 12:57 | 2024-02-05 12:57 | suction tree not working | (TC 11/02/2023 12:57:17 PM, Huang, Jessica) picked up the broken suction tree from OR ordered the new regulator and replaced it tested it, working fine sent it back to OR
| 201608 | CM | CMPLT | CMPLT | 2024-02-05 13:51 | 2024-02-05 13:51 | biomed meeting | biomed meeting
| 201609 | PA | CMPLT | CMPLT | 2024-02-05 13:54 | 2024-02-12 14:16 | Repair parts for Zeus | (TC 02/05/2024 02:00:16 PM, Scero, Tomlyn) Ordered parts on ereq (TC 02/12/2024 02:15:04 PM, Scero, Tomlyn) Parts received and taken to RT workshop
64069 | 201610 | CM | CMPLT | CMPLT | 2024-02-05 14:41 | 2024-02-05 14:41 | Removing board | (TC 02/05/2024 02:40:43 PM, Rhuland, Courtney) Removed communication board from device as requested by Kamran
| 201611 | CM | CMPLT | CMPLT | 2024-02-05 15:05 | 2024-02-05 15:26 | Meeting | (TC 02/05/2024 03:04:58 PM, Patel, Rutvik) Staff Meeting
| 201612 | MEP | CMPLT | CMPLT | 2024-02-05 15:49 | 2024-02-05 15:49 | New Hire Orientation | Attended New Hire Orientation
1062338 | 201613 | CM | CMPLT | CMPLT | 2024-02-05 15:51 | 2024-02-05 15:53 | E5.0 on screen | (TC 02/05/2024 03:52:29 PM, Rendon Lopez , Elizabeth) Confirmed issue, replaced temperature probe. Tested-OK.
1055954 | 201614 | CM | CMPLT | CMPLT | 2024-02-06 08:04 | 2024-02-06 08:04 | battery is not holding the charge. | unit shipped to Masimo. Sn.2000060488 has been replaced to Sn.2000185479 on Feb.06.2024. programmed. configured. tested bp, spo2 and temp. replaced temp probe. functional.
| 201615 | CM | CMPLT | CMPLT | 2024-02-06 09:54 | 2024-02-08 13:59 | OR 11 - Service call | (TC 02/06/2024 09:53:54 AM, Scero, Tomlyn) Went to OR 11 to ask what was happening. Nurse stated that the Davinci monitor was working fine, but the Olympus monitors on the ceiling were turning off and then back on every once and awhile. they asked if someone could come back between cases to take a look at them. I called Jessica at Bridgepoint to ask if this is something she's seen before. She said it could be the ESU causing interference with the monitors. She suggested they move the ESU further away and make sure it's not plugged in too close to the monitors as well. (TC 02/08/2024 01:57:35 PM, Scero, Tomlyn) Kam and I returned to the OR to check but they were still in a case. We could hear the ESU being used, we observed the monitors and didnt see any changes. OR11 didn't end up calling us back either. Notified Jessica of the issue. | OR ceiling monitors turning off and then back on
1069951 | 201616 | CM | CMPLT | CMPLT | 2024-02-06 10:02 | 2024-02-05 10:02 | incoming inspection and installation. | (TC 02/06/2024 10:02:34 AM, Chukri, Samir) Incoming Inspection and istalled work with Qwyynn - delivered to SH biomed by Perry . (TC 02/05/2024 10:04:58 AM, Ferreira, Qwynn) installed the loaner MRI ventilator.
| 201617 | CM | CMPLT | CMPLT | 2024-02-06 10:14 | 2024-02-06 10:14 | worked with NICU RT | worked with NICU RT
| 201618 | RDS | CMPLT | CMPLT | 2024-02-06 11:16 | 2024-02-07 07:48 | Weekly check for ED
| 201619 | RDS | CMPLT | CMPLT | 2024-02-06 11:16 | 2024-02-26 14:20 | Weekly check for ICU
1065115 | 201620 | CM | CMPLT | CMPLT | 2024-02-06 12:40 | 2024-02-06 12:42 | UroCap not connecting | (TC 02/06/2024 12:40:00 PM, Rhuland, Courtney) Troubleshooted device to connect UroCap Connected device with USB Bluetooth connectivity not working Talked to manager who said they only use device connected Connected both UroCap devices to both Aquarius
| 201621 | CM | CMPLT | CMPLT | 2024-02-06 12:44 | 2024-02-06 12:45 | Exam light not turning on | (TC 02/06/2024 12:43:59 PM, Rhuland, Courtney) Reporter said the bulb was burnt out Replaced bulb but device still did not turn on Power cord was not seated properly Reseated power cord and put old bulb back in Device working as expected
1059533 | 201622 | CM | CMPLT | CMPLT | 2024-02-06 13:18 | 2024-02-06 13:20 | PICiX kept restarting and needed rebuild | (TC 02/06/2024 01:18:21 PM, Canjimavadivel, Kamallesh ) PICiX was rebuilt after usb was removed. Patched to latest OS. Took PC back to PACU and reinstalled as PACU3S. Brought spare back to biomed shop.
1052978 | 201623 | CM | CMPLT | CMPLT | 2024-02-06 13:21 | 2024-02-06 13:29 | Not sending to power chart | (TC 02/06/2024 01:21:28 PM, Canjimavadivel, Kamallesh ) Verified ip address was correct and tested connection from VSM. Sent saved reading and send was successful. Pulled logs and informed staff monitor was ok to use and report if issue reoccurs.
1067975 | 201624 | CM | CMPLT | CMPLT | 2024-02-06 13:32 | 2024-02-14 08:41 | MSH ER 3rd (Main) Floor CONTROL PANEL ISSUES INSTALL PALM CONTROL TRACKBALL | (TC 02/06/2024 02:07:43 PM, Whittington, Scott) MSH ER 3rd (Main) Floor INSTALL PALM CONTROL TRACKBALL (TC 02/10/2024 01:11:10 PM, Whittington, Scott) MSH ER 3rd (Main) Floor INSTALL PALM CONTROL TRACKBALL CONTROL PANEL ISSUES, REPAIR AND INSTALL UPGRADE TO PALMCONTROL. WIRELESS ACCES SETUP, AWAITING USB WIRELESS DONGLE (TC 02/10/2024 01:11:10 PM, Whittington, Scott) MSH ER 3rd (Main) Floor INSTALL PALM CONTROL TRACKBALL CONTROL PANEL ISSUES, REPAIR AND INSTALL UPGRADE TO PALMCONTROL. WIRELESS ACCES SETUP, AWAITING USB WIRELESS DONGLE | MSH ER 3rd (Main) Floor CONTROL PANEL ISSUES INSTALL PALM CONTROL TRACKBALL MSH ER 3rd (Main) Floor INSTALL PALM CONTROL TRACKBALL
1068221 | 201625 | CM | CMPLT | CMPLT | 2024-02-06 13:33 | 2024-02-08 15:32 | MSH US PRISM ROOM 6 INSTALL WIRELESS | (TC 02/06/2024 01:33:33 PM, Whittington, Scott) MSH US PRISM ROOM 6 INSTALL WIRELESS (TC 02/08/2024 03:30:06 PM, Whittington, Scott) MSH US PRISM ROOM 6 INSTALL WIRELESS SETUP WIRELESS ABILTY WAIT FOR PW INSTALL MAC: C8:9E:43:A0:53:F2 | MSH US PRISM ROOM 6 INSTALL WIRELESS MSH US PRISM ROOM 6 INSTALL WIRELESS SETUP WIRELESS ABILTY WAIT FOR PW INSTALL MAC: C8:9E:43:A0:53:F2
1067437 | 201626 | FC | CMPLT | CMPLT | 2024-02-06 13:41 | 2024-02-06 13:41 | QC Checklist | (TC 02/06/2024 01:40:43 PM, Rhuland, Courtney) QC checklist complete Returned to 11S
1069406 | 201627 | CM | CMPLT | CMPLT | 2024-02-06 13:57 | 2024-02-06 13:57 | new incoming inspection
1069337 | 201628 | CM | CMPLT | CMPLT | 2024-02-06 13:58 | 2024-02-06 13:58 | new incoming inspection
1069407 | 201629 | CM | CMPLT | CMPLT | 2024-02-06 14:00 | 2024-02-06 14:00 | new incoming inspection
1069408 | 201630 | CM | CMPLT | CMPLT | 2024-02-06 14:03 | 2024-02-06 14:03 | new incoming inspection
1069409 | 201631 | CM | CMPLT | CMPLT | 2024-02-06 14:03 | 2024-02-06 14:03 | new incoming inspection
1069410 | 201632 | CM | CMPLT | CMPLT | 2024-02-06 14:04 | 2024-02-06 14:04 | new incoming inspection
1069411 | 201633 | CM | CMPLT | CMPLT | 2024-02-06 14:05 | 2024-02-06 14:05 | new incoming inspection
64085 | 201634 | CM | CMPLT | CMPLT | 2024-02-06 14:10 | 2024-02-07 07:48 | Replaced battery | (TC 02/06/2024 02:10:05 PM, Canjimavadivel, Kamallesh ) Batteries were no longer charging after power shutdown, had not been replaced in over 2 years. Replaced batteries
531113 | 201635 | CM | CMPLT | CMPLT | 2024-02-06 14:29 | 2024-02-06 14:31 | Buttons are not working | (TC 02/06/2024 02:29:16 PM, Rhuland, Courtney) Device turns on but none of the buttons are functional Brought it back to biomed to troubleshoot Reseated membrane cable and all buttons worked except the cycle button Looks like the connection on the membrane switch is damaged at the cycle pin No replacement parts are available and device will be retired
1046203 | 201636 | CM | CMPLT | CMPLT | 2024-02-06 14:45 | 2024-02-06 14:45 | Cannot save exam | (TC 02/06/2024 02:45:13 PM, Fan, Tak) No power on system. Outlet in reading room has no power. Apparently the system died earlier in the reading room after the system was brought back from a TEE. System booted up with saved images. Exam not associated with patient. Fellow placed a new order for TEE. Associated the exam with order.
| 201637 | CM | CMPLT | CMPLT | 2024-02-06 15:09 | 2024-02-06 15:11 | Scope guide not working | (TC 02/06/2024 03:08:38 PM, Canjimavadivel, Kamallesh ) Went up to quickly troubleshoot in between cases, not able to resolve. WAs told it is not urgent anbd can be looked at after cases are done for the day. Looked at manual for guide to understand how it works
| 201638 | CM | CMPLT | CMPLT | 2024-02-07 07:45 | 2024-02-07 07:53 | Rounds - NICU | (TC 02/07/2024 07:52:33 AM, Rhuland, Courtney) NICU rounds complete
| 201639 | CM | CMPLT | CMPLT | 2024-02-07 07:53 | 2024-02-07 08:00 | Rounds - RESUS | (TC 02/07/2024 07:59:47 AM, Rhuland, Courtney) RESUS rounds complete
1060001 | 201640 | CM | CMPLT | CMPLT | 2024-02-07 09:32 | 2024-02-07 09:50 | Not reading | (TC 02/07/2024 09:36:34 AM, Canjimavadivel, Kamallesh ) V6 not picking up correctly. Replaced chest lead set
66500 | 201641 | CM | CMPLT | CMPLT | 2024-02-07 10:11 | 2024-02-07 10:12 | Recorder not printing | (TC 02/06/2024 10:11:03 AM, Rendon Lopez , Elizabeth) Checking recorder in room, Cleaned, run self test, NFF. Ran on user mode, working no issues
1069462 | 201642 | CM | CMPLT | CMPLT | 2024-02-07 10:22 | 2024-02-07 10:38 | Loaner removed from service, to be returned to Olympus | (TC 02/07/2024 10:36:08 AM, Cortez, John) Inspected tubing for signs of fluid, looked ok Shipped to olympus
| 201643 | CM | CMPLT | CMPLT | 2024-02-07 10:25 | 2024-02-07 10:25 | daily checks | daily checks
66126 | 201644 | CM | CMPLT | CMPLT | 2024-02-07 10:26 | 2024-02-07 10:26 | holder is broken | replaced broken holder. tested. functional.
1066216 | 201645 | CM | CMPLT | CMPLT | 2024-02-07 10:27 | 2024-02-07 10:27 | alarming | tested. functional.
66454 | 201646 | CM | CMPLT | CMPLT | 2024-02-07 10:28 | 2024-02-07 10:28 | broken | tested. functional. Humidifier holder was broken. replaced.
| 201647 | CM | CMPLT | CMPLT | 2024-02-07 10:42 | 2024-02-15 13:01 | Formlabs Order
1062173 | 201648 | CM | CMPLT | CMPLT | 2024-02-07 11:15 | 2024-02-07 11:22 | "unable to adjust bed" | (TC 02/07/2024 11:15:24 AM, Cortez, John) Turned on/off battery switch Seems to work normally, all control functions tested ok. Probably just needed a full power cycle
| 201649 | CM | CMPLT | CMPLT | 2024-02-07 14:01 | 2024-02-07 14:01 | meeting | meeting
| 201650 | CM | CMPLT | CMPLT | 2024-02-07 14:02 | 2024-02-07 14:02 | worked with Elizabeth | worked with Elizabeth
1069959 | 201651 | CM | CMPLT | CMPLT | 2024-02-07 14:18 | 2024-02-08 07:31 | Incoming inspection | (TC 02/07/2024 02:29:00 PM, Rhuland, Courtney) Asset tagged and functional check - OK
1069960 | 201652 | CM | CMPLT | CMPLT | 2024-02-07 14:18 | 2024-02-08 07:32 | Incoming inspection | (TC 02/07/2024 02:30:00 PM, Rhuland, Courtney) Asset tagged and functional check - OK
1069961 | 201653 | CM | CMPLT | CMPLT | 2024-02-07 14:18 | 2024-02-08 07:32 | Incoming inspection | (TC 02/07/2024 02:31:00 PM, Rhuland, Courtney) Asset tagged and functional check - OK
1069962 | 201654 | CM | CMPLT | CMPLT | 2024-02-07 14:18 | 2024-02-08 07:33 | Incoming inspection | (TC 02/07/2024 02:32:00 PM, Rhuland, Courtney) Asset tagged and functional check - OK
1069963 | 201655 | CM | CMPLT | CMPLT | 2024-02-07 14:18 | 2024-02-08 07:34 | Incoming inspection | (TC 02/07/2024 02:33:00 PM, Rhuland, Courtney) Asset tagged and functional check - OK
1069964 | 201656 | CM | CMPLT | CMPLT | 2024-02-07 14:18 | 2024-02-08 07:34 | Incoming inspection | (TC 02/07/2024 02:33:00 PM, Rhuland, Courtney) Asset tagged and functional check - OK
1069965 | 201657 | CM | CMPLT | CMPLT | 2024-02-07 14:18 | 2024-02-08 07:35 | Incoming inspection | (TC 02/07/2024 02:34:00 PM, Rhuland, Courtney) Asset tagged and functional check - OK
| 201658 | CM | CMPLT | CMPLT | 2024-02-07 14:44 | 2024-02-07 14:47 | Department Tour | Feb 6 - Department Tour - NICU, ICU, Endoscopy, GI Feb 7 - Department Tour - OR
| 201659 | CM | CMPLT | CMPLT | 2024-02-07 14:50 | 2024-02-07 14:52 | Meeting | Meeting with Adeel for TAAAC and CMMS
1069411 | 201660 | CM | CANCL | CANCL | 2024-02-07 14:54 | 2024-02-07 14:54 | new incoming inspection
1069413 | 201661 | CM | CMPLT | CMPLT | 2024-02-07 14:55 | 2024-02-07 14:55 | new incoming inspection
1069414 | 201662 | CM | CMPLT | CMPLT | 2024-02-07 14:55 | 2024-02-07 14:55 | new incoming inspection
1069412 | 201663 | CM | CMPLT | CMPLT | 2024-02-07 14:56 | 2024-02-07 14:56 | new incoming inspection
1069415 | 201664 | CM | CMPLT | CMPLT | 2024-02-07 14:57 | 2024-02-07 14:57 | new incoming inspection
1069417 | 201665 | CM | CMPLT | CMPLT | 2024-02-07 14:58 | 2024-02-07 14:58 | new incoming inspection
1069416 | 201666 | CM | CMPLT | CMPLT | 2024-02-07 14:59 | 2024-02-07 14:59 | new incoming inspection
1069418 | 201667 | CM | CMPLT | CMPLT | 2024-02-07 14:59 | 2024-02-07 14:59 | new incoming inspection
1069419 | 201668 | CM | CMPLT | CMPLT | 2024-02-07 15:00 | 2024-02-07 15:00 | new incoming inspection
1069420 | 201669 | CM | CMPLT | CMPLT | 2024-02-07 15:01 | 2024-02-07 15:01 | new incoming inspection
| 201670 | CM | CMPLT | CMPLT | 2024-02-07 15:02 | 2024-02-06 15:03 | Meeting | Irrigation/scope tubing demo Meeting with Boston Sci Rep, Endo manager and others
1069421 | 201671 | CM | CMPLT | CMPLT | 2024-02-07 15:03 | 2024-02-07 15:03 | new incoming inspection
1069945 | 201672 | CM | CMPLT | CMPLT | 2024-02-07 15:05 | 2024-02-07 15:04 | new incoming inspection
1069946 | 201673 | CM | CMPLT | CMPLT | 2024-02-07 15:06 | 2024-02-07 15:06 | new incoming inspection
1069482 | 201674 | CM | CMPLT | CMPLT | 2024-02-07 15:14 | 2024-02-07 15:14 | new incoming inspection
1069483 | 201675 | CM | CMPLT | CMPLT | 2024-02-07 15:17 | 2024-02-07 15:17 | new incoming inspection
1050572 | 201676 | CM | CMPLT | CMPLT | 2024-02-07 15:20 | 2024-02-06 15:25 | Scope not detecting | Call re: not wotking. Went to check in between cases. Could not determine issue and was asked to come back after the case. Went again later. Was informed to come back again. Scheduled repair by EOD. Found unit was not detecting scope at the right plate position and was out of calibration. Contacted Olympus rep and worked with him to calibrate the device. Issue resolved. Returned to service.
1069484 | 201677 | CM | CMPLT | CMPLT | 2024-02-07 15:24 | 2024-02-07 15:24 | new incoming inspection
1069485 | 201678 | CM | CMPLT | CMPLT | 2024-02-07 15:25 | 2024-02-07 15:25 | new incoming inspection
1069486 | 201679 | CM | CMPLT | CMPLT | 2024-02-07 15:27 | 2024-02-07 15:27 | new incoming inspection
1069940 | 201680 | CM | CMPLT | CMPLT | 2024-02-07 15:28 | 2024-02-07 15:28 | new incoming inspection
1069941 | 201681 | CM | CMPLT | CMPLT | 2024-02-07 15:29 | 2024-02-07 15:29 | new incoming inspection
1069944 | 201682 | CM | CMPLT | CMPLT | 2024-02-07 15:32 | 2024-02-07 15:32 | new incoming inspection
| 201683 | CM | CMPLT | CMPLT | 2024-02-07 15:42 | 2024-02-07 15:42 | Phone pick up and set up | Phone pick up from IT and set up.
| 201684 | CM | CMPLT | CMPLT | 2024-02-07 15:43 | 2024-02-07 15:43 | hug me vac not working | (TC 11/01/2023 03:42:41 PM, Huang, Jessica) picked up the defective hug me vac from the OR desk and confirmed it cant be repaired. disposed the item and informed the OR team
1067480 | 201685 | CM | CMPLT | CMPLT | 2024-02-07 15:55 | 2024-02-07 15:55 | error message internal hardware failure | (TC 11/07/2023 03:55:02 PM, Huang, Jessica) contacted the tech support disconnected the battery and reconnected it working fine
1061318 | 201686 | CM | CMPLT | CMPLT | 2024-02-07 15:57 | 2024-02-07 15:57 | not working | (TC 11/07/2023 03:57:00 PM, Huang, Jessica) replaced the battery and charged it tested it, working fine
541271 | 201687 | CM | CMPLT | CMPLT | 2024-02-07 15:59 | 2024-02-07 15:59 | not working | (TC 11/07/2023 03:58:55 PM, Huang, Jessica) replaced the battery, BP cuff and the temperature probe
1049059 | 201688 | CM | CMPLT | CMPLT | 2024-02-07 16:04 | 2024-02-07 16:04 | broken | (TC 11/23/2023 04:03:45 PM, Huang, Jessica) picked up the device from the OR desk and requested the replacement device from the vendor sent the replacement unit to OR
541284 | 201689 | CM | CMPLT | CMPLT | 2024-02-07 16:36 | 2024-02-07 16:36 | mutiple issues | (TC 12/05/2023 04:36:09 PM, Huang, Jessica) the device was sent down because of defective power adapter requested the part number and quote from the vendor replaced it and noticed the power button was not working properly requested the part number and quote from the vendor and replaced it passed calibration and informed PCM
541303 | 201690 | CM | CMPLT | CMPLT | 2024-02-07 16:42 | 2024-02-07 16:42 | removed from service
541174 | 201691 | CM | CMPLT | CMPLT | 2024-02-07 16:43 | 2024-02-07 16:43 | removed from service
1066985 | 201692 | CM | CMPLT | CMPLT | 2024-02-07 16:46 | 2024-02-07 16:46 | unable to establish network communications. radio out of network range | (TC 12/19/2023 04:45:45 PM, Huang, Jessica) contacted the tech support regarding the error message disconnected the battery and reconnected it didnt fix the issue shipped the device for warranty repair
120015 | 201693 | CM | CMPLT | CMPLT | 2024-02-07 16:54 | 2024-02-07 16:54 | not working | (TC 12/05/2023 04:54:06 PM, Huang, Jessica) replaced the temperature probe and the battery
541029 | 201694 | CM | CMPLT | CMPLT | 2024-02-07 16:55 | 2024-02-07 16:55 | not working | (TC 12/05/2023 04:55:15 PM, Huang, Jessica) replaced the battery and temperature probe
51730 | 201695 | CM | CMPLT | CMPLT | 2024-02-07 16:57 | 2024-02-07 16:57 | removed from service | (TC 12/07/2023 04:57:10 PM, Huang, Jessica) the unit is not repairable removed from service
541051 | 201696 | CM | CMPLT | CMPLT | 2024-02-07 16:59 | 2024-02-07 16:59 | removed from service
541022 | 201697 | CM | CMPLT | CMPLT | 2024-02-07 17:00 | 2024-02-07 17:00 | removed from service
1059618 | 201698 | CM | CMPLT | CMPLT | 2024-02-07 17:03 | 2024-02-07 17:03 | not working | (TC 12/13/2023 05:03:07 PM, Huang, Jessica) received the call in the OR room and went there checked all the cables and everything was fine brought down to Biomed and just turned the knob on the side working fine brought back to OR
1059944 | 201699 | CM | CMPLT | CMPLT | 2024-02-07 17:06 | 2024-02-07 17:06 | light not working | (TC 12/14/2023 05:05:55 PM, Huang, Jessica) the user complained the light was not working checked it working fine and informed Tanya
1046635 | 201700 | CM | CMPLT | CMPLT | 2024-02-07 17:08 | 2024-02-07 17:08 | damaged cord on the foot pedal | (TC 12/14/2023 05:08:21 PM, Huang, Jessica) replaced the foot pedal
| 201701 | CM | CMPLT | CMPLT | 2024-02-07 17:10 | 2024-02-07 17:10 | defective ophthalmoscope from MBU
50362 | 201702 | CM | CMPLT | CMPLT | 2024-02-07 17:12 | 2024-02-07 17:12 | not working | (TC 12/18/2023 05:11:41 PM, Huang, Jessica) the scale is out of support and cant be repaired removed from service and informed the users
| 201703 | CM | CMPLT | CMPLT | 2024-02-07 17:14 | 2024-02-07 17:14 | suction tree not working | (TC 12/20/2023 05:14:30 PM, Huang, Jessica) picked the suction tree from the OR desk and confirmed the regulator was not working replaced the regulator and sent it back to OR
1047489 | 201704 | CM | CMPLT | CMPLT | 2024-02-07 17:19 | 2024-02-07 17:19 | not working | (TC 12/20/2023 05:18:50 PM, Huang, Jessica) tested it on the Biomed spare monitor working fine and sent it back to the unit
1050242 | 201705 | CM | CMPLT | CMPLT | 2024-02-07 17:20 | 2024-02-07 17:20 | not working | (TC 12/20/2023 05:20:11 PM, Huang, Jessica) tested it on the Biomed spare monitor working fine and sent it back to the unit
66040 | 201706 | CM | CMPLT | CMPLT | 2024-02-08 07:51 | 2024-02-08 07:52 | Not connecting to Wifi | (TC 02/08/2024 07:50:53 AM, Rhuland, Courtney) Powered on PCU and wifi connected Verified connectivity using ASM software No fault found Functional check - OK
1064418 | 201707 | CM | CMPLT | CMPLT | 2024-02-08 08:48 | 2024-04-03 11:05 | Broken Rear flap cable port . | (TC 02/08/2024 08:48:14 AM, Chukri, Samir) went with Qwynn, checked the incubator and Qwynn will order the Rear flap assembly. (TC 02/08/2024 08:58:51 AM, Ferreira, Qwynn) Went with Samir to the MRI incubator, rear flap was broken - going to order new parts (TC 04/03/2024 11:13:35 AM, Ferreira, Qwynn) part arrived - replaced the part
501888 | 201708 | CM | CMPLT | CMPLT | 2024-02-08 09:14 | 2024-02-08 09:14 | Trackball not tracking well | (TC 02/07/2024 09:13:50 AM, Fan, Tak) Cleaned out ball and roller assembly.
1050876 | 201709 | CM | CMPLT | CMPLT | 2024-02-08 09:28 | 2024-02-08 09:29 | Battery not holding charge | (TC 02/08/2024 09:28:08 AM, Rhuland, Courtney) Battery status was not showing on the screen Replaced battery with spare and battery status showed Functional check - OK Charged battery
| 201710 | CM | CMPLT | CMPLT | 2024-02-08 11:16 | 2024-02-08 11:16 | daily check | daily check
| 201711 | CM | CMPLT | CMPLT | 2024-02-08 11:18 | 2024-02-08 11:18 | meeting | meeting
1055579 | 201712 | CM | CMPLT | CMPLT | 2024-02-08 11:21 | 2024-02-08 11:21 | battery not charging | installed battery in Radical 7. Charged unit. Testing battery life.
1055580 | 201713 | CM | CMPLT | CMPLT | 2024-02-08 11:23 | 2024-02-08 11:23 | battery not charging | installed battery in Radical 7. Charged unit. Testing battery life.
| 201714 | CM | CMPLT | CMPLT | 2024-02-08 12:52 | 2024-02-26 08:28 | Freight Charge for return of the 4 surfaces mattress | (TC 02/08/2024 12:57:36 PM, Chukri, Samir) ordered a return PO (TC 02/26/2024 08:25:17 AM, Chukri, Samir) received and emailed the PO copy to Aleksandar from Stryker
| 201715 | CM | CMPLT | CMPLT | 2024-02-08 13:48 | 2024-02-08 13:56 | Setting up IT closet 12N | (TC 02/08/2024 01:53:08 PM, Rhuland, Courtney) Aided Kamran in installing switches in the IT closet on 12N
1063000 | 201716 | PM | CMPLT | CMPLT | 2024-02-08 14:32 | 2024-02-08 14:37 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/07/2024 02:36:17 PM, Cortez, John) Completed cleaning cycle
| 201717 | CM | CMPLT | CMPLT | 2024-02-08 14:45 | 2024-02-08 14:50 | CCU requesting spare telepack, Telepack 7 is MIA | (TC 02/08/2024 02:45:27 PM, Cortez, John) Put spare telepack (BIO1) into CCU pool and removed telepack 7. Tested connection to PICIX in CCU BIO1 returned next day, Tomi put back tele7 into pool
| 201718 | CM | CMPLT | CMPLT | 2024-02-08 15:03 | 2024-02-08 15:33 | "PACU bay D28 patient showing up in wrong bed for powerchart" | (TC 02/07/2024 03:03:01 PM, Cortez, John) Went to PACU to verify it's not a Philips issue, and if it's not related to DHCP issues that were currently occuring. Nurse said it was working after they closed/opened powerchart application. Told charge nurse to let us know if the issue comes back
501067 | 201719 | CM | CMPLT | CMPLT | 2024-02-09 07:28 | 2024-02-09 07:29 | Trackball not moving up and down | (TC 02/08/2024 11:28:00 AM, Fan, Tak) Confirmed problem. Resistance coming from the 45 degree small bearing. Removed trackball assembly, cleaned the bearings, reassembled. Problem resolved.
501584 | 201720 | CM | CMPLT | CMPLT | 2024-02-09 07:31 | 2024-02-09 07:32 | No ext video | (TC 02/08/2024 10:31:00 AM, Fan, Tak) Not connected in back.
501066 | 201721 | CM | CMPLT | CMPLT | 2024-02-09 07:33 | 2024-02-09 07:34 | Moving system | (TC 02/08/2024 12:33:00 PM, Fan, Tak) System was swapped with spare for PMs. Swapped back.
68259 | 201722 | CM | CMPLT | CMPLT | 2024-02-09 08:03 | 2024-02-15 08:39 | Msmt Ext Malfunction | (TC 02/08/2024 08:30:00 AM, Canjimavadivel, Kamallesh ) MSL port on FMS rack was corroded. Needed to swap out rack. Found spare in biomed shop and checked compatibility. Swapped out with defective rack. (TC 02/12/2024 08:16:10 AM, Canjimavadivel, Kamallesh ) Called vendor for quote. Placed order for MSL connecter board (TC 02/15/2024 08:37:32 AM, Canjimavadivel, Kamallesh ) Replaced MSL downboard, tested ok
| 201723 | TEDU | CMPLT | CMPLT | 2024-02-09 08:14 | 2024-02-09 08:19 | Draeger V500/VN500 Training | (TC 02/05/2024 04:00:00 PM, Baghirzade, Nara) Travel to Philadelphia/Sellersville (TC 02/06/2024 04:30:00 PM, Baghirzade, Nara) Draeger V500/VN500 Training (TC 02/07/2024 04:30:00 PM, Baghirzade, Nara) Draeger V500/VN500 Training (TC 02/08/2024 03:00:00 PM, Baghirzade, Nara) Travel to Toronto from Philadelphia/Sellersville
1067992 | 201724 | CM | CMPLT | CMPLT | 2024-02-09 08:28 | 2024-02-26 13:09 | MSH US RM 8 CANON PRISM SERVICE LIC & REQ PALM CONTROLLER TRKBL | (TC 02/09/2024 08:27:51 AM, Whittington, Scott) MSH US RM 8 CANON PRISM SERVICE LIC & REQ PALM CONTROLLER TRKBL (TC 02/12/2024 01:18:16 PM, Whittington, Scott) MSH US RM 8 CANON PRISM SERVICE LIC 7 REQ PALM CONTROLLER TRKBL WIRELESS PREP REQUIRES WIRELESS USB DONGLE | MSH US RM 8 CANON PRISM SERVICE LIC & REQ PALM CONTROLLER TRKBL
500847 | 201725 | CM | CMPLT | CMPLT | 2024-02-09 08:30 | 2024-02-26 13:08 | MSH US RM 1 APLIO XG MSH LIC & DFRG | (TC 02/09/2024 08:29:44 AM, Whittington, Scott) MSH US RM 1 APLIO XG MSH LIC & DFRG | MSH US RM 1 APLIO XG MSH LIC & DFRG
500857 | 201726 | CM | CMPLT | CMPLT | 2024-02-09 08:30 | 2024-02-26 13:08 | MSH US RM 3 APLIO XG MSH LIC & DFRG | (TC 02/09/2024 08:30:34 AM, Whittington, Scott) MSH US RM 3 APLIO XG MSH LIC & DFRG | MSH US RM 3 APLIO XG MSH LIC & DFRG
| 201727 | OC | CMPLT | CMPLT | 2024-02-09 08:31 | 2024-02-20 09:58 | On Call - Feb 9-16 | (TC 02/09/2024 08:30:42 AM, Scero, Tomlyn) Daily Checks (TC 02/12/2024 07:59:36 AM, Scero, Tomlyn) Daily checks (TC 02/13/2024 01:07:08 PM, Scero, Tomlyn) daily checks (TC 02/14/2024 11:08:32 AM, Scero, Tomlyn) Daily checks (TC 02/14/2024 02:00:00 AM, Scero, Tomlyn) After hours call CCU called with issues seeing their telemetry pack on patient in 15M. I helped her troubleshoot over the phone at first, changed telepacks, tried on herself. I looged into citrix to monitor the tele overview station as well. Tele signal going in and out as the patient moves around. We eventually decided to move the patient to another room as we suspected there to be an issue with the closest AP to room 15-367. (TC 02/15/2024 10:12:58 AM, Scero, Tomlyn) Daily checks
501888 | 201728 | CM | CMPLT | CMPLT | 2024-02-09 08:32 | 2024-03-01 13:07 | MSH US RM 2 APLIO 500 MSH LIC & DFRG | (TC 02/09/2024 08:31:39 AM, Whittington, Scott) MSH US RM 2 APLIO 500 MSH LIC & DFRG | MSH US RM 2 APLIO 500 MSH LIC & DFRG
1055712 | 201729 | CM | CMPLT | CMPLT | 2024-02-09 08:42 | 2024-02-09 08:45 | Retiring defective UPS | (TC 02/09/2024 08:44:06 AM, Baghirzade, Nara) UPS defective. Self-test failed. Confirmed with Kamran to retire it and dispose of it in e-waste.
1047867 | 201730 | CM | CMPLT | CMPLT | 2024-02-09 08:46 | 2024-02-09 13:05 | Cracked top case | (TC 02/08/2024 01:46:00 PM, Rhuland, Courtney) Replaced top case Reassembled device (TC 02/09/2024 01:02:20 PM, Rhuland, Courtney) When performing motor drive and occlusion test, device failed Clutch was worn out Replaced clutch greased leadscrew All tests passed post repair See procedure for results
1064593 | 201731 | CM | CMPLT | CMPLT | 2024-02-09 09:44 | 2024-02-09 09:46 | Remote clip broken | (TC 02/07/2024 09:44:39 AM, Rendon Lopez , Elizabeth) Confirmed issue, coordinator had purchased new clip. threading is broken off from remote and unable to install new clip, notified coordinator, covered open holes.
502088 | 201732 | CM | ACTIV | ISSUE | 2024-02-09 09:47 |  | Trying to delete the hermes network node | (TC 02/09/2024 09:47:28 AM, Paz, Ramon (Tino)) Trying to delete the network node 'hermes', however, the system does not allow.
| 201733 | CM | ACTIV | ISSUE | 2024-02-09 09:48 |  | L&D no asset WOs | (TC 02/08/2024 09:47:55 AM, Rendon Lopez , Elizabeth) OR3 lights are moving during surgery. Adjusted resistance screws on ceiling lights, tested-OK (TC 02/09/2024 09:57:04 AM, Rendon Lopez , Elizabeth) Nitrous device in L&D 7 and 10 having issues. Inspected devices, talked to clinical team , NFF (TC 02/14/2024 08:40:23 AM, Rendon Lopez , Elizabeth) OR 3 screen being held by tape. Adjusted screws, working. Same issue in OR 5, adjusted holding screws on all screens (TC 02/27/2024 01:33:47 PM, Rendon Lopez , Elizabeth) Rounds, meet with coordinators (TC 02/29/2024 08:26:29 AM, Rendon Lopez , Elizabeth) Philips monitor SW upgrade and config prep (TC 03/04/2024 08:57:20 AM, Rendon Lopez , Elizabeth) Philips monitors software upgrade (TC 03/05/2024 08:57:20 AM, Rendon Lopez , Elizabeth) Philips monitors software upgrade (TC 03/06/2024 08:57:20 AM, Rendon Lopez , Elizabeth) Philips monitors software upgrade (TC 03/18/2024 10:29:50 AM, Rendon Lopez , Elizabeth) OR 5 lights wondering. Adjusted screws, all good (TC 03/18/2024 12:09:03 PM, Rendon Lopez , Elizabeth) Testing fetal monitor Avalon CL for upgrade (TC 04/02/2024 11:36:08 AM, Rendon Lopez , Elizabeth) OR 2 lights wondering. Adjusted screws-Okay (TC 04/08/2024 12:52:04 PM, Rendon Lopez , Elizabeth) Repairing fetal monitor recorder (TC 06/06/2024 11:13:14 AM, Rendon Lopez , Elizabeth) Repairing 2 fetal monitor recorders (TC 06/13/2024 10:59:50 AM, Rendon Lopez , Elizabeth) 15M wall suction issues, testing different devices (TC 06/11/2024 02:30:22 PM, Rendon Lopez , Elizabeth) Checking wireless transducers (TC 06/19/2024 01:32:16 PM, Rendon Lopez , Elizabeth) CADD Pump SW update (TC 07/23/2024 03:58:47 PM, Rendon Lopez , Elizabeth) Checking Nitronox units, replacing missing demand valves (TC 07/30/2024 03:47:29 PM, Rendon Lopez , Elizabeth) Fetal procedure room, philips monitor move to fit AA desktop computer request (TC 08/12/2024 02:25:31 PM, Rendon Lopez , Elizabeth) Collecting all wireless trasnsducers for SW upgrade (TC 08/15/2024 09:47:34 AM, Rendon Lopez , Elizabeth) Wireless transducer SW upgrade (TC 08/16/2024 09:47:34 AM, Rendon Lopez , Elizabeth) Wireless transducer SW upgrade (TC 08/19/2024 03:21:02 PM, Rendon Lopez , Elizabeth) Wireless transducers and base SW upgrade (TC 08/21/2024 08:58:45 AM, Rendon Lopez , Elizabeth) Wireless transducer and base SW upgrade. (TC 08/20/2024 08:58:45 AM, Rendon Lopez , Elizabeth) Wireless transducer and base SW upgrade. (TC 09/10/2024 09:45:45 AM, Rendon Lopez , Elizabeth) Wireless transducer base station upgrade
1055561 | 201734 | CM | CMPLT | CMPLT | 2024-02-09 10:32 | 2024-02-09 11:32 | Not picking up contractions welll | (TC 02/09/2024 11:31:54 AM, Rendon Lopez , Elizabeth) Tested, NFF
67744 | 201735 | CM | CMPLT | NFF | 2024-02-09 12:30 | 2024-02-09 12:30 | Reading blood pressure high | (TC 02/09/2024 12:29:45 PM, Rhuland, Courtney) Went over and tested device using patient simulator All readings were accurate No fault found Functional - OK
1055732 | 201736 | PM | CMPLT | CMPLT | 2024-02-09 16:00 | 2024-02-09 16:03 | PM ACTIVE-MICROSTREAM EXTENSION | (TC 02/09/2024 04:00:27 PM, Cortez, John) Calibration completed, PM passed
1056424 | 201737 | CM | CMPLT | CMPLT | 2024-02-09 16:18 | 2024-02-09 16:26 | "missing power cord, battery not charging" | (TC 02/09/2024 04:18:28 PM, Cortez, John) Added spare power cord and charged overnight, Battery seems to hold charge, Added BP hose and Adult cuff All parameters tested ok Secured power cord to roll stand
| 201738 | RDS | CMPLT | CMPLT | 2024-02-12 07:50 | 2024-02-12 09:08 | Weekly check for ED
68268 | 201739 | CM | CMPLT | CMPLT | 2024-02-12 08:10 | 2024-04-04 09:11 | Msmt Ext Malfunction | (TC 02/12/2024 08:17:12 AM, Canjimavadivel, Kamallesh ) MSL connector corroded. Called vendor for quote. Placed order for MSL connector board (TC 04/04/2024 09:10:23 AM, Canjimavadivel, Kamallesh ) Replaced MSL down board. Connected to monitor in shop, error cleared
66683 | 201740 | CM | CMPLT | CMPLT | 2024-02-12 09:19 | 2024-02-12 09:32 | Missing parts | (TC 02/12/2024 09:31:06 AM, Rhuland, Courtney) Replaced missing X-ray tray with spare Replaced missing corner grommet Calibrated RH Functional check - OK
1054772 | 201741 | CM | CMPLT | CMPLT | 2024-02-12 09:53 | 2024-02-12 09:55 | BP not working, cuff won't inflate | (TC 02/12/2024 09:53:30 AM, Canjimavadivel, Kamallesh ) Connector piece was damaged causing leak. Replaced connector end. Tested ok
1047858 | 201742 | CM | CMPLT | CMPLT | 2024-02-12 10:08 | 2024-02-15 14:37 | System failure | (TC 02/12/2024 10:09:20 AM, Rhuland, Courtney) System error "primary audible alarm BGND test" message present upon startup Cleared error with biomed password Error repeated during tests and opened device to reseat interconnect cable and speaker cable Bottom case screw hold was broken when opened Reaching out to ICU medical to get part price (TC 02/15/2024 02:35:58 PM, Rhuland, Courtney) Replaced bottom case Performed various test infusions with no issue Functional check - OK
1050360 | 201743 | CM | CMPLT | CMPLT | 2024-02-12 10:24 | 2024-02-12 10:25 | Tubing clip broken | (TC 02/12/2024 10:24:11 AM, Canjimavadivel, Kamallesh ) Replaced damaged tube set. Also replaced second tube set as it appeared worn out. Tested ok
| 201744 | CM | CMPLT | CMPLT | 2024-02-12 10:41 | 2024-02-15 12:48 | Ordering medfusion cases | (TC 02/12/2024 10:47:17 AM, Rhuland, Courtney) Called ICU medical and got part prices Placed order for top cases for Medfusion 4000
65993 | 201745 | CM | CMPLT | CMPLT | 2024-02-12 11:20 | 2024-02-12 11:22 | Channel error | (TC 02/12/2024 11:20:45 AM, Rhuland, Courtney) channel error 354.6770 apparent when attatched to PCU Reseated and flashed syringe Check-in complete Error cleared
1059938 | 201746 | CM | CMPLT | CMPLT | 2024-02-12 11:33 | 2024-02-12 14:52 | Arm mount broken | (TC 02/12/2024 02:50:20 PM, Canjimavadivel, Kamallesh ) Replaced arm mount with extra mount from storage. Tested reading and noticed screen was flickering. Opened up and found ribbon cable was slighlty crimped by cover. Researted cable and closed back up. Flickering no longer present. To be monitored for reoccurance.
| 201747 | CM | CMPLT | CMPLT | 2024-02-12 11:45 | 2024-02-22 13:44 | HERMES NETWORK CONNECTIVITY TESTING AND SETUP | (TC 02/12/2024 11:44:44 AM, Paz, Ramon (Tino)) HERMES setup and testing for SHS Siemens and ERGO connectivity testing to HERMES server. It is not working at the moment. Port 10505 could be blocked. (TC 02/22/2024 01:43:08 PM, Paz, Ramon (Tino)) Connected VG aquisition station to HERMES. Also worked on Entegra computer.
1068220 | 201748 | CM | CMPLT | CMPLT | 2024-02-12 13:24 | 2024-03-01 13:07 | MSH US RM 7 SERVICE LICENCE & WIRELESS PREP | (TC 02/12/2024 01:28:13 PM, Whittington, Scott) MSH US RM 7 SERVICE LICENCE & WIRELESS PREP | MSH US RM 7 SERVICE LICENCE & WIRELESS PREP
65989 | 201749 | CM | CMPLT | CMPLT | 2024-02-12 14:47 | 2024-02-13 08:18 | Channel error | (TC 02/12/2024 02:47:53 PM, Rhuland, Courtney) Channel error 354.6770 Replaced rear case and right handle Flashed pump Error cleared Pressure sensor failed check-in and calibration (TC 02/13/2024 08:17:20 AM, Rhuland, Courtney) Reseated pressure sensor Calibrated and checked in pump All tests passed Refer to ASM for details
1066476 | 201750 | CM | CMPLT | CMPLT | 2024-02-12 15:00 | 2024-02-12 15:00 | replaced the HEPA filter | (TC 02/07/2024 02:59:56 PM, Huang, Jessica) got a call from OR desk and was asked to replace the filter
1061554 | 201751 | CM | CMPLT | CMPLT | 2024-02-12 15:02 | 2024-02-12 15:02 | replaced the HEPA filter | (TC 02/08/2024 03:02:32 PM, Huang, Jessica) got the call from Maureen and was asked to replace the filter by the support service supervisor
1066475 | 201752 | CM | CMPLT | CMPLT | 2024-02-12 15:04 | 2024-02-12 15:04 | replaced the HEPA filter | (TC 02/08/2024 03:04:20 PM, Huang, Jessica) got the call from Maureen and was asked to replace the filter by the support service supervisor
1066474 | 201753 | CM | CMPLT | CMPLT | 2024-02-12 15:05 | 2024-02-12 15:05 | replaced the HEPA filter | (TC 02/08/2024 03:05:17 PM, Huang, Jessica) got the call from Maureen and was asked to replace the filter by the support service supervisor
1058606 | 201754 | CM | CMPLT | CMPLT | 2024-02-12 15:10 | 2024-02-12 15:10 | replaced the HEPA filter | (TC 02/08/2024 03:09:41 PM, Huang, Jessica) got the call from Maureen and was asked to replace the filter by the support service supervisor
1056665 | 201755 | CM | CMPLT | CMPLT | 2024-02-12 15:11 | 2024-02-12 15:11 | zoom feature not working | (TC 01/31/2024 03:11:07 PM, Huang, Jessica) got the text message from Karmela and was requested to check the microscope. contacted Steve and was checking the device problem not found and informed Karmela
1049023 | 201756 | CM | CMPLT | CMPLT | 2024-02-12 15:14 | 2024-02-12 15:14 | not working | (TC 02/07/2024 03:13:53 PM, Huang, Jessica) picked up the device from OR and contacted Stryker for the replacement unit SN 190267716 was replaced by SN 230452744 in Feb 2024
| 201757 | CM | CMPLT | CMPLT | 2024-02-12 15:16 | 2024-02-12 15:16 | broeken suction device | (TC 02/07/2024 03:16:01 PM, Huang, Jessica) picked up the broken device from OR desk and replaced the regulator tested it, working fine and sent it back to OR
1053863 | 201758 | CM | CMPLT | CMPLT | 2024-02-12 15:17 | 2024-02-12 15:20 | Repair | "not delivering shock current". Not familiar with the unit. Went through documentation to understand the unit functionality. Cleaned unit, cable and lead wires. Checked functionality and current delivery. Tested ok. No issue found. Returned to OR.
| 201759 | CM | CMPLT | CMPLT | 2024-02-12 15:22 | 2024-02-12 15:23 | Meeting | Weekly Biomed Huddle
| 201760 | CM | CMPLT | CMPLT | 2024-02-12 15:24 | 2024-02-12 15:25 | Label Printer Set up and Biomed Label Design | Set up both label printers in computer. Downloaded and Installed driver and installed SW. Created Biomed inspection labels.
1062999 | 201761 | PM | CMPLT | CMPLT | 2024-02-12 16:24 | 2024-02-12 16:25 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/12/2024 04:24:35 PM, Cortez, John) Completed cleaning cycle
68346 | 201762 | CM | CMPLT | CMPLT | 2024-02-12 16:26 | 2024-02-12 16:31 | "error message on screen" | (TC 02/12/2024 04:25:52 PM, Cortez, John) Ran with test set and didn't come across any issues. Check error logs and the most recent errors are E13. E13 is related to pressure transducer. Ran test 4 to verify pressure values of sensor. Values were not within +- .2 mmgh (error tolerance stated in manual), more like within +- 2 mmgh. Calibrated using test 3 but got the same results. Will return to service. Changed one of the hoses with a spare.
1064418 | 201763 | CM | CANCL | DUP | 2024-02-13 08:55 |  | Rear Flap Door broken
| 201764 | CM | CMPLT | CMPLT | 2024-02-13 08:59 | 2024-02-13 09:00 | Damaged power cord for 11S Wound VAC (VAC ULTA) | (TC 02/13/2024 08:59:09 AM, Baghirzade, Nara) Replaced power cord only, adapter is functional. Checked with another wound VAC and the power supply is functional.
| 201765 | CM | CMPLT | CMPLT | 2024-02-13 09:16 | 2024-02-13 09:19 | Receoved safety issue report re: 4th fl. mezzanine AED | (TC 02/13/2024 09:15:42 AM, Baghirzade, Nara) Received following e-mail from security, Alyssa, and patient relations: Hello Wendy, I hope this email finds you well. Please see the safety concern outlined below for your review, follow-up and response. Please feel free to reach out to Julia should you require any additional information or clarification. Communicator (Visitor): Julia Dearing-Vollett Location: 4th floor, Pre-Admission Unit Tel: 416 460 1663 Email: jmdv25@gmail.com Communicator Perspective: -“I am a visitor waiting for my partner who is having a procedure; -I am on the 4th floor at the moment, opposite workstation 486, and noticed an empty AED box under the AED signage; -There are 2 other AED's on crash carts within my sight; -I am kindly suggesting to either remove the signage and box, or to put one of the crash carts under the sign; -I am including photos here to help; -Please feel free to contact me if you have any questions or concerns;” Desired Outcome: -Visitor suggested to either remove the signage and box, or to put one of the crash carts under the sign; Thank you for your time, follow-up and attention to this matter. (TC 02/13/2024 09:17:45 AM, Baghirzade, Nara) Confirmed with Kam that the AED has been relocated from the 4th fl mezzanine wall cabinet to the second crash cart. Followed up with Alyssa and Security regarding this as well as looped Steve Arai as RT has plans to relocate the wall cabinet.
1056125 | 201766 | PM | CMPLT | CMPLT | 2024-02-13 09:35 | 2024-02-13 10:01 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/13/2024 10:00:23 AM, Canjimavadivel, Kamallesh ) Cleaned hose end and valves Cleaning cycle complete
1056000 | 201767 | CM | CMPLT | CMPLT | 2024-02-13 09:38 | 2024-02-13 09:41 | Battery replacement | (TC 02/13/2024 09:40:42 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
| 201768 | CM | CMPLT | CMPLT | 2024-02-13 10:37 | 2024-02-29 09:31 | Masimo patient cable order for new X3 | (TC 02/13/2024 10:39:07 AM, Scero, Tomlyn) Called Masimo for the price. 6 are needed for bedside but I'll order 8 so they have 2 spares on hand. (TC 02/29/2024 09:30:43 AM, Scero, Tomlyn) emailed MohawkMedbuy to help locate order. Order received.
1057279 | 201769 | CM | CMPLT | CMPLT | 2024-02-13 11:06 | 2024-02-26 10:01 | Root monitor does not turn on; red flashing LED on the on/off button - replacement unit under warranty | (TC 02/13/2024 11:06:27 AM, Baghirzade, Nara) Plugged root to AC, switched it on. The screen stays on for about 5 seconds and then shuts off, with the on/off button flashing. Tried to turn it on another time, had the same outcome. Checked service manual and only option is to send it to Masimo for service. Requested service & RMA from Masimo for replacement unit. (TC 02/14/2024 01:17:30 PM, Baghirzade, Nara) SC 1003169532 (TC 02/16/2024 08:09:49 AM, Baghirzade, Nara) Received RMA #39120 KJV #417752 (TC 02/26/2024 09:59:03 AM, Baghirzade, Nara) PO 7500202175 Received replacement Root
| 201770 | RDS | CMPLT | CMPLT | 2024-02-13 11:29 | 2024-02-13 11:29 | weekly check for 14S and 16N | weekly check for 14S and 16N
1062029 | 201771 | CM | CMPLT | CMPLT | 2024-02-13 11:34 | 2024-02-13 11:34 | elevating height of bed not working properly | Checked. Functional. Talked to front desk and nurse. Checked together one more time. Functional.
| 201772 | CM | CMPLT | CMPLT | 2024-02-13 11:36 | 2024-02-13 11:36 | meeting | meeting
62701 | 201773 | CM | CMPLT | CMPLT | 2024-02-13 13:08 | 2024-02-13 13:11 | Intermittent signal loss. | (TC 02/13/2024 01:08:17 PM, Scero, Tomlyn) Confirmed bad signal when PIM was moved or squeezed. Opened PIM and found 4 loose pins. Resoldered the pins in place and sinal remained stable when PIM was moved. replaced patient data cable as it wouldn't sit in the PIM properly. Confirmed ECG with simulator
| 201774 | CM | CMPLT | CMPLT | 2024-02-13 14:10 | 2024-02-13 14:10 | worked with RT | worked with RT
| 201775 | CM | CMPLT | CMPLT | 2024-02-13 14:11 | 2024-02-13 14:11 | worked with Nara | worked with Nara
1055941 | 201776 | CM | CMPLT | CMPLT | 2024-02-13 14:11 | 2024-02-13 14:14 | Reported that the NIBP button is not working | (TC 02/13/2024 02:11:24 PM, Baghirzade, Nara) Tested device first in the morning, NIBP button was working. Took multiple blood pressure measurements and no issues with the button. Left the unit to charge and tested again in the afternoon; NIBP button is still functional. Took 3 successful measurements with no issues. Screen was dirty, cleaned the screen with wipes to see if it affected the screen. Pressed the NIBP buttons from different angles, no issues. NIBP, Temp, and SpO2 are all functional. Followed up with Jessica.
67566 | 201777 | CM | CMPLT | CMPLT | 2024-02-13 14:18 | 2024-02-13 14:19 | Found unit on the incoming repair shelf | (TC 02/13/2024 02:17:36 PM, Baghirzade, Nara) Old filac thermometers, end of support no parts. Could not get it to turn on even with new batteries. Confirmed with Elizabeth to retire this unit as L&D received new replacement thermometers.
1057278 | 201778 | CM | CMPLT | CMPLT | 2024-02-13 14:35 | 2024-02-13 14:38 | Battery replacement | (TC 02/13/2024 02:37:29 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
63234 | 201779 | CM | CMPLT | CMPLT | 2024-02-13 14:40 | 2024-07-24 09:26 | Cracked near threading for clamp | (TC 02/13/2024 02:54:33 PM, Cortez, John) Will send to Cardinal Health Requested no-charge PO (TC 02/14/2024 02:27:55 PM, Cortez, John) shipped to cardinal health WO-00099814 (TC 07/24/2024 09:25:39 AM, Cortez, John) Ran test infusion Incoming passed
| 201780 | CM | CMPLT | CMPLT | 2024-02-13 14:59 | 2024-02-13 15:00 | IVF Walkthrough | (TC 02/13/2024 02:58:49 PM, Rhuland, Courtney) Showed Rutvik around IVF clinic and introduced him to clinical staff
66704 | 201781 | CM | CMPLT | CMPLT | 2024-02-14 07:45 | 2024-02-14 07:48 | Front wall door was off the hinge, broken door latch and missing grommet | (TC 02/14/2024 07:45:33 AM, Baghirzade, Nara) Re-placed the south door back on the hinge, was missing the cover, replaced the cover on the snap (snap cover was missing). Replaced the SW latch. Replaced 1 missing grommets. Functional tests passed and completed. Scale is functional. Tilt is functional. Warmer is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.2 %RH Incubator is fully functional.
65919 | 201782 | PA | CMPLT | CMPLT | 2024-02-14 07:55 | 2024-02-23 10:38 | Ordering plunger assembly | (TC 02/14/2024 07:56:53 AM, Baghirzade, Nara) SC 1003168918 (TC 02/23/2024 10:37:54 AM, Baghirzade, Nara) PO 7202362669
1070003 | 201783 | CM | CMPLT | CMPLT | 2024-02-14 08:56 | 2024-02-14 09:10 | Incoming inspection | (TC 02/14/2024 09:09:41 AM, Rhuland, Courtney) Asset tagged and functional check - OK
68027 | 201784 | CM | CMPLT | CMPLT | 2024-02-14 09:16 | 2024-02-14 09:18 | Not working | (TC 02/14/2024 09:16:02 AM, Canjimavadivel, Kamallesh ) Attempted to calibrate and kept failing. Cleaned probe tip and calibrated, passed.
| 201785 | CM | CMPLT | CMPLT | 2024-02-14 09:34 | 2024-02-21 08:42 | batteries for stock | (TC 02/14/2024 09:37:09 AM, Scero, Tomlyn) ordered batteries for cabinet from SHS catalog (TC 02/21/2024 08:41:22 AM, Scero, Tomlyn) batteries received and put in the cabinet.
68865 | 201786 | CM | CMPLT | CMPLT | 2024-02-14 09:57 | 2024-02-14 09:59 | "Not charging" | (TC 02/13/2024 09:57:13 AM, Cortez, John) Confirmed, power cord defective Replaced cord and conditioned battery batt capacity: 4095 mAhrs
1054688 | 201787 | CM | CMPLT | NFF | 2024-02-14 10:23 | 2024-02-14 10:25 | "clogged downstream" | (TC 02/14/2024 10:23:21 AM, Cortez, John) Ran test infusion, unable to replicate issue
1065715 | 201788 | CM | CMPLT | CMPLT | 2024-02-14 10:30 | 2024-03-27 14:50 | "reads error, alarms will not run" | (TC 02/14/2024 10:31:02 AM, Cortez, John) Checked logs, 130.602 keyboard failure 133.608 backup speaker failure Ran keypad and alarm test, and both passed
1062275 | 201789 | CM | CMPLT | CMPLT | 2024-02-14 11:15 | 2024-02-14 14:26 | Mattress pump connection broke off | (TC 02/14/2024 02:24:33 PM, Scero, Tomlyn) The connector on the mattress came off when the mattress pump was pulled off. I reattached it to the mattress using the hose that was there. I pushed it in as far as it would go to create a good seal.
| 201790 | CM | CMPLT | CMPLT | 2024-02-14 11:16 | 2024-02-26 12:59 | AP15M06 not connecting to telepack | (TC 02/14/2024 11:16:02 AM, Scero, Tomlyn) Continuing from the call I got last night. Grabbed a telepack from CCU and my simulator to check the signal levels around the AP and room that wasn't working last night. Confirmed that the signal in that area was week at times and not consistent with what it should be (TC 02/14/2024 11:20:43 AM, Scero, Tomlyn) Working with Kamran to troubleshoot why AP15M06 isn't working. AP showing unregistered. In the IT closet I unplugged the AP from its port and reconnected it. No change as of yet. (TC 02/15/2024 03:00:40 PM, Scero, Tomlyn) Worked with Kamran to locate AP in IT closet. Reset AP by unpugging cable from ports 1, 6, 7, 8, 9. Took a telepak up to 15M under the AP and made sure it connected ok. Successful. (TC 02/26/2024 12:58:08 PM, Scero, Tomlyn) Unplugged remaining APs from the switch and plugged them back in. Checked on the PIIC system that it shows all APs available now.
| 201791 | CON | CMPLT | CMPLT | 2024-02-14 12:57 | 2024-02-14 12:59 | Compiled list of laser types MSH | (TC 02/14/2024 12:56:38 PM, Baghirzade, Nara) Compiled list of laser types at MSH along with wavelength ranges and optical density requirements for safety glasses/goggles. Purpose of the document is to identify vendors for safety goggles, discussed with Maisha and requested that she do the research for the vendors.
| 201792 | CM | CMPLT | CMPLT | 2024-02-14 13:33 | 2024-02-14 13:33 | worked with RT | worked with RT
1054214 | 201793 | CM | CMPLT | CMPLT | 2024-02-14 13:36 | 2024-02-14 13:37 | "shows "sec" after hitting "recall" button" | (TC 02/14/2024 01:35:55 PM, Cortez, John) Turned off security setting Tested ok
1053487 | 201794 | CM | CMPLT | NFF | 2024-02-14 13:39 | 2024-02-14 13:40 | "showed "monitor mode" only for 1 min, no result" | (TC 02/14/2024 01:39:21 PM, Cortez, John) Unable to replicate issue, Nurse might have hit mode button by accident Was able to get a spot check reading normally
| 201795 | CM | CMPLT | CMPLT | 2024-02-14 14:05 | 2024-02-14 14:07 | NICU Rounds | (TC 02/14/2024 02:04:51 PM, Rhuland, Courtney) NICU rounds completed
| 201796 | CM | CMPLT | CMPLT | 2024-02-14 14:15 | 2024-02-14 14:16 | RESUS rounds | (TC 02/14/2024 02:16:03 PM, Rhuland, Courtney) RESUS weekly rounds complete
| 201797 | CM | CMPLT | CMPLT | 2024-02-14 14:34 | 2024-02-16 15:43 | Requesting PO for Getinge Servo U/training | (TC 02/14/2024 02:34:15 PM, Cortez, John) Received quote for Servo U/N training on Apr 11/12 $5,375.00 (TC 02/16/2024 03:41:28 PM, Cortez, John) Requested, received PO and sent to Julien from Getinge. Registration confirmed for Apr 11/12
| 201798 | RDS | CMPLT | CMPLT | 2024-02-14 14:46 | 2024-02-14 14:48 | Weekly check for 10N/10S, Feb 12-16 2024 | (TC 02/14/2024 02:47:20 PM, Cortez, John) Nothing notable
| 201799 | RDS | CMPLT | CMPLT | 2024-02-14 14:48 | 2024-02-14 14:50 | Weekly check for ICU, Feb 12-16 2024 | (TC 02/14/2024 02:48:34 PM, Cortez, John) Nothing notable
1053348 | 201800 | CM | CMPLT | CMPLT | 2024-02-14 15:51 | 2024-03-01 13:10 | MSH BI BIOPSY ARM CAL ERROR | (TC 02/14/2024 04:07:05 PM, Whittington, Scott) MSH BI BIOPSY ARM CAL ERROR (TC 02/26/2024 01:54:59 PM, Whittington, Scott) REPLACEMENT BIOPSY DECK RECEIVED TO R3EPLACE DECK WITH FAAULTY LOCKING MECH (TC 02/29/2024 01:31:32 PM, Whittington, Scott) MSH BI BIOPSY ARM CAL ERROR From: Kim, So Young <SoYoung.Kim@sinaihealth.ca> Sent: February 29, 2024 11:53 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] H3 Ref# Hi Scott Left side lateral arm QC failing ref#:01464993 Thanks So Young | MSH BI BIOPSY ARM CAL ERROR
| 201801 | CM | CMPLT | CMPLT | 2024-02-14 15:53 | 2024-02-26 13:10 | MSH BI PHANTOM (PINK) BROKEN NYLON SCREWS | (TC 02/14/2024 03:53:27 PM, Whittington, Scott) MSH BI PHANTOM (PINK) BROKEN NYLON SCREWS (TC 02/14/2024 03:53:37 PM, Dolinar, John) MSH BI PHANTOM (PINK) BROKEN NYLON SCREWS | MSH BI PHANTOM (PINK) BROKEN NYLON SCREWS
68942 | 201802 | CM | CMPLT | CMPLT | 2024-02-15 08:38 | 2024-02-15 08:39 | Battery replace alarm | (TC 02/15/2024 08:38:23 AM, Rhuland, Courtney) Replaced and conditioned battery Display and keyboard tests passed Functional check - OK Electrical safety - OK
| 201803 | CM | CMPLT | CMPLT | 2024-02-15 08:51 | 2024-02-15 10:53 | Draeger ventilator training expenses | (TC 02/15/2024 09:00:11 AM, Baghirzade, Nara) Started the process of submitting expense claims. (TC 02/15/2024 10:52:49 AM, Baghirzade, Nara) Submitted all expenses during the training.
| 201804 | CM | CMPLT | CMPLT | 2024-02-15 08:52 | 2024-02-23 09:06 | panda warmer blender overhaul | (TC 02/15/2024 08:53:06 AM, Rendon Lopez , Elizabeth) Sending out Panda warmer blenders for overhaul Received
| 201805 | CM | CMPLT | CMPLT | 2024-02-15 09:43 | 2024-02-15 09:45 | Inventory of OR lasers with Jessica | (TC 02/15/2024 09:43:17 AM, Baghirzade, Nara) Taking inventory of OR lasers to compare what is on TMS. One laser missing: CO2 - Clarion Ultrapulse Encore - Asset #: 62125 Luisa from OR will look for it and confirm if its there.
| 201806 | CM | CANCL | CANCL | 2024-02-15 09:55 |  | ordering four table remotes
67959 | 201807 | CM | CMPLT | CMPLT | 2024-02-15 10:27 | 2024-02-15 10:48 | Broken bed tilt assembly | (TC 02/15/2024 10:27:27 AM, Rhuland, Courtney) Repaired bed tilt assembly Greased bed tilt Tightened NE/NW latches Calibrated RH Functional check - OK
1057938 | 201808 | CM | CMPLT | CMPLT | 2024-02-15 10:39 | 2024-04-02 13:33 | ordering remotes | (TC 02/15/2024 01:33:24 PM, Huang, Jessica) all the spare remotes were gone in OR 12 and infromed Maurren had to send the spare remote from Biomed and ordered new remotes (TC 02/22/2024 01:44:55 PM, Huang, Jessica) received the remotes and sent them to OR Maureen suggested to keep them in her office
62125 | 201809 | CM | CMPLT | CMPLT | 2024-02-15 10:55 | 2024-02-15 10:56 | Confirmation that laser is no longer in OR | (TC 02/15/2024 10:54:46 AM, Baghirzade, Nara) Luisa confirmed with Jessica that CO2 laser is no longer on site/in OR. Confirmed with Jessica to cancel all active PM work orders and retire the asset.
67913 | 201810 | CM | CMPLT | CMPLT | 2024-02-15 12:44 | 2024-02-15 12:47 | Missing door | (TC 02/15/2024 12:43:53 PM, Rhuland, Courtney) Noticed incubator in NICU that was missing a door During rounds this week, a door was found and brought to biomed from 18 storage room Brought door up from biomed and replaced it in unit
| 201811 | CM | CMPLT | CMPLT | 2024-02-15 13:30 | 2024-02-22 09:44 | Ordered New workstation for Nar | (TC 02/15/2024 01:29:42 PM, Chukri, Samir) Ordered biomed workstation to Nara (TC 02/22/2024 09:41:06 AM, Chukri, Samir) Received a new workshop and arranged to remove and disposal the old desk .
| 201812 | CM | CMPLT | CMPLT | 2024-02-15 13:50 | 2024-02-15 13:51 | Monitor not working in bay 8 | (TC 02/15/2024 01:50:16 PM, Canjimavadivel, Kamallesh ) MSL cable not correctly plugged into FMS rack, reseated connection and X3 displayin on monitor.
1067992 | 201813 | CM | CMPLT | CMPLT | 2024-02-15 15:52 | 2024-02-26 13:11 | MSH US RM 8 PRISM PROBE WITH CONNECTOR END GROMMET PULLED OUT | (TC 02/15/2024 03:51:58 PM, Whittington, Scott) MSH US RM 8 PRISM PROBE WITH CONNECTOR END GROMMET PULLED OUT | MSH US RM 8 PRISM PROBE WITH CONNECTOR END GROMMET PULLED OUT
1059828 | 201814 | CM | CMPLT | CMPLT | 2024-02-16 06:41 | 2024-02-16 06:42 | Cleaning | (TC 02/15/2024 01:00:00 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
| 201815 | CM | CMPLT | CMPLT | 2024-02-16 06:43 | 2024-02-16 06:48 | Working on the recall assigned on onerecall | (TC 02/15/2024 02:00:00 PM, Baghirzade, Nara) Looking over at recalls assigned. 1: Medfusion; 10 anomalies that are affecting the medfusion 4000s with specific software versions. Confirmed that the NICU medfusions are on software V 1.6.0 and are impacted by the recall. Followed up with Samir and decided to email vendor icumed to setup a webinar session to upgrade the software and firmware to help mitigate the anomalies. Reached out to Simon from icumed for session bookings. 2. INVOS neo sensors are affected, there are 2 that are in Dr. Deshpande's office. Contacted Dr. Deshpande to inquire about sensors and if she had purchased any, also forwarded the recall notification to her. She confirmed, so I emailed her secretary to set up a time for her to let me in her office so I can check the batch number of the neo sensors. Will visit on Friday Feb 16.
| 201816 | CM | CMPLT | CMPLT | 2024-02-16 06:49 | 2024-02-16 06:51 | NICU RT repairs | (TC 02/15/2024 11:00:00 AM, Baghirzade, Nara) Met with Rex to go over any devices that require repairs in NICU RT. Gave him the aerogen plus that was dropped off in biomed. Box on his desk had sensors and accessories that were broken however let him know that they need to purchase new sensors for replacement. Discussed CMAC blades with him, asked him to flag any old (black handle) CMAC blades, so we can keep track of how many of the old ones we have as they are end of support. Will meet with a CE to discuss this being purchased through capital. Asked Rex to find out if we can get a list of all serial numbers for the CMAC blades.
| 201817 | PA | CMPLT | CMPLT | 2024-02-16 06:54 | 2024-02-20 12:06 | Ordering sterile water for NICU altrix cleaning | (TC 02/16/2024 06:54:52 AM, Baghirzade, Nara) SC 1003171070 (TC 02/20/2024 12:05:07 PM, Baghirzade, Nara) Order # 4150272101
1059923 | 201818 | CM | CMPLT | CMPLT | 2024-02-16 07:39 | 2024-04-02 13:21 | the left touch screen is not responing | The left touch screen is not working and contacted conmed to request a loaner device. (TC 02/16/2024 09:18:01 AM, Baghirzade, Nara) SC 1003171241 (TC 02/22/2024 07:33:27 AM, Baghirzade, Nara) PO 7500202287 Loaner device received. Jessica is working on it. (TC 02/12/2024 01:22:38 PM, Huang, Jessica) got the call in OR 2 and saying the device was not working properly; however, the patient was on the table and the clinical team didnt want to be interupted. went back after the case was completed and noticed the left touch screen was not working. contacted Conmed for tech support and did some troubleshooting on the device. did not fix the issue; however, the device could be used because the right side was working. informed Maureen and was told not to remove the device unitl the loaner device was available. contacted Conmed again and was informed the sales rep Joe would be onsite to have a look at the device. (TC 02/15/2024 01:28:55 PM, Huang, Jessica) Joe Brown from Conmed was onsite and took him to the room. he checked the device and informed the device would need to be sent for repair. (TC 03/14/2024 01:30:27 PM, Huang, Jessica) got the device back and sent it back to OR 2 sent the loaner device back to the vendor and informed the OR team
1059938 | 201819 | CM | CMPLT | CMPLT | 2024-02-16 08:18 | 2024-02-16 09:03 | Screen flickering | (TC 02/16/2024 08:54:41 AM, Canjimavadivel, Kamallesh ) Removed back cover of display and reseated all connections. No visible damage on any cables. Turned on and flickering no longer present
| 201820 | CM | CMPLT | CMPLT | 2024-02-16 09:05 | 2024-02-16 09:11 | Confirming medical device recall regarding Medtronic Reusable Sensor Cable for PM7100 | (TC 02/16/2024 09:07:10 AM, Baghirzade, Nara) Reviewed the lot numbers of the PMAC71RSC reusable sensor cable for PM7100 (QTY: 4); confirmed that the lot numbers are from 2023-03-29 and are not affected by the medical device recall. Followed up with Dr. Deshpande. (TC 02/16/2024 09:09:49 AM, Rhuland, Courtney) Reviewed the lot numbers of the PMAC71RSC reusable sensor cable for PM7100 (QTY: 4); confirmed that the lot numbers are from 2023-03-29 and are not affected by the medical device recall.
1055772 | 201821 | CM | CMPLT | CMPLT | 2024-02-16 09:18 | 2024-02-16 09:20 | Out of paper | (TC 02/16/2024 09:18:17 AM, Rhuland, Courtney) Brought an extra roll up from biomed The paper was not empty but jammed Reseated the paper feed and tested the output Also intstructed Mary (RESUS nurse) to wait until transmission is complete when performing 30J test
502088 | 201822 | CM | CMPLT | CMPLT | 2024-02-16 09:34 | 2024-02-16 09:34 | INSTALLED QC SOURCES | (TC 02/16/2024 09:34:01 AM, Paz, Ramon (Tino)) Installed QC sources. Tested OK. Updated all labels.
502089 | 201823 | CM | CMPLT | CMPLT | 2024-02-16 09:36 | 2024-02-16 09:36 | INSTALLED QC SOURCES | (TC 02/16/2024 09:36:08 AM, Paz, Ramon (Tino)) Installed new QC sources. Tested OK. Updated all labels.
1063746 | 201824 | CM | CMPLT | CMPLT | 2024-02-16 10:18 | 2024-02-16 10:19 | Cleaning | (TC 02/16/2024 10:18:23 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1062389 | 201825 | CM | CMPLT | CMPLT | 2024-02-16 10:22 | 2024-02-16 10:23 | User reported that there is a functionality error; motion controls not working | (TC 02/16/2024 10:22:24 AM, Baghirzade, Nara) Staff reported bed not working however could not find an issue with the bed, did not see any error when powering on the bed. Checked all motion controls including both side panels and patient panels. Power cycled the bed again and no errors. Bed is functional.
1062505 | 201826 | CM | CMPLT | CMPLT | 2024-02-16 10:24 | 2024-02-16 10:24 | Functional check | (TC 02/16/2024 10:23:47 AM, Baghirzade, Nara) Checked all motion controls including both side panels and patient panels. Power cycled the bed, no errors or issues upon boot up. Bed is functional.
| 201827 | CM | CMPLT | CMPLT | 2024-02-16 10:53 | 2024-02-16 10:54 | Full team meeting | (TC 02/16/2024 10:53:18 AM, Baghirzade, Nara) Full team meeting
| 201828 | CM | CMPLT | CMPLT | 2024-02-16 11:36 | 2024-02-16 11:37 | 1-1 meeting with Adeel | (TC 02/16/2024 11:36:32 AM, Baghirzade, Nara) 1-1 meeting with Adeel
| 201829 | CM | CMPLT | CMPLT | 2024-02-16 13:08 | 2024-02-16 13:12 | Recurring issue with the NICU defibrillator not checking in on the dashboard | (TC 02/16/2024 01:08:06 PM, Baghirzade, Nara) Consistently the NICU crash cart defibrillator did not check in on the dashboard. Met with Janet to let her know. Reviewed the tip sheet that is attached to the crash cart and noted the discrepancy, it is incomplete. There is no instruction to wait for the data transmission to be completed before shutting off the defibrillator. Created a revised tip sheet and added the extra step: "10. Confirm that “transmission completed” is displayed at the top left corner of display (DO NOT SHUT OFF THE DEFIBRILLATOR BEFORE SEEING THIS MESSAGE)" Shared it with NICU management, requested they update both NICU and RESUS crash carts and circulate new tip sheet with the NICU staff. Confirmed with Samir that the tip sheet was acceptable.
| 201830 | RDS | CMPLT | CMPLT | 2024-02-16 13:12 | 2024-02-16 13:14 | Weekly rounds: NICU | (TC 02/16/2024 01:12:04 PM, Baghirzade, Nara) Weekly rounds completed in NICU & RESUS. See procedure for results.
| 201831 | RDS | CMPLT | CMPLT | 2024-02-16 13:14 | 2024-02-16 13:15 | Weekly Rounds: 11 South | (TC 02/16/2024 01:14:26 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 201832 | RDS | CMPLT | CMPLT | 2024-02-16 13:16 | 2024-02-16 13:17 | Weekly Rounds: 11 North | (TC 02/16/2024 01:15:40 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 201833 | CM | CMPLT | CMPLT | 2024-02-16 15:08 | 2024-02-16 15:09 | Meeting | Monthly division meeting Meeting with AA and SC. Created TOR for CMMS committee.
| 201834 | CM | CMPLT | CMPLT | 2024-02-16 15:11 | 2024-02-16 15:12 | Clean Up in 3D Printing room | Organized 3D printer room. Mounted power bar in wooden dash board
1054537 | 201835 | PM | CMPLT | CMPLT | 2024-02-16 15:24 | 2024-02-16 15:26 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/16/2024 03:24:20 PM, Cortez, John) Completed cleaning cycle During rinse, got an error saying that hose 3 wasn't being detected. Reseated port and ran altrix for another 10 minutes. Ran ok.
65590 | 201836 | CM | CMPLT | NFF | 2024-02-16 15:32 | 2024-02-16 15:33 | "Shows "direct mode" " | (TC 02/16/2024 03:31:58 PM, Cortez, John) Unable to replicate issue. User must have switched modes by accident, or left temp probe idle too long.
1051447 | 201837 | CM | CMPLT | CMPLT | 2024-02-16 15:37 | 2024-02-16 15:39 | "showed error after 1 min" | (TC 02/16/2024 03:37:09 PM, Cortez, John) Was sent down from 10S, however asset was retired. Unsure how they obtained retired device. Confirmed issue, thermometer unable to get reading. Tried different probes/cleaning contacts. Will retire again and discard device.
| 201838 | RDS | CMPLT | CMPLT | 2024-02-20 08:26 | 2024-02-20 08:26 | Weekly check for 14S and 16N | Weekly check for 14S and 16N
65898 | 201839 | CM | CMPLT | CMPLT | 2024-02-20 08:36 | 2024-02-20 08:38 | Broken latch and missing SN barcode sticker | (TC 02/20/2024 08:37:44 AM, Baghirzade, Nara) Replaced latch assembly. Reprinted the SN barcode sticker. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1054521 | 201840 | CM | CMPLT | CMPLT | 2024-02-20 09:39 | 2024-02-20 10:09 | V3 not working | (TC 02/20/2024 09:39:33 AM, Canjimavadivel, Kamallesh ) Replaced V1-V3 chest lead set. Repalced Leads V7,V8,V9
65952 | 201841 | CM | CMPLT | CMPLT | 2024-02-20 09:48 | 2024-02-23 11:17 | 354.6670 error when connecting to PCU | (TC 02/20/2024 09:51:58 AM, Baghirzade, Nara) Re-seated the pressure sensor, put device back together, connected to the PCU and still getting the same error: 354.6770. Replaced the pressure sensor. When replacing the pressure sensor, noted that the rear case and left handle are broken. Will replace both parts. (TC 02/23/2024 11:15:52 AM, Baghirzade, Nara) Replaced rear case. Replaced left handle. Reprinted asset tag. Calibrated pressure sensor. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. | 354.6770 pressure sensor error. Left handle and rear case broken.
| 201842 | RDS | CMPLT | CMPLT | 2024-02-20 09:59 | 2024-02-20 13:54 | Weekly checks in CCU | (TC 02/20/2024 01:53:42 PM, Scero, Tomlyn) weekly checks ok. 30J tests complete.
68687 | 201843 | CM | CMPLT | CMPLT | 2024-02-20 10:14 | 2024-02-20 10:15 | Does not switch on when operating on battery | (TC 02/20/2024 10:14:20 AM, Baghirzade, Nara) Battery replaced. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 93 mohms Leakage current: 74.4 uA
1061212 | 201844 | CM | CMPLT | CMPLT | 2024-02-20 10:33 | 2024-02-20 10:36 | VMC Error | (TC 02/20/2024 10:34:25 AM, Reid, Emmerson ) VMC needed to be restarted. Also showed Rutvic around the OR and PACU
| 201845 | RDS | CMPLT | CMPLT | 2024-02-20 10:36 | 2024-02-20 11:00 | Weekly Checks For PACU | (TC 02/20/2024 10:59:45 AM, Reid, Emmerson ) All good
68765 | 201846 | CM | CMPLT | CMPLT | 2024-02-20 11:04 | 2024-02-20 11:06 | PCU not switching on, AC LED does not light up when plugged in to AC | (TC 02/20/2024 11:04:15 AM, Baghirzade, Nara) Replaced front case as per recall. Connected PCU to AC, AC LED switches on however the device battery is defective. Replaced battery. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 185 mohms Leakage current: 75.1 uA
68963 | 201847 | CM | CMPLT | CMPLT | 2024-02-20 11:06 | 2024-02-20 11:08 | User did not report what the issue/errors are with the device | (TC 02/20/2024 11:06:20 AM, Baghirzade, Nara) Downloaded error log, did not see any issue with the device. Did not note any significant errors with the PCU; only error was the watchdog timer and that was cleared with a power cycle of the PCU. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 88 mohms Leakage current: 77.0 uA
1057550 | 201848 | CM | CMPLT | CMPLT | 2024-02-20 13:17 | 2024-03-21 14:32 | SPO2 port plugged | (TC 02/20/2024 01:16:50 PM, Scero, Tomlyn) Contacted philips customer service portal to create a case. Asking for price of replacement board. (TC 02/23/2024 01:58:16 PM, Scero, Tomlyn) Emailed Philips to get a price quote for the connector block. Ordered on ereq. (TC 03/21/2024 02:30:57 PM, Scero, Tomlyn) Connector plate replaced and parameters tested. repair complete.
| 201849 | CM | CMPLT | CMPLT | 2024-02-20 13:32 | 2024-02-20 13:32 | Meeting | Biomed meeting
| 201850 | CM | CMPLT | CMPLT | 2024-02-20 13:33 | 2024-02-20 13:33 | worked with RT | worked with RT
| 201851 | CM | CMPLT | CMPLT | 2024-02-20 13:33 | 2024-02-20 13:33 | worked on TMS | worked on TMS
1056128 | 201852 | CM | CMPLT | CMPLT | 2024-02-20 14:01 | 2024-02-20 14:02 | Signal cutting | (TC 02/20/2024 02:00:58 PM, Rendon Lopez , Elizabeth) Confirme issue. Previously repaired, all transducer. Retired
1049210 | 201853 | CM | CMPLT | CMPLT | 2024-02-20 14:03 | 2024-02-20 14:03 | Physical damage | (TC 02/20/2024 02:02:58 PM, Rendon Lopez , Elizabeth) Confirme issue. Previously repaired, all transducer. Retired
| 201854 | CM | CMPLT | CMPLT | 2024-02-20 14:10 | 2024-02-20 14:11 | Editing claim submission for Draeger ventilator training | (TC 02/20/2024 02:10:31 PM, Baghirzade, Nara) Adeel requested to review the submission as one item was over the limit; revised the claim and submitted the claim back on SAP concur.
1064732 | 201855 | CM | CMPLT | CMPLT | 2024-02-20 14:14 | 2024-05-17 10:57 | E210 Software version mismatch | (TC 02/12/2024 01:11:06 PM, Huang, Jessica) got the call in OR 1 saying the olympus 4k system was not working properly. restarted the device, disconnected the camera head and reconnected it. Error code 201 was showing on the display. Informed Irina to get the 4k moible cart. (TC 02/13/2024 01:16:15 PM, Huang, Jessica) Contacted Olympus for tech support and was informed the device needed to be sent for repair. Informed the OR management team and requested the loaner device (TC 02/13/2024 01:16:15 PM, Huang, Jessica) Contacted Olympus for tech support and was informed the device needed to be sent for repair. Informed the OR management team and requested the loaner device got the PO and forwarded it to the vendor (TC 02/22/2024 01:18:23 PM, Huang, Jessica) received the loaner and installed it in OR 1 (TC 03/18/2024 01:19:30 PM, Huang, Jessica) followed up with Olympus to see the repair status and was informed the device is being worked on one of the boards needed to be replaced and still waiting for the part (TC 05/09/2024 10:56:01 AM, Huang, Jessica) received the repair quote from Olympus and forwarded it to Reeba for approval (TC 05/17/2024 10:56:33 AM, Huang, Jessica) brought the device to OR 1 and sent the loaner device back to the company
1063299 | 201856 | CM | CMPLT | CMPLT | 2024-02-20 14:28 | 2024-02-20 14:29 | Functional Check | (TC 02/20/2024 02:28:15 PM, Rhuland, Courtney) Physical inspection passed. Functional tests passed. Test weight 10lb, measured 10lb Test weight 20lb, measured 20lb No calibration needed. Scale is fully functional.
1063125 | 201857 | CM | CMPLT | CMPLT | 2024-02-20 14:38 | 2024-02-20 14:43 | Functional Inspection | (TC 02/20/2024 02:40:18 PM, Rhuland, Courtney) Device is in good physical condition. Set 120/80 bpm 100; measured 119/81 bpm 100 Set 80/40 bpm 100; measured 79/40 bpm 100 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
1063124 | 201858 | CM | CMPLT | CMPLT | 2024-02-20 14:38 | 2024-02-20 14:45 | Functional Inspection | (TC 02/20/2024 02:44:04 PM, Rhuland, Courtney) Device is in good physical condition. Set 120/80 bpm 100; measured 118/82 bpm 100 Set 80/40 bpm 100; measured 81/40 bpm 100 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
531120 | 201859 | CM | CMPLT | CMPLT | 2024-02-20 14:38 | 2024-02-20 14:49 | Functional Inspection | (TC 02/20/2024 02:48:08 PM, Rhuland, Courtney) Device is in good physical condition. Set 120/80 bpm 100; measured 119/80 bpm 100 Set 80/40 bpm 100; measured 78/41 bpm 101 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
531121 | 201860 | CM | CMPLT | CMPLT | 2024-02-20 14:38 | 2024-02-20 14:50 | Functional Inspection | (TC 02/20/2024 02:49:49 PM, Rhuland, Courtney) Device is in good physical condition. Set 120/80 bpm 100; measured 119/80 bpm 100 Set 80/40 bpm 100; measured 80/40 bpm 100 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
66531 | 201861 | CM | CMPLT | CMPLT | 2024-02-20 14:45 | 2024-02-20 14:47 | Not turning on | (TC 02/20/2024 02:45:48 PM, Rendon Lopez , Elizabeth) Removed from L&D room 3. Troubleshooting, replaced API board and switch board. Configured SN, tested, Working.
| 201862 | CM | CMPLT | CMPLT | 2024-02-20 14:55 | 2024-03-04 13:30 | Tools sorting and purchase | Feb 20 - Went through available tools amd Complied the list of tools needed to order from Mcmaster carr, Digikey and Newark Canada. Feb 23 - Created new shopping carts #1003175207 and #1003175643 and sent it for approval. Feb 28 - Followed up with vendors. Was informed that newark put our PO on hold due to two past payment overdue. Reached out to MMC to have the issue resolved and po processed. Awaiting confirmation Mar 01 - Followed up and escalated issue with MMC. After a lot of back and forth communication, was instructed that vendor will be paid by next week. PO released by the vendor. Awaiting delivery. Mar 04 - Received. Organized in the work area.
1070054 | 201863 | CM | CMPLT | CMPLT | 2024-02-20 15:01 | 2024-02-20 15:04 | Functional Inspection | (TC 02/20/2024 03:00:55 PM, Rhuland, Courtney) Device is in good physical condition. Set 120/80 bpm 100; measured 120/81 bpm 100 Set 80/40 bpm 100; measured 79/40 bpm 100 Cuffs in good condition, no leaks or damage. Battery ok. All functional tests passed.
501107 | 201864 | CM | CMPLT | CMPLT | 2024-02-20 15:17 | 2024-02-26 13:14 | MSH OPG CEOU ROOM 3 SITE 553759 REF# 0123149935 FUZZY IMAGE | (TC 02/20/2024 03:17:23 PM, Dolinar, John) MSH OPG CEOU ROOM 3 SITE 553759 REF# 0123149935 FUZZY IMAGE (TC 02/22/2024 02:18:27 PM, Fan, Tak) Swapped with spare. (TC 02/22/2024 02:18:56 PM, Fan, Tak) Jillian replaced parts. | MSH OPG CEOU ROOM 3 SITE 553759 REF# 0123149935 FUZZY IMAGE
1057302 | 201865 | CM | CMPLT | CMPLT | 2024-02-20 15:20 | 2024-09-25 09:34 | MSH OPG SPP D7 NETWORK | (TC 02/20/2024 03:19:48 PM, Dolinar, John) MSH OPG SPP D7 NETWORK NOW WIRELESS ONLY NW CABLE AVAILABLE IF NEEDED INFORMED THAT CARDIOLOGISTS NOW ON STAFF AT MSH NO NEED TO TXFR STUDIES TO HSC (TC 02/21/2024 01:03:35 PM, Whittington, Scott) MSH OPG SPP D7 REACTIVATED WIRELESS PROVIDED LONGER ETHERNET CABLE | MSH OPG SPP D7 NETWORK
| 201866 | CM | CMPLT | CMPLT | 2024-02-20 15:25 | 2024-02-20 15:25 | Missing hose connection | (TC 02/20/2024 03:24:35 PM, Rhuland, Courtney) Replaced missing connector on sphyg hose Functional check - OK
1049904 | 201867 | CM | CMPLT | CMPLT | 2024-02-20 15:27 | 2024-02-20 15:28 | Broken pins on top battery charger | (TC 02/20/2024 03:27:02 PM, Fan, Tak) Replaced top charger. (TC 02/20/2024 03:27:02 PM, Dolinar, John) Broken pins on top battery charger (TC 02/20/2024 03:27:02 PM, Fan, Tak) MSH R64 Replaced top charger. (TC 02/20/2024 03:27:02 PM, Dolinar, John) MSH R64 Broken pins on top battery charger | Broken pins on top battery charger
| 201868 | CM | CMPLT | CMPLT | 2024-02-21 07:50 | 2024-02-21 07:52 | Creating assets for bassinets | (TC 02/21/2024 07:50:09 AM, Rhuland, Courtney) Created all assets for new fleet of bassinets When bassinets arrived, noted no unique identifier All bassinets have same batch number as serial number Decided not to physically tag bassinets, only in TMS for inventory purposes Assets 1070056-1070078
500857 | 201869 | CM | CMPLT | CMPLT | 2024-02-21 08:01 | 2024-02-26 13:13 | MSH 5 US RM 3 APLIO XG UNABLE TO SAVE IMAGES | (TC 02/21/2024 08:01:17 AM, Whittington, Scott) MSH 5 US RM 3 APLIO XG UNABLE TO SAVE IMAGES (TC 02/21/2024 08:01:40 AM, Whittington, Scott) MSH 5 US RM 3 APLIO XG UNABLE TO SAVE IMAGES HDD DRIVE PARTITION D REQUIRED CHECK DISK RESTORED ABILITY TO STORE IMAGES (TC 02/21/2024 08:01:17 AM, Whittington, Scott) MSH 5 US RM 3 APLIO XG UNABLE TO SAVE IMAGES | MSH 5 US RM 3 APLIO XG UNABLE TO SAVE IMAGES
| 201870 | RDS | CMPLT | CMPLT | 2024-02-21 08:32 | 2024-02-21 10:49 | Weekly check for ED
1057383 | 201871 | CM | CMPLT | CMPLT | 2024-02-21 09:49 | 2024-02-26 13:12 | MSH 525-2 CR ROOM CASSETTES WONT FEED | (TC 02/21/2024 01:27:19 PM, Dolinar, John) MSH 525-2 CR ROOM CASSETTES WONT FEED K4187-3265 REF# HARSH 930312 (TC 02/21/2024 01:27:19 PM, Dolinar, John) MSH 525-2 CR ROOM CASSETTES WONT FEED K4187-3265 REF# HARSH 930312 Changed UPS batteries (TC 02/22/2024 12:06:17 PM, Whittington, Scott) MAH CR READER CASSETTES WONT FEED SEAN INSTALLED REAAR SENSOR FORM HIS MOBILE STOCK WILL ORDER WHEELS AND REAR SENSOR | MSH 525-2 CR ROOM CASSETTES WONT FEED K4187-3265 REF# HARSH 930312 Changed UPS batteries
| 201872 | CM | CMPLT | CMPLT | 2024-02-21 10:02 | 2024-02-23 10:32 | Moving workspaces in biomed | (TC 02/23/2024 10:30:58 AM, Baghirzade, Nara) Completed move of all stuff and cleaned the old workspace.
| 201873 | CM | CMPLT | CMPLT | 2024-02-21 10:02 | 2024-02-21 10:03 | Meeting with CEs
1055580 | 201874 | CM | CMPLT | CMPLT | 2024-02-21 10:08 | 2024-03-05 13:05 | battery not holding charge | battery not holding charge. shipping 4 Masimo Root for battery replacement/repair. Asset#1055580(sn.200060501), 1055953(2000060408), 1055573(2000060500), 1055579(2000060521) Received 4 x replacement unit from Masimo. PO#7500202419
1057547 | 201875 | CM | CMPLT | CMPLT | 2024-02-21 10:21 | 2024-02-21 10:24 | SPO2 malfunction | (TC 02/21/2024 10:21:11 AM, Scero, Tomlyn) Confirmed SPO2 malunction even with a new cable and sensor. I ordered a new spo2 board but instaed of waiting for it to arrive, I used the board from 1057550 because it also has a damaged spo2 port. After the SPO2 board was replaced, no error occured and the issue was resolved. SPO2 reading ok.
1055036 | 201876 | CM | CMPLT | CMPLT | 2024-02-21 11:22 | 2024-02-21 11:23 | Battery replacement | (TC 02/21/2024 11:22:27 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
69126 | 201877 | CM | CMPLT | CMPLT | 2024-02-21 11:22 | 2024-02-21 11:26 | Broken door latch | (TC 02/21/2024 11:22:21 AM, Rhuland, Courtney) Door latch assembly was broken and door closed sensor was not registering Replaced door handle with spare Calibrated pressure and rate Performed check-in All tests passed Refer to ASM for details
1054982 | 201878 | CM | CMPLT | CMPLT | 2024-02-21 11:26 | 2024-02-21 11:27 | User reported that "cords" are missing | (TC 02/21/2024 11:26:06 AM, Baghirzade, Nara) NIBP hose and SpO2 cables are there. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional. Placed a y-splitter for safetynet.
| 201879 | CM | CMPLT | CMPLT | 2024-02-21 12:12 | 2024-02-21 12:12 | worked with RT | worked with RT
| 201880 | CM | CMPLT | CMPLT | 2024-02-21 12:32 | 2024-02-21 12:32 | Communications committee meeting
| 201881 | CM | CMPLT | CMPLT | 2024-02-21 12:36 | 2024-02-21 12:38 | NICU Rounds | (TC 02/21/2024 12:35:55 PM, Rhuland, Courtney) Rounds completed for NICU
1057384 | 201882 | CM | CMPLT | CMPLT | 2024-02-21 13:29 | 2024-03-04 11:26 | MSH 525-2 CR ROOM CASSETTES WONT FEED | (TC 02/21/2024 01:29:34 PM, Dolinar, John) MSH 525-2 CR ROOM CASSETTES WONT FEED (TC 02/22/2024 12:03:55 PM, Whittington, Scott) MSH 525-2 CR ROOM CASSETTES WONT FEED SITE K4187-3354 SEAN INSTALLED LAMPS FROM LOCAL STOCK WILL ORDER WHEELS AND REAR SENSOR | MSH 525-2 CR ROOM CASSETTES WONT FEED SITE K4187-3354 OEM came on site to repair
| 201883 | CM | CMPLT | CMPLT | 2024-02-21 13:33 | 2024-02-21 13:33 | worked at 14S | worked at 14S
| 201884 | CM | CMPLT | CMPLT | 2024-02-21 13:34 | 2024-02-21 13:34 | worked on Masimo | worked on Masimo
1055598 | 201885 | CM | CMPLT | CMPLT | 2024-02-21 13:35 | 2024-03-05 13:03 | not communicating with Root | not communicating with Root. Received replacement unit from Masimo
1057384 | 201886 | CM | CMPLT | CMPLT | 2024-02-21 13:48 | 2024-03-04 11:27 | CASSETTES WONT LOAD |  | CASSETTES WONT LOAD Ticket# 930330 OEM came in to repair.
68377 | 201887 | CM | CMPLT | CMPLT | 2024-02-21 13:51 | 2024-02-21 14:20 | RT/AA said pictogram was weird | (TC 02/21/2024 02:15:37 PM, Scero, Tomlyn) Sven called and told me that the pictogram in the upper left hand corner looked "really weird". I pulled the device into OR12 and ran it through the self test and all the different vent modes including auto mode. I made sure before starting testing to assemble AGM correctly ie; attach sample line, lock left side of breathing system, attached filter to the breathing hose, and attached SGA hose to water trap. emailed Peter to let him know the device is working with no issues found.
| 201888 | CM | CMPLT | CMPLT | 2024-02-21 14:13 | 2024-02-22 14:12 | Desk assembly
1070099 | 201889 | CM | CMPLT | CMPLT | 2024-02-21 14:22 | 2024-02-21 16:13 | Incoming Inspection | (TC 02/21/2024 04:07:25 PM, Patel, Rutvik) Unit was brought in by Endo department in conjustion with Steris Rep David Murphy and was put into clinical use. Perofrmed EST. Tested ok. However, unit does not possess ESA mark or apporval. Reached out to sales rep for confirmation. Until then, recommended not to use the unit clinically. Inspection failed. Removed unit from video tower and placed original olympus unit back. (TC 02/22/2024 12:38:57 PM, Patel, Rutvik) Unit was replaced with a different one with approval mark.Updated TMS asset info. Completed functional check and EST. Tested ok. Affixed asset tag and inspection label. Assisted rep to install the unit in Procedure room.
1056911 | 201890 | CM | CMPLT | CMPLT | 2024-02-21 14:44 | 2024-02-21 14:46 | Microscope not balanced | (TC 02/21/2024 02:44:30 PM, Reid, Emmerson ) Microscope was not balanced. Went up and balanced it. Working now.
1061364 | 201891 | CM | CMPLT | CMPLT | 2024-02-21 15:28 | 2024-02-21 15:27 | Needs brake replacement. | Replaced.
| 201892 | CM | CMPLT | CMPLT | 2024-02-21 16:16 | 2024-02-21 16:40 | Meeting - Endo Steris Tryout | Meeting - Endo Steris CO2 insufflator and irrigation pump Tryout
1059778 | 201893 | FC | CMPLT | CMPLT | 2024-02-22 07:32 | 2024-02-22 07:33 | QC Checklist | (TC 02/22/2024 07:32:23 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
67957 | 201894 | CM | CMPLT | CMPLT | 2024-02-22 07:44 | 2024-02-26 14:21 | Ordering hood | (TC 02/22/2024 07:46:48 AM, Baghirzade, Nara) SC 1003174026 (TC 02/26/2024 02:20:38 PM, Baghirzade, Nara) PO 7202366522
1046952 | 201895 | CM | CMPLT | CMPLT | 2024-02-22 07:50 | 2024-02-27 13:03 | MSH ICU SPARQ NOT BOOTING | (TC 02/22/2024 09:09:00 AM, Whittington, Scott) MSH ICU SPARQ NOT BOOTING (TC 02/26/2024 02:43:52 PM, Whittington, Scott) MSH ICU SPARQ NOT BOOTING site 72633784 REF# 0123179199 ALEX PN# 4535 617 59491 BUNGEE CORD PROBE CABLE HOLDER (TC 02/27/2024 11:28:48 AM, Fan, Tak) MSH ICU SPARQ NOT BOOTING TAK COPIED VER 3.0 SW AND PUT IN SYSTEM (ORIGINAL 3.01) BOOTED OK (TC 02/22/2024 08:02:00 PM, Whittington, Scott) MSH ICU SPARQ NOT BOOTING site 72633784 REF# 0123179199 ALEX PN# 4535 617 59491 BUNGEE CORD PROBE CABLE HOLDER (TC 02/27/2024 01:06:24 PM, Whittington, Scott) MSH ICU SPARQ NOT BOOTING site 72633784 REF# 0123179199 ALEX PN# 4535 617 59491 BUNGEE CORD PROBE CABLE HOLDER MARKED NO CONTRACT | MSH ICU SPARQ NOT BOOTING site 72633784 REF# 0123179199 ALEX PN# 4535 617 59491 BUNGEE CORD PROBE CABLE HOLDER MARKED NO CONTRACT
53081 | 201896 | CM | CMPLT | CMPLT | 2024-02-22 08:21 | 2024-02-22 08:23 | Light not turning on | (TC 02/22/2024 08:21:20 AM, Rhuland, Courtney) Checked bulb and noticed bulb was inserted incorrectly Bulb installed was incorrect voltage/wattage Replaced bulb with correct one Functional check - OK
63859 | 201897 | CM | CMPLT | CMPLT | 2024-02-22 08:28 | 2024-02-22 08:30 | Fallen off wall | (TC 02/22/2024 08:28:27 AM, Rhuland, Courtney) Removed bracket from basket Reattached sphyg to bracket and reshaped bend with pliers Reattached to wall Functional check - OK
| 201898 | CM | CMPLT | CMPLT | 2024-02-22 08:44 | 2024-02-26 09:02 | Bunnell HFV spare parts | (TC 02/22/2024 08:50:49 AM, Rendon Lopez , Elizabeth) Ordering spare parts for Bunnell HFV and patient boxes Received
66500 | 201899 | CM | CMPLT | CMPLT | 2024-02-22 08:53 | 2024-02-22 08:56 | Unable to associate to Fetalink | (TC 02/21/2024 08:54:56 AM, Rendon Lopez , Elizabeth) In house IT checked it, unable to associate monitor. Checked connections on fetal monitor, still unable to associate, sent issue to cerner to resolve
1054372 | 201900 | CM | CMPLT | CMPLT | 2024-02-22 09:02 | 2024-02-22 09:17 | Broken temp probe | (TC 02/22/2024 09:16:38 AM, Rhuland, Courtney) Temp probe was destroyed Replaced probe and probe well Functional check - OK
66470 | 201901 | CM | CMPLT | CMPLT | 2024-02-22 10:00 | 2024-02-22 15:35 | not passing device check, making noise | Collected logs and sent them to Draeger technical support. Log VN500 ASFH-0006 Log review for Error Possible causes Remedies Based on log entries I suggest 1 000768.32839 Unable to build up pressure in system during an IST. Possible causes 1. Leakage in expiratory valve. 2. Hose connection between inspiratory and expiratory port is missing. 3. M1.3 not calibrated. 4. Calibration valve V6 or V8 faulty. Remedies Connect the device to the HIT software and start the test mode. In the HIT software, select Tools menu and Self-test submenu, and display the ISTs of V500. Perform the "Airway Pressure Measurement" IST. In the event of an error, perform IST "ZeroingAllFlowSensors" and then repeat IST "Airway Pressure Measurement". Build up pressure in the system for leakage test. To do so, open the product page "PI V500 / PI Exp. valve". Switch expiratory valve test to "ON" position. Adjust an inspiratory flow of approx. 20 L/min and an expiratory valve pressure of approx. 20 mbar. Check pressure display of inspiratory and expiratory pressure sensors. Test for leakage as necessary. When a V500 or VN500 when experience issues where you need to verify the M1.3 is working as expected, you will want to run the following tests. 1. Set ventilator IP address – 169.254.0.1 2. Set ventilator Subnet Mask – 255.255.255.0 3. Disable Laptop Wifi 4. If using local laptop firewall disable the firewall 5. Set laptop IP address – 169.254.0.2 6. Set laptop Subnet Mask – 255.255.255.0 7. On Ventilator a. System Set Up i. System 1. Service a. Enter Password i. V500 – 8270 ii. VN500 – 8272 b. Enter touch key c. Service Utilities i. Test Mode Key – On (will highlight when laptop is connected). 8. On Laptop a. Service Connect Desktop i. Enter Evita V500 or Babylog VN500 ii. Select and Start HIT software iii. Type in the IP Address of the ventilator 169.254.0.1 and connect. 9. At this point the Test Mode Key should highlight if the connection is successful. 1. Device Pages a. M1.3 Gas Metering Module i. Valves & Sensors M1.3 1. M1.3 Control a. In this page you can set i. Inspiratory flow – 0 to 180 L/min ii. O2 Concentration – 21 to 100% b. You can read out the following: i. Valve Current V1 (Air) – A ii. Valve Current V2 (O2) – A iii. Valve Current V3 (LPSV) – A iv. Flow S1 (Air) – L/min v. Flow S2 (O2) – L/min vi. Flow S3 (LPSV) – L/min vii. Inspiratory Flow – L/min viii. S7 Tank Pressure – mBar ix. O2 Monitoring – Vol % x. O2 Calculated – decimal xi. Inspiratory Pressure – mBar 2. To test the M1.3 always use either 21% or 100% so you know what you should be reading at S1 and S2. 3. As an example: a. Set Inspiratory Flow to 10 L/min b. O2 Concentration to 21% 4. You should read the following: a. S1 = approximately 11 L/min (Set inspiratory flow, plus approximately 1 L/min for flow to Pato and Purge flow to expiratory pressure sensor. b. S2 = 0 L/min c. S3 = 10 L/min d. Inspiratory Flow = 10 L/min e. O2 Monitoring = 21% f. O2 Calculated = 21.0 g. Tank Pressure = XX mBar 5. You can change set inspiratory flow from 1 to 180 L/min. a. S1 = set inspiratory flow plus approximately 1 L/min b. S2 = 0 L/min c. S3 = Set Inspiratory Flow d. Inspiratory Flow = Set Inspiratory Flow e. O2 Monitoring = 21% f. O2 Calculated = 21.0 6. Now change O2 Concentration to 100% and set Inspiratory Flow to 10 L/min 7. You should read the following: a. S2 = approximately 11 L/min (Set inspiratory flow, plus approximately 1 L/min for flow to Pato and Purge flow to expiratory pressure sensor. b. S1 = 0 L/min c. S3 = 10 L/min d. Inspiratory Flow = 10 L/min e. O2 Monitoring = 100% f. O2 Calculated = 100.0 g. Tank Pressure = XX mBar 8. You can change set inspiratory flow from 1 to 180 L/min. a. S2 = set inspiratory flow plus approximately 1 L/min b. S1 = 0 L/min c. S3 = Set Inspiratory Flow d. Inspiratory Flow = Set Inspiratory Flow e. O2 Monitoring = 100% f. O2 Calculated = 100.0 9. You can also verify Inspiratory Flow and O2 Concentration via an external meter by either: a. Connecting the meter to the output of the M1.3 Module. b. Connecting to the Inspiratory Nozzle i. In this case before you can measure at this point you will have to close the Safety Valve. 1. Device Pages a. M15 Valve Drives i. M15.2 Solenoid – Set from Open to Closed. 10. By doing this, you can check for correct flows, pressures all the way through the M1.3 and out the Inspiratory Nozzle. From here you can determine if you have leaking Proportional Valves, LPSV, leakages at the tank, Safety Valve, Regulators in the Nebulizer, or Inspiratory Nozzle. Replaced flow sensor cable. Did the calibrations and tests as per Draeger. Run unit for 2 weeks. Did device check every morning during 2 weeks. Functional. Informed Rex. Returned back to service.
1070107 | 201902 | CM | CMPLT | CMPLT | 2024-02-22 12:53 | 2024-02-22 13:00 | Incoming inspection | (TC 02/22/2024 12:57:40 PM, Rhuland, Courtney) Unboxed and assembled flowstand and flowstar system Performed test flow and accuracy verification All tests passed Brought back to MKUWC (TC 02/22/2024 01:49:47 PM, Rhuland, Courtney) Assembled commode and set up device in unit Showed physicians the basics of the device Removed old device
1067012 | 201903 | CM | CMPLT | CMPLT | 2024-02-22 13:22 | 2024-02-22 13:25 | V1-V3 broken | (TC 02/22/2024 01:23:01 PM, Scero, Tomlyn) ECG tech called and asked for a replacement V1-V3 cable. Assembled the cable and delivered it to the ECG department
1070110 | 201904 | ACC | CMPLT | CMPLT | 2024-02-22 13:26 | 2024-02-22 13:28 | Incoming Inpspection | (TC 02/22/2024 01:25:47 PM, Patel, Rutvik) Unit was brought in by Endo department in conjuction with Steris Rep David Murphy (with no PO issued) with intend to put into clinical use. Assinged asset tag and recorded into TMS. Completed functional check and EST. Tested ok. Affixed asset tag and inspection label. Assisted rep to install the unit in Procedure room.
502088 | 201905 | CM | CMPLT | CMPLT | 2024-02-22 13:46 | 2024-02-22 13:46 | Crystal Hydration Detector 2 | (TC 02/22/2024 01:45:36 PM, Paz, Ramon (Tino)) Performed an image evaluation and concluded that Detector 2 is hydrated causing image artifacts. System is still usable. Inquiring on a replacement detector.
1051031 | 201906 | CM | CMPLT | CMPLT | 2024-02-22 14:25 | 2024-02-22 14:25 | Touch screen not working mid exam | (TC 02/21/2024 02:24:42 PM, Fan, Tak) Checked system, no problem with touch screen found. Checked usb cable for panel. Clear test exam that was in the "No Archive". Changed time due to daylight savings.
| 201907 | OC | CMPLT | CMPLT | 2024-02-22 14:44 | 2024-02-23 15:31 | On-call Feb 16-22, 2024 | (TC 02/17/2024 02:44:06 PM, Cortez, John) Received call from ER saying Y51 monitor was not appearing on central monitor. Logged into biomed ssupport PC and confirmed issue. Suggested to Allison (nurse) to try reseating MSL cable. Also suggested to swap X3s with one in empty room if possible. She said she would try it and get back to me if unsuccessful. Monitored Surveillance station and saw Y51 came back up. Issue resolved and did not need to come in. (TC 02/22/2024 03:22:19 PM, Cortez, John) Wrong number, 10N asking about issue with microwave/power outage
1064593 | 201908 | CM | CMPLT | CMPLT | 2024-02-22 17:25 | 2024-02-22 17:26 | Table doesn't move when locks taken off | (TC 02/22/2024 05:25:11 PM, Rendon Lopez , Elizabeth) Confirmed issue, one caster is rusted and not moving. Advised to call in Medaris
1050672 | 201909 | CM | CMPLT | CMPLT | 2024-02-23 08:48 | 2024-02-23 10:44 | Screen flickering | (TC 02/23/2024 10:41:45 AM, Canjimavadivel, Kamallesh ) Confirmed issue when device turned on. Removed back cover of display and reseated all cables. Issue resolved, all functions ok.
1056214 | 201910 | CM | CMPLT | CMPLT | 2024-02-23 09:30 | 2024-02-23 09:31 | "temp probe cover stuck" | (TC 02/23/2024 09:29:55 AM, Cortez, John) Replaced temp probe well with spare (TC 02/23/2024 09:30:34 AM, Cortez, John) All parameters tested ok
65883 | 201911 | CM | CMPLT | CMPLT | 2024-02-23 09:56 | 2024-02-23 09:56 | Broken handle | (TC 02/23/2024 09:55:57 AM, Rhuland, Courtney) Replaced cracked handle Check-in complete All tests passed
| 201912 | CM | CMPLT | CMPLT | 2024-02-23 09:57 | 2024-10-08 15:42 | BD Alaris pump parts order, for biomed stock | (TC 02/23/2024 10:24:13 AM, Cortez, John) Ordered spare parts for BD pump upgrades in March (TC 02/27/2024 11:17:59 AM, Cortez, John) Most of items were delivered. 10 sensors still incoming (TC 02/28/2024 01:55:17 PM, Cortez, John) Inquired about price for rear cases, ordered 5 (TC 03/01/2024 02:32:28 PM, Cortez, John) Ordered x100 syringe rear cases and left handles (TC 03/01/2024 04:11:41 PM, Cortez, John) Rest of IUIs arrived (TC 07/22/2024 09:57:57 AM, Cortez, John) Followed up with vendor, x41 rear cases outstanding, x100 left handles shipped They will be shipping rear cases soon (TC 10/08/2024 03:41:48 PM, Cortez, John) rest of rear cases (x25) arrived
64188 | 201913 | CM | CMPLT | CMPLT | 2024-02-23 10:32 | 2024-02-23 10:33 | Removal of MX800 | (TC 02/23/2024 10:32:37 AM, Baghirzade, Nara) Removed MX800, rack, arm, and mount from the area behind the main central at 17 Murray. No X2 found on the rack.
67148 | 201914 | CM | CMPLT | CMPLT | 2024-02-23 10:34 | 2024-02-23 10:40 | Temperature reading not working | (TC 02/23/2024 10:37:56 AM, Canjimavadivel, Kamallesh ) Temp reading not picking up. Checked with calibration plug, still not working. Opened up and inspected temp module board. No visible issues found. Reseated ribbon cable, did not resolve issue. Device no longer supported and spare temp module not available in our stock parts. Device to be retired and kept for parts.
68904 | 201915 | CM | CMPLT | CMPLT | 2024-02-23 10:44 | 2024-02-26 14:29 | Not charging | (TC 02/23/2024 10:43:54 AM, Rhuland, Courtney) Battery completely depleted and would not turn on Replaced battery but device was not detecting AC power Replaced power cord Conditioning battery (TC 02/26/2024 02:28:36 PM, Rhuland, Courtney) Verified battery conditioning Hardware tests passed Functional check - OK Electrical safety - OK
501133 | 201916 | CM | CMPLT | CMPLT | 2024-02-23 11:51 | 2024-02-27 13:07 | MSH 5 US RM 5 WENT TO SLEEP DURING EXAM, WILL NOT BOOT UP,GOES TO SERVICE LOG ON | (TC 02/23/2024 12:00:19 PM, Whittington, Scott) MSH 5 US RM 5 WENT TO SLEEP DURING EXAM, WILL NOT BOOT UP,GOES TO SERVICE LOG ON SITE 553814 REF# PAUL 0123169775 (TC 02/27/2024 01:07:10 PM, Whittington, Scott) MSH 5 US RM 5 WENT TO SLEEP DURING EXAM, WILL NOT BOOT UP,GOES TO SERVICE LOG ON SITE 553814 REF# PAUL 0123169775 | MSH 5 US RM 5 WENT TO SLEEP DURING EXAM, WILL NOT BOOT UP,GOES TO SERVICE LOG ON SITE 553814 REF# PAUL 0123169775 _____________________________________________ From: Whittington, Scott Sent: February 23, 2024 12:00 PM To: Thompson, Jillian <jillian.thompson@philips.com> Cc: Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca> Subject: MSH 5 US RM 5 WENT TO SLEEP DURING EXAM, WILL NOT BOOT UP,GOES TO SERVICE LOG ON SITE 553814 REF# PAUL 0123169775 Hi Jillian MSH 5 US RM 5 WENT TO SLEEP DURING EXAM, WILL NOT BOOT UP,GOES TO SERVICE LOG ON SITE 553814 REF# PAUL 0123169775 Thanks, Scott 416-402-2462
1065848 | 201917 | CM | CMPLT | CMPLT | 2024-02-23 12:46 | 2024-04-23 07:37 | Channel error | (TC 02/23/2024 12:47:03 PM, Rhuland, Courtney) Sensor broken low error in error log Failed calibration Replaced patient side sensor Failed check-in Failed calibration Investigating if batch of pressure sensors are defective
66701 | 201918 | CM | CMPLT | CMPLT | 2024-02-23 13:43 | 2024-02-23 13:44 | Unable to associate to fetalink | (TC 02/23/2024 01:43:12 PM, Rendon Lopez , Elizabeth) Checked monitor confirmed issue. Request nurse to place IT ticket with cerner
66474 | 201919 | CM | CMPLT | CMPLT | 2024-02-23 13:44 | 2024-03-01 08:28 | Failing patient circuit leakage test | (TC 02/23/2024 01:44:45 PM, Rendon Lopez , Elizabeth) Tried 4 new circuits and pots, replaced exp valve, insp valve, same issue. Refer to WO 202007
1059524 | 201920 | PM | CMPLT | CMPLT | 2024-02-23 13:52 | 2024-02-23 13:58 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 02/21/2024 01:52:03 PM, Rendon Lopez , Elizabeth) Inspected, cleaned, PM parts replaced (TC 02/22/2024 01:52:03 PM, Rendon Lopez , Elizabeth) Calibration done. Performance verification done (TC 02/23/2024 01:52:03 PM, Rendon Lopez , Elizabeth) PM tasks completed, returned to RT NICU
1061316 | 201921 | CM | CMPLT | CMPLT | 2024-02-23 13:55 | 2024-02-23 14:14 | CD burning
| 201922 | RDS | CMPLT | CMPLT | 2024-02-23 13:55 | 2024-02-23 13:58 | Weekly Rounds: 11 North | (TC 02/23/2024 01:55:32 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 201923 | RDS | CMPLT | CMPLT | 2024-02-23 13:58 | 2024-02-23 13:59 | Weekly Rounds: 11 South | (TC 02/23/2024 01:58:05 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 201924 | RDS | CMPLT | CMPLT | 2024-02-23 13:59 | 2024-02-23 14:00 | Weekly Rounds: NICU & RESUS | (TC 02/23/2024 01:59:20 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
| 201925 | CM | CMPLT | CMPLT | 2024-02-23 14:02 | 2024-02-23 14:02 | Meeting with Dr. Deshpande, Laura, and Kamran | (TC 02/23/2024 02:01:37 PM, Baghirzade, Nara) Meeting regarding the ICM+ software download and materials required for the study
1047096 | 201926 | CM | CMPLT | CMPLT | 2024-02-23 14:03 | 2024-02-23 14:09 | Does not turn on | (TC 02/23/2024 02:06:02 PM, Rhuland, Courtney) Device would not turn on Checked power cord continuity and power cord was fine Checked fuses in power supply - OK Keypad found to be faulty Replaced front case Replaced male IUI connector Functional check - OK
| 201927 | CM | CMPLT | CMPLT | 2024-02-23 16:25 | 2024-02-23 16:50 | OR PM Meeting | OR PM Meeting with AA, JH, ER.
| 201928 | CM | CANCL | DUP | 2024-02-23 16:52 |  | Scope Guide Cable broken
| 201929 | CM | CMPLT | CMPLT | 2024-02-23 20:14 | 2024-02-26 11:50 | MSH OPG SPP VOLUSON CONNECTION TO HSC & CONF SETUP | (TC 02/23/2024 08:14:05 PM, Whittington, Scott) MSH OPG SPP VOLUSON CONNECTION TO HSC (TC 02/23/2024 08:14:05 PM, Whittington, Scott) MSH OPG SPP VOLUSON CONNECTION TO HSC | MSH OPG SPP VOLUSON CONNECTION TO HSC & CONF SETUP
32510 | 201930 | CM | CMPLT | CMPLT | 2024-02-23 20:30 | 2024-02-19 08:32 | MSH DR1 TABLE NOT OPERATING | (TC 02/16/2024 11:30:00 PM, Whittington, Scott) MSH DR1 TABLE NOT OPERATING | MSH DR1 TABLE NOT OPERATING
1070114 | 201931 | ACC | CMPLT | CMPLT | 2024-02-26 08:26 | 2024-03-12 11:48 | MSH 5 US CANON PRISM PROBE 99C23X4146 PVT-681MVL 3D TV ACCEPTANCE | (TC 02/26/2024 10:49:42 AM, Whittington, Scott) PROBE 99C23X4146 PVT-681MVL 3D TV acceptance (TC 02/26/2024 10:50:01 AM, Dolinar, John) PROBE 99C23X4146 PVT-681MVL 3D TV ACCEPTANCE | PROBE 99C23X4146 PVT-681MVL 3D TV
| 201932 | RDS | CMPLT | CMPLT | 2024-02-26 08:29 | 2024-02-26 08:29 | Weekly check for 14S and 17M | 16N is moving to 17M today
1062603 | 201933 | CM | CMPLT | CMPLT | 2024-02-26 08:38 | 2024-04-02 13:33 | needs a new replacement unit | (TC 02/28/2024 01:01:35 PM, Huang, Jessica) SN 171159150 was replaced by SN 230654631 in Feb 2024
| 201934 | RDS | CMPLT | CMPLT | 2024-02-26 08:55 | 2024-02-26 08:56 | Weekly Checks For PACU | (TC 02/26/2024 08:56:07 AM, Reid, Emmerson ) Rounds complete, also went around OR
| 201935 | RDS | CMPLT | CMPLT | 2024-02-26 09:00 | 2024-02-27 09:55 | Weekly checks in CCU | (TC 02/27/2024 09:54:33 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed
63347 | 201936 | CM | CMPLT | CMPLT | 2024-02-26 09:04 | 2024-02-26 09:19 | Frayed lead wires | (TC 02/26/2024 09:04:37 AM, Baghirzade, Nara) Replaced lead wires: V4-V6 and V7. Connected to ECG simulator, all functional. Inspected all other leads; in good condition.
1060002 | 201937 | CM | CMPLT | CMPLT | 2024-02-26 09:11 | 2024-02-26 10:58 | V4-V6 not working | (TC 02/26/2024 09:11:10 AM, Canjimavadivel, Kamallesh ) Connected to simulator and was able to replciate issue. Replaced V4-V6 chest lead, tested ok. Replaced damaged keyboard cover
1055042 | 201938 | CM | CMPLT | CMPLT | 2024-02-26 09:36 | 2024-02-26 09:36 | Battery replacement | (TC 02/26/2024 09:36:04 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
65129 | 201939 | CM | CMPLT | CMPLT | 2024-02-26 09:44 | 2024-03-01 15:08 | Ordering Dose Cords | (TC 02/26/2024 09:44:25 AM, Reid, Emmerson ) Looked at 3 CADD pumps and all needed their dose cord replaced. Reached out to the vendor for a quote.
1057279 | 201940 | CM | CMPLT | CMPLT | 2024-02-26 10:02 | 2024-02-26 10:16 | Acceptance and configuration of new Root | (TC 02/26/2024 10:14:41 AM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055042 | 201941 | CM | CMPLT | CMPLT | 2024-02-26 10:25 | 2024-02-26 10:29 | Repurposed spare standalone Radical 7 | (TC 02/26/2024 10:27:11 AM, Baghirzade, Nara) Repurposed a spare standalone Radical 7 for Mother & Baby unit (old location was 12S). Update all asset information on TMS. Completed battery upgrade. Completed functional testing of SpO2: Set 100% = Measured 99% Set 98% = Measured 98% Set 90% = Measured 90% HR set to 80 BPM = measured 80 BPM Followed up and confirmed with Jessica.
| 201942 | CM | CMPLT | CMPLT | 2024-02-26 11:09 | 2024-02-26 11:09 | meeting | biomed meeitng
| 201943 | CM | CMPLT | CMPLT | 2024-02-26 11:10 | 2024-02-26 11:10 | worked on Masimo | worked on Masimo
1053214 | 201944 | CM | CMPLT | CMPLT | 2024-02-26 12:57 | 2024-02-26 13:01 | CMOS battery replacement and EST | (TC 02/26/2024 12:56:53 PM, Baghirzade, Nara) Replaced CMOS battery . Updated time (was off just by half a minute); booted the CNS down and then back up, time and date retained. Electrical safety results: power cord resistance = 53mohms current leakage = 47.0 uA Returned CNS machine to procedure room in NICU. Update Carol, Lesley, and Amy (OUCH LAB, YORKU).
| 201945 | CM | CMPLT | CMPLT | 2024-02-26 13:02 | 2024-02-26 13:02 | BD Upgrade Workflow meeting | (TC 02/26/2024 01:02:00 PM, Baghirzade, Nara) BD Upgrade Workflow meeting
| 201946 | CM | CMPLT | CMPLT | 2024-02-26 13:03 | 2024-02-29 13:08 | 11S Stryker beds missing headboards | (TC 02/26/2024 01:02:44 PM, Baghirzade, Nara) Sri sent a message at the end of the day on Friday reporting 23 headboards missing in 11S. Today (Monday. Feb 26th) went on the floor to follow up with Sri, she is not sure why or how they went missing. Went to every single room in 11S to confirm the number of missing headboards and which rooms are missing them. Confirmed 3 headboards missing in rooms: 1105B 1108B 1111C Went to the storage spaces on 11S, staff lounge, and checked behind all curtains in the rooms, could not find any of the missing headboards. Followed up with Sri with my findings and ask her to inquire with the SAs as it must have been removed during cleaning. Followed up with Naz and waiting to follow up with Becky to see if they know where else I can look for the headboards. (TC 02/29/2024 01:06:35 PM, Baghirzade, Nara) Becky was able to locate 2 headboards that were placed in the P3A storage room. Dropped it off with Sri and let her know that if the missing headboard doesn't turn up or if more headboards go missing, 11S unit will have to purchase the headboards from Stryker.
| 201947 | CM | CMPLT | CMPLT | 2024-02-26 14:01 | 2024-02-26 14:01 | biomed meeting | biomed meeting
| 201948 | CM | CMPLT | CMPLT | 2024-02-26 14:12 | 2024-02-26 14:12 | Biomed huddle | (TC 02/26/2024 02:11:38 PM, Baghirzade, Nara) Biomed huddle
| 201949 | CM | CMPLT | CMPLT | 2024-02-26 14:12 | 2024-09-10 14:09 | ICM+ Study Setup | (TC 02/26/2024 02:12:25 PM, Baghirzade, Nara) Dr. Deshpande forwarded the registration key for ICM+ software. Registered the key on her laptop. Connected to MX800 with MIB RS232 port; working with Kamran on simulating a patient. (TC 02/27/2024 09:30:59 AM, Baghirzade, Nara) Successfully configured the data communication between the MX800 and ICM+ software; we were able to get waveforms for ECG, ABP, and RESP. Some glitches in the software and usability issues which can be addressed when clinical training occurs. Kamran followed up with Dr. Deshpande regarding the successful connection. Next step is for her to bring the INVOS monitor to integrate with the ICM+ software. (TC 02/27/2024 11:49:47 AM, Baghirzade, Nara) Set up with Kamran, Laura, and Dr. Deshpande in NICU. (TC 03/14/2024 02:24:12 PM, Baghirzade, Nara) Tested INVOS and Philips connections to the laptop; created signal sources and configuration profile. Walked through step-by-step process for Kamran's documentation. Labelled on the laptop the USB ports for connections and labelled both cables. (TC 03/19/2024 08:53:52 AM, Baghirzade, Nara) Instructions for connecting the devices to the laptop and creating the configuration for the studies (also attached as a Word document to this WO): 1. Connect Philips cable to the Philips labelled USB port on the laptop. Connect the other end to the back of the MX800 MIB RS232 port. 2. Connect the INVOS cable to the Invos labelled USB port on the laptop. Connect the other end to the Invos device. 3. Launch ICM+ software on the laptop. 4. Select new patient and then select "generic" project name. 5. Input patient information on the new data acquisition session form. 6. At the bottom of the new data acquisition session form, click on "signal sources". 7. In the "configured devices with proprietary digital output protocols" section, click "add". 8. Click on "device type" and select "Philips Intellivue" 9. Confirm that the COM port is: COM1, Baud rate: 115200, and interface type is: RS232. 10. Click on Click on "device type" and select "INVOS" 11. Confirm that the COM port is: COM4, Baud rate: 9600, and interface type is: RS232. 12. In the "configured modalities to be collected" section, click on "add". 13. Configure the device to the parameters the study requires. For example: -Select device: Philips -Click on "select wave" and choose the waveforms corresponding to the parameter the study wants to measure. -Click ok to finish setting up the parameters that are to be acquired from Philips. -Select device: INVOS -Click on "select wave" and choose the waveforms/parameters corresponding to the parameter the study wants to measure. -Click ok to finish setting up the parameters that are to be acquired from INVOS. 14. At the bottom of the sample configuration dialog, click "save". (It will automatically route to the Configs folder). 15. Name the signal source and save it. 16. At the bottom of the sample configuration dialog, click "ok". 17. This will take you back to the new data acquisition session form (patient information window). Near the bottom of the window, locate "data acquisition/analysis configuration profile" and click on the 3 little boxes. 18. Locate the configs folder and select the configuration that was created. 19. Click ok at the bottom of the window and the study will start. (TC 03/19/2024 09:18:11 AM, Baghirzade, Nara) Instructions for connecting the devices to the laptop and creating the configuration for the studies (also attached as a Word document to this WO): 1. Connect Philips cable to the Philips labelled USB port on the laptop. Connect the other end to the back of the MX800 MIB RS232 port. 2. Connect the INVOS cable to the Invos labelled USB port on the laptop. Connect the other end to the serial port ( 9 pins connector in the bottom of docking station ) in INVOS device. 3. Launch ICM+ software on the laptop. 4. Select new patient and then select "generic" project name. 5. Input patient information on the new data acquisition session form. 6. At the bottom of the new data acquisition session form, click on "signal sources". 7. In the "configured devices with proprietary digital output protocols" section, click "add". 8. Click on "device type" and select "Philips Intellivue" 9. Confirm that the COM port is: COM1, Baud rate: 115200, and interface type is: RS232. 10. Click on “communication test” to test the connection. Ensure the connection test has passed. 11. Click on Click on "device type" and select "INVOS" 12. Confirm that the COM port is: COM4, Baud rate: 9600, and interface type is: RS232. 13. Click on “communication test” to test the connection. Ensure the connection test has passed. 14. In the "configured modalities to be collected" section, click on "add". 15. Configure the device to the parameters the study requires. For example: -Select device: Philips -Click on "select wave" and choose the waveforms corresponding to the parameter the study wants to measure. -Click ok to finish setting up the parameters that are to be acquired from Philips. -Select device: INVOS -Click on "select wave" and choose the waveforms/parameters corresponding to the parameter the study wants to measure. -Click ok to finish setting up the parameters that are to be acquired from INVOS. 16. At the bottom of the sample configuration dialog, click "save". (It will automatically route to the Configs folder). 17. Name the signal source and save it. 18. At the bottom of the sample configuration dialog, click "ok". 19. This will take you back to the new data acquisition session form (patient information window). Near the bottom of the window, locate "data acquisition/analysis configuration profile" and click on the 3 little boxes. 20. Locate the configs folder and select the configuration that was created. 21. Click ok at the bottom of the window and the study will start. (TC 03/19/2024 10:34:27 AM, Baghirzade, Nara) Met with Thaiani to go over all the steps to start and configure a study. Went over the instructions and had Thaiani connect the cables and configure the signal sources and create a configuration file. She took notes of all steps. (TC 04/09/2024 01:04:01 PM, Rhuland, Courtney) Training on the ICM+ software (TC 04/09/2024 01:05:04 PM, Rhuland, Courtney) set up testing of the system with simulator (TC 04/24/2024 12:51:42 PM, Rhuland, Courtney) Set up 2 more laptops for the study (TC 04/25/2024 01:37:01 PM, Rhuland, Courtney) Labeled laptop and cables Cable tied ethernet connector to USB connector (TC 04/26/2024 01:22:53 PM, Rhuland, Courtney) Confirmed setup with Philips monitor and INVOS monitor with new laptop (TC 04/29/2024 01:44:59 PM, Rhuland, Courtney) Set up final laptop with cables Labeled all cables During testing, INVOS monitor would not connect tried troubleshooting in unit with research team Tested with different INVOS monitor and connection passed Brought monitor and laptop down to biomed and found in service mode that the serial port had been turned off Turned port on and ICM+ was able to connect with INVOS monitor Returned monitor and laptop to research team (TC 04/30/2024 01:05:39 PM, Rhuland, Courtney) Labeled all cables with associated laptops Created specific profiles based on laptops so connection issues do not arise (TC 05/28/2024 10:55:55 AM, Baghirzade, Nara) Laptop 1 and Laptop 3 were both not connecting to the INVOS. INVOS ports not working (not opening); disabled and enabled the ports, restarted the laptop after disabling and then enabling, and disconnected the cable from INVOS; also adjusted the workflow for thaiani: prepare the devices and open the ICM+ software, once in the signal source part, then connect the USB cables to the laptop. (TC 05/30/2024 11:54:40 AM, Baghirzade, Nara) ICM was not communicating with the INVOS. Adjusted the workflow again for the research staff and confirmed with Thaiani. Also adjusted the port settings to hardware as per the manual for Invos. Adjusted the password for one of the devices. (TC 05/31/2024 01:29:51 PM, Baghirzade, Nara) Researching different cables and options for the INVOS connection to the laptop. Followed up with Kamran and confirmed the type of cable that the research team should order. Found an option that might work but will need to measure the connector on the INVOS first before purchasing. Will reconvene with Thaiani regarding this matter. (TC 06/03/2024 12:05:00 PM, Baghirzade, Nara) Invos was problematic over the weekend; RTs were unable to connect the invos to the software. In first room, was able to connect the invos without any issues by connecting it last on the device end. Second room, the gender changer adapter was on a 7100 invos which was causing a connectivity issue. Gender changer adapter was tightly screwed in so had to use a flathead to remove it. Connected the second invos to the laptop without any issue. Went over the connectivity steps with Thaiani and Laura. Third Invos was the older model; Invos 5100. The gender changer adapter is meant for this model only. Connected it to the laptop but was having an issue connecting to ICM. Looked into the output configuration on the invos and found that "PC link" needs to be selected and output must be set to the first one (three options in total). Demonstrated this to Thaiani and Laura, and had them configure as well for practice. Followed up with Dr. Deshpande and Laura let us know that Dr. Deshpande has already ordered the startech adapter as suggested by Dr. Smielewski. Followed up with Kamran regarding this and updated Courtney as well. (TC 06/07/2024 02:20:56 PM, Rhuland, Courtney) Troubleshooting INVOS monitor to find out why it disconnects intermittently. Connection procedure is very specific in terms of connecting order. Issue is still present with the new cables that were ordered. (TC 09/10/2024 02:09:01 PM, Baghirzade, Nara) Followed up Thaiani and Phoebe regarding any issues with ICM+; no issues reported. Followed up with Kamran to close this WO.
65264 | 201950 | CM | CMPLT | CMPLT | 2024-02-26 14:24 | 2024-02-26 14:24 | broken control knob | replaced broken rotary knob. tested. Functional.
1048472 | 201951 | CM | CMPLT | CMPLT | 2024-02-26 14:29 | 2024-03-04 11:13 | MSH R59 Tube cover melted. | (TC 02/26/2024 02:30:32 PM, Dolinar, John) MSH R59 Tube cover melted. Ordered parts (TC 02/29/2024 01:39:42 PM, Whittington, Scott) MSH R59 Tube cover melted. Ordered parts PARTS ARRIVED TGH MOVED TO MSH (TC 03/04/2024 11:15:00 AM, Dolinar, John) Parts replaced. | MSH R59 Tube cover melted. Ordered parts . Replaced cover
69018 | 201952 | CM | CMPLT | CMPLT | 2024-02-26 14:44 | 2024-02-27 13:16 | Bent power cord | (TC 02/26/2024 02:45:40 PM, Rhuland, Courtney) Replaced bent power cord Letting device charge overnight (TC 02/27/2024 01:15:45 PM, Rhuland, Courtney) Check in complete All tests passed Refer to ASM for details
| 201953 | CM | CMPLT | CMPLT | 2024-02-26 15:04 | 2024-02-27 09:58 | QC check not passing on AVOXimeter 1000E | (TC 02/26/2024 09:55:20 AM, Scero, Tomlyn) Sandra brought unit down and asked if I could disassemble and clean the sensor as the QC test isn't passing. Downloaded and read the service manual online to learn the proper cleaning proceedure. Disassembled and cleaned the optical unit. Tested the QC after reassembling and passed. Reset the date and time. Returned the unit to the Cath Lab.
1070135 | 201954 | CM | CMPLT | CMPLT | 2024-02-26 15:08 | 2024-02-26 16:02 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 80 mOhm; 77 uA. Tested ok. Installed unit in endoscopy.
1070136 | 201955 | CM | CMPLT | CMPLT | 2024-02-26 15:17 | 2024-02-26 15:17 | Incoming Inpsection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 90 mOhm; 78 uA. Tested ok. Installed unit in endoscopy.
1070137 | 201956 | CM | CMPLT | CMPLT | 2024-02-26 15:20 | 2024-04-11 15:20 | Incoming Inspection | Feb 26 - Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 100 mOhm; 76 uA. Tested ok. Apr 11 - Delivered unit to Endo and is being stored on a sheft.
1070138 | 201957 | CM | CMPLT | CMPLT | 2024-02-26 15:25 | 2024-02-26 15:26 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 120 mOhm; 77 uA. Tested ok. Installed unit in endoscopy.
1070139 | 201958 | CM | CMPLT | CMPLT | 2024-02-26 15:38 | 2024-02-26 15:38 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 90 mOhm; 78 uA. Tested ok. Installed unit in endoscopy.
1070140 | 201959 | CM | CMPLT | CMPLT | 2024-02-26 15:45 | 2024-02-26 15:45 | Incoming Inspection | Feb 26 - Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 160 mOhm; 78 uA. Tested ok. Apr 11 - Delivered unit to Endo and is being stored on a sheft.
502048 | 201960 | CM | CMPLT | CMPLT | 2024-02-26 16:40 | 2024-02-26 16:40 | Move unit back into Rm2 Hall C |  | Rewired Tested OK
1063331 | 201961 | CM | CMPLT | CMPLT | 2024-02-26 16:42 | 2024-02-26 16:42 | Move unit into Rm3 Hall C |  | Rewired Tested Ok
| 201962 | CM | CMPLT | CMPLT | 2024-02-26 16:43 | 2024-02-26 16:43 | Site coverage
67952 | 201963 | CM | CMPLT | CMPLT | 2024-02-27 07:31 | 2024-03-01 09:30 | Missing parts | (TC 02/27/2024 08:10:39 AM, Rhuland, Courtney) Came down missing drawer tray Noticed hood was cracked Placed order for new hood (TC 02/29/2024 02:18:23 PM, Rhuland, Courtney) Replaced hood (TC 03/01/2024 09:29:11 AM, Rhuland, Courtney) Calibrated RH Functional check - OK
1063310 | 201964 | CM | CMPLT | CMPLT | 2024-02-27 08:24 | 2024-04-02 12:46 | micro-manipulator needs to be repaired | (TC 03/04/2024 12:46:08 PM, Huang, Jessica) Dr. Witterick was using the micromanipulator during the case and noticed it was not working properly. Coordinated with Melissa from NOI and sent it for repair. got it back and informed all the users. Dr. Mentiro was using it on March 25th.,
| 201965 | PA | CMPLT | CMPLT | 2024-02-27 08:38 | 2024-03-07 07:33 | Battery modules for servo-u | (TC 03/05/2024 09:04:15 AM, Canjimavadivel, Kamallesh ) Followed up with Fan regarding shopping cart (TC 03/07/2024 07:32:03 AM, Canjimavadivel, Kamallesh ) Received batteries
| 201966 | CM | CMPLT | CMPLT | 2024-02-27 09:47 | 2024-04-16 08:52 | BD Pump Remediation Plan | (TC 02/27/2024 09:47:38 AM, Chukri, Samir) Meeting with Barb 10s (TC 02/27/2024 01:51:59 PM, Chukri, Samir) M&B Pump Remediation (TC 03/12/2024 11:41:13 AM, Chukri, Samir) talking to Nara and Courteny regarding the IUI connection issue. try to solve the cleaning problem to avoide the BD reject to perform the PM. (TC 03/14/2024 09:24:33 AM, Chukri, Samir) discuss regarding the Alaris pump return for IUI cleaning, Rutvik will help Nara and courtney for this week. (TC 03/25/2024 02:31:12 PM, Chukri, Samir) searching for the Alaris pump SN 13994457 location and emailed to Mirza Helpdisk. (TC 03/26/2024 02:47:07 PM, Chukri, Samir) working with Courtney searching for BD pump with same serial number contact Mirza to help us. we can't find this pump. (TC 04/09/2024 02:00:00 PM, Chukri, Samir) meet with Maisha and BD for the clean plan and completed and remove the tables and dispoasle parts and batteries by emmerson, received email from BD for the return parts (TC 04/16/2024 08:51:17 AM, Chukri, Samir) Completed.
68762 | 201967 | CM | CMPLT | CMPLT | 2024-02-27 10:16 | 2024-02-27 10:19 | Damaged power cord | (TC 02/27/2024 10:16:24 AM, Baghirzade, Nara) Replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 221 mohms Leakage current: 68.0 uA
1061210 | 201968 | CM | CMPLT | CMPLT | 2024-02-27 10:55 | 2024-02-27 10:56 | Light not balanced | (TC 02/27/2024 10:55:15 AM, Reid, Emmerson ) Went up with Rutvik and was able to rebalance the light. also introduced him to Maureen.
1064351 | 201969 | CM | CMPLT | CMPLT | 2024-02-27 11:02 | 2024-02-28 13:12 | MSH ICU SPARQ POD C SITE # 98117064 CLICKING BETWEEN 2D & CPA NO RESPONSE FROM OP PANEL | (TC 02/27/2024 11:01:49 AM, Whittington, Scott) MSH ICU SPARQ POD C SITE # 98117064 CLICKING BETWEEN 2D & CPA NO RESPONSE FROM OP PANEL HARD REBOOT APPEARS TO BE OPERATING NORMALLY CONSULTED JILLIAN T WHILE SHE IS ON SITE | MSH ICU SPARQ POD C SITE # 98117064 CLICKING BETWEEN 2D & CPA NO RESPONSE FROM OP PANEL
| 201970 | CM | CMPLT | CMPLT | 2024-02-27 11:51 | 2024-02-27 11:51 | Meeting with NOI re: service expectations | (TC 02/27/2024 11:51:04 AM, Baghirzade, Nara) Meeting with NOI re: service expectations
1051085 | 201971 | CM | CMPLT | CMPLT | 2024-02-27 12:44 | 2024-02-27 12:49 | BP low and sounds like it's leaking | (TC 02/27/2024 12:47:30 PM, Scero, Tomlyn) Tested BP and confirmed issue. Replaced the air hose and BP cuff. NIBP no longer leaking, accurate BP was taken.
| 201972 | CM | CMPLT | CMPLT | 2024-02-27 12:57 | 2024-02-27 12:57 | meeting | vent meeting
| 201973 | CM | CMPLT | CMPLT | 2024-02-27 12:58 | 2024-02-27 12:58 | biomed meeting | biomed meeting
| 201974 | CM | CMPLT | CMPLT | 2024-02-27 12:59 | 2024-02-27 12:59 | worked with RT | Worked with RT
1070144 | 201975 | CM | CMPLT | CMPLT | 2024-02-27 13:31 | 2024-02-27 13:31 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 80 mOhm; 78 uA. Tested ok. Installed unit in endoscopy.
1070145 | 201976 | CM | CMPLT | CMPLT | 2024-02-27 13:42 | 2024-02-27 13:42 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 100 mOhm; 77 uA. Tested ok. Installed unit in endoscopy.
1070146 | 201977 | CM | CMPLT | CMPLT | 2024-02-27 13:52 | 2024-04-11 13:53 | Incoming Inspection | Feb 27 - Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 100 mOhm; 77 uA. Tested ok. Apr 11 - Delivered unit to Endo and is being stored on a sheft.
1070147 | 201978 | CM | CMPLT | CMPLT | 2024-02-27 13:58 | 2024-04-11 13:58 | Incoming Inspection | Feb 27 - Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 120 mOhm; 78 uA. Tested ok. Apr 11 - Delivered unit to Endo and is being stored on a sheft.
1070148 | 201979 | CM | CMPLT | CMPLT | 2024-02-27 14:14 | 2024-02-27 14:15 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 140 mOhm; 77 uA. Tested ok. Installed unit in endoscopy.
1070149 | 201980 | CM | CMPLT | CMPLT | 2024-02-27 14:22 | 2024-02-27 14:24 | Incoming Inspection | Feb 27 - Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 160 mOhm; 79 uA. Tested ok.
67957 | 201981 | CM | CMPLT | CMPLT | 2024-02-27 14:59 | 2024-02-27 15:11 | Disposal and replacement of hoods | (TC 02/27/2024 03:09:38 PM, Baghirzade, Nara) Removed components off of old hood. Disposed of the old hood and any components that were broken. Replaced new hood and transferred existing components to the new hood. Replaced south latch.
| 201982 | ACC | CMPLT | CMPLT | 2024-02-27 15:04 | 2024-03-15 12:23 | MSH OPG SPP SAMSUNG US HERA W10 DEMO | (TC 02/27/2024 03:03:51 PM, Whittington, Scott) MSH OPG SPP SAMSUNG US HERA W10 DEMO (TC 03/08/2024 02:31:00 PM, Whittington, Scott) MSH OPG SPP SAMSUNG US HERA W10 DEMO ACCEPTANCE AND PROGRAMMING SN# 82A3M3HRA00001T (TC 03/08/2024 02:33:28 PM, Dolinar, John) MSH OPG SPP SAMSUNG US HERA W10 DEMO (TC 03/15/2024 12:22:41 PM, Whittington, Scott) MSH OPG SPP SAMSUNG US HERA W10 DEMO SYSTEM REMOVED (TC 03/15/2024 12:23:00 PM, Dolinar, John) MSH OPG SPP SAMSUNG US HERA W10 DEMO SYSTEM REMOVED | MSH OPG SPP SAMSUNG US HERA W10 DEMO
| 201983 | PA | CMPLT | CMPLT | 2024-02-27 15:07 | 2024-05-07 14:19 | No charge PO request for trial equipment | (TC 02/28/2024 02:12:52 PM, Canjimavadivel, Kamallesh ) Received no charge PO and forwarded to Becky for next steps
| 201984 | CM | CMPLT | CMPLT | 2024-02-27 15:11 | 2024-02-27 15:11 | PM update meeting | (TC 02/27/2024 03:10:57 PM, Baghirzade, Nara) PM update meeting
| 201985 | CM | CMPLT | CMPLT | 2024-02-27 15:14 | 2024-02-27 15:14 | Transferring NICU defib back to NICU | (TC 02/27/2024 03:13:38 PM, Baghirzade, Nara) Helped Courtney bring back the NICU defib and swap it with the spare biomed X-series that was on the NICU crash cart. Updated biomed team that biomed X-series is back on the cart.
| 201986 | CM | CMPLT | CMPLT | 2024-02-27 15:18 | 2024-02-27 15:46 | Meeting with 11N re: BD pump upgrade
| 201987 | PA | CMPLT | CMPLT | 2024-02-28 06:54 | 2024-02-29 13:18 | Ordering sterile water for NICU altrix cleaning | (TC 02/28/2024 06:54:19 AM, Baghirzade, Nara) SC 1003178125 (TC 02/29/2024 01:17:21 PM, Baghirzade, Nara) Order #: 4150272682
68874 | 201988 | CM | CMPLT | CMPLT | 2024-02-28 07:10 | 2024-02-28 07:12 | Battery defective; will not stay on | (TC 02/28/2024 07:10:38 AM, Baghirzade, Nara) Battery replaced. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 180 mohms Leakage current: 73.0 uA
| 201989 | CM | CMPLT | CMPLT | 2024-02-28 08:29 | 2024-03-01 15:19 | Hotline warmer tank gasket for stock | received
| 201990 | CM | CMPLT | CMPLT | 2024-02-28 08:48 | 2024-02-28 08:53 | NICU Rounds | (TC 02/28/2024 08:50:40 AM, Rhuland, Courtney) Weekly rounds completed Issue with booms in rooms: 48 & 30 Contacting Draeger for support
| 201991 | CM | CMPLT | CMPLT | 2024-02-28 08:53 | 2024-02-28 08:55 | RESUS Rounds | (TC 02/28/2024 08:54:23 AM, Rhuland, Courtney) RESUS weekly rounds complete
66694 | 201992 | CM | CMPLT | CMPLT | 2024-02-28 10:02 | 2024-03-12 15:46 | Unable to connect to CL and wireless transducers, red light | (TC 02/27/2024 10:02:35 AM, Rendon Lopez , Elizabeth) Confirmed issue. Removed from room and replaced with spare monitor. (TC 02/28/2024 10:02:35 AM, Rendon Lopez , Elizabeth) Troubleshooting, replaced connector from spare monitor, still not working. Suspect defective API board, order replacement (TC 03/11/2024 03:45:11 PM, Rendon Lopez , Elizabeth) Replaced API board, same issue. Rev G software not compatible with CL wireless station, placed old API back, swap monitor wirh Biomed spare 66517. kept monitor for spare, API board as spare
| 201993 | CM | CMPLT | CMPLT | 2024-02-28 10:38 | 2024-02-28 10:38 | Medfusion webinar | (TC 02/28/2024 10:37:59 AM, Baghirzade, Nara) Medfusion webinar for software/firmware upgrade.
| 201994 | CM | CMPLT | CMPLT | 2024-02-28 10:39 | 2024-02-29 08:50 | Booms in NICU with stuck buttons/defects | (TC 02/28/2024 10:38:50 AM, Baghirzade, Nara) Found 2 booms in NICU with stuck buttons. Tried to unstick the buttons but could not get it to work. Followed up with Samir re: issue. Courtney to contact Draeger for service. Will assess all NICU booms for defects. (TC 02/29/2024 08:46:39 AM, Baghirzade, Nara) Checked most of the NICU rooms including some isolation rooms. Found following rooms with defective booms (depressed buttons and broken ethernet ports): - Room 19 (depressed button) - Room 23 (depressed button) - Room 30 (broken ethernet port) - Room 48 (depressed button) - Room 53A (broken ethernet port) - Room 75 (broken ethernet port) Courtney and I will follow up with Naz regarding corrective maintenance for the rooms, specifically with the brakes issue.
1065735 | 201995 | CM | CMPLT | CMPLT | 2024-02-28 10:44 | 2024-02-28 10:45 | PCU does not switch on; failed front case | (TC 02/28/2024 10:44:09 AM, Baghirzade, Nara) No AC LED/PCU does not switch on. Replaced front case as per recall. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 120 mohms Leakage current: 63.0 uA
| 201996 | CM | CMPLT | CMPLT | 2024-02-28 10:45 | 2024-03-01 14:54 | ECG lead clips needed ASAP | (TC 02/28/2024 10:53:35 AM, Scero, Tomlyn) Emailed philips for an updated quote. Ordered 5 bags for ECG department. (TC 03/01/2024 02:52:58 PM, Scero, Tomlyn) clips received and delivered to Dianne
1060462 | 201997 | CM | CMPLT | CMPLT | 2024-02-28 10:50 | 2024-02-29 15:54 | Broken Cable | (TC 02/23/2024 02:18:54 PM, Patel, Rutvik) Call re: not working. Went to check. Found Scope guide cable connector broken. Same length cable not available. Removed broken cable and repalced it with a shorter cable. Calibrated the scope guide position. Tested ok. Will order new cable soon. (TC 02/28/2024 02:19:30 PM, Patel, Rutvik) Contacted sales rep for part# and Pricing twice. No response. Reached out to olympus canada. Info Received. Created SAP Req. SC #1003178392 Requested approval. Received. Awaiting delivery. (TC 02/29/2024 03:53:28 PM, Patel, Rutvik) Cable received. Coordinated and schedule the replacement. Routed and repalced scopeguide cable. Calibrated scope position. Tested ok. Returned to service.
1063746 | 201998 | PM | CMPLT | CMPLT | 2024-02-28 10:52 | 2024-02-28 10:53 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/28/2024 10:52:22 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1058543 | 201999 | PM | CMPLT | CMPLT | 2024-02-28 11:02 | 2024-02-28 11:06 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 02/28/2024 11:03:20 AM, Reid, Emmerson ) Previous WO closed by accident. Battery Replaced, PM Complete. | Every year needs dust cleaning, every two years the battery is changed.
1058543 | 202000 | PM | ACTIV | ISSUE | 2024-02-28 11:02 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible |  | Every year needs dust cleaning, every two years the battery is changed.
1051826 | 202001 | PM | CMPLT | CMPLT | 2024-02-28 11:09 | 2024-02-28 11:10 | PM MAJOR - Annual PM for: UPS | (TC 02/28/2024 11:09:57 AM, Reid, Emmerson ) Replaced Battery, PM Complete.
| 202002 | CM | CMPLT | CMPLT | 2024-02-28 11:14 | 2024-05-09 14:06 | Ordering UPS Batteries | (TC 02/28/2024 11:18:46 AM, Reid, Emmerson ) Placed the order on erec (TC 05/09/2024 02:05:55 PM, Reid, Emmerson ) Recieved and put away
1047839 | 202003 | CM | CMPLT | CMPLT | 2024-02-28 11:24 | 2024-02-28 11:25 | Software upgrade to V1.6.5 | (TC 02/28/2024 11:24:05 AM, Baghirzade, Nara) Software upgrade to V1.6.5 compeleted successfully.
67957 | 202004 | CM | CMPLT | CMPLT | 2024-02-28 11:31 | 2024-02-28 11:32 | Functional check post hood replacement | (TC 02/28/2024 11:31:42 AM, Baghirzade, Nara) Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 76.7 %RH Incubator is fully functional.
| 202005 | CM | CMPLT | CMPLT | 2024-02-28 12:59 | 2024-02-28 12:59 | meeting | meeting
| 202006 | CM | CMPLT | CMPLT | 2024-02-28 13:00 | 2024-02-28 13:00 | worked with RT | worked with RT
66474 | 202007 | CM | CMPLT | CMPLT | 2024-02-28 13:02 | 2024-03-04 13:45 | not passing patient circuit test | can't run unit in any mode except O2 therapy. Vent and humidifier options are disabled on cockpit. It passed device check. C500 had been repaired in 07.2021. Before this error there were pressure errors on 23th Feb. 2024. Collected logs and sent to Drageger. 1.tried putting an image on HDD with Cosego USB and then used Infinity installer. But got failure 0769.32930 2. tried putting an image on HDD with Cosego USB and then used Infinity installer. But got failure 0769.32930 3. tried putting an image on HDD with Cosego USB and then used Infinity installer but this time I tried a full installation instead of cockpit only. It worked. Copied config from another unit. Passed pre-use check. Started venting. Second day still working without any error. Third day worked without error. Returned back to service.
1053734 | 202008 | FC | CMPLT | CMPLT | 2024-02-28 13:06 | 2024-02-28 13:58 | Functional check after use | (TC 02/28/2024 01:57:17 PM, Scero, Tomlyn) Functional checks complete and passed after use.
| 202009 | CM | CMPLT | CMPLT | 2024-02-28 13:30 | 2024-02-28 13:30 | Meeting with adeel
1059828 | 202010 | PM | CMPLT | CMPLT | 2024-02-28 13:30 | 2024-02-28 14:18 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/28/2024 02:17:43 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 202011 | CM | CMPLT | CMPLT | 2024-02-28 13:31 | 2024-02-28 13:31 | Capital meeting | (TC 02/28/2024 01:30:45 PM, Baghirzade, Nara) Capital meeting
1066929 | 202012 | CM | CMPLT | CMPLT | 2024-02-28 13:33 | 2024-02-28 13:35 | No Video | (TC 02/28/2024 01:33:30 PM, Patel, Rutvik) CAll re: no video. Went to check. Found the monitor not being able to duplicate the primary dispaly. Checked and secured all connection. Found the monitor was set to incorrect input. Fixed it. Tested ok.
1053982 | 202013 | CM | CMPLT | CMPLT | 2024-02-28 13:41 | 2024-02-28 13:43 | Display cable broken. | (TC 02/28/2024 01:41:30 PM, Patel, Rutvik) Found video cable connector (BNC Connector) broken. Replaced it. Checked functionality. Tested ok. Returned to service.
67872 | 202014 | CM | CMPLT | CMPLT | 2024-02-28 13:45 | 2024-02-28 13:47 | Not working | (TC 02/28/2024 01:45:52 PM, Rendon Lopez , Elizabeth) Testing on user mode, NFF. Performance verification done-OK
1047833 | 202015 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-11 13:04 | Software upgrade to V1.6.5 | (TC 06/11/2024 01:04:06 PM, Rhuland, Courtney) Software upgraded to 1.6.5
1047840 | 202016 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-04-10 13:45 | Software upgrade to V1.6.5 | (TC 04/10/2024 01:43:28 PM, Rhuland, Courtney) Software upgraded to V1.6.5
1047841 | 202017 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-28 13:25 | Software upgrade to V1.6.5 | (TC 06/28/2024 01:24:48 PM, Rhuland, Courtney) Software upgraded to V1.6.5
1047842 | 202018 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-02-28 13:56 | Software upgrade to V1.6.5 | (TC 02/28/2024 01:55:55 PM, Rhuland, Courtney) Software upgrade to V1.6.5 compeleted successfully.
1047843 | 202019 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-03-27 08:22 | Software upgrade to V1.6.5 | (TC 03/27/2024 08:20:40 AM, Rhuland, Courtney) Software upgraded to V1.6.5
1047844 | 202020 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-04-11 08:49 | Software upgrade to V1.6.5 | (TC 04/11/2024 08:48:10 AM, Rhuland, Courtney) Software upgraded to V1.6.5
1047845 | 202021 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-14 14:53 | Software upgrade to V1.6.5 | (TC 06/14/2024 02:52:42 PM, Rhuland, Courtney) Software upgraded to V1.6.5
1047846 | 202022 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-03-28 09:12 | Software upgrade to V1.6.5 | (TC 03/28/2024 09:11:36 AM, Rhuland, Courtney) Software upgraded to V1.6.5
1047847 | 202023 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-07-23 14:03 | Software upgrade to V1.6.5 | (TC 07/23/2024 02:02:40 PM, Baghirzade, Nara) Software upgrade to V1.6.5 completed using DL4000
1047848 | 202024 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-07-04 10:33 | Software upgrade to V1.6.5 | (TC 07/04/2024 10:32:42 AM, Rhuland, Courtney) Upgraded software to 1.6.5
1047849 | 202025 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-05-08 12:47 | Software upgrade to V1.6.5 | (TC 05/08/2024 12:46:54 PM, Baghirzade, Nara) Software upgrade to V1.6.5 completed successfully.
1047850 | 202026 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-03-01 11:06 | Software upgrade to V1.6.5 | (TC 03/01/2024 11:06:00 AM, Rhuland, Courtney) Software upgrade to V1.6.5 completed successfully.
1047851 | 202027 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-02-29 09:49 | Software upgrade to V1.6.5 | (TC 02/29/2024 09:48:35 AM, Baghirzade, Nara) Software upgrade to V1.6.5 completed successfully.
1047852 | 202028 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-03-04 13:48 | Software upgrade to V1.6.5 | (TC 03/04/2024 01:48:06 PM, Rhuland, Courtney) Software upgrade to V1.6.5 completed successfully.
1047853 | 202029 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-08-27 10:37 | Software upgrade to V1.6.5 | (TC 08/27/2024 10:36:58 AM, Rhuland, Courtney) Software successfully upgraded to V1.6.5
1047854 | 202030 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-07-08 13:19 | Software upgrade to V1.6.5 | (TC 07/08/2024 01:19:07 PM, Rhuland, Courtney) Software updated to V1.6.5
1047855 | 202031 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-05-06 14:08 | Software upgrade to V1.6.5 | (TC 05/06/2024 02:07:55 PM, Rhuland, Courtney) Located pump in NICU Software upgrade to V1.6.5 completed successfully.
1047856 | 202032 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-19 14:15 | Software upgrade to V1.6.5 | (TC 06/19/2024 02:14:46 PM, Rhuland, Courtney) Software upgraded to V1.6.5
1056125 | 202033 | PM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-03-05 08:31 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 02/28/2024 01:55:18 PM, Canjimavadivel, Kamallesh ) Cleaned hoses ends and ports. Cleaning cycle complete
1047857 | 202034 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-07-09 11:14 | Software upgrade to V1.6.5 | (TC 07/09/2024 11:13:56 AM, Rhuland, Courtney) Software upgraded to V1.6.5
1047858 | 202035 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-07 08:41 | Software upgrade to V1.6.5 | (TC 06/07/2024 08:41:03 AM, Rhuland, Courtney) Successfuly upgraded software to V1.6.5.
1047859 | 202036 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-14 14:26 | Software upgrade to V1.6.5 | (TC 06/14/2024 02:26:09 PM, Rhuland, Courtney) Software upgraded to V1.6.5
1047860 | 202037 | CM | CANCL | CANCL | 2024-02-28 13:55 | 2024-08-13 12:21 | Software upgrade to V1.6.5
1047861 | 202038 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-14 09:28 | Software upgrade to V1.6.5 | (TC 06/14/2024 09:27:39 AM, Rhuland, Courtney) Software upgraded to V1.6.5
1047862 | 202039 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-14 14:53 | Software upgrade to V1.6.5 | (TC 06/14/2024 02:53:16 PM, Rhuland, Courtney) Software upgraded to V1.6.5
1047863 | 202040 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-05 13:33 | Software upgrade to V1.6.5 | (TC 06/05/2024 01:33:25 PM, Rhuland, Courtney) Software upgrade to V1.6.5 completed successfully.
1047864 | 202041 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-06-07 09:40 | Software upgrade to V1.6.5 | (TC 06/07/2024 09:39:52 AM, Rhuland, Courtney) Software successfully upgraded to V1.6.5.
1047865 | 202042 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-05-08 12:44 | Software upgrade to V1.6.5 | (TC 05/08/2024 12:43:39 PM, Baghirzade, Nara) Software upgrade to V1.6.5 completed successfully.
1047866 | 202043 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-02-29 09:47 | Software upgrade to V1.6.5 | (TC 02/29/2024 09:47:08 AM, Baghirzade, Nara) Software upgrade to V1.6.5 completed successfully.
1047867 | 202044 | CM | CMPLT | CMPLT | 2024-02-28 13:55 | 2024-05-10 11:05 | Software upgrade to V1.6.5 | (TC 05/10/2024 11:04:43 AM, Rhuland, Courtney) Software upgrade to V1.6.5 completed successfully. Functional check - OK
1066036 | 202045 | CM | CMPLT | CMPLT | 2024-02-28 14:02 | 2024-02-28 14:04 | Functional check | (TC 02/28/2024 02:02:29 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
69507 | 202046 | CM | CMPLT | CMPLT | 2024-02-28 14:12 | 2024-02-28 14:15 | Functional check | (TC 02/28/2024 02:14:23 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
69531 | 202047 | CM | CMPLT | CMPLT | 2024-02-28 14:19 | 2024-02-28 14:22 | Functional check | (TC 02/28/2024 02:21:31 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1065923 | 202048 | CM | CMPLT | CMPLT | 2024-02-28 14:24 | 2024-02-28 14:25 | Functional check | (TC 02/28/2024 02:24:13 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1070156 | 202049 | CM | CMPLT | CMPLT | 2024-02-28 14:36 | 2024-04-11 14:36 | Incoming Inspection | Feb 28 - Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 160 mOhm; 78 uA. Tested ok. Apr 11 - Delivered unit to Endo and is being stored on a sheft.
1070157 | 202050 | CM | CMPLT | CMPLT | 2024-02-28 14:41 | 2024-02-28 14:41 | Incoming Inspection | Feb 28 - Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 100 mOhm; 77 uA. Tested ok. Apr 11 - Delivered unit to Endo and is being stored on a sheft.
1058175 | 202051 | CM | CMPLT | CMPLT | 2024-02-28 15:06 | 2024-02-28 15:07 | Broken | (TC 02/28/2024 03:06:05 PM, Reid, Emmerson ) Was not broken, needed to change the batteries and update the library.
| 202052 | CON | CMPLT | CMPLT | 2024-02-29 07:50 | 2024-03-11 09:26 | CCU Upgrade project | (TC 02/28/2024 07:50:20 AM, Scero, Tomlyn) Kick off Zoom meeting. Phiips to come on site Tuesday March 5th (TC 02/29/2024 02:10:40 PM, Scero, Tomlyn) Met with Dianne informally in CCU to discuss configuration changes (TC 03/05/2024 01:41:46 PM, Scero, Tomlyn) worked with Philips, completed SW upgrades on all 6 MX700s and swapped the X3s. Removed X2s from the CCU pool on PICIX. Showed nursing staff how to use the X3 (TC 03/08/2024 10:32:20 AM, Scero, Tomlyn) Downloaded both MX700 and X3 configuration files to the T drive (TC 03/08/2024 09:24:16 AM, Scero, Tomlyn) All X2s and CO2 microstream extension modules "retired" from TMS and stored on the shelf in Biomed
66474 | 202053 | CM | CMPLT | CMPLT | 2024-02-29 08:04 | 2024-03-08 10:54 | ordering parts for VN500 vents | received parts
1065889 | 202054 | CM | CMPLT | CMPLT | 2024-02-29 08:41 | 2024-02-29 08:46 | Functional check | (TC 02/29/2024 08:45:35 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
69364 | 202055 | CM | CMPLT | CMPLT | 2024-02-29 09:04 | 2024-02-29 09:04 | Functional check | (TC 02/29/2024 09:03:53 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1065867 | 202056 | CM | CMPLT | CMPLT | 2024-02-29 09:19 | 2024-02-29 09:20 | Functional check | (TC 02/29/2024 09:19:28 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
| 202057 | CM | CMPLT | CMPLT | 2024-02-29 09:22 | 2024-02-29 09:24 | Garbage Disposal | Garbage Disposal for Steris Order. Brought in garbage bin. Cut all boxes and disposed all packing materials.
69672 | 202058 | CM | CMPLT | CMPLT | 2024-02-29 09:22 | 2024-02-29 09:22 | Functional check | (TC 02/29/2024 09:22:13 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
67369 | 202059 | PM | CMPLT | CMPLT | 2024-02-29 11:00 | 2024-04-12 08:12 | PM-MAJOR - Annual PM For: Air-oxygen Blender | shipped (TC 04/12/2024 08:11:33 AM, Ferreira, Qwynn) Functional Test - OK Tested at 21%, 40%, 60%, 80%, 100%
1066919 | 202060 | CM | CMPLT | CMPLT | 2024-02-29 11:05 | 2024-04-23 09:50 | "can't use scale" | (TC 02/28/2024 11:05:19 AM, Cortez, John) Confirmed issue, might be related to "scale uncalibrated" message Marlene also mentioned they tried weighing patient below minimum required weight for bed (225 lbs) Called Stryker and they described process of calibrating scale. Will attempt later. (TC 02/29/2024 04:03:27 PM, Cortez, John) Reviewed service manual for scale calibration
| 202061 | CM | CMPLT | CMPLT | 2024-02-29 11:29 | 2024-02-29 11:30 | Setting up space for BD software upgrade | (TC 02/29/2024 11:29:21 AM, Baghirzade, Nara) Moving tables, chair, and BD parts boxes to 14th floor space for the BD software upgrade with Courtney and Maisha (TC 02/29/2024 11:29:21 AM, Baghirzade, Nara) Moving tables, chair, and BD parts boxes to 14th floor space for the BD software upgrade with Courtney and Maisha
| 202062 | CM | CMPLT | CMPLT | 2024-02-29 11:58 | 2024-02-29 11:58 | Incubator Delivery and Task Overview meeting | (TC 02/29/2024 11:57:36 AM, Baghirzade, Nara) Incubator Delivery and Task Overview meeting with Naz and Courtney.
1063369 | 202063 | CM | CMPLT | CMPLT | 2024-02-29 13:11 | 2024-04-19 11:35 | Requires calibration - sending to CHS for calibration | (TC 02/29/2024 03:00:03 PM, Baghirzade, Nara) Contacted CHS for a service request to send unit for calibration. CHS coordinator sent over the request form and requested a no charge PO for the calibration and loaner device. (TC 02/29/2024 03:01:08 PM, Baghirzade, Nara) SC 1003180188 (TC 03/01/2024 03:14:19 PM, Baghirzade, Nara) Forwaded the loaner agreement and the response from the vendor to Sarabeth; she prefers to send the device for calibration when we are guaranteed a loaner as their need for a bladder scanner is high on 11N. Relayed this information to the vendor and let him know that I will send our bladder scanner once we have a loaner. (TC 03/20/2024 10:20:38 AM, Baghirzade, Nara) Received loaner; created asset on TMS as loaner. PO 7500202840 (TC 03/20/2024 10:26:09 AM, Baghirzade, Nara) Sent device for calibraton. case# CAS-04629-Y3Z8R9 KJV Courier # 423362 (TC 04/19/2024 11:31:42 AM, Baghirzade, Nara) Received bladder scanner; attached service report. Placed bladder scanner back on the rolling stand. Completed a caliscan for the probe. Returned to 11 North
502453 | 202064 | CM | CMPLT | CMPLT | 2024-02-29 13:34 | 2024-02-29 13:35 | MSH OPG SPP-3525(Rm D1) GE TO C2 | (TC 02/29/2024 01:34:49 PM, Whittington, Scott) MSH OPG SPP-3525(Rm D1) GE TO C2 | MSH OPG SPP-3525(Rm D1) GE TO C2
| 202065 | CM | CMPLT | CMPLT | 2024-02-29 14:06 | 2024-02-29 14:09 | NICU Pump Remediation Plan Meeting | (TC 02/29/2024 02:08:45 PM, Baghirzade, Nara) NICU Pump Remediation Plan Meeting (TC 02/29/2024 02:08:45 PM, Baghirzade, Nara) NICU Pump Remediation Plan Meeting
| 202066 | CM | CMPLT | CMPLT | 2024-02-29 14:28 | 2024-02-29 14:28 | Servo-N tranining | Servo-N tranining
| 202067 | CM | CMPLT | CMPLT | 2024-02-29 14:29 | 2024-02-29 14:29 | meeting | meeting
| 202068 | CM | CMPLT | CMPLT | 2024-02-29 14:46 | 2024-05-07 09:02 | X3 batteries, bulk order for Biomed | (TC 02/29/2024 02:51:16 PM, Scero, Tomlyn) placed order for 30 batteries from philips (TC 05/07/2024 09:01:02 AM, Scero, Tomlyn) batteries were received and stored on the shelf | batteries for stock
69656 | 202069 | CM | CMPLT | CMPLT | 2024-02-29 15:09 | 2024-02-29 15:09 | Functional check | (TC 02/29/2024 03:08:44 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1061202 | 202070 | CM | CMPLT | CMPLT | 2024-02-29 15:16 | 2024-02-29 15:17 | OR wall monitor not working for 4K scope | (TC 02/29/2024 03:15:59 PM, Reid, Emmerson ) The 4K camera was plugged into the wrong port of the boom. Working now.
| 202071 | CM | CMPLT | CMPLT | 2024-02-29 15:56 | 2024-03-05 09:26 | Biomed Work Area set up | Biomed Work area set up. Cleaned and vacuumed NB's cubicle. Drilled holes and Mounted Peg board for tools. Organized tools. Modified work bench height. Moved and arranged my stuff to the new work area.
| 202072 | CM | CMPLT | CMPLT | 2024-02-29 15:59 | 2024-02-29 16:00 | Endoscopy Emergency Video Cart Organization | (TC 02/29/2024 03:58:41 PM, Patel, Rutvik) As per manager's request, Clenaed up the cart. Relocated the power bar and equipment on the cart. Organized and sorted out hanging cables. Checked functionality. Tested ok.
50517 | 202073 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-04-04 09:50 | PM-MAJOR - Annual PM For: Vacuum Extractors, Obstetric | (TC 04/04/2024 09:49:50 AM, Rendon Lopez , Elizabeth) Performance verification-Ok
62781 | 202074 | PM | ACTIV | ISSUE | 2024-03-01 02:14 |  | PM-MAJOR - Annual PM For: Vacuum Pumps, Central
63874 | 202075 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-09-05 07:27 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/05/2024 07:27:13 AM, Baghirzade, Nara) PM completed. Replaced power cord. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 70.9 uA Power cord resistance = 85 mohms
64078 | 202076 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-03-27 14:55 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 03/27/2024 02:53:50 PM, Rhuland, Courtney) Repaired broken door latch Repaired bed tilt assembly PM completed Calibrated: LV and main voltage Scale Functional check - OK Electrical safety - OK
65123 | 202077 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-05-22 09:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/22/2024 09:45:58 AM, Reid, Emmerson ) PM Complete, see procedure for details. (TC 05/22/2024 09:45:58 AM, Reid, Emmerson ) PM Complete, see procedure for details.
65125 | 202078 | PM | ACTIV | ISSUE | 2024-03-01 02:14 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65135 | 202079 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-05-22 14:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/22/2024 02:44:24 PM, Reid, Emmerson ) PM Complete, see procedure for details.
65154 | 202080 | PM | CMPLT | PMFAIL | 2024-03-01 02:14 | 2024-05-13 13:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/13/2024 01:58:06 PM, Reid, Emmerson ) Unit needs to be sent out for repair.
65263 | 202081 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-03-28 15:03 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced air filter. Run battery check. 25 minutes and 230 minutes. Tested. Functional. SW2.51 Leakage: 85uA and 93 ohm.
65416 | 202082 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-05-06 14:01 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/06/2024 02:00:38 PM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 81.3% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 56.2 uA Power cord resistance = 231 mohms
65420 | 202083 | PM | ACTIV | ISSUE | 2024-03-01 02:14 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
65891 | 202084 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-03-15 13:34 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 01:33:25 PM, Baghirzade, Nara) PM completed by BD.
65902 | 202085 | PM | ACTIV | ISSUE | 2024-03-01 02:14 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65972 | 202086 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-04-17 10:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
66057 | 202087 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-03-18 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66454 | 202088 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-04-04 15:12 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced air filter. Run battery check. 25 minutes and 180 minutes. Tested. Functional. SW2.51 Leakage: 89uA and 92 ohm.
66455 | 202089 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-03-20 13:41 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced air filter. Run battery check. 25 minutes and 230 minutes. Tested. Functional. SW2.51 Leakage: 83uA and 91 ohm.
66761 | 202090 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-04-05 08:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/05/2024 08:38:49 AM, Reid, Emmerson ) PM Complete, see procedure
66767 | 202091 | PM | ACTIV | ISSUE | 2024-03-01 02:14 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66819 | 202092 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-07-18 09:59 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/18/2024 09:59:13 AM, DeGuzman, Leila) PM completed. See procedures for details.
67432 | 202093 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-04-29 13:03 | PM-MAJOR - Annual PM For: Vacuum Pumps, Central | (TC 04/29/2024 01:02:29 PM, Chukri, Samir) completed the functional check and replaced the battery on 2022
67763 | 202094 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-05-06 10:00 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/06/2024 09:59:40 AM, Reid, Emmerson ) PM Complete, see procedure. Had to run volume test twice.
67919 | 202095 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-05-06 12:29 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/06/2024 12:12:58 PM, Rhuland, Courtney) Replaced side door hinges (x2) Replaced porthole door spring Tightened porthole door hinges Replaced exam light Calibrated: Main and REF voltage Scale RH Functional check - OK Set 35C & 75% humidity Read 35C & 75.7% humidity Electrical safety - OK
67969 | 202096 | PM | CMPLT | CMPLT | 2024-03-01 02:14 | 2024-05-07 09:42 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/07/2024 09:40:41 AM, Rhuland, Courtney) PM completed Calibrated: Scale Main & REF voltage Functional check - OK Set 35C, 75% humidity Read 35C, 74.7% humidity Electrical safety - OK Ground resistance - 0.287mOhms Leakage current - 83.7uA
67970 | 202097 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-10 13:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/10/2024 01:01:25 PM, Rhuland, Courtney) 9V rechargeable battery replaced System calibration completed Scale calibrated RH calibrated 2 porthole doors replaced Functional check - OK Electrical safety - OK Resistance - 62mOhms Leakage current - 18.9uA
68545 | 202098 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-10 15:39 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 03/13/2024 11:48:28 AM, Rendon Lopez , Elizabeth) Removed AGM from OR# 3 for PM (TC 04/08/2024 03:21:08 PM, Rendon Lopez , Elizabeth) Replaced parts for 1 year PM, cleaned. PM procedures started (TC 04/10/2024 03:21:08 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
68546 | 202099 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-10-22 11:17 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 10/01/2024 01:09:01 PM, Scero, Tomlyn) AGM swapped out of OR5 and brought downstairs for annual PM (TC 10/09/2024 01:21:53 PM, Scero, Tomlyn) Replaced all PM parts including batteries. Self test complete. (TC 10/17/2024 11:16:31 AM, Scero, Tomlyn) Continued with PM test as per IPM-L (TC 10/22/2024 11:17:12 AM, Scero, Tomlyn) All tests have been complete and passed as per the attched IPM-L. Device also used to test the function of diva module 68385 that was in the shop for repair. See the attached results sheet for test results.
68582 | 202100 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-12 09:42 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
68587 | 202101 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-11-21 09:02 | PM-MAJOR-Annual PM for Servo-U Ventilator | (TC 11/18/2024 02:50:03 PM, Canjimavadivel, Kamallesh ) PM performed by vendor, see attached service report. Waiting for batteries. (TC 11/21/2024 09:01:44 AM, Canjimavadivel, Kamallesh ) Received batteries and installed
68594 | 202102 | PM | CMPLT | PMFAIL | 2024-03-01 02:15 | 2024-03-20 14:25 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 02:24:19 PM, Baghirzade, Nara) Broken rear case and requires full calibration.
68599 | 202103 | PM | CMPLT | PMFAIL | 2024-03-01 02:15 | 2024-03-21 15:52 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed by BD
68603 | 202104 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-12 10:04 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
68629 | 202105 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-25 10:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68631 | 202106 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-05-27 13:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/24/2024 10:56:35 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 108 mohms Leakage current: 77.8 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 05/27/2024 01:51:51 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3883 mamphr PM completed.
68635 | 202107 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-05-03 09:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/03/2024 09:03:45 AM, Baghirzade, Nara) PM completed by BD.
68645 | 202108 | PM | CMPLT | PMFAIL | 2024-03-01 02:15 | 2024-05-21 13:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/21/2024 01:09:52 PM, Baghirzade, Nara) Rate accuracy failed during PM. Rate was reading high. Will calibrate and PM.
68661 | 202109 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-14 15:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68677 | 202110 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-20 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68687 | 202111 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-03 13:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:33:50 PM, Rhuland, Courtney) PM completed by BD
68710 | 202112 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-07-17 11:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/17/2024 11:00:22 AM, DeGuzman, Leila) PM completed, all tests passed. Cleaned device. Electrical safety Ground resistance: 233 mOhm Leakage current: 41.9 uA Downloaded error and event log. Cleared error and event log.
68716 | 202113 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-11 15:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68719 | 202114 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-27 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68724 | 202115 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 10:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68727 | 202116 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-08-16 13:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68732 | 202117 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-08-16 13:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68736 | 202118 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-26 13:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68775 | 202119 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-02 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68777 | 202120 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-12 09:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:26:15 AM, Rendon Lopez , Elizabeth) PM completed by BD
68786 | 202121 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-03 13:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:23:35 PM, Rhuland, Courtney) PM completed by BD
68802 | 202122 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-02 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68814 | 202123 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-18 12:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68817 | 202124 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-12 13:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68825 | 202125 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68846 | 202126 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-05-23 12:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/23/2024 12:17:17 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 183 mohms Leakage current: 75.9 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68853 | 202127 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68866 | 202128 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-27 14:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68868 | 202129 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-05-06 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/06/2024 01:56:25 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 40 mohms Leakage current: 75.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68877 | 202130 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-08-16 14:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68886 | 202131 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-26 12:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68888 | 202132 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-03 14:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:07:00 PM, Rhuland, Courtney) PM completed by BD
68894 | 202133 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-20 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68897 | 202134 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-02 10:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68911 | 202135 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-09-10 08:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68934 | 202136 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-24 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68942 | 202137 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-25 10:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68955 | 202138 | PM | CANCL | CANCL | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68958 | 202139 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-21 09:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68963 | 202140 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-08-16 14:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68965 | 202141 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-24 14:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68987 | 202142 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-11 14:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68991 | 202143 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-14 15:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68996 | 202144 | PM | CANCL | DUP | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69014 | 202145 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-27 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69026 | 202146 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-12 11:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69042 | 202147 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-27 10:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69045 | 202148 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-19 21:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69058 | 202149 | PM | CMPLT | PMFAIL | 2024-03-01 02:15 | 2024-03-15 09:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69078 | 202150 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-27 10:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69098 | 202151 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-14 14:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:08:39 PM, Huang, Jessica) replaced the male iui connector and completed PM
69107 | 202152 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-14 14:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69113 | 202153 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-14 07:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69136 | 202154 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-03 11:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69141 | 202155 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-13 09:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:45:45 AM, Cortez, John) Female IUI replaced PM completed
69143 | 202156 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69148 | 202157 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-08-16 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69155 | 202158 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-13 13:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69157 | 202159 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-12 08:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69206 | 202160 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69213 | 202161 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-07-03 13:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:50:31 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered
69237 | 202162 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-12 10:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69243 | 202163 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-27 11:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69248 | 202164 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69297 | 202165 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-07-03 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:44:58 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered
69321 | 202166 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-18 15:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/18/2024 03:06:16 PM, Reid, Emmerson ) ASM PM Complete.
69322 | 202167 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-14 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 10:34:06 AM, Huang, Jessica) replaced the male iui connector and completed PM
69334 | 202168 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-08-09 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
69343 | 202169 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-21 18:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69349 | 202170 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-28 11:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2024 02:33:32 PM, Reid, Emmerson ) Replaced the front case
69365 | 202171 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-14 14:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:01:25 PM, Huang, Jessica) replaced the female iui connector and completed PM
69373 | 202172 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-08-09 13:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 01:00:29 PM, Patel, Rutvik) by BD
69378 | 202173 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-25 14:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69382 | 202174 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-19 20:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69404 | 202175 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-21 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69410 | 202176 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-09 15:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69440 | 202177 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-15 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69460 | 202178 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-18 13:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69552 | 202179 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-15 13:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69599 | 202180 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-27 10:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69606 | 202181 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-13 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69608 | 202182 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69618 | 202183 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69626 | 202184 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69660 | 202185 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-15 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69664 | 202186 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-19 13:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69669 | 202187 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 10:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69672 | 202188 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68149 | 202189 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: Testers
67356 | 202190 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
67364 | 202191 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
500846 | 202192 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:51 | MSH 5 US XG PM - Semi-Annual | (TC 03/12/2024 11:34:16 AM, Whittington, Scott) MSH 5 US XG PM - Semi-Annual | MSH 5 US XG PM - Semi-Annual
500847 | 202193 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:51 | MSH 5 US XG PM - Semi-Annual | (TC 03/12/2024 11:35:09 AM, Whittington, Scott) MSH 5 US XG PM - Semi-Annual | MSH 5 US XG PM - Semi-Annual
500850 | 202194 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-15 12:14 | MSH 5 RPCU IRIUS XG "CLYDE" PM - Semi-Annual | (TC 03/12/2024 11:36:35 AM, Whittington, Scott) MSH 5 RPCU IRIUS XG "CLYDE" PM - Semi-Annual (TC 03/15/2024 12:13:48 PM, Dolinar, John) MSH 5 RPCU IRIUS XG "CLYDE" PM - Semi-Annual | MSH 5 RPCU IRIUS XG "CLYDE" PM - Semi-Annual
32510 | 202195 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:47 | MSH DR1 PM - Semi-Annual | (TC 04/17/2024 09:46:51 AM, Whittington, Scott) MSH DR1 PM - Semi-Annual | MSH DR1 PM - Semi-Annual
32510 | 202196 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:45 | MSH DR1 HARP TRGT APR 21 Semi Annual | (TC 04/16/2024 08:06:46 PM, Whittington, Scott) MSH DR1 HARP TRGT APR 21 Semi Annual pre harp (TC 04/17/2024 09:45:14 AM, Whittington, Scott) MSH DR1 HARP TRGT APR 21 Semi Annual (TC 04/17/2024 11:45:00 AM, Whittington, Scott) MSH DR1 HARP TRGT APR 21 Semi Annual | MSH DR1 HARP TRGT APR 21 Semi Annual
32540 | 202197 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-08 15:34 | PM - Monthly | PM completed Mechanical PM done sc-30448 TOS phantom a(-994.7,9.9) b(330.9,13.8) c(133.1,14.3) d(96.1,14.0) e(-106.1,12.0) f(0.8,12.3) -passed all phantom tests
32560 | 202198 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-03 14:49 | PM - Monthly | Mechanical PM done sc-1396276 TOS phantom a(-990.0,9.6) b(332.7,13.3) c(132.8,13.0) d(93.2,12.9) e(-105.0,12.7) f(0.1,14.8) -passed all phantom tests
32580 | 202199 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-03 14:55 | PM - Monthly | PM completed Mechanical PM done sc-215210 TOS phantom a(-993.6,11.1) b(335.3,16.1) c(133.2,15.2) d(97.2,14.2) e(-106.8,14.0) f(0.7,14.9) -passed all phantom tests
500732 | 202200 | PM | CANCL | CANCL | 2024-03-01 02:15 | 2024-09-24 11:58 | PM- SCHEDULED SEMI-ANNUAL | APRIL 26 2024 MSH-20-470 (Sports Clinic) APLIO XG REMOVED FROM CLINICAL USE IN SPORTS CLINIC
500800 | 202201 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-09 17:11 | MSH BIUS XG PM - Semi-Annual | (TC 03/13/2024 09:34:37 AM, Whittington, Scott) MSH BIUS XG PM - Semi-Annual (TC 04/09/2024 03:39:38 PM, Dolinar, John) MSH BIUS XG PM - Semi-Annual | MSH BIUS XG PM - Semi-Annual
500801 | 202202 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-09 17:11 | MSH BIUS XG PM - Semi-Annual | (TC 03/13/2024 09:35:06 AM, Whittington, Scott) MSH BIUS XG PM - Semi-Annual (TC 04/09/2024 03:40:04 PM, Dolinar, John) MSH BIUS XG PM - Semi-Annual | MSH BIUS XG PM - Semi-Annual
500802 | 202203 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-09 17:11 | MSH BIUS XG PM - Semi-Annual | (TC 03/13/2024 09:35:36 AM, Whittington, Scott) MSH BIUS XG PM - Semi-Annual (TC 04/09/2024 03:40:29 PM, Dolinar, John) MSH BIUS XG PM - Semi-Annual | MSH BIUS XG PM - Semi-Annual
500857 | 202204 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:49 | MSH BIUS XG PM - Semi-Annual | (TC 03/13/2024 09:36:13 AM, Whittington, Scott) MSH BIUS XG PM - Semi-Annual | MSH BIUS XG PM - Semi-Annual
500858 | 202205 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:49 | MSH-5-US PM - Semi-Annual | (TC 03/13/2024 03:17:24 PM, Whittington, Scott) MSH-5-US PM - Semi-Annual | MSH-5-US PM - Semi-Annual
500868 | 202206 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:49 | MSH-5-US PM - Semi-Annual | (TC 03/13/2024 03:17:57 PM, Whittington, Scott) MSH-5-US PM - Semi-Annual | MSH-5-US PM - Semi-Annual
500869 | 202207 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:50 | MSH-5-US PM - Semi-Annual | (TC 03/13/2024 03:18:32 PM, Whittington, Scott) MSH-5-US PM - Semi-Annual | MSH-5-US PM - Semi-Annual
501647 | 202208 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-26 15:17 | PM - Annual | (TC 04/26/2024 03:17:06 PM, Green, Paul) From: Green, Paul <Paul.Green@uhn.ca> Sent: April 26, 2024 3:16 PM To: Yang, Helen - Sinai Health System <Helen.Yang@sinaihealth.ca>; Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca> Cc: Paz, Ramon <Ramon.Paz@uhn.ca> Subject: MSH Veris MRI monitor on 2 Hi Helen/Jason, We checked out the NIBP on the Veris monitor on 2. With the replaced tubing, the NIBP is ok. The old tubing had a leak and we will try to repair it. We did a PM on the Veris monitor while we were there. Have a great weekend. Regards, Paul Paul Green, CET Senior Technologist Medical Engineering, Imaging Service Group – UHN, MSH & WCH Toronto General Hospital, TGH-1-NU160 200 Elizabeth Street, Toronto ON, M5G 2C4 Imaging Service Support: 437-688-9325 Cell: (416) 919-7245 Email: paul.green@uhn.ca
501888 | 202209 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-17 09:50 | MSH-5-US PM - Semi-Annual | (TC 03/13/2024 03:18:58 PM, Whittington, Scott) MSH-5-US PM - Semi-Annual | MSH-5-US PM - Semi-Annual
1046035 | 202210 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-05-02 14:26 | PM Major | (TC 05/02/2024 10:23:00 AM, Fan, Tak) PM completed.
1046035 | 202211 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-19 13:28 | MSH ER DRX EVO E5-02(ER) HARP - Semi Annual | (TC 04/19/2024 01:27:41 PM, Dolinar, John) MSH ER DRX EVO E5-02(ER) HARP - Semi Annual (TC 04/19/2024 01:27:56 PM, Fan, Tak) MSH ER DRX EVO E5-02(ER) HARP - Semi Annual (TC 04/19/2024 01:27:41 PM, Dolinar, John) MSH ER DRX EVO E5-02(ER) HARP - Semi Annual | MSH ER DRX EVO E5-02(ER) HARP - Semi Annual
1046448 | 202212 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-26 10:17 | Annual PM | (TC 04/22/2024 10:16:22 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached results
1046449 | 202213 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-26 10:29 | Annual PM | (TC 04/23/2024 10:28:13 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached results
1046450 | 202214 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-26 10:27 | Annual PM | (TC 04/23/2024 10:26:44 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Attached results
1046787 | 202215 | PM | ACTIV | READY | 2024-03-01 02:15 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1046788 | 202216 | PM | ACTIV | READY | 2024-03-01 02:15 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1047830 | 202217 | PM | ACTIV | ISSUE | 2024-03-01 02:15 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
1048529 | 202218 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-03-25 09:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 09:38:26 AM, Rhuland, Courtney) Calibration and PM completed All tests passed Refer to ASM for details
1048598 | 202219 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-18 13:19 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/18/2024 01:18:20 PM, Rhuland, Courtney) PM completed Refer to procedure for details
1048599 | 202220 | PM | CNP | ENL | 2024-03-01 02:15 | 2024-09-30 07:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048600 | 202221 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-19 13:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/19/2024 01:05:14 PM, Rhuland, Courtney) PM completed Batteries replaced Refer to procedure for details
1048601 | 202222 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-22 14:35 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/22/2024 02:34:58 PM, Rhuland, Courtney) Batteries replaced PM completed Refer to procedure for details
1048604 | 202223 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-06-10 13:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 04/24/2024 02:05:36 PM, Rhuland, Courtney) Screen was found to be broken 14S uses the pump and agreed to cover the cost of repair (TC 06/10/2024 01:57:40 PM, Rhuland, Courtney) Returned from Smiths medical updated drug library to paliative Replaced batteries
1048605 | 202224 | PM | CNP | ENL | 2024-03-01 02:15 | 2024-09-30 07:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1048628 | 202225 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-19 09:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:05:33 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048630 | 202226 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-19 09:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048631 | 202227 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-04-19 09:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048633 | 202228 | PM | CMPLT | CMPLT | 2024-03-01 02:15 | 2024-05-03 09:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1048635 | 202229 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD (TC 04/16/2024 09:20:04 AM, Huang, Jessica) replaced the female iui connector completed PM
1048638 | 202230 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:05:09 AM, Huang, Jessica) replaced the female iui connector completed PM
1048640 | 202231 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 09:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048644 | 202232 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 11:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 03:50:54 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048657 | 202233 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:02:14 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048661 | 202234 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 11:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:13:26 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048665 | 202235 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 03:49:28 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048672 | 202236 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 11:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048673 | 202237 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 11:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 05:21:14 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048675 | 202238 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 11:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048680 | 202239 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 13:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048683 | 202240 | PM | CANCL | CANCL | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048684 | 202241 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 13:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:07:56 AM, Huang, Jessica) replaced the female iui connector completed PM
1048685 | 202242 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 13:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048689 | 202243 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048696 | 202244 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 13:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1050251 | 202245 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-13 14:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:25:46 PM, Baghirzade, Nara) PM completed by BD.
1050577 | 202246 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-18 10:44 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | (TC 04/15/2024 08:55:03 AM, Rendon Lopez , Elizabeth) 6000hr PM due. Disassembled, inspected, replacing PM parts (TC 04/16/2024 10:42:55 AM, Rendon Lopez , Elizabeth) Finished parts replacement. Completed calibration. (TC 04/17/2024 10:42:55 AM, Rendon Lopez , Elizabeth) Functional tests done. PM completed, all tests passed. Checklist sent to Bunnell for update
1050578 | 202247 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-18 10:46 | PM-MAJOR - Annual PM For: Wisper Jet Patient Box | (TC 04/15/2024 08:53:37 AM, Rendon Lopez , Elizabeth) Disassambled, cleaned, replaced PM parts, reassembled (TC 04/16/2024 10:45:10 AM, Rendon Lopez , Elizabeth) Calibration tests completed (TC 04/17/2024 10:45:10 AM, Rendon Lopez , Elizabeth) Functional tests completed. PM completed, all tests passed. Sent checklist to Bunnell for updated records
1050623 | 202248 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-13 13:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 01:23:06 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1053349 | 202249 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-15 17:09 | MSH BI 1274 H2 HARP TRGT APR 1 SEMI-ANNUAL MSH BI 1274 |  | MSH BI 1274 H2 HARP TRGT APR 1 SEMI-ANNUAL MSH BI 1274 HARP completed.
1053452 | 202250 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM MAJOR-Stretchers, Mobile, Hospital, Patient Transfer
1054411 | 202251 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-07-25 14:46 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 07/25/2024 02:45:54 PM, Canjimavadivel, Kamallesh ) Verified calibration, batteries ok.
1054448 | 202252 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1054460 | 202253 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-07-04 10:30 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 07/04/2024 10:30:00 AM, Canjimavadivel, Kamallesh ) Replaced and lubricated o-rings Cleaned fan fliter All tests pass
1054462 | 202254 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-07-11 10:34 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 07/11/2024 10:33:56 AM, Canjimavadivel, Kamallesh ) O-rings replaced Filter cleaned All tests pass PM complete
1054464 | 202255 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054465 | 202256 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054466 | 202257 | PM | CMPLT | PMFAIL | 2024-03-01 02:16 | 2024-09-04 13:57 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/04/2024 01:56:56 PM, Cortez, John) error code, "CPR switch failed" no pendant turn assist fail-left bottom rail installed nurse call cable
1054467 | 202258 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054468 | 202259 | PM | CMPLT | PMFAIL | 2024-03-01 02:16 | 2024-07-12 15:26 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 07/12/2024 03:25:39 PM, Canjimavadivel, Kamallesh ) Started PM and found damaged component. Contacted tech support and was able to get required part number and service manual. Received quote for part.
1054469 | 202260 | PM | CMPLT | PMFAIL | 2024-03-01 02:16 | 2024-09-05 11:29 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/05/2024 11:29:14 AM, Cortez, John) PM fail, turn assist-left head rail , right bottom rail sensor installed nurse call cable
1054470 | 202261 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054471 | 202262 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054473 | 202263 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054474 | 202264 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054475 | 202265 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1054479 | 202266 | PM | CMPLT | PMFAIL | 2024-03-01 02:16 | 2024-09-04 14:04 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/04/2024 02:03:36 PM, Cortez, John) PM failed left siderail screen not responsive, took a while to startup, couldn't calibrate screen
1054541 | 202267 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
1054613 | 202268 | PM | CANCL | CANCL | 2024-03-01 02:16 |  | PM MAJOR-Power Systems, Uninterruptible | Battery just changed and UPS verified Feb 2 2023. See WO 201578. PM rescheduled
1054617 | 202269 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-08 14:03 | PM MAJOR-Power Systems, Uninterruptible | (TC 05/08/2024 02:03:01 PM, Scero, Tomlyn) battery replaced. PM tests complete and passed.
1054806 | 202270 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
1055103 | 202271 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-04 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055122 | 202272 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055181 | 202273 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-14 08:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055642 | 202274 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1055643 | 202275 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-08-14 16:06 | PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 08/14/2024 04:05:19 PM, Cortez, John) 70 mOhms, 22 uAmps Pendant missing clip PM passed
1055644 | 202276 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1055645 | 202277 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1055646 | 202278 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric
1055647 | 202279 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1056137 | 202280 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-07-26 08:30 | PM MAJOR-Power Systems, Uninterruptible | (TC 07/26/2024 08:30:00 AM, Canjimavadivel, Kamallesh ) Replaced battery
1057846 | 202281 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-09-17 11:45 | PM-MAJOR - Annual PM For: Insufflators | (TC 07/31/2024 11:45:25 AM, Huang, Jessica) sent the device to conmed for PM from OR 9 installed it in OR 8 after it was received
1058172 | 202282 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-22 09:13 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/22/2024 09:13:10 AM, Reid, Emmerson ) Replaced batteries. PM Complete, see procedure for details.
1058173 | 202283 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058174 | 202284 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-21 15:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/21/2024 03:21:26 PM, Reid, Emmerson ) PM complete, see procedure for details.
1058177 | 202285 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-22 08:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/22/2024 08:39:14 AM, Reid, Emmerson ) PM Complete, see procedure for details. Also replaced the batteries (TC 05/22/2024 08:39:14 AM, Reid, Emmerson ) PM Complete, see procedure for details. Also replaced the batteries (TC 05/22/2024 08:39:14 AM, Reid, Emmerson ) PM Complete, see procedure for details. Also replaced the batteries
1058180 | 202286 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-06 10:19 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/06/2024 10:18:15 AM, Reid, Emmerson ) PM Complete, see procedure for details.
1058187 | 202287 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-06-14 09:18 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/14/2024 09:17:45 AM, Reid, Emmerson ) PM complete, see procedure for details.
1058192 | 202288 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-13 11:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/13/2024 11:21:10 AM, Reid, Emmerson ) PM complete, see procedure for details.
1058195 | 202289 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058229 | 202290 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-15 14:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058235 | 202291 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-28 11:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1058261 | 202292 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-12 13:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058275 | 202293 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-08 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/08/2024 02:42:40 PM, Rhuland, Courtney) PM completed by BD Pushed network configuration and data set
1058286 | 202294 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-12 09:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:35:15 AM, Rendon Lopez , Elizabeth) PM completed by BD
1059447 | 202295 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059461 | 202296 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059465 | 202297 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059869 | 202298 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
1059876 | 202299 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-05-23 12:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/23/2024 12:24:18 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1059883 | 202300 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-11 19:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059884 | 202301 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-08-16 10:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1059980 | 202302 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-11-12 15:16 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 11/12/2024 03:16:07 PM, Alvarado, Karla) PM Completed See procedure for details
1059981 | 202303 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1059982 | 202304 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1059983 | 202305 | PM | CANCL | CANCL | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 03/01/2024 04:29:52 PM, Cortez, John) PM completed recently, will close WO
1059984 | 202306 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1060164 | 202307 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-01 02:16 | Sonesta 3 Exam Table
1063193 | 202308 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-17 09:51 | MSH 5 RPCU IRUS XG "BONNIE" PM - Semi-Annual | (TC 03/12/2024 11:38:01 AM, Whittington, Scott) MSH 5 RPCU IRUS XG "BONNIE" PM - Semi-Annual | MSH 5 RPCU IRUS XG "BONNIE" PM - Semi-Annual
1063372 | 202309 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Elecric |  | Bed under contract ( Annual -PM ) performed by Vendor ( Hill-Rom )
1064964 | 202310 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-19 16:43 | HARP | HARP
1064964 | 202311 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-18 16:44 | PM - Semi-Annual | PM
1065463 | 202312 | PM | ACTIV | ISSUE | 2024-03-01 02:16 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
1065513 | 202313 | PM | CANCL | CANCL | 2024-03-01 02:16 | 2024-10-02 07:51 | PM-MAJOR - Annual PM For: Waste-Disposal Units
1065514 | 202314 | PM | CANCL | CANCL | 2024-03-01 02:16 | 2024-10-02 07:55 | PM-MAJOR - Annual PM For: Waste-Disposal Units
1065516 | 202315 | PM | CANCL | CANCL | 2024-03-01 02:16 | 2024-10-02 07:46 | PM-MAJOR - Annual PM For: Waste-Disposal Units
1066474 | 202316 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-10-02 08:14 | PM-MAJOR - Annual PM For: Waste-Disposal Units
1066475 | 202317 | PM | CANCL | CANCL | 2024-03-01 02:16 | 2024-10-02 08:12 | PM-MAJOR - Annual PM For: Waste-Disposal Units
1066476 | 202318 | PM | CANCL | CANCL | 2024-03-01 02:16 | 2024-10-02 08:09 | PM-MAJOR - Annual PM For: Waste-Disposal Units
1066756 | 202319 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-07-16 10:00 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 07/16/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Replaced o-rings and lubricated All tests pass PM complete
1066757 | 202320 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-07-11 14:04 | PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 07/11/2024 02:04:08 PM, Canjimavadivel, Kamallesh ) Replaced o-rings Cleaned filter All tests pass PM complete
1066770 | 202321 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-07-22 12:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/22/2024 12:14:31 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
1066771 | 202322 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-04-16 08:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066772 | 202323 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-03-18 11:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066917 | 202324 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR-Annual PM For: Bariatric Bed
1066918 | 202325 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-08-26 11:09 | PM-MAJOR-Annual PM For: Bariatric Bed | (TC 08/26/2024 11:08:48 AM, Canjimavadivel, Kamallesh ) Completed by Stryker
1066919 | 202326 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-09-05 07:52 | PM-MAJOR-Annual PM For: Bariatric Bed | Completed by Stryker
1066920 | 202327 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-09-05 07:53 | PM-MAJOR-Annual PM For: Bariatric Bed | Completed by Stryker
1066921 | 202328 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-08-26 11:11 | PM-MAJOR-Annual PM For: Bariatric Bed | (TC 08/26/2024 11:11:07 AM, Canjimavadivel, Kamallesh ) Completed by Stryker
1066922 | 202329 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-08-26 11:05 | PM-MAJOR-Annual PM For: Bariatric Bed | (TC 08/26/2024 11:05:06 AM, Canjimavadivel, Kamallesh ) Completed by Stryker
1067048 | 202330 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-09-04 09:34 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 09/04/2024 09:33:48 AM, Cortez, John) PM complete
1067049 | 202331 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-09-04 09:28 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 09/04/2024 09:28:20 AM, Cortez, John) PM complete, pendant missing clip
1067050 | 202332 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
1067051 | 202333 | PM | CMPLT | CMPLT | 2024-03-01 02:16 | 2024-09-05 10:25 | PM-MAJOR- Factory PM for: ICU Bed Electric | (TC 09/05/2024 10:25:05 AM, Cortez, John) PM complete Did have "battery voltage less than..." error, but cleared after restart and didn't return during PM
1067052 | 202334 | PM | ACTIV | READY | 2024-03-01 02:16 |  | PM-MAJOR- Factory PM for: ICU Bed Electric
| 202335 | CM | CMPLT | CMPLT | 2024-03-01 07:42 | 2024-04-04 09:38 | Ordering BP cuff for NICU | (TC 03/01/2024 07:42:08 AM, Rhuland, Courtney) NICU has begun receiving shippments of backordered cuffs but are still not receiving size 1 Spectramedx has size 1 cuffs in stock and were requested to hold them over
1062084 | 202336 | CM | CMPLT | CMPLT | 2024-03-01 08:42 | 2024-03-05 15:42 | Error message, scale needs to be calibrated | (TC 02/29/2024 08:44:20 AM, Rendon Lopez , Elizabeth) Confirmed issue, error message. Removed bed from room and swap with other bed. Contacted Mark that was on site. Scale is now calibrated, bed is working
1056125 | 202337 | CM | CMPLT | CMPLT | 2024-03-01 08:50 | 2024-03-05 09:03 | Water level low | (TC 03/01/2024 08:50:03 AM, Canjimavadivel, Kamallesh ) Received email for ED that Altrix water level was low. Went up and found reservoir almost empty. Topped up with water and spoke to David to find out more details on how the water was low (TC 03/05/2024 08:58:48 AM, Canjimavadivel, Kamallesh ) Followed up with David, low water was due to nursing error.
| 202338 | CM | CMPLT | CMPLT | 2024-03-01 09:31 | 2024-04-25 14:51 | Ordering Medfusion parts for stock | (TC 03/01/2024 09:49:57 AM, Rhuland, Courtney) Placed order for medfusion parts for biomed stock SC 1003180519
1070166 | 202339 | CM | CMPLT | CMPLT | 2024-03-01 09:59 | 2024-03-01 09:59 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 130 mOhm; 77 uA. Tested ok. Installed unit in endoscopy.
| 202340 | CM | CMPLT | CMPLT | 2024-03-01 10:02 | 2024-03-01 10:03 | Thulium Fiber Laser Meeting | (TC 03/01/2024 10:02:53 AM, Baghirzade, Nara) Thulium Fiber Laser Meeting
1065746 | 202341 | CM | CMPLT | CMPLT | 2024-03-01 10:04 | 2024-03-01 10:04 | Functional check | (TC 03/01/2024 10:03:51 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
69387 | 202342 | CM | CMPLT | CMPLT | 2024-03-01 10:04 | 2024-03-01 10:05 | Functional check | (TC 03/01/2024 10:04:39 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1070167 | 202343 | CM | CMPLT | CMPLT | 2024-03-01 10:05 | 2024-03-01 10:04 | Incoming Inspection | Received unit. Unboxed and Assembled it. Completed inspection and electrical safety. 110 mOhm; 76 uA. Tested ok. Installed unit in endoscopy.
69444 | 202344 | CM | CMPLT | CMPLT | 2024-03-01 10:24 | 2024-03-01 10:24 | Functional check | (TC 03/01/2024 10:23:58 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1050259 | 202345 | CM | CMPLT | CMPLT | 2024-03-01 10:26 | 2024-03-01 10:26 | Functional check | (TC 03/01/2024 10:25:40 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
68798 | 202346 | CM | CMPLT | CMPLT | 2024-03-01 10:42 | 2024-03-01 10:43 | Functional check | (TC 03/01/2024 10:42:00 AM, Baghirzade, Nara) Downloaded error logs - checked for any battery failure errors. No errors found. Checked battery status; capacity is within nominal range. Switched PCU in regular user mode; battery run time is good. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 140 mohms Leakage current: 74.5 uA
1058276 | 202347 | CM | CMPLT | CMPLT | 2024-03-01 10:43 | 2024-03-01 10:44 | Functional check | (TC 03/01/2024 10:43:31 AM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 124 mohms Leakage current: 67.3 uA
| 202348 | CM | CMPLT | CMPLT | 2024-03-01 11:04 | 2024-03-01 11:04 | worked on TMS | worked on TMS
| 202349 | RDS | CMPLT | CMPLT | 2024-03-01 12:09 | 2024-03-01 12:11 | Weekly Rounds: 11 South | (TC 03/01/2024 12:09:55 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 202350 | CM | CMPLT | CMPLT | 2024-03-01 12:11 | 2024-03-01 12:13 | Weekly Rounds: 11 North | (TC 03/01/2024 12:11:18 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 202351 | RDS | CMPLT | CMPLT | 2024-03-01 12:13 | 2024-03-01 12:15 | Weekly Rounds: NICU & RESUS | (TC 03/01/2024 12:12:56 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
| 202352 | CM | CMPLT | CMPLT | 2024-03-01 12:15 | 2024-03-01 12:17 | Confirmation of baud rates on MX800 monitors | (TC 03/01/2024 12:15:18 PM, Baghirzade, Nara) Kamran requested to confirm the baud rates on MX800 monitors in NICU for Dr. Poorva's study for brain monitoring/trauma and Carol's baby AI study. Checked MX800 in all empty rooms (8 rooms total) and confirmed that the baud rate is 115200; baud rates work for both studies. Followed up with Kamran.
1059763 | 202353 | CM | CMPLT | CMPLT | 2024-03-01 13:45 | 2024-03-01 16:04 | MSH 15 L&D TE-7 D WILL NOT BOOT | (TC 03/01/2024 03:58:02 PM, Dolinar, John) MSH 15 L&D TE-7 D WILL NOT BOOT STAND POWER SUPPLY FAULTY 20. Adapter 022-000187-00 2121 Trolley power adapter 19 V 150 W 022-000194-00 | MSH 15 L&D TE-7 D WILL NOT BOOT STAND POWER SUPPLY FAULTY 20. Adapter 022-000187-00 2121 Trolley power adapter 19 V 150 W 022-000194-00
64995 | 202354 | CM | CMPLT | CMPLT | 2024-03-01 13:55 | 2024-04-12 11:12 | Created by Failed PM WO # 198132 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 03/01/2024 01:55:30 PM, Rendon Lopez , Elizabeth) Testing pump with different settings, same over delivering.Contacted company for RMA. Pump is over 7 years old, repair is $1000 (TC 03/05/2024 03:37:25 PM, Rendon Lopez , Elizabeth) Receved RMA# 318651 To be sent out (TC 04/12/2024 11:11:32 AM, Rendon Lopez , Elizabeth) Received pump, see attached report. Incoming inspection, all tests passed | (TC 02/29/2024 01:54:35 PM, Rendon Lopez , Elizabeth) performing PM, pumps is overdelivering. Tried different cassettes, same issue
| 202355 | CM | CMPLT | CMPLT | 2024-03-01 14:25 | 2024-03-01 14:27 | Otoscope not working | (TC 03/01/2024 02:24:46 PM, Rhuland, Courtney) Light was burnt out on ophthalmoscope Nurse had tried replacing bulb but inserted bulb upside down and was stuck Replaced head with spare Brought broken head to biomed and was able to remove bulb with pliers Functional check - OK
1047851 | 202356 | CM | CMPLT | CMPLT | 2024-03-01 14:28 | 2024-03-04 10:08 | worn clutch | (TC 03/01/2024 02:54:39 PM, Rhuland, Courtney) Took device apart and replaced clutch Greased leadscrew Replaced plunger head assembly (TC 03/04/2024 10:06:27 AM, Rhuland, Courtney) Completed rebuild Motor not running error present Lead screw nut was too tight - loosened and error cleared Performed PM - All tests passed Refer to procedure for details
1062999 | 202357 | CM | CMPLT | CMPLT | 2024-03-01 14:34 | 2024-03-01 14:48 | "equipment malfunction" | (TC 03/01/2024 02:33:54 PM, Cortez, John) Confirmed issue, unit doesn't boot up normally, requires service. Found error codes 85 and 191 in log. UIB and CCB components may be related Swapped CCB, error remains. Swapped 9V battery, boots to regular menu but in service mode asks user to "unplug for 10 seconds" Swapped back original 9V battery, to see if it fails to boot, but same result as new 9V battery DC output on serial cable at 21.5 V Swapped back original CCB. Ran warmer for 1.5 hours, didn't come across issues Will send back to ICU, but made not of error (TC 03/01/2024 02:47:24 PM, Canjimavadivel, Kamallesh ) Used spare Altrix for parts
1049062 | 202358 | CM | CMPLT | CMPLT | 2024-03-01 14:42 | 2024-03-05 10:19 | While using an error code came up | (TC 03/05/2024 10:18:21 AM, Reid, Emmerson ) Ran for an hour twice and no errors came up. Sending back
| 202359 | MEP | CMPLT | CMPLT | 2024-03-01 14:56 | 2024-03-01 15:03 | Meeting - 12 N Walk through with RPO | Meeting with NB. Walk through of 12N with ROP team. Tour of 11N,11S with NB for Bed location positioning.
1047866 | 202360 | CM | CMPLT | CMPLT | 2024-03-01 15:08 | 2024-05-15 12:51 | Plunger was hard to move | (TC 03/01/2024 03:08:12 PM, Baghirzade, Nara) Plunger was difficult to move. Disassembled the device, found that an O-ring seal was caught between the plunger and case, causing it to get stuck. Removed o-ring from device. Assembled the device and when doing a functional check, the device was failing the occlusion test, giving the error "coarse travel/check plunger or clutch". Will inspect the clutch for wear. (TC 04/26/2024 10:28:03 AM, Baghirzade, Nara) Replaced clutch entirely. Assembled the pump back together; plunger was still hard to move and was loud when moving. Disassembled the pump. lubricated the leadscrew, assembled the pump again. Still finding the plunger hard to move. Will keep looking for the cause of this issue. (TC 04/30/2024 10:18:41 AM, Baghirzade, Nara) Disassembled the syringe pump; cleaned and greased the guide rod. Ensured that the data cable was in place in the black notch. Put the device back together and there were no errors or difficulties moving the plunger. Completed a full calibration of the syringe size and position, force sensor, plunger position, and syringe size sensor. When running the pump for the occlusion test, the first 30 mL had no issue however when it got down to the last 30 mL, before the occlusion even occurs, there was an error message "syringe plunger not in place". Ran the digital sensor test as per the manual instructions however could not determine anythihng from that. Ran the test again and had the same result. Repeated calibration for the force sensor, ran the test again and had the same result. Replaced the plungerhead assembly, ran the test again and had the same result. Will keep looking for the root cause. (TC 05/07/2024 01:17:41 PM, Baghirzade, Nara) Disassembled the device again; replaced the plunger cable. When assembling the device, the plunger cable was not seated correctly. Will need to disassemble the device and re-seat the plunger cable. (TC 05/15/2024 12:50:04 PM, Baghirzade, Nara) Disassembled the entire device, noticed that the plunger cable doesn't sit properly and its preventing the plunger from moving properly. Assembled the device. Ran a functional test and no errors or issues observed. Occlusion is functional.
| 202361 | CM | CMPLT | CMPLT | 2024-03-01 15:16 | 2024-03-01 15:16 | 0.25 | (TC 03/01/2024 03:15:37 PM, Baghirzade, Nara) Completed the daily on-call checks. Sent email with details to biomedical engineering for: -Philips -Masimo -Zoll defibrillators -Stryker beds
| 202362 | CM | CMPLT | CMPLT | 2024-03-01 15:17 | 2024-03-01 15:18 | Meeting | Meeting with Steris rep with regards to Anita Chin's request to move around the equipment on Emergency video cart. Relocated printer. Found CO2 hose is too long. Devid murphy to find an alternative length hose.
| 202363 | CM | CMPLT | CMPLT | 2024-03-01 15:36 | 2024-04-03 16:08 | misc biomed - Mar 2024 | (TC 03/01/2024 03:36:12 PM, Cortez, John) Talked with Ivy again about power bar usage in IR rooms. provided ordering info for medical grade power bars (patient care areas) (TC 03/05/2024 04:27:06 PM, Cortez, John) Medication cart sent down from 10N. Noticed mouse, keybord cables were damaged, Brought cart over to IT and they replaced accessories. Sent back to 10N (TC 03/04/2024 04:28:20 PM, Cortez, John) huddle (TC 03/06/2024 11:56:05 AM, Cortez, John) infant scale meeting (TC 03/07/2024 04:26:53 PM, Cortez, John) systems meeting (TC 03/06/2024 04:29:21 PM, Cortez, John) Contacted Rita from Vaughan clinic about scale in Triage B, as well as the need for a wider scale for larger patients. Looped in Naz (TC 03/12/2024 10:24:23 AM, Cortez, John) Checked Bariatric bed in ICU pod B room 22, keeps alarming. User had to acknowledge on-screen prompt to confirm set bed width (TC 03/12/2024 09:27:45 AM, Cortez, John) 10N had issue with suction regulator in room 44. After troubleshooting, confirmed regulator and suction from wall was ok. Nurse replaced hose/tube and suction improved (TC 03/12/2024 09:29:53 AM, Cortez, John) drove to Vaughan clinic (TC 03/12/2024 09:30:22 AM, Cortez, John) Visited Vaughan for infant scale delivery and to look at issue with scale in Triage B. Height bar was loose, tightened. Infant scale delivery was cancelled, Vendor mixed up delivery day. Pushed to Mar 18. Talked with Rita about need for wider scale for larger patients (TC 03/13/2024 11:01:47 AM, Cortez, John) Closing WOs for pumps on TMS (TC 03/14/2024 04:38:12 PM, Cortez, John) TMS BD workorders (TC 03/15/2024 05:10:44 PM, Cortez, John) TMS WOs (TC 03/19/2024 02:13:16 PM, Cortez, John) Looked into APs marked as unregistered Reseating patch cables from switch didn't resolve issue (TC 03/19/2024 10:09:07 AM, Cortez, John) travel to Vaughan (TC 03/20/2024 04:07:03 PM, Cortez, John) Went with Kamran to look at unregistered AP issue (TC 03/20/2024 04:50:28 PM, Cortez, John) Retrieved spare pumps for Barb (10S) (TC 03/20/2024 04:08:10 PM, Cortez, John) Went back to clear unregistered AP list (TC 03/22/2024 04:08:40 PM, Cortez, John) 1 on 1 (TC 03/21/2024 04:10:55 PM, Cortez, John) Dropped off hyfrecator to Vaughan (TC 03/22/2024 04:11:20 PM, Cortez, John) TMS BD WOs (TC 03/25/2024 01:29:18 PM, Cortez, John) TMS BD WOs (TC 03/25/2024 01:29:51 PM, Cortez, John) replace pump IUIs (TC 03/26/2024 04:11:58 PM, Cortez, John) Replaced IUIs for BD (TC 03/27/2024 04:13:55 PM, Cortez, John) PM meeting (TC 03/28/2024 02:27:52 PM, Cortez, John) Measurement ext. malf in Y42 emerg, Room busy, will return to try to swap FMS rack
1046059 | 202364 | PS | ACTIV | ISSUE | 2024-03-01 15:53 |  | 2024 Alaris Software Recall
1046072 | 202365 | PS | CANCL | CANCL | 2024-03-01 15:53 | 2024-10-03 12:54 | 2024 Alaris Software Recall
1046073 | 202366 | PS | CANCL | CANCL | 2024-03-01 15:53 | 2024-10-03 12:54 | 2024 Alaris Software Recall
1046075 | 202367 | PS | CANCL | CANCL | 2024-03-01 15:53 | 2024-10-03 12:55 | 2024 Alaris Software Recall
1047003 | 202368 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:15 | 2024 Alaris Software Recall
1047004 | 202369 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:22 | 2024 Alaris Software Recall
1047006 | 202370 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:27 | 2024 Alaris Software Recall
1047007 | 202371 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:32 | 2024 Alaris Software Recall
1047010 | 202372 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:39 | 2024 Alaris Software Recall
1047011 | 202373 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:47 | 2024 Alaris Software Recall
1047014 | 202374 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:57 | 2024 Alaris Software Recall
1047015 | 202375 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 13:07 | 2024 Alaris Software Recall
1047034 | 202376 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-21 13:56 | 2024 Alaris Software Recall
1047035 | 202377 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-19 13:51 | 2024 Alaris Software Recall
1047036 | 202378 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-05-16 11:35 | 2024 Alaris Software Recall | (TC 05/16/2024 11:34:34 AM, Baghirzade, Nara) Software upgrade completed.
1047037 | 202379 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-28 11:17 | 2024 Alaris Software Recall
1047038 | 202380 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-02 14:45 | 2024 Alaris Software Recall
1047039 | 202381 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-14 13:56 | 2024 Alaris Software Recall | (TC 03/06/2024 01:56:40 PM, Huang, Jessica) completed by BD
1047040 | 202382 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-11-06 13:39 | 2024 Alaris Software Recall | (TC 11/06/2024 01:39:13 PM, Baghirzade, Nara) Software upgraded.
1047041 | 202383 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-15 14:28 | 2024 Alaris Software Recall | Completed by BD
1047042 | 202384 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-04 15:57 | 2024 Alaris Software Recall | Recall completed by BD
1047043 | 202385 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-07-03 13:34 | 2024 Alaris Software Recall
1047044 | 202386 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-15 08:45 | 2024 Alaris Software Recall
1047045 | 202387 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-11 19:07 | 2024 Alaris Software Recall
1047046 | 202388 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-28 11:18 | 2024 Alaris Software Recall | Completed
1047047 | 202389 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-12 10:46 | 2024 Alaris Software Recall | (TC 03/12/2024 10:45:43 AM, Rendon Lopez , Elizabeth) Recall completed
1047048 | 202390 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-27 11:21 | 2024 Alaris Software Recall | Completed
1047049 | 202391 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 13:23 | 2024 Alaris Software Recall
1047050 | 202392 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 13:29 | 2024 Alaris Software Recall
1047062 | 202393 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-20 14:53 | 2024 Alaris Software Recall
1047069 | 202394 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-26 12:15 | 2024 Alaris Software Recall | Completed
1047095 | 202395 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-13 09:26 | 2024 Alaris Software Recall | Recall completed
1047096 | 202396 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-02 11:47 | 2024 Alaris Software Recall
1047097 | 202397 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-20 14:56 | 2024 Alaris Software Recall
1047098 | 202398 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-08-13 07:42 | 2024 Alaris Software Recall
1047099 | 202399 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-20 14:59 | 2024 Alaris Software Recall
1047100 | 202400 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-20 15:01 | 2024 Alaris Software Recall
1047101 | 202401 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-13 09:29 | 2024 Alaris Software Recall | Recall completed
1047102 | 202402 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-20 15:03 | 2024 Alaris Software Recall
1047104 | 202403 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-12 12:57 | 2024 Alaris Software Recall | Completed by BD
1048522 | 202404 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-28 13:11 | 2024 Alaris Software Recall | (TC 03/28/2024 01:10:26 PM, Rhuland, Courtney) Recall completed by BD
1048523 | 202405 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-07-31 09:16 | 2024 Alaris Software Recall
1048524 | 202406 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-17 10:55 | 2024 Alaris Software Recall | Completed by BD
1048525 | 202407 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-22 12:00 | 2024 Alaris Software Recall | (TC 03/22/2024 11:59:22 AM, Baghirzade, Nara) Software recall completed by BD.
1048526 | 202408 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-10 14:50 | 2024 Alaris Software Recall | (TC 04/10/2024 02:49:47 PM, Rhuland, Courtney) Recall completed by BD
1048527 | 202409 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-14 14:47 | 2024 Alaris Software Recall | (TC 03/14/2024 02:46:23 PM, Baghirzade, Nara) Software recall completed by BD.
1048528 | 202410 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-07-02 10:17 | 2024 Alaris Software Recall | (TC 07/02/2024 10:16:45 AM, Rhuland, Courtney) Syringe flashed and software updated
1048529 | 202411 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-22 13:43 | 2024 Alaris Software Recall | (TC 03/22/2024 01:43:01 PM, Rhuland, Courtney) Recall completed by BD
1048530 | 202412 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-17 10:27 | 2024 Alaris Software Recall | Completed by BD
1048531 | 202413 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-03-13 11:28 | 2024 Alaris Software Recall | (TC 03/13/2024 03:53:27 PM, Patel, Rutvik) BD returned unit indicating Rear cover need replacement. Cleaned unit. Found fluid ingress inside until when I removed rear case. Cleaned everything and air dried it. Repaced rear cover and S/N Sticker. My computer does not have ASM. The unit will be PMed by the pump team.
1048627 | 202414 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:36 | 2024 Alaris Software Recall | Completed by BD
1048628 | 202415 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:37 | 2024 Alaris Software Recall | Completed by BD
1048629 | 202416 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:38 | 2024 Alaris Software Recall | Completed by BD
1048630 | 202417 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:39 | 2024 Alaris Software Recall | Completed by BD
1048631 | 202418 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:43 | 2024 Alaris Software Recall | Completed by BD
1048632 | 202419 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-22 08:26 | 2024 Alaris Software Recall | Completed by BD
1048633 | 202420 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-05-03 09:19 | 2024 Alaris Software Recall | Completed by BD.
1048634 | 202421 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:49 | 2024 Alaris Software Recall | Completed by BD
1048635 | 202422 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:50 | 2024 Alaris Software Recall | Completed by BD
1048636 | 202423 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:50 | 2024 Alaris Software Recall | Completed by BD
1048637 | 202424 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:51 | 2024 Alaris Software Recall | Completed by BD
1048638 | 202425 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:53 | 2024 Alaris Software Recall | Completed by BD
1048639 | 202426 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:55 | 2024 Alaris Software Recall | Completed by BD
1048640 | 202427 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:57 | 2024 Alaris Software Recall | Completed by BD
1048641 | 202428 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 09:59 | 2024 Alaris Software Recall | Completed by BD
1048642 | 202429 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 10:00 | 2024 Alaris Software Recall | Completed by BD
1048643 | 202430 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:03 | 2024 Alaris Software Recall | Completed by BD
1048644 | 202431 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:05 | 2024 Alaris Software Recall | Completed by BD
1048645 | 202432 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:06 | 2024 Alaris Software Recall | Completed by BD
1048646 | 202433 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:07 | 2024 Alaris Software Recall | Completed by BD
1048647 | 202434 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:08 | 2024 Alaris Software Recall | Completed by BD
1048648 | 202435 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:09 | 2024 Alaris Software Recall | Completed by BD
1048649 | 202436 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:10 | 2024 Alaris Software Recall | Completed by BD
1048650 | 202437 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:11 | 2024 Alaris Software Recall | Completed by BD
1048651 | 202438 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:12 | 2024 Alaris Software Recall | Completed by BD
1048652 | 202439 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:13 | 2024 Alaris Software Recall | Completed by BD
1048653 | 202440 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:14 | 2024 Alaris Software Recall | Completed by BD
1048654 | 202441 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:14 | 2024 Alaris Software Recall | Completed by BD
1048655 | 202442 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:15 | 2024 Alaris Software Recall | Completed by BD
1048656 | 202443 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:16 | 2024 Alaris Software Recall | Completed by BD
1048657 | 202444 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:17 | 2024 Alaris Software Recall | Completed by BD
1048658 | 202445 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:18 | 2024 Alaris Software Recall | Completed by BD
1048659 | 202446 | PS | ACTIV | ISSUE | 2024-03-01 15:53 |  | 2024 Alaris Software Recall
1048660 | 202447 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:19 | 2024 Alaris Software Recall | Completed by BD
1048661 | 202448 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:20 | 2024 Alaris Software Recall | Completed by BD
1048662 | 202449 | PS | ACTIV | ISSUE | 2024-03-01 15:53 |  | 2024 Alaris Software Recall
1048663 | 202450 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:34 | 2024 Alaris Software Recall | Completed by BD
1048664 | 202451 | PS | ACTIV | ISSUE | 2024-03-01 15:53 |  | 2024 Alaris Software Recall
1048665 | 202452 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:34 | 2024 Alaris Software Recall | Completed by BD
1048666 | 202453 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:35 | 2024 Alaris Software Recall | Completed by BD
1048667 | 202454 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:36 | 2024 Alaris Software Recall | Completed by BD
1048668 | 202455 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:39 | 2024 Alaris Software Recall | Completed by BD
1048669 | 202456 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:41 | 2024 Alaris Software Recall | Completed by BD
1048670 | 202457 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:42 | 2024 Alaris Software Recall | Completed by BD
1048671 | 202458 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:44 | 2024 Alaris Software Recall | Completed by BD
1048672 | 202459 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:44 | 2024 Alaris Software Recall | Completed by BD
1048673 | 202460 | PS | CMPLT | CMPLT | 2024-03-01 15:53 | 2024-04-19 11:45 | 2024 Alaris Software Recall | Completed by BD
1048674 | 202461 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 11:46 | 2024 Alaris Software Recall | Completed by BD
1048675 | 202462 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 11:47 | 2024 Alaris Software Recall | Completed by BD
1048676 | 202463 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:27 | 2024 Alaris Software Recall | Completed by BD
1048677 | 202464 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:28 | 2024 Alaris Software Recall | Completed by BD
1048678 | 202465 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:29 | 2024 Alaris Software Recall | Completed by BD
1048679 | 202466 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:30 | 2024 Alaris Software Recall | Completed by BD
1048680 | 202467 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:31 | 2024 Alaris Software Recall | Completed by BD
1048681 | 202468 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:31 | 2024 Alaris Software Recall | Completed by BD
1048682 | 202469 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:32 | 2024 Alaris Software Recall | Completed by BD
1048683 | 202470 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:33 | 2024 Alaris Software Recall | Completed by BD
1048684 | 202471 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:34 | 2024 Alaris Software Recall | Completed by BD
1048685 | 202472 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:35 | 2024 Alaris Software Recall | Completed by BD
1048686 | 202473 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:36 | 2024 Alaris Software Recall | Completed by BD
1048687 | 202474 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:37 | 2024 Alaris Software Recall | Completed by BD
1048688 | 202475 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:38 | 2024 Alaris Software Recall | Completed by BD
1048689 | 202476 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:39 | 2024 Alaris Software Recall | Completed by BD
1048690 | 202477 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:40 | 2024 Alaris Software Recall | Completed by BD
1048691 | 202478 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:40 | 2024 Alaris Software Recall | Completed by BD
1048692 | 202479 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:41 | 2024 Alaris Software Recall | Completed by BD
1048693 | 202480 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:42 | 2024 Alaris Software Recall | Completed by BD
1048694 | 202481 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:44 | 2024 Alaris Software Recall | Completed by BD
1048695 | 202482 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:45 | 2024 Alaris Software Recall | Completed by BD
1048696 | 202483 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 13:47 | 2024 Alaris Software Recall | Completed by BD
1048697 | 202484 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1048698 | 202485 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-22 17:11 | 2024 Alaris Software Recall | completed BD
1048699 | 202486 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1048700 | 202487 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-22 17:17 | 2024 Alaris Software Recall | completed by BD
1048701 | 202488 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1050035 | 202489 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-26 14:41 | 2024 Alaris Software Recall | Completed
1050039 | 202490 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-19 13:53 | 2024 Alaris Software Recall
1050045 | 202491 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 07:51 | 2024 Alaris Software Recall
1050240 | 202492 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-26 14:46 | 2024 Alaris Software Recall | Completed
1050244 | 202493 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 09:30 | 2024 Alaris Software Recall | Recall completed
1050247 | 202494 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 07:54 | 2024 Alaris Software Recall | Completed by BD.
1050249 | 202495 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-26 12:16 | 2024 Alaris Software Recall | Completed
1050250 | 202496 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 10:27 | 2024 Alaris Software Recall
1050251 | 202497 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 14:25 | 2024 Alaris Software Recall | (TC 03/13/2024 02:25:11 PM, Baghirzade, Nara) Software recall completed by BD.
1050252 | 202498 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-15 11:29 | 2024 Alaris Software Recall | Completed by BD
1050253 | 202499 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-15 11:28 | 2024 Alaris Software Recall | Completed by BD
1050256 | 202500 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 18:12 | 2024 Alaris Software Recall
1050257 | 202501 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 07:57 | 2024 Alaris Software Recall | Completed by BD.
1050258 | 202502 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-14 07:57 | 2024 Alaris Software Recall | Completed by BD
1050259 | 202503 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 07:59 | 2024 Alaris Software Recall | Completed by BD.
1050260 | 202504 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 18:15 | 2024 Alaris Software Recall
1050263 | 202505 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-27 11:22 | 2024 Alaris Software Recall | Completed
1050264 | 202506 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-06-17 13:31 | 2024 Alaris Software Recall | (TC 06/17/2024 01:30:39 PM, Rendon Lopez , Elizabeth) Software updated
1050265 | 202507 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-15 14:29 | 2024 Alaris Software Recall | Completed by BD
1050266 | 202508 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:52 | 2024 Alaris Software Recall
1050267 | 202509 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-27 11:24 | 2024 Alaris Software Recall | Completed
1050269 | 202510 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 18:18 | 2024 Alaris Software Recall
1050270 | 202511 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 11:01 | 2024 Alaris Software Recall
1050271 | 202512 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-23 09:36 | 2024 Alaris Software Recall
1050272 | 202513 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1050273 | 202514 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-15 14:31 | 2024 Alaris Software Recall | Completed by BD
1050274 | 202515 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-16 13:06 | 2024 Alaris Software Recall | (TC 04/16/2024 01:05:48 PM, Baghirzade, Nara) Software recall completed by BD.
1050614 | 202516 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-25 13:45 | 2024 Alaris Software Recall | (TC 03/25/2024 01:44:21 PM, Baghirzade, Nara) Software recall completed by BD.
1050615 | 202517 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-05 12:17 | 2024 Alaris Software Recall | (TC 04/05/2024 12:17:12 PM, Rhuland, Courtney) Recall completed by BD
1050616 | 202518 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-10 14:52 | 2024 Alaris Software Recall | (TC 04/10/2024 02:51:15 PM, Rhuland, Courtney) Recall completed by BD
1050617 | 202519 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 14:42 | 2024 Alaris Software Recall | (TC 03/13/2024 02:41:23 PM, Rhuland, Courtney) Recall completed by BD
1050618 | 202520 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-25 13:20 | 2024 Alaris Software Recall | (TC 03/25/2024 01:20:10 PM, Baghirzade, Nara) Software recall completed by BD.
1050619 | 202521 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:04 | 2024 Alaris Software Recall | (TC 03/15/2024 02:03:40 PM, Rhuland, Courtney) Recall completed by BD
1050620 | 202522 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-07-30 13:31 | 2024 Alaris Software Recall | Software updated.
1050622 | 202523 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-14 10:40 | 2024 Alaris Software Recall | Recall completed
1050623 | 202524 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 13:24 | 2024 Alaris Software Recall | (TC 03/13/2024 01:23:36 PM, Baghirzade, Nara) Software recall completed by BD.
1050773 | 202525 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 14:42 | 2024 Alaris Software Recall | Completed by BD
1051030 | 202526 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-17 08:54 | 2024 Alaris Software Recall | (TC 04/17/2024 08:53:52 AM, Rhuland, Courtney) Recall completed by BD
1051369 | 202527 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-30 14:00 | 2024 Alaris Software Recall
1051370 | 202528 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:39 | 2024 Alaris Software Recall
1051371 | 202529 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-26 10:22 | 2024 Alaris Software Recall
1051372 | 202530 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-22 13:25 | 2024 Alaris Software Recall
1051373 | 202531 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-28 11:21 | 2024 Alaris Software Recall | Completed
1051374 | 202532 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-15 11:26 | 2024 Alaris Software Recall | Completed by BD
1051375 | 202533 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-27 14:43 | 2024 Alaris Software Recall
1051376 | 202534 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-18 15:09 | 2024 Alaris Software Recall
1051377 | 202535 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-26 12:18 | 2024 Alaris Software Recall | Completed
1051378 | 202536 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-17 08:50 | 2024 Alaris Software Recall | Completed by BD
1055094 | 202537 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-03 14:19 | 2024 Alaris Software Recall | (TC 04/03/2024 02:18:31 PM, Rhuland, Courtney) Recall completed by BD
1055097 | 202538 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:58 | 2024 Alaris Software Recall
1055102 | 202539 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 08:03 | 2024 Alaris Software Recall | Completed by BD.
1055103 | 202540 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-07-31 09:50 | 2024 Alaris Software Recall
1055104 | 202541 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-14 10:25 | 2024 Alaris Software Recall | (TC 03/04/2024 10:25:25 AM, Huang, Jessica) BD completed
1055105 | 202542 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 13:00 | 2024 Alaris Software Recall | Completed by BD
1055106 | 202543 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-02 11:57 | 2024 Alaris Software Recall
1055107 | 202544 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-17 08:51 | 2024 Alaris Software Recall | Completed by BD
1055108 | 202545 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 08:05 | 2024 Alaris Software Recall | Completed by BD.
1055109 | 202546 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-02 11:58 | 2024 Alaris Software Recall
1055110 | 202547 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-18 15:10 | 2024 Alaris Software Recall
1055111 | 202548 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 08:08 | 2024 Alaris Software Recall | Completed by BD.
1055112 | 202549 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-13 08:09 | 2024 Alaris Software Recall | Completed by BD.
1055113 | 202550 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 10:31 | 2024 Alaris Software Recall
1055114 | 202551 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 09:49 | 2024 Alaris Software Recall | Software recall completed by BD.
1055115 | 202552 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 09:46 | 2024 Alaris Software Recall | Recall completed
1055116 | 202553 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 09:50 | 2024 Alaris Software Recall | Completed by BD.
1055117 | 202554 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-14 15:48 | 2024 Alaris Software Recall | Recall completed
1055118 | 202555 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-25 10:54 | 2024 Alaris Software Recall
1055119 | 202556 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 09:32 | 2024 Alaris Software Recall | Recall completed
1055120 | 202557 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 09:44 | 2024 Alaris Software Recall | Recall completed
1055121 | 202558 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-03 14:21 | 2024 Alaris Software Recall | (TC 04/03/2024 02:20:39 PM, Rhuland, Courtney) Recall completed by BD
1055122 | 202559 | PS | CNP | ENL | 2024-03-01 15:54 | 2024-08-16 15:31 | 2024 Alaris Software Recall
1055123 | 202560 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 09:58 | 2024 Alaris Software Recall | Completed by BD.
1055124 | 202561 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-20 15:07 | 2024 Alaris Software Recall
1055125 | 202562 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-13 09:42 | 2024 Alaris Software Recall | Recall completed
1055126 | 202563 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 10:00 | 2024 Alaris Software Recall | Completed by BD.
1055127 | 202564 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-11 15:13 | 2024 Alaris Software Recall | Completed by BD
1055128 | 202565 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 10:02 | 2024 Alaris Software Recall | Completed by BD.
1055155 | 202566 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-26 10:24 | 2024 Alaris Software Recall
1055156 | 202567 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 10:04 | 2024 Alaris Software Recall | Completed by BD.
1055157 | 202568 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 15:00 | 2024 Alaris Software Recall
1055158 | 202569 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-20 10:48 | 2024 Alaris Software Recall
1055159 | 202570 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-15 14:31 | 2024 Alaris Software Recall | Completed by BD
1055160 | 202571 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-26 10:25 | 2024 Alaris Software Recall
1055161 | 202572 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-26 13:31 | 2024 Alaris Software Recall
1055162 | 202573 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-02 14:48 | 2024 Alaris Software Recall
1055163 | 202574 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-10 10:08 | 2024 Alaris Software Recall | Completed
1055164 | 202575 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1055165 | 202576 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 08:51 | 2024 Alaris Software Recall
1055166 | 202577 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 10:15 | 2024 Alaris Software Recall
1055167 | 202578 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-20 10:40 | 2024 Alaris Software Recall
1055168 | 202579 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 10:07 | 2024 Alaris Software Recall | Completed by BD.
1055169 | 202580 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 10:09 | 2024 Alaris Software Recall | Completed by BD.
1055170 | 202581 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1055171 | 202582 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-20 10:44 | 2024 Alaris Software Recall
1055172 | 202583 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 08:55 | 2024 Alaris Software Recall
1055173 | 202584 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 10:18 | 2024 Alaris Software Recall
1055174 | 202585 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-04 15:22 | 2024 Alaris Software Recall | Recall completed by BD
1055175 | 202586 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-07-10 10:17 | 2024 Alaris Software Recall | (TC 07/10/2024 10:17:08 AM, DeGuzman, Leila) Device came down with an issue channel error. Error code: 200.5040 meaning to Reflash the pump. Asked Courtney how to reflash the pump. Performed flush pump. Software updated.
1055176 | 202587 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 09:00 | 2024 Alaris Software Recall
1055177 | 202588 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-09 10:10 | 2024 Alaris Software Recall
1055178 | 202589 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 10:11 | 2024 Alaris Software Recall | Completed by BD.
1055179 | 202590 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 10:49 | 2024 Alaris Software Recall | (TC 03/12/2024 10:48:41 AM, Rendon Lopez , Elizabeth) Recall completed
1055180 | 202591 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 10:47 | 2024 Alaris Software Recall | (TC 03/12/2024 10:47:06 AM, Rendon Lopez , Elizabeth) Recall completed
1055181 | 202592 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-14 08:03 | 2024 Alaris Software Recall | Completed by BD
1055743 | 202593 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 08:18 | 2024 Alaris Software Recall | (TC 03/21/2024 08:17:44 AM, Rhuland, Courtney) Recall completed by BD
1055745 | 202594 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-20 13:51 | 2024 Alaris Software Recall | (TC 03/20/2024 01:50:33 PM, Baghirzade, Nara) Software recall completed by BD.
1055746 | 202595 | PS | CMPLT | NFF | 2024-03-01 15:54 | 2024-03-21 08:27 | 2024 Alaris Software Recall | (TC 03/21/2024 08:26:18 AM, Rhuland, Courtney) Recall completed by BD
1055747 | 202596 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-25 09:36 | 2024 Alaris Software Recall | (TC 03/25/2024 09:35:41 AM, Rhuland, Courtney) Recall completed by BD
1055748 | 202597 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-07-31 09:53 | 2024 Alaris Software Recall
1055749 | 202598 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-07-31 09:54 | 2024 Alaris Software Recall
1055750 | 202599 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 08:28 | 2024 Alaris Software Recall | (TC 03/21/2024 08:27:45 AM, Rhuland, Courtney) Recall completed by BD
1055751 | 202600 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-10 14:54 | 2024 Alaris Software Recall | (TC 04/10/2024 02:53:34 PM, Rhuland, Courtney) Recall completed by BD
1055838 | 202601 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-19 09:45 | 2024 Alaris Software Recall | Completed by BD
1055839 | 202602 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-22 08:31 | 2024 Alaris Software Recall | Completed by BD
1055840 | 202603 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-22 08:30 | 2024 Alaris Software Recall | Completed by BD
1055841 | 202604 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-22 08:28 | 2024 Alaris Software Recall | Completed by BD
1055842 | 202605 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-22 08:29 | 2024 Alaris Software Recall | Completed by BD
1057388 | 202606 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:41 | 2024 Alaris Software Recall
1057389 | 202607 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-02 14:46 | 2024 Alaris Software Recall
1057390 | 202608 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-15 14:33 | 2024 Alaris Software Recall | Completed by BD
1057391 | 202609 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-11 19:09 | 2024 Alaris Software Recall
1057392 | 202610 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 11:02 | 2024 Alaris Software Recall
1057393 | 202611 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:43 | 2024 Alaris Software Recall
1058163 | 202612 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-18 13:12 | 2024 Alaris Software Recall
1058164 | 202613 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-14 08:07 | 2024 Alaris Software Recall | Completed by BD
1058221 | 202614 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1058222 | 202615 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1058223 | 202616 | PS | ACTIV | ISSUE | 2024-03-01 15:54 |  | 2024 Alaris Software Recall
1058224 | 202617 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-08-16 10:16 | 2024 Alaris Software Recall | Completed by BD.
1058225 | 202618 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 10:20 | 2024 Alaris Software Recall
1058226 | 202619 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 10:22 | 2024 Alaris Software Recall
1058227 | 202620 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-19 13:55 | 2024 Alaris Software Recall
1058228 | 202621 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-04-26 10:26 | 2024 Alaris Software Recall
1058229 | 202622 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:46 | 2024 Alaris Software Recall
1058230 | 202623 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 18:20 | 2024 Alaris Software Recall
1058231 | 202624 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-15 14:56 | 2024 Alaris Software Recall
1058232 | 202625 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-21 18:23 | 2024 Alaris Software Recall
1058233 | 202626 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-20 10:50 | 2024 Alaris Software Recall
1058234 | 202627 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-20 10:58 | 2024 Alaris Software Recall
1058235 | 202628 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-28 11:22 | 2024 Alaris Software Recall | Completed
1058236 | 202629 | PS | CMPLT | CMPLT | 2024-03-01 15:54 | 2024-03-12 11:03 | 2024 Alaris Software Recall
1058237 | 202630 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-19 13:58 | 2024 Alaris Software Recall
1058238 | 202631 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-12 09:16 | 2024 Alaris Software Recall | Completed by BD
1058239 | 202632 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-15 15:02 | 2024 Alaris Software Recall
1058240 | 202633 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-20 11:15 | 2024 Alaris Software Recall
1058241 | 202634 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-14 14:44 | 2024 Alaris Software Recall | (TC 03/08/2024 02:44:19 PM, Huang, Jessica) completed by BD
1058242 | 202635 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-30 14:02 | 2024 Alaris Software Recall
1058243 | 202636 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 10:22 | 2024 Alaris Software Recall
1058244 | 202637 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 10:24 | 2024 Alaris Software Recall
1058245 | 202638 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-14 08:09 | 2024 Alaris Software Recall | Completed by BD
1058246 | 202639 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:18 | 2024 Alaris Software Recall | Completed by BD.
1058247 | 202640 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:20 | 2024 Alaris Software Recall | Completed by BD.
1058248 | 202641 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-19 14:00 | 2024 Alaris Software Recall
1058249 | 202642 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-15 15:06 | 2024 Alaris Software Recall
1058250 | 202643 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 10:25 | 2024 Alaris Software Recall
1058251 | 202644 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-12 09:17 | 2024 Alaris Software Recall | Completed by BD
1058252 | 202645 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-12 09:19 | 2024 Alaris Software Recall | Completed by BD
1058253 | 202646 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 14:50 | 2024 Alaris Software Recall
1058254 | 202647 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:22 | 2024 Alaris Software Recall | Completed by BD.
1058255 | 202648 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-19 14:04 | 2024 Alaris Software Recall
1058256 | 202649 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-14 08:11 | 2024 Alaris Software Recall | Completed by BD
1058257 | 202650 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:29 | 2024 Alaris Software Recall | Completed by BD.
1058258 | 202651 | PS | ACTIV | ISSUE | 2024-03-01 15:55 |  | 2024 Alaris Software Recall
1058259 | 202652 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-14 13:54 | 2024 Alaris Software Recall | (TC 03/06/2024 01:54:05 PM, Huang, Jessica) completed by BD
1058260 | 202653 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:34 | 2024 Alaris Software Recall | Completed by BD.
1058261 | 202654 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:03 | 2024 Alaris Software Recall | Completed by BD
1058262 | 202655 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-27 14:45 | 2024 Alaris Software Recall
1058263 | 202656 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-27 14:46 | 2024 Alaris Software Recall
1058264 | 202657 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 09:33 | 2024 Alaris Software Recall | (TC 03/12/2024 09:32:29 AM, Rendon Lopez , Elizabeth) Recall completed
1058265 | 202658 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-15 10:16 | 2024 Alaris Software Recall | Recall completed
1058266 | 202659 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 09:34 | 2024 Alaris Software Recall | (TC 03/12/2024 09:34:10 AM, Rendon Lopez , Elizabeth) Recall completed
1058267 | 202660 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-13 09:55 | 2024 Alaris Software Recall | Recall completed
1058268 | 202661 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:36 | 2024 Alaris Software Recall | Completed by BD.
1058269 | 202662 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-21 10:33 | 2024 Alaris Software Recall
1058270 | 202663 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-13 09:57 | 2024 Alaris Software Recall | Recall completed
1058271 | 202664 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-11 15:14 | 2024 Alaris Software Recall | Completed by BD
1058272 | 202665 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-27 14:48 | 2024 Alaris Software Recall
1058273 | 202666 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-17 10:50 | 2024 Alaris Software Recall | Completed by BD
1058274 | 202667 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-06-06 09:50 | 2024 Alaris Software Recall | (TC 06/06/2024 09:49:52 AM, Rendon Lopez , Elizabeth) Software upgrade completed
1058275 | 202668 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-05-08 14:44 | 2024 Alaris Software Recall
1058276 | 202669 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:37 | 2024 Alaris Software Recall | Completed by BD.
1058277 | 202670 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 15:12 | 2024 Alaris Software Recall
1058278 | 202671 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 12:00 | 2024 Alaris Software Recall
1058279 | 202672 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-13 14:27 | 2024 Alaris Software Recall | (TC 03/13/2024 02:26:39 PM, Baghirzade, Nara) Software recall completed by BD.
1058280 | 202673 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:40 | 2024 Alaris Software Recall | Completed by BD.
1058281 | 202674 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-13 10:00 | 2024 Alaris Software Recall | Recall completed
1058282 | 202675 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-11 15:16 | 2024 Alaris Software Recall | Completed by BD
1058283 | 202676 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:41 | 2024 Alaris Software Recall | Completed by BD.
1058284 | 202677 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:09 | 2024 Alaris Software Recall | Completed by BD
1058285 | 202678 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:43 | 2024 Alaris Software Recall | Completed by BD.
1058286 | 202679 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 09:36 | 2024 Alaris Software Recall | (TC 03/12/2024 09:35:50 AM, Rendon Lopez , Elizabeth) Recall completed
1058287 | 202680 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-15 11:25 | 2024 Alaris Software Recall | Completed by BD
1058288 | 202681 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-11 15:17 | 2024 Alaris Software Recall | Completed by BD
1058289 | 202682 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-13 10:01 | 2024 Alaris Software Recall | Recall completed
1059873 | 202683 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-15 10:37 | 2024 Alaris Software Recall | Recall completed
1059874 | 202684 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:46 | 2024 Alaris Software Recall | Completed by BD.
1059875 | 202685 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:49 | 2024 Alaris Software Recall | Completed by BD.
1059876 | 202686 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-05-22 11:35 | 2024 Alaris Software Recall | (TC 05/22/2024 11:34:40 AM, Baghirzade, Nara) Software upgrade completed.
1059877 | 202687 | PS | CNP | ENL | 2024-03-01 15:55 | 2024-08-16 15:31 | 2024 Alaris Software Recall
1059878 | 202688 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-05 10:35 | 2024 Alaris Software Recall | (TC 04/05/2024 10:33:54 AM, Rhuland, Courtney) Recall completed by BD
1059879 | 202689 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:54 | 2024 Alaris Software Recall | Completed by BD.
1059880 | 202690 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:56 | 2024 Alaris Software Recall | Completed by BD.
1059882 | 202691 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 10:27 | 2024 Alaris Software Recall
1059883 | 202692 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-11 19:16 | 2024 Alaris Software Recall
1059884 | 202693 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 10:58 | 2024 Alaris Software Recall | Completed by BD.
1064641 | 202694 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 12:09 | 2024 Alaris Software Recall
1064642 | 202695 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 12:03 | 2024 Alaris Software Recall
1064643 | 202696 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 12:06 | 2024 Alaris Software Recall
1064644 | 202697 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 12:08 | 2024 Alaris Software Recall
1064645 | 202698 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-09 11:22 | 2024 Alaris Software Recall
1064646 | 202699 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-09 11:18 | 2024 Alaris Software Recall
1064647 | 202700 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 11:01 | 2024 Alaris Software Recall | Completed by BD.
1064648 | 202701 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-07-31 07:35 | 2024 Alaris Software Recall | (TC 07/31/2024 07:34:40 AM, Baghirzade, Nara) Software upgrade completed.
1064649 | 202702 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 14:51 | 2024 Alaris Software Recall
1064650 | 202703 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 14:56 | 2024 Alaris Software Recall
1064651 | 202704 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 14:55 | 2024 Alaris Software Recall
1064652 | 202705 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 14:53 | 2024 Alaris Software Recall
1065040 | 202706 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-23 10:11 | 2024 Alaris Software Recall | (TC 04/23/2024 10:10:54 AM, Rhuland, Courtney) Recall completed by BD
1065041 | 202707 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:04 | 2024 Alaris Software Recall | (TC 04/04/2024 01:03:43 PM, Rhuland, Courtney) Recall completed by BD
1065045 | 202708 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-23 10:10 | 2024 Alaris Software Recall | (TC 04/23/2024 10:09:34 AM, Rhuland, Courtney) Recall completed by BD
1065046 | 202709 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-03 14:26 | 2024 Alaris Software Recall | (TC 04/03/2024 02:25:52 PM, Rhuland, Courtney) Recall completed by BD
1065047 | 202710 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-05 14:53 | 2024 Alaris Software Recall | (TC 04/05/2024 02:52:56 PM, Rhuland, Courtney) Recall completed by BD
1065048 | 202711 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-03 14:24 | 2024 Alaris Software Recall | (TC 04/03/2024 02:24:11 PM, Rhuland, Courtney) Recall completed by BD
1065049 | 202712 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 08:56 | 2024 Alaris Software Recall
1065050 | 202713 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 08:16 | 2024 Alaris Software Recall
1065051 | 202714 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 08:57 | 2024 Alaris Software Recall
1065052 | 202715 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-05 14:52 | 2024 Alaris Software Recall | (TC 04/05/2024 02:51:49 PM, Rhuland, Courtney) Recall completed by BD
1065053 | 202716 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-03 14:22 | 2024 Alaris Software Recall | (TC 04/03/2024 02:21:52 PM, Rhuland, Courtney) Recall completed by BD
1065054 | 202717 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-03 14:23 | 2024 Alaris Software Recall | (TC 04/03/2024 02:23:04 PM, Rhuland, Courtney) Recall completed by BD
1065055 | 202718 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-23 10:13 | 2024 Alaris Software Recall | (TC 04/23/2024 10:12:29 AM, Rhuland, Courtney) Recall completed by BD
1065056 | 202719 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-05 14:51 | 2024 Alaris Software Recall | (TC 04/05/2024 02:50:32 PM, Rhuland, Courtney) Recall completed by BD
1065057 | 202720 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-05 14:50 | 2024 Alaris Software Recall | (TC 04/05/2024 02:49:30 PM, Rhuland, Courtney) Recall completed by BD
1065058 | 202721 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 08:50 | 2024 Alaris Software Recall
1065059 | 202722 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:21 | 2024 Alaris Software Recall
1065060 | 202723 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:22 | 2024 Alaris Software Recall
1065061 | 202724 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:22 | 2024 Alaris Software Recall | (TC 04/04/2024 01:21:11 PM, Rhuland, Courtney) Recall completed by BD
1065062 | 202725 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:37 | 2024 Alaris Software Recall
1065063 | 202726 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 10:33 | 2024 Alaris Software Recall
1065064 | 202727 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:24 | 2024 Alaris Software Recall
1065065 | 202728 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 14:58 | 2024 Alaris Software Recall
1065066 | 202729 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 10:35 | 2024 Alaris Software Recall
1065067 | 202730 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:33 | 2024 Alaris Software Recall
1065068 | 202731 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:21 | 2024 Alaris Software Recall
1065069 | 202732 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 10:32 | 2024 Alaris Software Recall
1065070 | 202733 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 10:34 | 2024 Alaris Software Recall
1065071 | 202734 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:14 | 2024 Alaris Software Recall | (TC 04/04/2024 01:14:00 PM, Rhuland, Courtney) Recall completed by BD
1065072 | 202735 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:27 | 2024 Alaris Software Recall
1065073 | 202736 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:29 | 2024 Alaris Software Recall
1065074 | 202737 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:25 | 2024 Alaris Software Recall | (TC 04/04/2024 01:24:08 PM, Rhuland, Courtney) Recall completed by BD
1065075 | 202738 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:35 | 2024 Alaris Software Recall
1065076 | 202739 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-07-19 13:23 | 2024 Alaris Software Recall | (TC 07/19/2024 01:22:53 PM, Baghirzade, Nara) Software upgrade completed.
1065077 | 202740 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:28 | 2024 Alaris Software Recall | (TC 04/04/2024 01:27:20 PM, Rhuland, Courtney) Recall completed by BD
1065078 | 202741 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:09 | 2024 Alaris Software Recall | (TC 04/04/2024 01:08:28 PM, Rhuland, Courtney) Recall completed by BD
1065079 | 202742 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:32 | 2024 Alaris Software Recall
1065080 | 202743 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:34 | 2024 Alaris Software Recall
1065081 | 202744 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:40 | 2024 Alaris Software Recall
1065082 | 202745 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:37 | 2024 Alaris Software Recall
1065083 | 202746 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:23 | 2024 Alaris Software Recall | (TC 04/04/2024 01:22:36 PM, Rhuland, Courtney) Recall completed by BD
1065084 | 202747 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:28 | 2024 Alaris Software Recall
1065085 | 202748 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:25 | 2024 Alaris Software Recall
1065086 | 202749 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:05 | 2024 Alaris Software Recall | (TC 04/04/2024 01:04:48 PM, Rhuland, Courtney) Recall completed by BD
1065087 | 202750 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:07 | 2024 Alaris Software Recall | (TC 04/04/2024 01:07:16 PM, Rhuland, Courtney) Recall completed by BD
1065088 | 202751 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:17 | 2024 Alaris Software Recall | (TC 04/04/2024 01:17:15 PM, Rhuland, Courtney) Recall completed by BD
1065089 | 202752 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:31 | 2024 Alaris Software Recall
1065090 | 202753 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:34 | 2024 Alaris Software Recall
1065091 | 202754 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-24 14:36 | 2024 Alaris Software Recall
1065092 | 202755 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-04 13:26 | 2024 Alaris Software Recall | (TC 04/04/2024 01:26:00 PM, Rhuland, Courtney) Recall completed by BD
1065686 | 202756 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:11 | 2024 Alaris Software Recall | Completed by BD
1065687 | 202757 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 09:45 | 2024 Alaris Software Recall
1065690 | 202758 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-11 15:18 | 2024 Alaris Software Recall | Completed by BD
1065691 | 202759 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-25 10:57 | 2024 Alaris Software Recall
1065692 | 202760 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-15 11:24 | 2024 Alaris Software Recall | Completed by BD
1065693 | 202761 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-15 08:27 | 2024 Alaris Software Recall | Completed by BD
1065694 | 202762 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 11:02 | 2024 Alaris Software Recall | Completed by BD.
1065695 | 202763 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 11:04 | 2024 Alaris Software Recall | Completed by BD.
1065696 | 202764 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-17 10:51 | 2024 Alaris Software Recall | Completed by BD
1065697 | 202765 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 13:14 | 2024 Alaris Software Recall
1065698 | 202766 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:14 | 2024 Alaris Software Recall | Completed by BD
1065699 | 202767 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-25 11:04 | 2024 Alaris Software Recall
1065700 | 202768 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:21 | 2024 Alaris Software Recall | Completed by BD
1065701 | 202769 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 15:15 | 2024 Alaris Software Recall
1065702 | 202770 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-25 11:06 | 2024 Alaris Software Recall
1065703 | 202771 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-17 08:52 | 2024 Alaris Software Recall | Completed by BD
1065704 | 202772 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 15:18 | 2024 Alaris Software Recall
1065705 | 202773 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 12:11 | 2024 Alaris Software Recall
1065706 | 202774 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 11:05 | 2024 Alaris Software Recall | Completed by BD.
1065707 | 202775 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-05-22 11:27 | 2024 Alaris Software Recall | (TC 05/22/2024 11:27:20 AM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
1065708 | 202776 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 11:07 | 2024 Alaris Software Recall | Completed by BD.
1065709 | 202777 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-15 08:24 | 2024 Alaris Software Recall | Completed by BD
1065710 | 202778 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:18 | 2024 Alaris Software Recall | Completed by BD
1065711 | 202779 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 11:08 | 2024 Alaris Software Recall | Completed by BD.
1065712 | 202780 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 15:20 | 2024 Alaris Software Recall
1065713 | 202781 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 13:19 | 2024 Alaris Software Recall
1065714 | 202782 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 15:22 | 2024 Alaris Software Recall
1065715 | 202783 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-27 14:51 | 2024 Alaris Software Recall
1065716 | 202784 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:23 | 2024 Alaris Software Recall | Completed by BD
1065717 | 202785 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-27 14:54 | 2024 Alaris Software Recall
1065718 | 202786 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-11 15:19 | 2024 Alaris Software Recall | Completed by BD
1065719 | 202787 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-20 15:11 | 2024 Alaris Software Recall
1065720 | 202788 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-02 12:13 | 2024 Alaris Software Recall
1065721 | 202789 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-26 08:58 | 2024 Alaris Software Recall
1065722 | 202790 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-25 11:09 | 2024 Alaris Software Recall
1065723 | 202791 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 15:25 | 2024 Alaris Software Recall
1065724 | 202792 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 13:23 | 2024 Alaris Software Recall
1065725 | 202793 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-16 11:20 | 2024 Alaris Software Recall | Completed by BD.
1065726 | 202794 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-12 11:08 | 2024 Alaris Software Recall
1065727 | 202795 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-12 13:26 | 2024 Alaris Software Recall | Completed by BD
1065728 | 202796 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-27 14:55 | 2024 Alaris Software Recall
1065729 | 202797 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-12 11:03 | 2024 Alaris Software Recall
1065730 | 202798 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-27 14:58 | 2024 Alaris Software Recall
1065731 | 202799 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-04-15 08:23 | 2024 Alaris Software Recall | Completed by BD
1065732 | 202800 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-12 11:02 | 2024 Alaris Software Recall
1065733 | 202801 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-12 11:01 | 2024 Alaris Software Recall
1065734 | 202802 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-12 10:59 | 2024 Alaris Software Recall
1065735 | 202803 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-08-12 10:58 | 2024 Alaris Software Recall
1065739 | 202804 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-15 15:09 | 2024 Alaris Software Recall
1065740 | 202805 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-21 18:25 | 2024 Alaris Software Recall
1065741 | 202806 | PS | CMPLT | CMPLT | 2024-03-01 15:55 | 2024-03-18 09:47 | 2024 Alaris Software Recall
1065742 | 202807 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:18 | 2024 Alaris Software Recall
1065743 | 202808 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 15:10 | 2024 Alaris Software Recall
1065744 | 202809 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:22 | 2024 Alaris Software Recall | Completed by BD.
1065745 | 202810 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 09:50 | 2024 Alaris Software Recall
1065746 | 202811 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:24 | 2024 Alaris Software Recall | Completed by BD.
1065747 | 202812 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:26 | 2024 Alaris Software Recall
1065748 | 202813 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 09:08 | 2024 Alaris Software Recall
1065749 | 202814 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 09:19 | 2024 Alaris Software Recall
1065750 | 202815 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:20 | 2024 Alaris Software Recall
1065751 | 202816 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 09:20 | 2024 Alaris Software Recall | Completed by BD
1065752 | 202817 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 08:14 | 2024 Alaris Software Recall | Completed by BD
1065753 | 202818 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 09:54 | 2024 Alaris Software Recall
1065754 | 202819 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 10:26 | 2024 Alaris Software Recall
1065755 | 202820 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 15:14 | 2024 Alaris Software Recall
1065756 | 202821 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 15:16 | 2024 Alaris Software Recall
1065757 | 202822 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:05 | 2024 Alaris Software Recall
1065758 | 202823 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 08:28 | 2024 Alaris Software Recall | Completed by BD
1065759 | 202824 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:27 | 2024 Alaris Software Recall | Completed by BD.
1065760 | 202825 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:29 | 2024 Alaris Software Recall
1065761 | 202826 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 11:18 | 2024 Alaris Software Recall
1065762 | 202827 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 09:22 | 2024 Alaris Software Recall | Completed by BD
1065763 | 202828 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:28 | 2024 Alaris Software Recall | Completed by BD.
1065764 | 202829 | PS | ACTIV | ISSUE | 2024-03-01 15:56 |  | 2024 Alaris Software Recall
1065765 | 202830 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:22 | 2024 Alaris Software Recall
1065766 | 202831 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 15:22 | 2024 Alaris Software Recall
1065767 | 202832 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-28 11:24 | 2024 Alaris Software Recall | Completed
1065768 | 202833 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-28 11:26 | 2024 Alaris Software Recall | Completed
1065769 | 202834 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-28 11:27 | 2024 Alaris Software Recall | Completed
1065770 | 202835 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:24 | 2024 Alaris Software Recall
1065771 | 202836 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 08:35 | 2024 Alaris Software Recall | Completed by BD
1065772 | 202837 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:06 | 2024 Alaris Software Recall
1065773 | 202838 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:26 | 2024 Alaris Software Recall
1065774 | 202839 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-26 10:37 | 2024 Alaris Software Recall
1065775 | 202840 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 09:25 | 2024 Alaris Software Recall
1065776 | 202841 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:34 | 2024 Alaris Software Recall | Completed by BD
1065777 | 202842 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 08:37 | 2024 Alaris Software Recall | Completed by BD
1065778 | 202843 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:28 | 2024 Alaris Software Recall
1065779 | 202844 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 11:20 | 2024 Alaris Software Recall
1065780 | 202845 | PS | CNP | ENL | 2024-03-01 15:56 | 2024-11-22 10:54 | 2024 Alaris Software Recall
1065781 | 202846 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-10 10:12 | 2024 Alaris Software Recall | Completed
1065782 | 202847 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 09:28 | 2024 Alaris Software Recall
1065783 | 202848 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:01 | 2024 Alaris Software Recall
1065784 | 202849 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 11:30 | 2024 Alaris Software Recall
1065785 | 202850 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 09:33 | 2024 Alaris Software Recall
1065786 | 202851 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:34 | 2024 Alaris Software Recall | Completed by BD.
1065787 | 202852 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:35 | 2024 Alaris Software Recall | Completed by BD
1065788 | 202853 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:30 | 2024 Alaris Software Recall
1065789 | 202854 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 08:46 | 2024 Alaris Software Recall | Completed by BD
1065790 | 202855 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:36 | 2024 Alaris Software Recall | Completed by BD.
1065791 | 202856 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:03 | 2024 Alaris Software Recall
1065792 | 202857 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 08:56 | 2024 Alaris Software Recall | Completed by BD
1065793 | 202858 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 09:43 | 2024 Alaris Software Recall
1065794 | 202859 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-11 19:32 | 2024 Alaris Software Recall
1065795 | 202860 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 10:28 | 2024 Alaris Software Recall
1065796 | 202861 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-10 10:13 | 2024 Alaris Software Recall | Completed
1065797 | 202862 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 11:34 | 2024 Alaris Software Recall
1065798 | 202863 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 08:59 | 2024 Alaris Software Recall | Completed by BD
1065799 | 202864 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 09:23 | 2024 Alaris Software Recall | Completed by BD
1065800 | 202865 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:06 | 2024 Alaris Software Recall
1065801 | 202866 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:17 | 2024 Alaris Software Recall
1065802 | 202867 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 09:04 | 2024 Alaris Software Recall | Completed by BD
1065803 | 202868 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 11:40 | 2024 Alaris Software Recall
1065804 | 202869 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:36 | 2024 Alaris Software Recall | Completed by BD
1065805 | 202870 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:08 | 2024 Alaris Software Recall
1065806 | 202871 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 09:24 | 2024 Alaris Software Recall | Completed by BD
1065807 | 202872 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:42 | 2024 Alaris Software Recall | Completed by BD.
1065808 | 202873 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:18 | 2024 Alaris Software Recall
1065809 | 202874 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:08 | 2024 Alaris Software Recall
1065810 | 202875 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:07 | 2024 Alaris Software Recall
1065811 | 202876 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:06 | 2024 Alaris Software Recall
1065812 | 202877 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-26 13:34 | 2024 Alaris Software Recall
1065813 | 202878 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 13:34 | 2024 Alaris Software Recall
1065814 | 202879 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 09:25 | 2024 Alaris Software Recall | Completed by BD
1065815 | 202880 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 13:50 | 2024 Alaris Software Recall
1065816 | 202881 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-07-03 13:44 | 2024 Alaris Software Recall
1065817 | 202882 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:37 | 2024 Alaris Software Recall | Completed by BD
1065818 | 202883 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:30 | 2024 Alaris Software Recall
1065819 | 202884 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:38 | 2024 Alaris Software Recall | Completed by BD
1065820 | 202885 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:08 | 2024 Alaris Software Recall
1065821 | 202886 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:21 | 2024 Alaris Software Recall
1065822 | 202887 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:40 | 2024 Alaris Software Recall | Completed by BD
1065823 | 202888 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:19 | 2024 Alaris Software Recall
1065824 | 202889 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:20 | 2024 Alaris Software Recall
1065825 | 202890 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:09 | 2024 Alaris Software Recall
1065826 | 202891 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:11 | 2024 Alaris Software Recall
1065827 | 202892 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:34 | 2024 Alaris Software Recall
1065828 | 202893 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 09:26 | 2024 Alaris Software Recall | Completed by BD
1065829 | 202894 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:31 | 2024 Alaris Software Recall
1065830 | 202895 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-26 10:38 | 2024 Alaris Software Recall
1065831 | 202896 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:21 | 2024 Alaris Software Recall
1065832 | 202897 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:44 | 2024 Alaris Software Recall | Completed by BD.
1065833 | 202898 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:36 | 2024 Alaris Software Recall
1065834 | 202899 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:41 | 2024 Alaris Software Recall | Completed by BD
1065835 | 202900 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:10 | 2024 Alaris Software Recall
1065836 | 202901 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-10 10:15 | 2024 Alaris Software Recall | Completed
1065837 | 202902 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:40 | 2024 Alaris Software Recall
1065838 | 202903 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 14:02 | 2024 Alaris Software Recall
1065839 | 202904 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:43 | 2024 Alaris Software Recall
1065840 | 202905 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:22 | 2024 Alaris Software Recall
1065841 | 202906 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:13 | 2024 Alaris Software Recall
1065842 | 202907 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-22 08:45 | 2024 Alaris Software Recall
1065843 | 202908 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 14:04 | 2024 Alaris Software Recall
1065844 | 202909 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:23 | 2024 Alaris Software Recall
1065845 | 202910 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:33 | 2024 Alaris Software Recall
1065846 | 202911 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:46 | 2024 Alaris Software Recall | Completed by BD.
1065847 | 202912 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:47 | 2024 Alaris Software Recall | Completed by BD.
1065848 | 202913 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-22 13:23 | 2024 Alaris Software Recall
1065849 | 202914 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:23 | 2024 Alaris Software Recall
1065850 | 202915 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:48 | 2024 Alaris Software Recall | Completed by BD.
1065851 | 202916 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 09:06 | 2024 Alaris Software Recall | Completed by BD
1065852 | 202917 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:09 | 2024 Alaris Software Recall
1065853 | 202918 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-28 11:29 | 2024 Alaris Software Recall | Completed
1065854 | 202919 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:27 | 2024 Alaris Software Recall
1065855 | 202920 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:10 | 2024 Alaris Software Recall
1065856 | 202921 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:28 | 2024 Alaris Software Recall
1065857 | 202922 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:33 | 2024 Alaris Software Recall
1065858 | 202923 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-13 08:10 | 2024 Alaris Software Recall
1065859 | 202924 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:36 | 2024 Alaris Software Recall
1065860 | 202925 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 14:06 | 2024 Alaris Software Recall
1065861 | 202926 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:38 | 2024 Alaris Software Recall
1065862 | 202927 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-04 13:31 | 2024 Alaris Software Recall | (TC 04/04/2024 01:30:44 PM, Rhuland, Courtney) Recall completed by BD
1065863 | 202928 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-17 10:17 | 2024 Alaris Software Recall | Completed by BD
1065864 | 202929 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 09:27 | 2024 Alaris Software Recall | Completed by BD
1065865 | 202930 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-28 11:31 | 2024 Alaris Software Recall | Completed
1065866 | 202931 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 14:08 | 2024 Alaris Software Recall
1065867 | 202932 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:50 | 2024 Alaris Software Recall | Completed by BD.
1065868 | 202933 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:14 | 2024 Alaris Software Recall
1065869 | 202934 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-20 14:42 | 2024 Alaris Software Recall
1065870 | 202935 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:36 | 2024 Alaris Software Recall
1065871 | 202936 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-10 10:17 | 2024 Alaris Software Recall | Completed
1065872 | 202937 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-24 14:42 | 2024 Alaris Software Recall
1065873 | 202938 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:17 | 2024 Alaris Software Recall
1065874 | 202939 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-22 13:28 | 2024 Alaris Software Recall
1065875 | 202940 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:24 | 2024 Alaris Software Recall
1065876 | 202941 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 10:39 | 2024 Alaris Software Recall
1065877 | 202942 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-27 11:26 | 2024 Alaris Software Recall | Completed
1065878 | 202943 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 11:22 | 2024 Alaris Software Recall
1065879 | 202944 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 11:52 | 2024 Alaris Software Recall | Completed by BD.
1065880 | 202945 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:26 | 2024 Alaris Software Recall
1065881 | 202946 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:25 | 2024 Alaris Software Recall
1065882 | 202947 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:30 | 2024 Alaris Software Recall
1065883 | 202948 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:11 | 2024 Alaris Software Recall
1065884 | 202949 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 16:57 | 2024 Alaris Software Recall | by BD
1065885 | 202950 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:32 | 2024 Alaris Software Recall
1065886 | 202951 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-10 10:19 | 2024 Alaris Software Recall | Completed
1065887 | 202952 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:36 | 2024 Alaris Software Recall
1065888 | 202953 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-28 11:33 | 2024 Alaris Software Recall | Completed
1065889 | 202954 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-14 14:02 | 2024 Alaris Software Recall | Recall completed
1065890 | 202955 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:31 | 2024 Alaris Software Recall
1065891 | 202956 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:37 | 2024 Alaris Software Recall
1065892 | 202957 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-05-24 09:51 | 2024 Alaris Software Recall
1065893 | 202958 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 10:18 | 2024 Alaris Software Recall | Completed by BD
1065894 | 202959 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:38 | 2024 Alaris Software Recall
1065895 | 202960 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:02 | 2024 Alaris Software Recall | Completed by BD.
1065896 | 202961 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-26 10:39 | 2024 Alaris Software Recall
1065897 | 202962 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:04 | 2024 Alaris Software Recall | Completed by BD.
1065898 | 202963 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-02 15:00 | 2024 Alaris Software Recall
1065899 | 202964 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:05 | 2024 Alaris Software Recall | Completed by BD.
1065900 | 202965 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 16:59 | 2024 Alaris Software Recall | by BD
1065901 | 202966 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-24 14:43 | 2024 Alaris Software Recall
1065902 | 202967 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 17:02 | 2024 Alaris Software Recall | by BD
1065903 | 202968 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-02 16:08 | 2024 Alaris Software Recall | Completed
1065904 | 202969 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-26 10:40 | 2024 Alaris Software Recall
1065905 | 202970 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:06 | 2024 Alaris Software Recall | Completed by BD.
1065906 | 202971 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:39 | 2024 Alaris Software Recall
1065907 | 202972 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:07 | 2024 Alaris Software Recall | Completed by BD.
1065908 | 202973 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 17:20 | 2024 Alaris Software Recall | by BD
1065909 | 202974 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:10 | 2024 Alaris Software Recall | Completed by BD.
1065910 | 202975 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-26 13:37 | 2024 Alaris Software Recall
1065911 | 202976 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:41 | 2024 Alaris Software Recall
1065912 | 202977 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 17:34 | 2024 Alaris Software Recall | by BD
1065913 | 202978 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 17:35 | 2024 Alaris Software Recall | by BD
1065914 | 202979 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:42 | 2024 Alaris Software Recall
1065915 | 202980 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 11:24 | 2024 Alaris Software Recall
1065916 | 202981 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:15 | 2024 Alaris Software Recall
1065917 | 202982 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:12 | 2024 Alaris Software Recall | Completed by BD.
1065918 | 202983 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-02 16:20 | 2024 Alaris Software Recall | Completed
1065919 | 202984 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 11:25 | 2024 Alaris Software Recall
1065920 | 202985 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:33 | 2024 Alaris Software Recall
1065921 | 202986 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:44 | 2024 Alaris Software Recall
1065922 | 202987 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:46 | 2024 Alaris Software Recall
1065923 | 202988 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:12 | 2024 Alaris Software Recall | Completed by BD.
1065924 | 202989 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 11:31 | 2024 Alaris Software Recall
1065925 | 202990 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 17:37 | 2024 Alaris Software Recall | by BD
1065926 | 202991 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:39 | 2024 Alaris Software Recall
1065927 | 202992 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-15 11:36 | 2024 Alaris Software Recall
1065928 | 202993 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-12 11:23 | 2024 Alaris Software Recall | Completed by BD
1065929 | 202994 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 17:52 | 2024 Alaris Software Recall | by BD
1065930 | 202995 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 08:47 | 2024 Alaris Software Recall
1065931 | 202996 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:42 | 2024 Alaris Software Recall | Completed by BD
1065932 | 202997 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:49 | 2024 Alaris Software Recall | Completed by BD
1065933 | 202998 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-22 13:34 | 2024 Alaris Software Recall
1065934 | 202999 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-22 13:37 | 2024 Alaris Software Recall
1065935 | 203000 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-12 11:12 | 2024 Alaris Software Recall
1065936 | 203001 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-19 14:17 | 2024 Alaris Software Recall
1065937 | 203002 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-09 17:54 | 2024 Alaris Software Recall | by BD
1065938 | 203003 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-21 18:42 | 2024 Alaris Software Recall
1065939 | 203004 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-08-16 12:13 | 2024 Alaris Software Recall | Completed by BD.
1065940 | 203005 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-03-18 10:34 | 2024 Alaris Software Recall
1065941 | 203006 | PS | CMPLT | CMPLT | 2024-03-01 15:56 | 2024-04-15 14:50 | 2024 Alaris Software Recall | Completed by BD
1065942 | 203007 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-17 10:18 | 2024 Alaris Software Recall | Completed by BD
1065943 | 203008 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:15 | 2024 Alaris Software Recall | Completed by BD.
1065944 | 203009 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 10:49 | 2024 Alaris Software Recall | Recall completed
1065945 | 203010 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-09 17:55 | 2024 Alaris Software Recall | by BD
1065946 | 203011 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:39 | 2024 Alaris Software Recall
1065947 | 203012 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-17 10:19 | 2024 Alaris Software Recall | Completed by BD
1065948 | 203013 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:16 | 2024 Alaris Software Recall | Completed by BD.
1065949 | 203014 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:17 | 2024 Alaris Software Recall | Completed by BD.
1065950 | 203015 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 11:12 | 2024 Alaris Software Recall
1065951 | 203016 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 16:22 | 2024 Alaris Software Recall | Completed
1065952 | 203017 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:42 | 2024 Alaris Software Recall
1065953 | 203018 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 13:39 | 2024 Alaris Software Recall
1065954 | 203019 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:18 | 2024 Alaris Software Recall | Completed by BD.
1065955 | 203020 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 08:49 | 2024 Alaris Software Recall
1065956 | 203021 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 08:51 | 2024 Alaris Software Recall
1065957 | 203022 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 10:36 | 2024 Alaris Software Recall
1065958 | 203023 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 13:43 | 2024 Alaris Software Recall
1065959 | 203024 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-10 10:21 | 2024 Alaris Software Recall | Completed
1065960 | 203025 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 08:57 | 2024 Alaris Software Recall
1065961 | 203026 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 18:43 | 2024 Alaris Software Recall
1065962 | 203027 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 13:51 | 2024 Alaris Software Recall
1065963 | 203028 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 10:37 | 2024 Alaris Software Recall
1065964 | 203029 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:19 | 2024 Alaris Software Recall | Completed by BD.
1065965 | 203030 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 14:02 | 2024 Alaris Software Recall
1065966 | 203031 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 10:44 | 2024 Alaris Software Recall
1065967 | 203032 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:44 | 2024 Alaris Software Recall
1065968 | 203033 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:21 | 2024 Alaris Software Recall | Completed by BD.
1065969 | 203034 | PS | ACTIV | ISSUE | 2024-03-01 15:57 |  | 2024 Alaris Software Recall
1065970 | 203035 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-26 13:40 | 2024 Alaris Software Recall
1065971 | 203036 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 10:52 | 2024 Alaris Software Recall
1065972 | 203037 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 10:56 | 2024 Alaris Software Recall
1065973 | 203038 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 14:09 | 2024 Alaris Software Recall
1065974 | 203039 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-12 11:25 | 2024 Alaris Software Recall | Completed by BD
1065975 | 203040 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:24 | 2024 Alaris Software Recall | Completed by BD.
1065976 | 203041 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 16:23 | 2024 Alaris Software Recall | Completed
1065977 | 203042 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:25 | 2024 Alaris Software Recall | Completed by BD.
1065978 | 203043 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 11:04 | 2024 Alaris Software Recall
1065979 | 203044 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-19 14:18 | 2024 Alaris Software Recall
1065980 | 203045 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 14:13 | 2024 Alaris Software Recall
1065981 | 203046 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 11:17 | 2024 Alaris Software Recall
1065982 | 203047 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 14:22 | 2024 Alaris Software Recall
1065983 | 203048 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:26 | 2024 Alaris Software Recall | Completed by BD.
1065984 | 203049 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 08:59 | 2024 Alaris Software Recall
1065985 | 203050 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 09:00 | 2024 Alaris Software Recall
1065986 | 203051 | PS | ACTIV | ISSUE | 2024-03-01 15:57 |  | 2024 Alaris Software Recall
1065987 | 203052 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:28 | 2024 Alaris Software Recall | Completed by BD.
1065988 | 203053 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:46 | 2024 Alaris Software Recall
1065989 | 203054 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-26 13:43 | 2024 Alaris Software Recall
1065990 | 203055 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:29 | 2024 Alaris Software Recall | Completed by BD.
1065991 | 203056 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:30 | 2024 Alaris Software Recall | Completed by BD.
1065992 | 203057 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-27 11:28 | 2024 Alaris Software Recall | Completed
1065993 | 203058 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-12 11:26 | 2024 Alaris Software Recall | Completed by BD
1065994 | 203059 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:31 | 2024 Alaris Software Recall | Completed by BD.
1065995 | 203060 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-09 17:57 | 2024 Alaris Software Recall | by BD
1065996 | 203061 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 16:24 | 2024 Alaris Software Recall | Completed
1065997 | 203062 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:32 | 2024 Alaris Software Recall | Completed by BD.
1065998 | 203063 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:32 | 2024 Alaris Software Recall | Completed by BD.
1065999 | 203064 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:49 | 2024 Alaris Software Recall
1066000 | 203065 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 16:32 | 2024 Alaris Software Recall | Completed
1066001 | 203066 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-16 10:58 | 2024 Alaris Software Recall | (TC 07/16/2024 10:57:58 AM, DeGuzman, Leila) Software updated.
1066002 | 203067 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:53 | 2024 Alaris Software Recall
1066003 | 203068 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-17 10:20 | 2024 Alaris Software Recall | Completed by BD
1066004 | 203069 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:33 | 2024 Alaris Software Recall | Completed by BD.
1066005 | 203070 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-26 13:46 | 2024 Alaris Software Recall
1066006 | 203071 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 11:18 | 2024 Alaris Software Recall
1066007 | 203072 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 10:29 | 2024 Alaris Software Recall
1066008 | 203073 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-26 10:43 | 2024 Alaris Software Recall
1066009 | 203074 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-26 10:44 | 2024 Alaris Software Recall
1066010 | 203075 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 18:45 | 2024 Alaris Software Recall
1066011 | 203076 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:54 | 2024 Alaris Software Recall
1066012 | 203077 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 16:26 | 2024 Alaris Software Recall | Completed
1066013 | 203078 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 10:29 | 2024 Alaris Software Recall
1066014 | 203079 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-05-16 11:34 | 2024 Alaris Software Recall | (TC 05/16/2024 11:34:01 AM, Baghirzade, Nara) Software upgrade completed.
1066015 | 203080 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 14:24 | 2024 Alaris Software Recall
1066016 | 203081 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:56 | 2024 Alaris Software Recall
1066017 | 203082 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-12 11:27 | 2024 Alaris Software Recall | Completed by BD
1066018 | 203083 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 14:31 | 2024 Alaris Software Recall
1066019 | 203084 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 11:59 | 2024 Alaris Software Recall
1066020 | 203085 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 09:01 | 2024 Alaris Software Recall
1066021 | 203086 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-04 13:32 | 2024 Alaris Software Recall | (TC 04/04/2024 01:32:11 PM, Rhuland, Courtney) Recall completed by BD
1066022 | 203087 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-16 08:38 | 2024 Alaris Software Recall | Completed by BD
1066023 | 203088 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-09 17:59 | 2024 Alaris Software Recall | by BD
1066024 | 203089 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 09:03 | 2024 Alaris Software Recall
1066025 | 203090 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-09 18:01 | 2024 Alaris Software Recall | by BD
1066026 | 203091 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:34 | 2024 Alaris Software Recall | Completed by BD.
1066027 | 203092 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 09:05 | 2024 Alaris Software Recall
1066028 | 203093 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:35 | 2024 Alaris Software Recall | Completed by BD.
1066029 | 203094 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-12 11:28 | 2024 Alaris Software Recall | Completed by BD
1066030 | 203095 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-26 10:45 | 2024 Alaris Software Recall
1066031 | 203096 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-09 18:13 | 2024 Alaris Software Recall | by BD
1066032 | 203097 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 12:02 | 2024 Alaris Software Recall
1066033 | 203098 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 14:34 | 2024 Alaris Software Recall
1066034 | 203099 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 16:27 | 2024 Alaris Software Recall | Completed
1066035 | 203100 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 12:03 | 2024 Alaris Software Recall
1066036 | 203101 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-16 12:36 | 2024 Alaris Software Recall | Completed by BD.
1066037 | 203102 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 12:05 | 2024 Alaris Software Recall
1066038 | 203103 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 15:02 | 2024 Alaris Software Recall
1066039 | 203104 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 09:07 | 2024 Alaris Software Recall
1066040 | 203105 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-16 08:37 | 2024 Alaris Software Recall | Completed by BD
1066041 | 203106 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-26 10:46 | 2024 Alaris Software Recall
1066042 | 203107 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 12:08 | 2024 Alaris Software Recall
1066043 | 203108 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 10:30 | 2024 Alaris Software Recall
1066044 | 203109 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 11:13 | 2024 Alaris Software Recall
1066045 | 203110 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-19 14:20 | 2024 Alaris Software Recall
1066046 | 203111 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 12:10 | 2024 Alaris Software Recall
1066047 | 203112 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-02 16:29 | 2024 Alaris Software Recall | Completed
1066048 | 203113 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 15:15 | 2024 Alaris Software Recall
1066049 | 203114 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 09:08 | 2024 Alaris Software Recall
1066050 | 203115 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-09 18:22 | 2024 Alaris Software Recall | by BD
1066051 | 203116 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 15:18 | 2024 Alaris Software Recall
1066052 | 203117 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-12 11:29 | 2024 Alaris Software Recall | Completed by BD
1066053 | 203118 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 15:24 | 2024 Alaris Software Recall
1066054 | 203119 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-12 11:30 | 2024 Alaris Software Recall | Completed by BD
1066055 | 203120 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 15:27 | 2024 Alaris Software Recall
1066056 | 203121 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 08:29 | 2024 Alaris Software Recall
1066057 | 203122 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-12 11:31 | 2024 Alaris Software Recall | Completed by BD
1066770 | 203123 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-22 12:09 | 2024 Alaris Software Recall | (TC 07/22/2024 12:09:25 PM, Rendon Lopez , Elizabeth) Recall completed
1066771 | 203124 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-16 08:36 | 2024 Alaris Software Recall | Completed by BD
1066772 | 203125 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-18 11:20 | 2024 Alaris Software Recall
65875 | 203126 | PS | CNP | ENL | 2024-03-01 15:57 | 2024-08-16 15:32 | 2024 Alaris Software Recall
65876 | 203127 | PS | CNP | ENL | 2024-03-01 15:57 | 2024-08-16 15:33 | 2024 Alaris Software Recall
65877 | 203128 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-19 09:32 | 2024 Alaris Software Recall
65878 | 203129 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-13 14:31 | 2024 Alaris Software Recall | (TC 03/13/2024 02:30:56 PM, Rhuland, Courtney) Recall completed by BD
65879 | 203130 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-18 12:56 | 2024 Alaris Software Recall | (TC 04/18/2024 12:55:17 PM, Baghirzade, Nara) Software recall completed by BD.
65880 | 203131 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-26 14:18 | 2024 Alaris Software Recall | (TC 03/26/2024 02:18:02 PM, Rhuland, Courtney) Recall completed by BD
65881 | 203132 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 12:45 | 2024 Alaris Software Recall | (TC 03/21/2024 12:44:35 PM, Baghirzade, Nara) Software recall completed by BD.
65882 | 203133 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-20 09:24 | 2024 Alaris Software Recall | (TC 03/20/2024 09:23:33 AM, Baghirzade, Nara) Software recall completed by BD.
65883 | 203134 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-30 11:54 | 2024 Alaris Software Recall | (TC 07/30/2024 11:54:09 AM, Baghirzade, Nara) Software upgrade completed.
65884 | 203135 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 09:58 | 2024 Alaris Software Recall
65885 | 203136 | PS | CNP | ENL | 2024-03-01 15:57 | 2024-08-16 15:33 | 2024 Alaris Software Recall
65886 | 203137 | PS | CNP | ENL | 2024-03-01 15:57 | 2024-08-16 15:33 | 2024 Alaris Software Recall
65887 | 203138 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-10 14:34 | 2024 Alaris Software Recall | (TC 04/10/2024 02:33:43 PM, Rhuland, Courtney) Recall completed by BD
65888 | 203139 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-21 08:34 | 2024 Alaris Software Recall | (TC 03/21/2024 08:33:14 AM, Rhuland, Courtney) Recall completed by BD
65889 | 203140 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-25 13:22 | 2024 Alaris Software Recall | (TC 03/25/2024 01:20:41 PM, Rhuland, Courtney) Recall completed by BD
65890 | 203141 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 12:49 | 2024 Alaris Software Recall | (TC 03/28/2024 12:48:30 PM, Rhuland, Courtney) Recall completed by BD
65891 | 203142 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 13:33 | 2024 Alaris Software Recall | (TC 03/15/2024 01:33:00 PM, Baghirzade, Nara) Software recall completed by BD.
65892 | 203143 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 10:01 | 2024 Alaris Software Recall
65893 | 203144 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-09 11:04 | 2024 Alaris Software Recall
65894 | 203145 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 11:07 | 2024 Alaris Software Recall
65895 | 203146 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-03 13:37 | 2024 Alaris Software Recall
65896 | 203147 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-11 13:57 | 2024 Alaris Software Recall | Recall completed by BD
65897 | 203148 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-14 13:23 | 2024 Alaris Software Recall | (TC 03/14/2024 01:23:05 PM, Baghirzade, Nara) Software recall completed by BD.
65898 | 203149 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 14:09 | 2024 Alaris Software Recall | (TC 03/12/2024 02:08:27 PM, Baghirzade, Nara) Software recall completed by BD.
65899 | 203150 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 09:21 | 2024 Alaris Software Recall | Completed by BD
65900 | 203151 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 13:41 | 2024 Alaris Software Recall | (TC 03/15/2024 01:41:15 PM, Baghirzade, Nara) Software recall completed by BD.
65901 | 203152 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-03 13:40 | 2024 Alaris Software Recall
65902 | 203153 | PS | CNP | ENL | 2024-03-01 15:57 | 2024-08-16 15:33 | 2024 Alaris Software Recall
65903 | 203154 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-20 09:17 | 2024 Alaris Software Recall | (TC 03/20/2024 09:17:21 AM, Baghirzade, Nara) Software recall completed by BD.
65904 | 203155 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-26 14:24 | 2024 Alaris Software Recall | (TC 03/26/2024 02:23:51 PM, Rhuland, Courtney) Recall completed by BD
65905 | 203156 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-10 14:37 | 2024 Alaris Software Recall | (TC 04/10/2024 02:37:02 PM, Rhuland, Courtney) Recall completed by BD
65906 | 203157 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 11:10 | 2024 Alaris Software Recall
65907 | 203158 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-25 13:47 | 2024 Alaris Software Recall | (TC 03/25/2024 01:46:49 PM, Baghirzade, Nara) Software recall completed by BD.
65908 | 203159 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 11:12 | 2024 Alaris Software Recall
65909 | 203160 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 12:51 | 2024 Alaris Software Recall | (TC 03/28/2024 12:50:18 PM, Rhuland, Courtney) Recall completed by BD
65910 | 203161 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-02 07:37 | 2024 Alaris Software Recall | (TC 08/02/2024 07:36:44 AM, Baghirzade, Nara) Software upgrade completed.
65911 | 203162 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-20 09:27 | 2024 Alaris Software Recall | (TC 03/20/2024 09:26:18 AM, Baghirzade, Nara) Software recall completed by BD.
65912 | 203163 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-20 09:21 | 2024 Alaris Software Recall | (TC 03/20/2024 09:20:20 AM, Baghirzade, Nara) Software recall completed by BD.
65913 | 203164 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 13:01 | 2024 Alaris Software Recall | (TC 03/28/2024 01:00:39 PM, Rhuland, Courtney) Recall completed by BD
65914 | 203165 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 14:14 | 2024 Alaris Software Recall | (TC 03/15/2024 02:13:19 PM, Rhuland, Courtney) Recall completed by BD
65915 | 203166 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 13:05 | 2024 Alaris Software Recall | (TC 03/28/2024 01:05:07 PM, Rhuland, Courtney) Recall completed by BD
65916 | 203167 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 12:57 | 2024 Alaris Software Recall | (TC 03/28/2024 12:56:43 PM, Rhuland, Courtney) Recall completed by BD
65917 | 203168 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-25 13:12 | 2024 Alaris Software Recall | (TC 03/25/2024 01:12:09 PM, Baghirzade, Nara) Software recall completed by BD.
65918 | 203169 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-14 07:18 | 2024 Alaris Software Recall | (TC 03/14/2024 07:17:18 AM, Rhuland, Courtney) Recall completed by BD
65919 | 203170 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-14 13:21 | 2024 Alaris Software Recall | (TC 03/14/2024 01:21:00 PM, Baghirzade, Nara) Software recall completed by BD.
65920 | 203171 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 12:55 | 2024 Alaris Software Recall | (TC 03/28/2024 12:54:58 PM, Rhuland, Courtney) Recall completed by BD
65921 | 203172 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-05 14:07 | 2024 Alaris Software Recall | (TC 04/05/2024 02:07:06 PM, Rhuland, Courtney) Recall completed by BD
65922 | 203173 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 14:11 | 2024 Alaris Software Recall | (TC 03/12/2024 02:10:43 PM, Baghirzade, Nara) Software recall completed by BD.
65923 | 203174 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-20 09:19 | 2024 Alaris Software Recall | (TC 03/20/2024 09:18:56 AM, Baghirzade, Nara) Software recall completed by BD.
65924 | 203175 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 13:07 | 2024 Alaris Software Recall | (TC 03/28/2024 01:06:33 PM, Rhuland, Courtney) Recall completed by BD
65925 | 203176 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 07:17 | 2024 Alaris Software Recall | (TC 07/31/2024 07:17:15 AM, Baghirzade, Nara) Software upgrade completed.
65926 | 203177 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-05 09:19 | 2024 Alaris Software Recall | (TC 04/05/2024 09:18:47 AM, Rhuland, Courtney) Recall completed by BD
65927 | 203178 | PS | CNP | ENL | 2024-03-01 15:57 | 2024-08-16 15:34 | 2024 Alaris Software Recall
65928 | 203179 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 11:16 | 2024 Alaris Software Recall
65929 | 203180 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-13 14:29 | 2024 Alaris Software Recall | (TC 03/13/2024 02:28:58 PM, Rhuland, Courtney) Recall completed by BD
65930 | 203181 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 13:40 | 2024 Alaris Software Recall | (TC 03/15/2024 01:39:38 PM, Baghirzade, Nara) Software recall completed by BD.
65931 | 203182 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-15 10:55 | 2024 Alaris Software Recall | Recall completed
65932 | 203183 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-19 10:40 | 2024 Alaris Software Recall | (TC 03/19/2024 10:39:08 AM, Rhuland, Courtney) Recall completed by BD
65933 | 203184 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-10 14:39 | 2024 Alaris Software Recall | (TC 04/10/2024 02:39:09 PM, Rhuland, Courtney) Recall completed by BD
65934 | 203185 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-10 14:42 | 2024 Alaris Software Recall | (TC 04/10/2024 02:41:28 PM, Rhuland, Courtney) Recall completed by BD
65935 | 203186 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-25 09:48 | 2024 Alaris Software Recall | (TC 03/25/2024 09:47:54 AM, Baghirzade, Nara) Software recall completed by BD.
65936 | 203187 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-15 11:53 | 2024 Alaris Software Recall | (TC 04/15/2024 11:52:23 AM, Baghirzade, Nara) Software recall completed by BD.
65937 | 203188 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 11:19 | 2024 Alaris Software Recall
65938 | 203189 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-07-31 09:48 | 2024 Alaris Software Recall | Software updated.
65939 | 203190 | PS | CNP | ENL | 2024-03-01 15:57 | 2024-08-16 15:35 | 2024 Alaris Software Recall
65940 | 203191 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 10:29 | 2024 Alaris Software Recall | (TC 03/22/2024 10:29:10 AM, Baghirzade, Nara) Software recall completed by BD.
65941 | 203192 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-01 13:18 | 2024 Alaris Software Recall
65942 | 203193 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-22 12:02 | 2024 Alaris Software Recall | (TC 03/22/2024 12:01:55 PM, Baghirzade, Nara) Software recall completed by BD.
65943 | 203194 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-04-17 13:41 | 2024 Alaris Software Recall | (TC 04/17/2024 01:40:37 PM, Baghirzade, Nara) Software recall completed by BD.
65944 | 203195 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-14 13:25 | 2024 Alaris Software Recall | (TC 03/14/2024 01:24:42 PM, Baghirzade, Nara) Software recall completed by BD.
65945 | 203196 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-08-01 13:09 | 2024 Alaris Software Recall
65946 | 203197 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-12 10:53 | 2024 Alaris Software Recall | (TC 03/12/2024 10:52:57 AM, Rendon Lopez , Elizabeth) Recall completed
65947 | 203198 | PS | CMPLT | CMPLT | 2024-03-01 15:57 | 2024-03-28 13:08 | 2024 Alaris Software Recall | (TC 03/28/2024 01:08:07 PM, Rhuland, Courtney) Recall completed by BD
65948 | 203199 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 12:12 | 2024 Alaris Software Recall | (TC 04/17/2024 12:11:44 PM, Baghirzade, Nara) Software recall completed by BD.
65949 | 203200 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-07-30 11:50 | 2024 Alaris Software Recall | (TC 07/30/2024 11:50:05 AM, Baghirzade, Nara) Software upgrade completed.
65950 | 203201 | PS | CNP | ENL | 2024-03-01 15:58 | 2024-08-16 15:35 | 2024 Alaris Software Recall
65951 | 203202 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-10 14:48 | 2024 Alaris Software Recall | (TC 04/10/2024 02:47:41 PM, Rhuland, Courtney) Recall completed by BD
65952 | 203203 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-05-02 14:54 | 2024 Alaris Software Recall
65953 | 203204 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-28 12:59 | 2024 Alaris Software Recall | (TC 03/28/2024 12:59:05 PM, Rhuland, Courtney) Recall completed by BD
65954 | 203205 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 14:11 | 2024 Alaris Software Recall | (TC 03/15/2024 02:11:09 PM, Rhuland, Courtney) Recall completed by BD
65955 | 203206 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 13:14 | 2024 Alaris Software Recall | (TC 03/25/2024 01:14:06 PM, Baghirzade, Nara) Software recall completed by BD.
65956 | 203207 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-20 11:32 | 2024 Alaris Software Recall | (TC 03/20/2024 11:15:30 AM, Reid, Emmerson ) Replaced front case and gasket. Ran Calibration, failed numerour times for the sizer at 30mm. After reopening it several times I was able to get it to work. Calibration passed. ASM PM Passed.
65957 | 203208 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 13:37 | 2024 Alaris Software Recall | (TC 03/15/2024 01:36:49 PM, Baghirzade, Nara) Software recall completed by BD.
65958 | 203209 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-05-02 10:46 | 2024 Alaris Software Recall
65959 | 203210 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 10:55 | 2024 Alaris Software Recall | (TC 03/18/2024 10:54:40 AM, Baghirzade, Nara) Software recall completed by BD.
65960 | 203211 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 15:54 | 2024 Alaris Software Recall | Recall completed by BD
65961 | 203212 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 08:04 | 2024 Alaris Software Recall | (TC 08/01/2024 08:03:56 AM, Baghirzade, Nara) Software upgrade completed.
65962 | 203213 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-30 13:33 | 2024 Alaris Software Recall
65963 | 203214 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-13 14:21 | 2024 Alaris Software Recall | (TC 03/13/2024 02:20:40 PM, Rhuland, Courtney) Recall completed by BD
65964 | 203215 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 13:06 | 2024 Alaris Software Recall
65965 | 203216 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 13:28 | 2024 Alaris Software Recall | (TC 03/14/2024 01:27:18 PM, Baghirzade, Nara) Software recall completed by BD.
65966 | 203217 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-19 09:05 | 2024 Alaris Software Recall | (TC 03/19/2024 09:04:12 AM, Rhuland, Courtney) Recall completed by BD
65967 | 203218 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 10:26 | 2024 Alaris Software Recall | Completed by BD
65968 | 203219 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 13:04 | 2024 Alaris Software Recall
65969 | 203220 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-13 13:10 | 2024 Alaris Software Recall | (TC 03/13/2024 01:09:21 PM, Baghirzade, Nara) Software recall completed by BD.
65970 | 203221 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-28 13:03 | 2024 Alaris Software Recall | (TC 03/28/2024 01:02:51 PM, Rhuland, Courtney) Recall completed by BD
65971 | 203222 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 10:51 | 2024 Alaris Software Recall | (TC 03/12/2024 10:50:53 AM, Rendon Lopez , Elizabeth) Recall completed
65972 | 203223 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 10:23 | 2024 Alaris Software Recall | Completed by BD
65973 | 203224 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-07-25 11:48 | 2024 Alaris Software Recall
65974 | 203225 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-07-03 13:32 | 2024 Alaris Software Recall
65975 | 203226 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-26 14:20 | 2024 Alaris Software Recall | (TC 03/26/2024 02:20:06 PM, Rhuland, Courtney) Recall completed by BD
65976 | 203227 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 07:38 | 2024 Alaris Software Recall | (TC 03/14/2024 07:37:54 AM, Rhuland, Courtney) Recall completed by BD
65977 | 203228 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-07-31 15:56 | 2024 Alaris Software Recall | Software update completed.
65978 | 203229 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 12:58 | 2024 Alaris Software Recall
65979 | 203230 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 10:24 | 2024 Alaris Software Recall | Completed by BD
65981 | 203231 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 12:54 | 2024 Alaris Software Recall
65982 | 203232 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 10:25 | 2024 Alaris Software Recall | Completed by BD
65983 | 203233 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-10 14:46 | 2024 Alaris Software Recall | (TC 04/10/2024 02:45:46 PM, Rhuland, Courtney) Recall completed by BD
65984 | 203234 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-20 09:25 | 2024 Alaris Software Recall | (TC 03/20/2024 09:24:46 AM, Baghirzade, Nara) Software recall completed by BD.
65985 | 203235 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-13 14:27 | 2024 Alaris Software Recall | (TC 03/13/2024 02:26:33 PM, Rhuland, Courtney) Recall completed by BD
65986 | 203236 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 13:17 | 2024 Alaris Software Recall | (TC 03/25/2024 01:16:40 PM, Baghirzade, Nara) Software recall completed by BD.
65987 | 203237 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-13 14:22 | 2024 Alaris Software Recall | (TC 03/13/2024 02:22:11 PM, Rhuland, Courtney) Recall completed by BD
65988 | 203238 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 08:30 | 2024 Alaris Software Recall | (TC 03/21/2024 08:29:45 AM, Rhuland, Courtney) Recall completed by BD
65989 | 203239 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-26 14:21 | 2024 Alaris Software Recall | (TC 03/26/2024 02:21:09 PM, Rhuland, Courtney) Recall completed by BD
65990 | 203240 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 10:54 | 2024 Alaris Software Recall | Completed by BD
65992 | 203241 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:24 | 2024 Alaris Software Recall | Completed by BD
65993 | 203242 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-05-10 13:38 | 2024 Alaris Software Recall
65994 | 203243 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 12:45 | 2024 Alaris Software Recall
65996 | 203244 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 13:58 | 2024 Alaris Software Recall
65997 | 203245 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 12:40 | 2024 Alaris Software Recall
65998 | 203246 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 12:31 | 2024 Alaris Software Recall
65999 | 203247 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:35 | 2024 Alaris Software Recall
66000 | 203248 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 11:15 | 2024 Alaris Software Recall | Recall completed by BD
66001 | 203249 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 08:15 | 2024 Alaris Software Recall
66002 | 203250 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-11 14:14 | 2024 Alaris Software Recall | Completed by BD
66003 | 203251 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 09:37 | 2024 Alaris Software Recall
66004 | 203252 | PS | CANCL | CANCL | 2024-03-01 15:58 | 2024-11-05 08:03 | 2024 Alaris Software Recall
66005 | 203253 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 08:12 | 2024 Alaris Software Recall
66006 | 203254 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 12:27 | 2024 Alaris Software Recall
66007 | 203255 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 08:10 | 2024 Alaris Software Recall
66008 | 203256 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 14:09 | 2024 Alaris Software Recall | (TC 03/15/2024 02:08:26 PM, Baghirzade, Nara) Software recall completed by BD.
66009 | 203257 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 12:21 | 2024 Alaris Software Recall
66010 | 203258 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-08 10:47 | 2024 Alaris Software Recall
66011 | 203259 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 13:55 | 2024 Alaris Software Recall | Recall completed
66012 | 203260 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-26 12:53 | 2024 Alaris Software Recall | Completed
66013 | 203261 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-11 14:32 | 2024 Alaris Software Recall | Completed by BD
66014 | 203262 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 10:27 | 2024 Alaris Software Recall | Recall completed
66015 | 203263 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 08:08 | 2024 Alaris Software Recall
66016 | 203264 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 08:03 | 2024 Alaris Software Recall
66017 | 203265 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-03 14:02 | 2024 Alaris Software Recall
66018 | 203266 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:18 | 2024 Alaris Software Recall
66019 | 203267 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 08:02 | 2024 Alaris Software Recall
66020 | 203268 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:58 | 2024 Alaris Software Recall
66021 | 203269 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 15:10 | 2024 Alaris Software Recall | Recall completed
66022 | 203270 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:09 | 2024 Alaris Software Recall | (TC 03/12/2024 09:08:49 AM, Rendon Lopez , Elizabeth) Recall completed
66023 | 203271 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:55 | 2024 Alaris Software Recall
66024 | 203272 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-05-23 12:23 | 2024 Alaris Software Recall | (TC 05/23/2024 12:23:11 PM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files. Battery replaced.
66025 | 203273 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:06 | 2024 Alaris Software Recall | (TC 03/12/2024 09:06:01 AM, Rendon Lopez , Elizabeth) Recall completed
66026 | 203274 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:32 | 2024 Alaris Software Recall
66027 | 203275 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 14:21 | 2024 Alaris Software Recall
66028 | 203276 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 09:46 | 2024 Alaris Software Recall
66029 | 203277 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:54 | 2024 Alaris Software Recall
66030 | 203278 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 12:23 | 2024 Alaris Software Recall
66031 | 203279 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:53 | 2024 Alaris Software Recall
66032 | 203280 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:50 | 2024 Alaris Software Recall
66033 | 203281 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 09:40 | 2024 Alaris Software Recall
66034 | 203282 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 09:41 | 2024 Alaris Software Recall
66035 | 203283 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:47 | 2024 Alaris Software Recall
66036 | 203284 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 13:25 | 2024 Alaris Software Recall
66037 | 203285 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:45 | 2024 Alaris Software Recall
66039 | 203286 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 11:20 | 2024 Alaris Software Recall | Completed by BD
66040 | 203287 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 11:26 | 2024 Alaris Software Recall | Recall completed by BD
66041 | 203288 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:43 | 2024 Alaris Software Recall
66042 | 203289 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 11:22 | 2024 Alaris Software Recall | Completed by BD
66043 | 203290 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-11 14:35 | 2024 Alaris Software Recall | Completed by BD
66044 | 203291 | PS | CNP | ENL | 2024-03-01 15:58 | 2024-08-16 15:35 | 2024 Alaris Software Recall
66045 | 203292 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 14:00 | 2024 Alaris Software Recall | Recall completed
66046 | 203293 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-27 13:50 | 2024 Alaris Software Recall
66047 | 203294 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:04 | 2024 Alaris Software Recall
66048 | 203295 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 13:45 | 2024 Alaris Software Recall | Recall completed
66049 | 203296 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:40 | 2024 Alaris Software Recall
66050 | 203297 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 09:31 | 2024 Alaris Software Recall
66051 | 203298 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 12:18 | 2024 Alaris Software Recall
66052 | 203299 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-02 09:25 | 2024 Alaris Software Recall | SW recall complete by BD
66053 | 203300 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:38 | 2024 Alaris Software Recall
66054 | 203301 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:36 | 2024 Alaris Software Recall
66055 | 203302 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 08:41 | 2024 Alaris Software Recall | Completed by BD
66056 | 203303 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-11 14:19 | 2024 Alaris Software Recall
66057 | 203304 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 14:11 | 2024 Alaris Software Recall
66058 | 203305 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-23 10:01 | 2024 Alaris Software Recall | (TC 04/23/2024 10:00:53 AM, Rhuland, Courtney) Recall completed by BD
66059 | 203306 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:34 | 2024 Alaris Software Recall
66060 | 203307 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 10:46 | 2024 Alaris Software Recall | Completed by BD
66061 | 203308 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:20 | 2024 Alaris Software Recall | Recall completed
66062 | 203309 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 10:28 | 2024 Alaris Software Recall | (TC 03/04/2024 10:28:02 AM, Huang, Jessica) BD completed
66063 | 203310 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:32 | 2024 Alaris Software Recall
66064 | 203311 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-24 14:46 | 2024 Alaris Software Recall
66065 | 203312 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-01 07:30 | 2024 Alaris Software Recall
66066 | 203313 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 12:15 | 2024 Alaris Software Recall
66067 | 203314 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:41 | 2024 Alaris Software Recall
66068 | 203315 | PS | CNP | ENL | 2024-03-01 15:58 | 2024-08-16 15:35 | 2024 Alaris Software Recall
66069 | 203316 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 14:18 | 2024 Alaris Software Recall
66070 | 203317 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:16 | 2024 Alaris Software Recall
68288 | 203318 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-11 14:37 | 2024 Alaris Software Recall | Completed by BD
68289 | 203319 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:50 | 2024 Alaris Software Recall
68290 | 203320 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-27 13:51 | 2024 Alaris Software Recall
68291 | 203321 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 09:44 | 2024 Alaris Software Recall
68292 | 203322 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 09:50 | 2024 Alaris Software Recall
68293 | 203323 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:21 | 2024 Alaris Software Recall
68294 | 203324 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 18:04 | 2024 Alaris Software Recall
68295 | 203325 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-05-03 12:57 | 2024 Alaris Software Recall | (TC 05/03/2024 12:56:27 PM, Baghirzade, Nara) Completed by BD
68296 | 203326 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-05-21 13:08 | 2024 Alaris Software Recall | (TC 05/21/2024 01:08:10 PM, Baghirzade, Nara) Software upgrade completed.
68297 | 203327 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 09:44 | 2024 Alaris Software Recall | Recall completed
68298 | 203328 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-03 10:50 | 2024 Alaris Software Recall | Completed
68299 | 203329 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 11:01 | 2024 Alaris Software Recall | Recall completed
68300 | 203330 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:27 | 2024 Alaris Software Recall
68301 | 203331 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 12:46 | 2024 Alaris Software Recall | Completed by BD.
68302 | 203332 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 09:00 | 2024 Alaris Software Recall | Completed by BD
68579 | 203333 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-19 09:37 | 2024 Alaris Software Recall | (TC 03/19/2024 09:36:20 AM, Rhuland, Courtney) Recall completed by BD
68580 | 203334 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 12:49 | 2024 Alaris Software Recall | Completed by BD.
68581 | 203335 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:39 | 2024 Alaris Software Recall | Completed by BD
68582 | 203336 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:46 | 2024 Alaris Software Recall | Completed by BD
68583 | 203337 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:30 | 2024 Alaris Software Recall | Completed by BD
68589 | 203338 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 12:52 | 2024 Alaris Software Recall | Completed by BD.
68590 | 203339 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 14:03 | 2024 Alaris Software Recall | (TC 03/15/2024 02:02:23 PM, Rhuland, Courtney) Recall completed by BD
68592 | 203340 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 13:19 | 2024 Alaris Software Recall | (TC 03/25/2024 01:18:41 PM, Baghirzade, Nara) Software recall completed by BD.
68593 | 203341 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-13 15:06 | 2024 Alaris Software Recall | (TC 03/13/2024 03:05:27 PM, Rhuland, Courtney) Recall completed by BD
68594 | 203342 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-20 14:24 | 2024 Alaris Software Recall | (TC 03/20/2024 02:23:55 PM, Baghirzade, Nara) Software recall completed by BD.
68595 | 203343 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-11 14:43 | 2024 Alaris Software Recall | Recall completed
68596 | 203344 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-20 09:22 | 2024 Alaris Software Recall | (TC 03/20/2024 09:21:35 AM, Baghirzade, Nara) Software recall completed by BD.
68597 | 203345 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 12:53 | 2024 Alaris Software Recall | Completed by BD.
68598 | 203346 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-20 13:56 | 2024 Alaris Software Recall
68599 | 203347 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 15:52 | 2024 Alaris Software Recall | Recall completed by BD
68600 | 203348 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 09:57 | 2024 Alaris Software Recall | Completed by BD
68601 | 203349 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 13:12 | 2024 Alaris Software Recall | (TC 03/21/2024 01:11:28 PM, Baghirzade, Nara) Software recall completed by BD.
68602 | 203350 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 10:00 | 2024 Alaris Software Recall | Completed by BD
68603 | 203351 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 10:04 | 2024 Alaris Software Recall | Completed by BD
68604 | 203352 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 12:47 | 2024 Alaris Software Recall | (TC 03/21/2024 12:46:27 PM, Baghirzade, Nara) Software recall completed by BD.
68605 | 203353 | PS | ACTIV | ISSUE | 2024-03-01 15:58 |  | 2024 Alaris Software Recall
68606 | 203354 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-07-30 13:20 | 2024 Alaris Software Recall | (TC 07/30/2024 01:19:37 PM, DeGuzman, Leila) Software updated.
68607 | 203355 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-21 13:58 | 2024 Alaris Software Recall | (TC 03/21/2024 01:57:49 PM, Baghirzade, Nara) Software recall completed by BD.
68608 | 203356 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-15 14:08 | 2024 Alaris Software Recall | (TC 03/15/2024 02:08:05 PM, Rhuland, Courtney) Recall completed by BD
68609 | 203357 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 12:55 | 2024 Alaris Software Recall | Completed by BD.
68610 | 203358 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-22 10:31 | 2024 Alaris Software Recall | (TC 03/22/2024 10:30:18 AM, Baghirzade, Nara) Software recall completed by BD.
68611 | 203359 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-12 10:58 | 2024 Alaris Software Recall | (TC 03/12/2024 10:57:25 AM, Rendon Lopez , Elizabeth) Recall completed
68612 | 203360 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 12:57 | 2024 Alaris Software Recall | Completed by BD.
68613 | 203361 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-03 13:26 | 2024 Alaris Software Recall | (TC 04/03/2024 01:25:36 PM, Rhuland, Courtney) Recall completed by BD
68614 | 203362 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 12:35 | 2024 Alaris Software Recall
68615 | 203363 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 12:59 | 2024 Alaris Software Recall | Completed by BD.
68616 | 203364 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-03 14:12 | 2024 Alaris Software Recall | (TC 04/03/2024 02:11:27 PM, Rhuland, Courtney) Recall completed by BD
68617 | 203365 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-08-16 13:01 | 2024 Alaris Software Recall | Completed by BD.
68618 | 203366 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-14 14:14 | 2024 Alaris Software Recall | Recall completed
68619 | 203367 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-03 13:23 | 2024 Alaris Software Recall | (TC 04/03/2024 01:22:35 PM, Rhuland, Courtney) Recall completed by BD
68620 | 203368 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-24 14:49 | 2024 Alaris Software Recall | Software upgrade completed by BD
68621 | 203369 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-17 08:45 | 2024 Alaris Software Recall | Completed by BD
68622 | 203370 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-26 12:08 | 2024 Alaris Software Recall | Completed
68623 | 203371 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 14:46 | 2024 Alaris Software Recall
68624 | 203372 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-12 15:34 | 2024 Alaris Software Recall | Completed by BD
68625 | 203373 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-18 14:49 | 2024 Alaris Software Recall
68626 | 203374 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-26 12:58 | 2024 Alaris Software Recall | Completed
68627 | 203375 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-15 11:35 | 2024 Alaris Software Recall | Completed by BD
68628 | 203376 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-04-23 10:07 | 2024 Alaris Software Recall | (TC 04/23/2024 10:07:07 AM, Rhuland, Courtney) Recall completed by BD
68629 | 203377 | PS | CMPLT | CMPLT | 2024-03-01 15:58 | 2024-03-25 10:49 | 2024 Alaris Software Recall
68630 | 203378 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-15 11:34 | 2024 Alaris Software Recall | Completed by BD
68631 | 203379 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-05-24 10:56 | 2024 Alaris Software Recall | (TC 05/24/2024 10:55:51 AM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68632 | 203380 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-23 10:23 | 2024 Alaris Software Recall
68633 | 203381 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-15 11:43 | 2024 Alaris Software Recall | Completed by BD
68634 | 203382 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 18:17 | 2024 Alaris Software Recall
68635 | 203383 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-05-03 09:04 | 2024 Alaris Software Recall | (TC 05/03/2024 09:04:11 AM, Baghirzade, Nara) Software recall completed by BD.
68636 | 203384 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:02 | 2024 Alaris Software Recall | Completed by BD.
68637 | 203385 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-26 09:06 | 2024 Alaris Software Recall
68638 | 203386 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-10 09:12 | 2024 Alaris Software Recall | Completed
68639 | 203387 | PS | CNP | ENL | 2024-03-01 15:59 | 2024-11-22 11:00 | 2024 Alaris Software Recall
68640 | 203388 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-22 14:08 | 2024 Alaris Software Recall
68641 | 203389 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-10 09:10 | 2024 Alaris Software Recall | Completed
68642 | 203390 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:08 | 2024 Alaris Software Recall | Completed by BD.
68643 | 203391 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-28 10:41 | 2024 Alaris Software Recall | Completed
68644 | 203392 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-04 12:52 | 2024 Alaris Software Recall | (TC 04/04/2024 12:51:22 PM, Rhuland, Courtney) Recall completed by BD
68645 | 203393 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-05-21 13:09 | 2024 Alaris Software Recall | (TC 05/21/2024 01:09:27 PM, Baghirzade, Nara) Software upgrade completed.
68646 | 203394 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 10:50 | 2024 Alaris Software Recall
68647 | 203395 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-26 09:57 | 2024 Alaris Software Recall
68648 | 203396 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 18:22 | 2024 Alaris Software Recall
68649 | 203397 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 10:30 | 2024 Alaris Software Recall | Completed
68650 | 203398 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-19 20:36 | 2024 Alaris Software Recall
68651 | 203399 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 12:01 | 2024 Alaris Software Recall
68652 | 203400 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 14:47 | 2024 Alaris Software Recall | (TC 04/03/2024 02:46:37 PM, Rhuland, Courtney) Recall completed by BD
68653 | 203401 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-05-16 11:31 | 2024 Alaris Software Recall | (TC 05/16/2024 11:30:41 AM, Baghirzade, Nara) Software upgrade completed
68654 | 203402 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-24 09:51 | 2024 Alaris Software Recall
68655 | 203403 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 14:08 | 2024 Alaris Software Recall
68656 | 203404 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 11:05 | 2024 Alaris Software Recall | Completed
68657 | 203405 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-09 12:18 | 2024 Alaris Software Recall | (TC 08/09/2024 12:18:12 PM, Patel, Rutvik) by BD
68658 | 203406 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-24 09:50 | 2024 Alaris Software Recall
68659 | 203407 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 13:55 | 2024 Alaris Software Recall
68660 | 203408 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 09:42 | 2024 Alaris Software Recall
68661 | 203409 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-14 15:02 | 2024 Alaris Software Recall
68662 | 203410 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-05-08 14:19 | 2024 Alaris Software Recall | Completed by BD.
68663 | 203411 | PS | ACTIV | ISSUE | 2024-03-01 15:59 |  | 2024 Alaris Software Recall
68664 | 203412 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 13:43 | 2024 Alaris Software Recall
68665 | 203413 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 14:14 | 2024 Alaris Software Recall
68666 | 203414 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 10:50 | 2024 Alaris Software Recall | Completed
68667 | 203415 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 09:59 | 2024 Alaris Software Recall | (TC 03/12/2024 09:58:34 AM, Rendon Lopez , Elizabeth) Recall completed
68668 | 203416 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-26 09:48 | 2024 Alaris Software Recall
68669 | 203417 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 11:06 | 2024 Alaris Software Recall | Completed
68670 | 203418 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-24 09:36 | 2024 Alaris Software Recall
68671 | 203419 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 13:57 | 2024 Alaris Software Recall
68672 | 203420 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 10:10 | 2024 Alaris Software Recall | Recall completed
68673 | 203421 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:09 | 2024 Alaris Software Recall | Completed by BD.
68674 | 203422 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:13 | 2024 Alaris Software Recall | Completed by BD.
68675 | 203423 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-20 14:41 | 2024 Alaris Software Recall
68676 | 203424 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-15 11:30 | 2024 Alaris Software Recall | Completed by BD
68677 | 203425 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-20 14:38 | 2024 Alaris Software Recall
68678 | 203426 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 10:28 | 2024 Alaris Software Recall
68679 | 203427 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 13:13 | 2024 Alaris Software Recall | Completed
68680 | 203428 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-14 15:34 | 2024 Alaris Software Recall | Recall completed
68681 | 203429 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 13:41 | 2024 Alaris Software Recall | Completed by BD
68682 | 203430 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-24 14:56 | 2024 Alaris Software Recall
68683 | 203431 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 15:10 | 2024 Alaris Software Recall | Completed by BD
68684 | 203432 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-19 09:35 | 2024 Alaris Software Recall | Completed by BD
68685 | 203433 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-14 10:04 | 2024 Alaris Software Recall | (TC 03/04/2024 10:04:11 AM, Huang, Jessica) BD completed
68686 | 203434 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:20 | 2024 Alaris Software Recall | (TC 04/03/2024 01:19:44 PM, Rhuland, Courtney) Recall completed by BD
68687 | 203435 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:35 | 2024 Alaris Software Recall | (TC 04/03/2024 01:34:22 PM, Rhuland, Courtney) Recall completed by BD
68688 | 203436 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-30 13:31 | 2024 Alaris Software Recall
68689 | 203437 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 14:18 | 2024 Alaris Software Recall
68690 | 203438 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 09:12 | 2024 Alaris Software Recall | (TC 03/12/2024 09:11:35 AM, Rendon Lopez , Elizabeth) Recall completed
68691 | 203439 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 15:06 | 2024 Alaris Software Recall
68692 | 203440 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:55 | 2024 Alaris Software Recall | (TC 04/03/2024 01:54:24 PM, Rhuland, Courtney) Recall completed by BD
68693 | 203441 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:13 | 2024 Alaris Software Recall | Completed by BD.
68694 | 203442 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 14:46 | 2024 Alaris Software Recall | Completed by BD
68696 | 203443 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 14:42 | 2024 Alaris Software Recall | Completed by BD
68697 | 203444 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 14:44 | 2024 Alaris Software Recall | Completed by BD
68698 | 203445 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 11:52 | 2024 Alaris Software Recall | (TC 03/15/2024 11:50:28 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 80 mohms Leakage current: 68.9 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. (TC 03/15/2024 11:50:28 AM, Baghirzade, Nara) Software recall completed by BD.
68699 | 203446 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 10:44 | 2024 Alaris Software Recall
68700 | 203447 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:15 | 2024 Alaris Software Recall | (TC 04/03/2024 01:14:54 PM, Rhuland, Courtney) Recall completed by BD
68701 | 203448 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 13:00 | 2024 Alaris Software Recall | Completed
68702 | 203449 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 10:51 | 2024 Alaris Software Recall
68703 | 203450 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 14:55 | 2024 Alaris Software Recall | Completed by BD
68704 | 203451 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 11:17 | 2024 Alaris Software Recall
68705 | 203452 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-13 14:24 | 2024 Alaris Software Recall | (TC 03/13/2024 02:23:37 PM, Baghirzade, Nara) Software recall completed by BD.
68706 | 203453 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 13:04 | 2024 Alaris Software Recall
68707 | 203454 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-13 08:52 | 2024 Alaris Software Recall
68708 | 203455 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 10:05 | 2024 Alaris Software Recall
68709 | 203456 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:20 | 2024 Alaris Software Recall | Completed by BD.
68710 | 203457 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-07-17 10:58 | 2024 Alaris Software Recall | Software updated. (TC 07/17/2024 01:37:03 PM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68711 | 203458 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 12:04 | 2024 Alaris Software Recall | Completed
68712 | 203459 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 14:56 | 2024 Alaris Software Recall | Completed by BD
68713 | 203460 | PS | CNP | ENL | 2024-03-01 15:59 | 2024-08-16 15:36 | 2024 Alaris Software Recall
68714 | 203461 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 10:19 | 2024 Alaris Software Recall | Recall completed
68715 | 203462 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 09:52 | 2024 Alaris Software Recall | Recall completed
68716 | 203463 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 15:05 | 2024 Alaris Software Recall | Completed by BD
68717 | 203464 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 10:36 | 2024 Alaris Software Recall
68718 | 203465 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:59 | 2024 Alaris Software Recall | (TC 04/03/2024 01:59:04 PM, Rhuland, Courtney) Recall completed by BD
68719 | 203466 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 13:55 | 2024 Alaris Software Recall
68720 | 203467 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:25 | 2024 Alaris Software Recall | Completed
68721 | 203468 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 11:06 | 2024 Alaris Software Recall
68722 | 203469 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-05 14:47 | 2024 Alaris Software Recall | (TC 04/05/2024 02:46:59 PM, Rhuland, Courtney) Recall completed by BD
68723 | 203470 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-17 08:48 | 2024 Alaris Software Recall | Completed by BD
68724 | 203471 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-17 10:48 | 2024 Alaris Software Recall | Completed by BD
68725 | 203472 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:22 | 2024 Alaris Software Recall | Completed by BD.
68726 | 203473 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 14:29 | 2024 Alaris Software Recall
68727 | 203474 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:23 | 2024 Alaris Software Recall | Completed by BD.
68728 | 203475 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:24 | 2024 Alaris Software Recall | Completed by BD.
68729 | 203476 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:35 | 2024 Alaris Software Recall | Completed
68730 | 203477 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 10:35 | 2024 Alaris Software Recall
68731 | 203478 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 15:08 | 2024 Alaris Software Recall | Completed by BD
68732 | 203479 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:25 | 2024 Alaris Software Recall | Completed by BD.
68733 | 203480 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 09:47 | 2024 Alaris Software Recall
68734 | 203481 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:27 | 2024 Alaris Software Recall | Completed by BD.
68735 | 203482 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 09:52 | 2024 Alaris Software Recall
68736 | 203483 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 13:06 | 2024 Alaris Software Recall | Completed
68737 | 203484 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 15:07 | 2024 Alaris Software Recall
68738 | 203485 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 14:34 | 2024 Alaris Software Recall
68739 | 203486 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 11:14 | 2024 Alaris Software Recall
68740 | 203487 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:29 | 2024 Alaris Software Recall | Completed by BD.
68741 | 203488 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-15 11:32 | 2024 Alaris Software Recall | Completed by BD
68742 | 203489 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 15:06 | 2024 Alaris Software Recall | Completed by BD
68743 | 203490 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-15 11:31 | 2024 Alaris Software Recall | Completed by BD
68744 | 203491 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-20 14:24 | 2024 Alaris Software Recall
68745 | 203492 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:31 | 2024 Alaris Software Recall | Completed by BD.
68746 | 203493 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-17 10:48 | 2024 Alaris Software Recall | Completed by BD
68747 | 203494 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 10:41 | 2024 Alaris Software Recall
68748 | 203495 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-20 14:50 | 2024 Alaris Software Recall
68749 | 203496 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-13 09:24 | 2024 Alaris Software Recall | Recall completed
68750 | 203497 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 09:56 | 2024 Alaris Software Recall
68751 | 203498 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:34 | 2024 Alaris Software Recall | Completed
68752 | 203499 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-13 13:20 | 2024 Alaris Software Recall | (TC 03/13/2024 01:19:50 PM, Baghirzade, Nara) Software recall completed by BD.
68753 | 203500 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 11:26 | 2024 Alaris Software Recall | Completed by BD
68754 | 203501 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:14 | 2024 Alaris Software Recall | (TC 04/03/2024 01:13:43 PM, Rhuland, Courtney) Recall completed by BD
68755 | 203502 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-17 09:42 | 2024 Alaris Software Recall | (TC 04/17/2024 09:41:20 AM, Rhuland, Courtney) Recall completed by BD
68756 | 203503 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 12:53 | 2024 Alaris Software Recall
68757 | 203504 | PS | CNP | ENL | 2024-03-01 15:59 | 2024-08-16 15:36 | 2024 Alaris Software Recall
68758 | 203505 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 10:21 | 2024 Alaris Software Recall
68759 | 203506 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-10 09:01 | 2024 Alaris Software Recall | (TC 04/10/2024 09:01:00 AM, Rhuland, Courtney) Recall completed by BD
68760 | 203507 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 11:24 | 2024 Alaris Software Recall
68761 | 203508 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 14:09 | 2024 Alaris Software Recall | (TC 04/03/2024 02:08:55 PM, Rhuland, Courtney) Recall completed by BD
68762 | 203509 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 11:15 | 2024 Alaris Software Recall
68763 | 203510 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:33 | 2024 Alaris Software Recall | Completed by BD.
68764 | 203511 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-30 13:25 | 2024 Alaris Software Recall
68765 | 203512 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:33 | 2024 Alaris Software Recall | Completed by BD
68766 | 203513 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 11:02 | 2024 Alaris Software Recall
68767 | 203514 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:32 | 2024 Alaris Software Recall | Completed by BD
68768 | 203515 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-05 14:42 | 2024 Alaris Software Recall | (TC 04/05/2024 02:41:54 PM, Rhuland, Courtney) Recall completed by BD
68769 | 203516 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-14 14:37 | 2024 Alaris Software Recall | Recall completed
68770 | 203517 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:12 | 2024 Alaris Software Recall | Completed
68771 | 203518 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-05 14:45 | 2024 Alaris Software Recall | (TC 04/05/2024 02:44:34 PM, Rhuland, Courtney) Recall completed by BD
68772 | 203519 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:13 | 2024 Alaris Software Recall | Completed
68773 | 203520 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:22 | 2024 Alaris Software Recall | Completed by BD
68774 | 203521 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:21 | 2024 Alaris Software Recall | Completed by BD
68775 | 203522 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 09:53 | 2024 Alaris Software Recall
68776 | 203523 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-20 14:20 | 2024 Alaris Software Recall
68777 | 203524 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 09:27 | 2024 Alaris Software Recall | (TC 03/12/2024 09:26:36 AM, Rendon Lopez , Elizabeth) Recall completed
68778 | 203525 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-17 08:44 | 2024 Alaris Software Recall | Completed by BD
68779 | 203526 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-13 09:16 | 2024 Alaris Software Recall | Recall completed
68780 | 203527 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 11:07 | 2024 Alaris Software Recall
68781 | 203528 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-24 14:53 | 2024 Alaris Software Recall
68782 | 203529 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-15 10:03 | 2024 Alaris Software Recall | Recall completed
68783 | 203530 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 09:30 | 2024 Alaris Software Recall | (TC 03/12/2024 09:29:59 AM, Rendon Lopez , Elizabeth) Recall completed
68784 | 203531 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 13:45 | 2024 Alaris Software Recall | Completed by BD
68785 | 203532 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 11:04 | 2024 Alaris Software Recall
68786 | 203533 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:24 | 2024 Alaris Software Recall | (TC 04/03/2024 01:23:57 PM, Rhuland, Courtney) Recall completed by BD
68787 | 203534 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:35 | 2024 Alaris Software Recall | Completed by BD.
68788 | 203535 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-07-03 13:57 | 2024 Alaris Software Recall
68789 | 203536 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 10:09 | 2024 Alaris Software Recall
68790 | 203537 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:37 | 2024 Alaris Software Recall | Completed by BD.
68791 | 203538 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 09:55 | 2024 Alaris Software Recall
68792 | 203539 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 11:39 | 2024 Alaris Software Recall | Completed by BD
68793 | 203540 | PS | CNP | ENL | 2024-03-01 15:59 | 2024-08-16 15:36 | 2024 Alaris Software Recall
68794 | 203541 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-05 14:44 | 2024 Alaris Software Recall | (TC 04/05/2024 02:43:18 PM, Rhuland, Courtney) Recall completed by BD
68795 | 203542 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 14:08 | 2024 Alaris Software Recall
68796 | 203543 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 12:03 | 2024 Alaris Software Recall | Completed
68797 | 203544 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 10:02 | 2024 Alaris Software Recall
68798 | 203545 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:39 | 2024 Alaris Software Recall | Completed by BD.
68799 | 203546 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 13:19 | 2024 Alaris Software Recall | Completed
68800 | 203547 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:36 | 2024 Alaris Software Recall | (TC 04/03/2024 01:36:12 PM, Rhuland, Courtney) Recall completed by BD
68801 | 203548 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 09:57 | 2024 Alaris Software Recall
68802 | 203549 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-02 10:46 | 2024 Alaris Software Recall
68803 | 203550 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 13:50 | 2024 Alaris Software Recall | Completed by BD
68804 | 203551 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-12 11:34 | 2024 Alaris Software Recall | Completed by BD
68805 | 203552 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:07 | 2024 Alaris Software Recall | Completed
68806 | 203553 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 14:37 | 2024 Alaris Software Recall
68807 | 203554 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 10:06 | 2024 Alaris Software Recall
68808 | 203555 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-07-03 14:00 | 2024 Alaris Software Recall
68809 | 203556 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-14 13:35 | 2024 Alaris Software Recall | Recall completed
68810 | 203557 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:17 | 2024 Alaris Software Recall | Completed by BD
68811 | 203558 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-14 15:11 | 2024 Alaris Software Recall
68812 | 203559 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:41 | 2024 Alaris Software Recall | Completed by BD.
68813 | 203560 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-24 14:50 | 2024 Alaris Software Recall
68814 | 203561 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 12:59 | 2024 Alaris Software Recall
68815 | 203562 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:43 | 2024 Alaris Software Recall | Completed by BD.
68816 | 203563 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-13 14:14 | 2024 Alaris Software Recall | (TC 03/13/2024 02:14:04 PM, Baghirzade, Nara) Software recall completed by BD.
68817 | 203564 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:16 | 2024 Alaris Software Recall | Completed by BD
68818 | 203565 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:45 | 2024 Alaris Software Recall | Completed by BD.
68819 | 203566 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:18 | 2024 Alaris Software Recall | Completed by BD
68820 | 203567 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 15:03 | 2024 Alaris Software Recall | Completed by BD
68821 | 203568 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:10 | 2024 Alaris Software Recall | Completed
68822 | 203569 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-18 14:35 | 2024 Alaris Software Recall
68823 | 203570 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 10:01 | 2024 Alaris Software Recall
68824 | 203571 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:45 | 2024 Alaris Software Recall | (TC 04/03/2024 01:44:14 PM, Rhuland, Courtney) Recall completed by BD
68825 | 203572 | PS | CNP | ENL | 2024-03-01 15:59 | 2024-08-16 15:36 | 2024 Alaris Software Recall
68826 | 203573 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:23 | 2024 Alaris Software Recall | Completed by BD
68827 | 203574 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:42 | 2024 Alaris Software Recall | (TC 04/03/2024 01:41:57 PM, Rhuland, Courtney) Recall completed by BD
68828 | 203575 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-11 15:05 | 2024 Alaris Software Recall | Completed by BD
68829 | 203576 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-12 14:20 | 2024 Alaris Software Recall | Completed by BD
68830 | 203577 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-07-03 13:55 | 2024 Alaris Software Recall
68831 | 203578 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-21 09:58 | 2024 Alaris Software Recall
68832 | 203579 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-27 14:23 | 2024 Alaris Software Recall
68833 | 203580 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 10:11 | 2024 Alaris Software Recall
68834 | 203581 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-26 14:20 | 2024 Alaris Software Recall | Completed
68835 | 203582 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-03-25 10:33 | 2024 Alaris Software Recall
68836 | 203583 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:51 | 2024 Alaris Software Recall | Completed by BD.
68837 | 203584 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-05 14:46 | 2024 Alaris Software Recall | (TC 04/05/2024 02:46:09 PM, Rhuland, Courtney) Recall completed by BD
68838 | 203585 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-08-16 13:53 | 2024 Alaris Software Recall | Completed by BD.
68839 | 203586 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-24 14:55 | 2024 Alaris Software Recall | Recall completed by BD
68840 | 203587 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:47 | 2024 Alaris Software Recall | (TC 04/03/2024 01:46:37 PM, Rhuland, Courtney) Recall completed by BD
68841 | 203588 | PS | CMPLT | CMPLT | 2024-03-01 15:59 | 2024-04-03 13:49 | 2024 Alaris Software Recall | (TC 04/03/2024 01:48:31 PM, Rhuland, Courtney) Recall completed by BD
68842 | 203589 | PS | CNP | ENL | 2024-03-01 16:00 | 2024-08-16 15:37 | 2024 Alaris Software Recall
68843 | 203590 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-17 08:47 | 2024 Alaris Software Recall | Completed by BD
68844 | 203591 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-21 10:13 | 2024 Alaris Software Recall
68845 | 203592 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 11:37 | 2024 Alaris Software Recall | Completed by BD
68846 | 203593 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-23 12:17 | 2024 Alaris Software Recall | (TC 05/23/2024 12:16:54 PM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68847 | 203594 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 13:02 | 2024 Alaris Software Recall
68848 | 203595 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-21 10:19 | 2024 Alaris Software Recall
68849 | 203596 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-30 13:29 | 2024 Alaris Software Recall
68850 | 203597 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 10:01 | 2024 Alaris Software Recall | (TC 03/04/2024 10:00:42 AM, Huang, Jessica) BD completed
68851 | 203598 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 09:52 | 2024 Alaris Software Recall | (TC 03/04/2024 09:52:21 AM, Huang, Jessica) BD completed it
68852 | 203599 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 11:41 | 2024 Alaris Software Recall | Completed by BD
68853 | 203600 | PS | CNP | ENL | 2024-03-01 16:00 | 2024-08-16 15:37 | 2024 Alaris Software Recall
68854 | 203601 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 14:19 | 2024 Alaris Software Recall | (TC 03/06/2024 02:18:53 PM, Huang, Jessica) completed by BD
68855 | 203602 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 14:35 | 2024 Alaris Software Recall
68856 | 203603 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-24 14:51 | 2024 Alaris Software Recall
68857 | 203604 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 14:53 | 2024 Alaris Software Recall | Recall completed
68858 | 203605 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 13:15 | 2024 Alaris Software Recall | (TC 03/13/2024 01:15:12 PM, Baghirzade, Nara) Software recall completed by BD.
68859 | 203606 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-31 13:11 | 2024 Alaris Software Recall | (TC 05/31/2024 01:10:36 PM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68860 | 203607 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 13:55 | 2024 Alaris Software Recall | Completed by BD.
68861 | 203608 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 13:08 | 2024 Alaris Software Recall
68862 | 203609 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-07-03 13:53 | 2024 Alaris Software Recall
68863 | 203610 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 15:16 | 2024 Alaris Software Recall
68864 | 203611 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 14:13 | 2024 Alaris Software Recall | (TC 03/13/2024 02:12:30 PM, Baghirzade, Nara) Software recall completed by BD.
68865 | 203612 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-15 09:22 | 2024 Alaris Software Recall | Recall completed
68866 | 203613 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-27 14:15 | 2024 Alaris Software Recall
68867 | 203614 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 14:16 | 2024 Alaris Software Recall | (TC 03/13/2024 02:15:35 PM, Baghirzade, Nara) Software recall completed by BD.
68868 | 203615 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-06 13:56 | 2024 Alaris Software Recall | (TC 05/06/2024 01:55:21 PM, Baghirzade, Nara) Software recall completed by BD.
68869 | 203616 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-19 09:32 | 2024 Alaris Software Recall | Completed by BD
68870 | 203617 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 14:29 | 2024 Alaris Software Recall | Completed
68871 | 203618 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-03 14:04 | 2024 Alaris Software Recall | (TC 04/03/2024 02:04:04 PM, Rhuland, Courtney) Recall completed by BD
68872 | 203619 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 13:57 | 2024 Alaris Software Recall | Completed by BD.
68873 | 203620 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-24 14:47 | 2024 Alaris Software Recall
68874 | 203621 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 13:58 | 2024 Alaris Software Recall | Completed by BD.
68875 | 203622 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-25 09:52 | 2024 Alaris Software Recall
68876 | 203623 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-24 14:47 | 2024 Alaris Software Recall
68877 | 203624 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:01 | 2024 Alaris Software Recall | Completed by BD.
68878 | 203625 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 11:56 | 2024 Alaris Software Recall | Completed
68879 | 203626 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 11:24 | 2024 Alaris Software Recall | Completed by BD
68880 | 203627 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-21 10:26 | 2024 Alaris Software Recall
68881 | 203628 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-02 11:44 | 2024 Alaris Software Recall
68882 | 203629 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 12:06 | 2024 Alaris Software Recall | Completed
68883 | 203630 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-21 10:24 | 2024 Alaris Software Recall
68884 | 203631 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 15:12 | 2024 Alaris Software Recall | Completed by BD
68885 | 203632 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-25 10:46 | 2024 Alaris Software Recall
68886 | 203633 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 12:11 | 2024 Alaris Software Recall | Completed
68887 | 203634 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-15 11:33 | 2024 Alaris Software Recall | Completed by BD
68888 | 203635 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-03 14:08 | 2024 Alaris Software Recall | (TC 04/03/2024 02:07:29 PM, Rhuland, Courtney) Recall completed by BD
68889 | 203636 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 09:41 | 2024 Alaris Software Recall | (TC 03/04/2024 09:40:55 AM, Huang, Jessica) BD completed it
68890 | 203637 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-02 11:23 | 2024 Alaris Software Recall
68891 | 203638 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-27 14:25 | 2024 Alaris Software Recall
68892 | 203639 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 12:46 | 2024 Alaris Software Recall
68893 | 203640 | PS | CNP | ENL | 2024-03-01 16:00 | 2024-08-16 15:37 | 2024 Alaris Software Recall
68894 | 203641 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-20 14:19 | 2024 Alaris Software Recall
68895 | 203642 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-22 11:35 | 2024 Alaris Software Recall | (TC 05/22/2024 11:35:23 AM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68896 | 203643 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 09:32 | 2024 Alaris Software Recall | (TC 03/04/2024 09:31:49 AM, Huang, Jessica) BD completed it
68897 | 203644 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-02 10:54 | 2024 Alaris Software Recall
68898 | 203645 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-22 11:25 | 2024 Alaris Software Recall | (TC 05/22/2024 11:25:19 AM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68899 | 203646 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:04 | 2024 Alaris Software Recall | Completed by BD.
68900 | 203647 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 14:59 | 2024 Alaris Software Recall | Completed by BD
68901 | 203648 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-03 09:25 | 2024 Alaris Software Recall
68902 | 203649 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-24 14:54 | 2024 Alaris Software Recall
68903 | 203650 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 15:02 | 2024 Alaris Software Recall
68904 | 203651 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-27 13:59 | 2024 Alaris Software Recall
68905 | 203652 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-15 10:40 | 2024 Alaris Software Recall | Recall completed
68906 | 203653 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 14:56 | 2024 Alaris Software Recall
68907 | 203654 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-03 14:02 | 2024 Alaris Software Recall | (TC 04/03/2024 02:01:24 PM, Rhuland, Courtney) Recall completed by BD
68908 | 203655 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-07-22 11:20 | 2024 Alaris Software Recall | Recall completed by BD
68909 | 203656 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:09 | 2024 Alaris Software Recall | Completed by BD.
68910 | 203657 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 14:43 | 2024 Alaris Software Recall
68911 | 203658 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-09-10 08:52 | 2024 Alaris Software Recall | Completed by BD.
68912 | 203659 | PS | CNP | ENL | 2024-03-01 16:00 | 2024-08-16 15:38 | 2024 Alaris Software Recall
68913 | 203660 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 09:12 | 2024 Alaris Software Recall | Recall completed
68914 | 203661 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 14:08 | 2024 Alaris Software Recall | (TC 03/13/2024 02:07:52 PM, Baghirzade, Nara) Software recall completed by BD.
68915 | 203662 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 14:46 | 2024 Alaris Software Recall | Completed by BD
68916 | 203663 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 13:14 | 2024 Alaris Software Recall | Completed
68917 | 203664 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-22 11:23 | 2024 Alaris Software Recall | (TC 05/22/2024 11:23:11 AM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68918 | 203665 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-03 13:18 | 2024 Alaris Software Recall | (TC 04/03/2024 01:17:49 PM, Rhuland, Courtney) Recall completed by BD
68919 | 203666 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 09:10 | 2024 Alaris Software Recall | Recall completed
68920 | 203667 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:11 | 2024 Alaris Software Recall | Completed by BD.
68921 | 203668 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 14:16 | 2024 Alaris Software Recall | Completed
68922 | 203669 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-25 10:20 | 2024 Alaris Software Recall
68923 | 203670 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 14:22 | 2024 Alaris Software Recall | (TC 03/13/2024 02:21:57 PM, Baghirzade, Nara) Software recall completed by BD.
68924 | 203671 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:12 | 2024 Alaris Software Recall | Completed by BD.
68925 | 203672 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-17 08:53 | 2024 Alaris Software Recall | (TC 04/17/2024 08:52:53 AM, Rhuland, Courtney) Recall completed by BD
68926 | 203673 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-03 13:58 | 2024 Alaris Software Recall | (TC 04/03/2024 01:57:31 PM, Rhuland, Courtney) Recall completed by BD
68927 | 203674 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:15 | 2024 Alaris Software Recall | Completed by BD.
68928 | 203675 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:17 | 2024 Alaris Software Recall | Completed by BD.
68929 | 203676 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 11:44 | 2024 Alaris Software Recall | Completed by BD
68930 | 203677 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 14:54 | 2024 Alaris Software Recall | Completed by BD
68931 | 203678 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 09:25 | 2024 Alaris Software Recall | (TC 03/12/2024 09:24:25 AM, Rendon Lopez , Elizabeth) Recall completed
68932 | 203679 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 14:38 | 2024 Alaris Software Recall | Completed
68933 | 203680 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 14:50 | 2024 Alaris Software Recall | Completed by BD
68934 | 203681 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-24 14:49 | 2024 Alaris Software Recall
68935 | 203682 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 11:51 | 2024 Alaris Software Recall | Completed
68936 | 203683 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:18 | 2024 Alaris Software Recall | Completed by BD.
68937 | 203684 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-03 13:22 | 2024 Alaris Software Recall | (TC 04/03/2024 01:21:16 PM, Rhuland, Courtney) Recall completed by BD
68938 | 203685 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 09:20 | 2024 Alaris Software Recall | Recall completed
68939 | 203686 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 14:26 | 2024 Alaris Software Recall
68940 | 203687 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 13:07 | 2024 Alaris Software Recall | Completed
68941 | 203688 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-02 09:39 | 2024 Alaris Software Recall
68942 | 203689 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-25 10:07 | 2024 Alaris Software Recall
68943 | 203690 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 15:00 | 2024 Alaris Software Recall
68944 | 203691 | PS | CMPLT | CMPLT | 2024-03-01 16:00 |  | 2024 Alaris Software Recall | Completed by BD.
68945 | 203692 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 12:54 | 2024 Alaris Software Recall | Completed
68946 | 203693 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 12:13 | 2024 Alaris Software Recall | Completed
68947 | 203694 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 09:28 | 2024 Alaris Software Recall | (TC 03/12/2024 09:27:50 AM, Rendon Lopez , Elizabeth) Recall completed
68948 | 203695 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:29 | 2024 Alaris Software Recall | Completed by BD.
68949 | 203696 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 11:46 | 2024 Alaris Software Recall | Completed by BD
68950 | 203697 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 14:51 | 2024 Alaris Software Recall | Completed by BD
68951 | 203698 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 12:56 | 2024 Alaris Software Recall | Completed
68952 | 203699 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 13:10 | 2024 Alaris Software Recall | Completed
68953 | 203700 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 09:17 | 2024 Alaris Software Recall | Recall completed
68954 | 203701 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 09:09 | 2024 Alaris Software Recall | Recall completed
68955 | 203702 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 14:48 | 2024 Alaris Software Recall | Completed by BD
68956 | 203703 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 11:54 | 2024 Alaris Software Recall | Completed
68957 | 203704 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-17 08:43 | 2024 Alaris Software Recall | Completed by BD
68958 | 203705 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-21 09:51 | 2024 Alaris Software Recall
68959 | 203706 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-12 13:42 | 2024 Alaris Software Recall | Completed by BD
68960 | 203707 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-02 10:22 | 2024 Alaris Software Recall
68961 | 203708 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 12:48 | 2024 Alaris Software Recall
68962 | 203709 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-25 09:58 | 2024 Alaris Software Recall
68963 | 203710 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:30 | 2024 Alaris Software Recall | Completed by BD.
68964 | 203711 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:31 | 2024 Alaris Software Recall | Completed by BD.
68965 | 203712 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-24 14:55 | 2024 Alaris Software Recall
68966 | 203713 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-27 14:21 | 2024 Alaris Software Recall
68967 | 203714 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:34 | 2024 Alaris Software Recall | Completed by BD.
68968 | 203715 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 14:27 | 2024 Alaris Software Recall | Completed
68969 | 203716 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-25 10:24 | 2024 Alaris Software Recall
68970 | 203717 | PS | CNP | ENL | 2024-03-01 16:00 | 2024-08-16 15:38 | 2024 Alaris Software Recall
68971 | 203718 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 14:45 | 2024 Alaris Software Recall | Recall completed
68972 | 203719 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-02 09:43 | 2024 Alaris Software Recall
68973 | 203720 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-18 13:09 | 2024 Alaris Software Recall
68974 | 203721 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 13:17 | 2024 Alaris Software Recall | Completed
68975 | 203722 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:35 | 2024 Alaris Software Recall | Completed by BD.
68976 | 203723 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 11:42 | 2024 Alaris Software Recall | Completed by BD
68977 | 203724 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-23 12:19 | 2024 Alaris Software Recall | (TC 05/23/2024 12:18:45 PM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68978 | 203725 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-21 10:08 | 2024 Alaris Software Recall
68979 | 203726 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-22 11:29 | 2024 Alaris Software Recall | (TC 05/22/2024 11:28:54 AM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
68980 | 203727 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 15:00 | 2024 Alaris Software Recall | Completed by BD
68981 | 203728 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:36 | 2024 Alaris Software Recall | Completed by BD.
68982 | 203729 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 12:01 | 2024 Alaris Software Recall | Completed
68983 | 203730 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-13 10:58 | 2024 Alaris Software Recall
68984 | 203731 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 14:59 | 2024 Alaris Software Recall
68985 | 203732 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-05-21 14:10 | 2024 Alaris Software Recall | (TC 05/21/2024 02:09:55 PM, Baghirzade, Nara) Completed software upgrade. Completed the transfer of the network configuration and dataset files.
68986 | 203733 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-03 13:39 | 2024 Alaris Software Recall | (TC 04/03/2024 01:38:40 PM, Rhuland, Courtney) Recall completed by BD
68987 | 203734 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 14:57 | 2024 Alaris Software Recall | Completed by BD
68988 | 203735 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-12 11:30 | 2024 Alaris Software Recall | Completed by BD
68989 | 203736 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:38 | 2024 Alaris Software Recall | Completed by BD.
68990 | 203737 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 13:03 | 2024 Alaris Software Recall | Completed
68991 | 203738 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-14 15:10 | 2024 Alaris Software Recall
68992 | 203739 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-12 13:43 | 2024 Alaris Software Recall | Completed by BD
68993 | 203740 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-21 09:49 | 2024 Alaris Software Recall
68994 | 203741 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-17 08:49 | 2024 Alaris Software Recall | Completed by BD
68995 | 203742 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 14:31 | 2024 Alaris Software Recall | Completed
68996 | 203743 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-11 15:01 | 2024 Alaris Software Recall | Completed by BD
68997 | 203744 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-03-26 14:18 | 2024 Alaris Software Recall | Completed
68998 | 203745 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-04-17 10:49 | 2024 Alaris Software Recall | Completed by BD
68999 | 203746 | PS | CMPLT | CMPLT | 2024-03-01 16:00 | 2024-08-16 14:40 | 2024 Alaris Software Recall | Completed by BD.
69000 | 203747 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 14:06 | 2024 Alaris Software Recall | (TC 04/03/2024 02:06:06 PM, Rhuland, Courtney) Recall completed by BD
69001 | 203748 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-11 14:43 | 2024 Alaris Software Recall | Completed by BD
69002 | 203749 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 14:27 | 2024 Alaris Software Recall
69003 | 203750 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:40 | 2024 Alaris Software Recall | Completed by BD.
69004 | 203751 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-18 14:31 | 2024 Alaris Software Recall
69005 | 203752 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 11:29 | 2024 Alaris Software Recall | Completed by BD
69006 | 203753 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:43 | 2024 Alaris Software Recall | Completed by BD.
69007 | 203754 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 10:28 | 2024 Alaris Software Recall
69008 | 203755 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:45 | 2024 Alaris Software Recall | Completed by BD.
69009 | 203756 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-11 14:39 | 2024 Alaris Software Recall | Completed by BD
69010 | 203757 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 11:27 | 2024 Alaris Software Recall | Completed by BD
69011 | 203758 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-05 14:40 | 2024 Alaris Software Recall | (TC 04/05/2024 02:39:48 PM, Rhuland, Courtney) Recall completed by BD
69012 | 203759 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-17 08:42 | 2024 Alaris Software Recall | Completed by BD
69013 | 203760 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-11 14:53 | 2024 Alaris Software Recall | Completed by BD
69014 | 203761 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 14:38 | 2024 Alaris Software Recall
69015 | 203762 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 14:10 | 2024 Alaris Software Recall
69016 | 203763 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-05-23 12:15 | 2024 Alaris Software Recall | (TC 05/23/2024 12:14:40 PM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
69017 | 203764 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 13:46 | 2024 Alaris Software Recall | Completed by BD
69018 | 203765 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-05 14:41 | 2024 Alaris Software Recall | (TC 04/05/2024 02:40:41 PM, Rhuland, Courtney) Recall completed by BD
69019 | 203766 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 09:47 | 2024 Alaris Software Recall | Recall completed
69020 | 203767 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-20 14:46 | 2024 Alaris Software Recall
69021 | 203768 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-20 14:44 | 2024 Alaris Software Recall
69022 | 203769 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 09:23 | 2024 Alaris Software Recall | Recall completed
69023 | 203770 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:47 | 2024 Alaris Software Recall | Completed by BD.
69024 | 203771 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 09:14 | 2024 Alaris Software Recall | Recall completed
69025 | 203772 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-26 16:47 | 2024 Alaris Software Recall
69026 | 203773 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 11:33 | 2024 Alaris Software Recall | Completed by BD
69027 | 203774 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 14:10 | 2024 Alaris Software Recall | (TC 03/13/2024 02:09:46 PM, Baghirzade, Nara) Software recall completed by BD.
69028 | 203775 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-18 14:33 | 2024 Alaris Software Recall
69029 | 203776 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 13:48 | 2024 Alaris Software Recall | Completed by BD
69030 | 203777 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-26 11:58 | 2024 Alaris Software Recall | Completed
69032 | 203778 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-28 10:26 | 2024 Alaris Software Recall | Completed
69033 | 203779 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 11:52 | 2024 Alaris Software Recall
69034 | 203780 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 17:58 | 2024 Alaris Software Recall
69036 | 203781 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 20:53 | 2024 Alaris Software Recall
69037 | 203782 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 14:57 | 2024 Alaris Software Recall | Completed by BD
69038 | 203783 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 09:04 | 2024 Alaris Software Recall | Completed by BD
69039 | 203784 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 15:08 | 2024 Alaris Software Recall | Completed by BD
69040 | 203785 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-05-31 13:13 | 2024 Alaris Software Recall | (TC 05/31/2024 01:12:52 PM, Baghirzade, Nara) Software upgrade completed via ASM.
69041 | 203786 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-15 14:21 | 2024 Alaris Software Recall | Completed by BD
69042 | 203787 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:35 | 2024 Alaris Software Recall | Completed
69043 | 203788 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 15:03 | 2024 Alaris Software Recall
69044 | 203789 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-28 10:13 | 2024 Alaris Software Recall | Completed
69045 | 203790 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 21:54 | 2024 Alaris Software Recall
69046 | 203791 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 13:54 | 2024 Alaris Software Recall
69047 | 203792 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 11:49 | 2024 Alaris Software Recall
69048 | 203793 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:48 | 2024 Alaris Software Recall | Completed by BD.
69049 | 203794 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 09:40 | 2024 Alaris Software Recall
69050 | 203795 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:50 | 2024 Alaris Software Recall | Completed by BD.
69051 | 203796 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:53 | 2024 Alaris Software Recall | Completed
69052 | 203797 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-15 14:19 | 2024 Alaris Software Recall | Completed by BD
69053 | 203798 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-09 11:04 | 2024 Alaris Software Recall
69054 | 203799 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:51 | 2024 Alaris Software Recall | Completed
69055 | 203800 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 14:28 | 2024 Alaris Software Recall
69056 | 203801 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 09:02 | 2024 Alaris Software Recall | Completed by BD
69057 | 203802 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 09:57 | 2024 Alaris Software Recall | (TC 03/12/2024 09:56:25 AM, Rendon Lopez , Elizabeth) Recall completed
69058 | 203803 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 09:34 | 2024 Alaris Software Recall | Recall completed
69059 | 203804 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-01 10:26 | 2024 Alaris Software Recall | Software update completed.
69060 | 203805 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-18 14:01 | 2024 Alaris Software Recall
69061 | 203806 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 21:05 | 2024 Alaris Software Recall
69062 | 203807 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 15:01 | 2024 Alaris Software Recall | Completed by BD
69063 | 203808 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 13:28 | 2024 Alaris Software Recall
69064 | 203809 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 20:59 | 2024 Alaris Software Recall
69065 | 203810 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:51 | 2024 Alaris Software Recall | Completed by BD.
69066 | 203811 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-11 18:51 | 2024 Alaris Software Recall
69067 | 203812 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 15:06 | 2024 Alaris Software Recall
69068 | 203813 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:44 | 2024 Alaris Software Recall | Completed
69069 | 203814 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-26 09:46 | 2024 Alaris Software Recall
69070 | 203815 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 10:46 | 2024 Alaris Software Recall
69071 | 203816 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 14:07 | 2024 Alaris Software Recall
69072 | 203817 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 21:52 | 2024 Alaris Software Recall
69073 | 203818 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 18:07 | 2024 Alaris Software Recall
69074 | 203819 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-20 10:09 | 2024 Alaris Software Recall
69075 | 203820 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-17 09:59 | 2024 Alaris Software Recall | Completed by BD
69076 | 203821 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 09:36 | 2024 Alaris Software Recall
69077 | 203822 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 21:05 | 2024 Alaris Software Recall
69078 | 203823 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:56 | 2024 Alaris Software Recall | Completed
69079 | 203824 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 15:10 | 2024 Alaris Software Recall
69080 | 203825 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 13:56 | 2024 Alaris Software Recall
69081 | 203826 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 10:46 | 2024 Alaris Software Recall
69082 | 203827 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-24 09:47 | 2024 Alaris Software Recall
69083 | 203828 | PS | ACTIV | ISSUE | 2024-03-01 16:01 |  | 2024 Alaris Software Recall
69084 | 203829 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 14:24 | 2024 Alaris Software Recall
69085 | 203830 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-15 14:18 | 2024 Alaris Software Recall | Completed by BD
69086 | 203831 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-09 15:27 | 2024 Alaris Software Recall | Completed
69087 | 203832 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-28 09:57 | 2024 Alaris Software Recall | Completed
69088 | 203833 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 14:49 | 2024 Alaris Software Recall | Recall completed
69089 | 203834 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 14:13 | 2024 Alaris Software Recall | Completed by BD
69090 | 203835 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-30 13:40 | 2024 Alaris Software Recall
69091 | 203836 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 10:58 | 2024 Alaris Software Recall
69092 | 203837 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 11:22 | 2024 Alaris Software Recall
69093 | 203838 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-15 11:44 | 2024 Alaris Software Recall | Completed by BD
69094 | 203839 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 10:55 | 2024 Alaris Software Recall
69095 | 203840 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 13:38 | 2024 Alaris Software Recall
69096 | 203841 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 21:10 | 2024 Alaris Software Recall
69097 | 203842 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-18 14:11 | 2024 Alaris Software Recall
69098 | 203843 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 14:08 | 2024 Alaris Software Recall | (TC 03/06/2024 02:08:01 PM, Huang, Jessica) completed by BD
69099 | 203844 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-06-18 13:54 | 2024 Alaris Software Recall | (TC 06/18/2024 01:54:18 PM, Rhuland, Courtney) Software upgraded
69100 | 203845 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-10 09:14 | 2024 Alaris Software Recall | Completed
69101 | 203846 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 15:00 | 2024 Alaris Software Recall
69102 | 203847 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 20:50 | 2024 Alaris Software Recall
69103 | 203848 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-05-02 14:25 | 2024 Alaris Software Recall | (TC 05/02/2024 02:24:45 PM, Baghirzade, Nara) Software recall completed by BD.
69104 | 203849 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:53 | 2024 Alaris Software Recall | Completed by BD.
69105 | 203850 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-24 09:49 | 2024 Alaris Software Recall
69106 | 203851 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 09:58 | 2024 Alaris Software Recall | Recall completed
69107 | 203852 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 14:56 | 2024 Alaris Software Recall
69108 | 203853 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 11:39 | 2024 Alaris Software Recall
69109 | 203854 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 21:01 | 2024 Alaris Software Recall
69110 | 203855 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:59 | 2024 Alaris Software Recall | Completed by BD
69111 | 203856 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:40 | 2024 Alaris Software Recall
69112 | 203857 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-28 10:24 | 2024 Alaris Software Recall | Completed
69113 | 203858 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 07:49 | 2024 Alaris Software Recall | Completed by BD
69114 | 203859 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 10:48 | 2024 Alaris Software Recall
69115 | 203860 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:55 | 2024 Alaris Software Recall | Completed by BD.
69116 | 203861 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 18:19 | 2024 Alaris Software Recall
69117 | 203862 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-10 09:18 | 2024 Alaris Software Recall | Completed
69118 | 203863 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:20 | 2024 Alaris Software Recall | Completed
69119 | 203864 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-09 10:39 | 2024 Alaris Software Recall
69120 | 203865 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 15:38 | 2024 Alaris Software Recall | Recall completed
69121 | 203866 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 21:31 | 2024 Alaris Software Recall
69122 | 203867 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:40 | 2024 Alaris Software Recall
69123 | 203868 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:56 | 2024 Alaris Software Recall | Completed by BD.
69124 | 203869 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 10:39 | 2024 Alaris Software Recall
69125 | 203870 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:07 | 2024 Alaris Software Recall | Completed
69126 | 203871 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 18:12 | 2024 Alaris Software Recall
69127 | 203872 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-05-21 13:12 | 2024 Alaris Software Recall | (TC 05/21/2024 01:11:40 PM, Baghirzade, Nara) Software upgrade completed.
69128 | 203873 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-07-29 08:42 | 2024 Alaris Software Recall | (TC 07/29/2024 08:42:15 AM, Baghirzade, Nara) Software upgrade completed.
69129 | 203874 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-23 10:20 | 2024 Alaris Software Recall
69130 | 203875 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-28 09:52 | 2024 Alaris Software Recall | Completed
69131 | 203876 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:40 | 2024 Alaris Software Recall
69132 | 203877 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-23 10:19 | 2024 Alaris Software Recall
69133 | 203878 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 14:41 | 2024 Alaris Software Recall | (TC 03/08/2024 02:41:05 PM, Huang, Jessica) completed by BD
69134 | 203879 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 14:44 | 2024 Alaris Software Recall | (TC 04/03/2024 02:44:06 PM, Rhuland, Courtney) Recall completed by BD
69135 | 203880 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 18:35 | 2024 Alaris Software Recall
69136 | 203881 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 11:00 | 2024 Alaris Software Recall | Completed
69137 | 203882 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 14:26 | 2024 Alaris Software Recall
69138 | 203883 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 09:03 | 2024 Alaris Software Recall | Completed by BD
69139 | 203884 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 13:58 | 2024 Alaris Software Recall
69140 | 203885 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 21:36 | 2024 Alaris Software Recall
69141 | 203886 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 09:47 | 2024 Alaris Software Recall
69142 | 203887 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 16:20 | 2024 Alaris Software Recall
69143 | 203888 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:41 | 2024 Alaris Software Recall
69144 | 203889 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-23 10:26 | 2024 Alaris Software Recall
69145 | 203890 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:41 | 2024 Alaris Software Recall
69146 | 203891 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:52 | 2024 Alaris Software Recall | Completed by BD
69147 | 203892 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 16:30 | 2024 Alaris Software Recall
69148 | 203893 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 14:59 | 2024 Alaris Software Recall | Completed by BD.
69149 | 203894 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:41 | 2024 Alaris Software Recall
69150 | 203895 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:45 | 2024 Alaris Software Recall | Completed by BD
69151 | 203896 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-26 12:02 | 2024 Alaris Software Recall
69152 | 203897 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-23 10:39 | 2024 Alaris Software Recall
69153 | 203898 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 19:59 | 2024 Alaris Software Recall
69154 | 203899 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 14:13 | 2024 Alaris Software Recall | (TC 03/06/2024 02:13:30 PM, Huang, Jessica) completed by BD
69155 | 203900 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 13:44 | 2024 Alaris Software Recall | Completed by BD
69156 | 203901 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-09 12:27 | 2024 Alaris Software Recall | (TC 08/09/2024 12:26:59 PM, Patel, Rutvik) by BD
69157 | 203902 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:48 | 2024 Alaris Software Recall | Completed by BD
69158 | 203903 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:51 | 2024 Alaris Software Recall | Completed by BD
69159 | 203904 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 11:10 | 2024 Alaris Software Recall | Completed
69160 | 203905 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:50 | 2024 Alaris Software Recall | Completed by BD
69161 | 203906 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 09:48 | 2024 Alaris Software Recall | (TC 03/12/2024 09:47:49 AM, Rendon Lopez , Elizabeth) Recall completed
69162 | 203907 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 11:08 | 2024 Alaris Software Recall | Completed
69163 | 203908 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 15:00 | 2024 Alaris Software Recall | Completed by BD.
69164 | 203909 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-26 09:21 | 2024 Alaris Software Recall
69165 | 203910 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 14:44 | 2024 Alaris Software Recall
69166 | 203911 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-26 09:22 | 2024 Alaris Software Recall
69167 | 203912 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-13 14:25 | 2024 Alaris Software Recall | Completed by BD
69168 | 203913 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-15 11:49 | 2024 Alaris Software Recall | Completed by BD
69169 | 203914 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 20:18 | 2024 Alaris Software Recall
69170 | 203915 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 11:27 | 2024 Alaris Software Recall | Recall completed by BD
69171 | 203916 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-16 15:01 | 2024 Alaris Software Recall | Completed by BD.
69172 | 203917 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:23 | 2024 Alaris Software Recall | Completed
69173 | 203918 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 11:28 | 2024 Alaris Software Recall
69174 | 203919 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-15 11:47 | 2024 Alaris Software Recall | Completed by BD
69175 | 203920 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:56 | 2024 Alaris Software Recall | Completed by BD
69176 | 203921 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-11 16:52 | 2024 Alaris Software Recall
69177 | 203922 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 15:09 | 2024 Alaris Software Recall
69178 | 203923 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-21 20:36 | 2024 Alaris Software Recall
69179 | 203924 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-12 08:54 | 2024 Alaris Software Recall | Completed by BD
69180 | 203925 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 14:47 | 2024 Alaris Software Recall
69181 | 203926 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-11 16:56 | 2024 Alaris Software Recall
69182 | 203927 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-05-24 09:49 | 2024 Alaris Software Recall | (TC 05/24/2024 09:48:30 AM, Baghirzade, Nara) Completed software upgrade using ASM.
69183 | 203928 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 20:16 | 2024 Alaris Software Recall
69184 | 203929 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-28 10:01 | 2024 Alaris Software Recall | Completed
69185 | 203930 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:41 | 2024 Alaris Software Recall
69186 | 203931 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-19 20:12 | 2024 Alaris Software Recall
69187 | 203932 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-09 12:37 | 2024 Alaris Software Recall | (TC 08/09/2024 12:37:25 PM, Patel, Rutvik) by BD
69188 | 203933 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-25 11:32 | 2024 Alaris Software Recall
69189 | 203934 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-02 13:49 | 2024 Alaris Software Recall
69190 | 203935 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:42 | 2024 Alaris Software Recall
69191 | 203936 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 11:15 | 2024 Alaris Software Recall
69192 | 203937 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-28 10:32 | 2024 Alaris Software Recall | Completed
69193 | 203938 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 09:39 | 2024 Alaris Software Recall
69194 | 203939 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 09:24 | 2024 Alaris Software Recall | Recall completed
69195 | 203940 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:42 | 2024 Alaris Software Recall
69196 | 203941 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-26 09:44 | 2024 Alaris Software Recall
69197 | 203942 | PS | CNP | ENL | 2024-03-01 16:01 | 2024-08-16 15:42 | 2024 Alaris Software Recall
69198 | 203943 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 11:19 | 2024 Alaris Software Recall | Completed
69199 | 203944 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-12 09:32 | 2024 Alaris Software Recall
69200 | 203945 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-09 14:53 | 2024 Alaris Software Recall | by BD
69201 | 203946 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 13:42 | 2024 Alaris Software Recall | Recall completed
69202 | 203947 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-05-16 11:34 | 2024 Alaris Software Recall | (TC 05/16/2024 11:33:27 AM, Baghirzade, Nara) Software upgrade completed.
69203 | 203948 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-24 09:41 | 2024 Alaris Software Recall
69204 | 203949 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-05-16 11:32 | 2024 Alaris Software Recall | (TC 05/16/2024 11:31:22 AM, Baghirzade, Nara) Software upgrade completed.
69205 | 203950 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-27 10:48 | 2024 Alaris Software Recall | Completed
69206 | 203951 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-17 09:58 | 2024 Alaris Software Recall | Completed by BD
69207 | 203952 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-18 14:51 | 2024 Alaris Software Recall
69208 | 203953 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-15 13:42 | 2024 Alaris Software Recall
69209 | 203954 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 14:42 | 2024 Alaris Software Recall | (TC 04/03/2024 02:42:16 PM, Rhuland, Courtney) Recall completed by BD
69210 | 203955 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-03-14 14:30 | 2024 Alaris Software Recall | (TC 03/08/2024 02:29:41 PM, Huang, Jessica) completed by BD
69211 | 203956 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-04-03 14:52 | 2024 Alaris Software Recall | (TC 04/03/2024 02:51:29 PM, Rhuland, Courtney) Recall completed by BD
69212 | 203957 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-08-09 14:55 | 2024 Alaris Software Recall | By BD
69213 | 203958 | PS | CMPLT | CMPLT | 2024-03-01 16:01 | 2024-07-03 13:51 | 2024 Alaris Software Recall
69214 | 203959 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:12 | 2024 Alaris Software Recall | Completed by BD
69215 | 203960 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-02 13:50 | 2024 Alaris Software Recall
69216 | 203961 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:11 | 2024 Alaris Software Recall | Completed by BD
69217 | 203962 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-26 09:27 | 2024 Alaris Software Recall
69218 | 203963 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 11:55 | 2024 Alaris Software Recall | Completed by BD
69219 | 203964 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-24 09:39 | 2024 Alaris Software Recall
69220 | 203965 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:16 | 2024 Alaris Software Recall | Completed by BD
69221 | 203966 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 11:01 | 2024 Alaris Software Recall
69222 | 203967 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-19 21:45 | 2024 Alaris Software Recall
69223 | 203968 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-03 14:48 | 2024 Alaris Software Recall | (TC 04/03/2024 02:47:32 PM, Rhuland, Courtney) Recall completed by BD
69224 | 203969 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-05-24 09:46 | 2024 Alaris Software Recall | (TC 05/24/2024 09:46:17 AM, Baghirzade, Nara) Completed the software upgrade using ASM.
69225 | 203970 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-13 15:05 | 2024 Alaris Software Recall | Completed by BD
69226 | 203971 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 09:48 | 2024 Alaris Software Recall
69227 | 203972 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 10:41 | 2024 Alaris Software Recall | Completed
69228 | 203973 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-13 13:41 | 2024 Alaris Software Recall | Completed by BD
69229 | 203974 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 11:38 | 2024 Alaris Software Recall | (TC 08/09/2024 11:40:23 AM, Patel, Rutvik) by BD
69230 | 203975 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 10:29 | 2024 Alaris Software Recall | Completed
69231 | 203976 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-02 16:02 | 2024 Alaris Software Recall
69232 | 203977 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 11:11 | 2024 Alaris Software Recall | Completed
69233 | 203978 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-12 09:09 | 2024 Alaris Software Recall | Completed by BD
69234 | 203979 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-17 10:53 | 2024 Alaris Software Recall | Completed by BD
69235 | 203980 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-10 09:19 | 2024 Alaris Software Recall | Completed
69236 | 203981 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 11:10 | 2024 Alaris Software Recall | Completed
69237 | 203982 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 10:59 | 2024 Alaris Software Recall
69238 | 203983 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-12 09:12 | 2024 Alaris Software Recall | Completed by BD
69239 | 203984 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 10:57 | 2024 Alaris Software Recall
69240 | 203985 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 10:57 | 2024 Alaris Software Recall | Completed
69241 | 203986 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-19 11:27 | 2024 Alaris Software Recall | Completed by BD
69242 | 203987 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-20 10:34 | 2024 Alaris Software Recall
69243 | 203988 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 11:19 | 2024 Alaris Software Recall | Completed
69244 | 203989 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-24 14:17 | 2024 Alaris Software Recall
69246 | 203990 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 10:43 | 2024 Alaris Software Recall | Completed
69247 | 203991 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 14:36 | 2024 Alaris Software Recall
69248 | 203992 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-17 10:16 | 2024 Alaris Software Recall | Completed by BD
69249 | 203993 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-19 20:22 | 2024 Alaris Software Recall
69250 | 203994 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-26 09:25 | 2024 Alaris Software Recall
69251 | 203995 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-02 13:40 | 2024 Alaris Software Recall
69252 | 203996 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-22 12:29 | 2024 Alaris Software Recall
69253 | 203997 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 13:02 | 2024 Alaris Software Recall | By BD
69254 | 203998 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 10:53 | 2024 Alaris Software Recall
69255 | 203999 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 09:30 | 2024 Alaris Software Recall
69256 | 204000 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-26 09:42 | 2024 Alaris Software Recall
69258 | 204001 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-21 20:46 | 2024 Alaris Software Recall
69259 | 204002 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 13:32 | 2024 Alaris Software Recall
69260 | 204003 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 14:42 | 2024 Alaris Software Recall | (TC 03/08/2024 02:42:03 PM, Huang, Jessica) completed by BD
69261 | 204004 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 14:42 | 2024 Alaris Software Recall
69262 | 204005 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 16:26 | 2024 Alaris Software Recall
69263 | 204006 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:23 | 2024 Alaris Software Recall | Completed by BD
69264 | 204007 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-05-08 14:19 | 2024 Alaris Software Recall | Completed by BD.
69265 | 204008 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-26 12:12 | 2024 Alaris Software Recall
69266 | 204009 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 14:28 | 2024 Alaris Software Recall
69267 | 204010 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-02 14:39 | 2024 Alaris Software Recall
69268 | 204011 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-12 09:11 | 2024 Alaris Software Recall | Completed by BD
69269 | 204012 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-02 14:40 | 2024 Alaris Software Recall
69270 | 204013 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 10:30 | 2024 Alaris Software Recall | Recall completed
69271 | 204014 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-11 18:57 | 2024 Alaris Software Recall
69272 | 204015 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-04 12:54 | 2024 Alaris Software Recall | (TC 04/04/2024 12:54:06 PM, Rhuland, Courtney) Recall completed by BD
69273 | 204016 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-10 09:39 | 2024 Alaris Software Recall | Completed
69274 | 204017 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 16:08 | 2024 Alaris Software Recall
69275 | 204018 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 09:58 | 2024 Alaris Software Recall
69276 | 204019 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 10:33 | 2024 Alaris Software Recall | Recall completed
69277 | 204020 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-10 09:36 | 2024 Alaris Software Recall | Completed
69278 | 204021 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 16:19 | 2024 Alaris Software Recall
69279 | 204022 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-19 13:22 | 2024 Alaris Software Recall
69280 | 204023 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-10 09:34 | 2024 Alaris Software Recall | Completed
69281 | 204024 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 08:42 | 2024 Alaris Software Recall
69282 | 204025 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 14:35 | 2024 Alaris Software Recall | (TC 03/08/2024 02:34:52 PM, Huang, Jessica) completed by BD
69283 | 204026 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-10 09:38 | 2024 Alaris Software Recall | Completed
69284 | 204027 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-13 10:07 | 2024 Alaris Software Recall
69285 | 204028 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-19 13:27 | 2024 Alaris Software Recall
69286 | 204029 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 09:53 | 2024 Alaris Software Recall
69287 | 204030 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:26 | 2024 Alaris Software Recall | Completed by BD
69288 | 204031 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-24 14:18 | 2024 Alaris Software Recall
69289 | 204032 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-12 09:10 | 2024 Alaris Software Recall | Completed by BD
69290 | 204033 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 11:13 | 2024 Alaris Software Recall | Completed
69291 | 204034 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 11:08 | 2024 Alaris Software Recall | Completed
69292 | 204035 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-25 15:06 | 2024 Alaris Software Recall
69293 | 204036 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 10:33 | 2024 Alaris Software Recall | Completed
69294 | 204037 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-09 15:52 | 2024 Alaris Software Recall | Completed
69296 | 204038 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 09:29 | 2024 Alaris Software Recall
69297 | 204039 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-07-03 13:46 | 2024 Alaris Software Recall
69298 | 204040 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-13 13:13 | 2024 Alaris Software Recall | (TC 03/13/2024 01:12:25 PM, Baghirzade, Nara) Software recall completed by BD.
69299 | 204041 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-09 15:58 | 2024 Alaris Software Recall | Completed
69300 | 204042 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-09 15:37 | 2024 Alaris Software Recall | Completed
69301 | 204043 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-25 11:42 | 2024 Alaris Software Recall
69302 | 204044 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-03 14:31 | 2024 Alaris Software Recall | (TC 04/03/2024 02:30:36 PM, Rhuland, Courtney) Recall completed by BD
69303 | 204045 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 11:51 | 2024 Alaris Software Recall | Completed by BD
69304 | 204046 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 09:59 | 2024 Alaris Software Recall | Completed
69305 | 204047 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-07-19 13:19 | 2024 Alaris Software Recall | (TC 07/19/2024 01:18:44 PM, Baghirzade, Nara) Software upgrade completed.
69306 | 204048 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 10:10 | 2024 Alaris Software Recall | Completed
69307 | 204049 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 12:24 | 2024 Alaris Software Recall | (TC 08/09/2024 12:23:45 PM, Patel, Rutvik) by bd
69308 | 204050 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-21 10:45 | 2024 Alaris Software Recall
69309 | 204051 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-16 15:05 | 2024 Alaris Software Recall | Completed by BD.
69310 | 204052 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 11:53 | 2024 Alaris Software Recall | Completed by BD
69311 | 204053 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 10:22 | 2024 Alaris Software Recall | Recall completed
69312 | 204054 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-26 11:59 | 2024 Alaris Software Recall
69313 | 204055 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 12:53 | 2024 Alaris Software Recall | (TC 08/09/2024 12:53:12 PM, Patel, Rutvik) by BD
69314 | 204056 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-05-23 14:33 | 2024 Alaris Software Recall | (TC 05/23/2024 02:32:32 PM, Rendon Lopez , Elizabeth) Completed SW update
69315 | 204057 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 10:01 | 2024 Alaris Software Recall | Completed
69316 | 204058 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 12:58 | 2024 Alaris Software Recall | (TC 08/09/2024 12:58:27 PM, Patel, Rutvik) by BD
69317 | 204059 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 10:19 | 2024 Alaris Software Recall | Completed
69318 | 204060 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-06-17 11:39 | 2024 Alaris Software Recall | Recall completed by BD
69319 | 204061 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-03 11:18 | 2024 Alaris Software Recall | Completed
69320 | 204062 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 10:08 | 2024 Alaris Software Recall | Recall completed
69321 | 204063 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-18 15:03 | 2024 Alaris Software Recall
69322 | 204064 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-04 10:35 | 2024 Alaris Software Recall | (TC 03/04/2024 10:34:52 AM, Huang, Jessica) BD completed
69323 | 204065 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-26 09:08 | 2024 Alaris Software Recall
69324 | 204066 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-13 10:44 | 2024 Alaris Software Recall
69325 | 204067 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-03 14:37 | 2024 Alaris Software Recall | (TC 04/03/2024 02:36:26 PM, Rhuland, Courtney) Recall completed by BD
69326 | 204068 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-03 14:41 | 2024 Alaris Software Recall | (TC 04/03/2024 02:41:08 PM, Rhuland, Courtney) Recall completed by BD
69327 | 204069 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 10:57 | 2024 Alaris Software Recall | Recall completed
69328 | 204070 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 09:55 | 2024 Alaris Software Recall | (TC 03/12/2024 09:54:38 AM, Rendon Lopez , Elizabeth) Recall completed
69329 | 204071 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-26 09:40 | 2024 Alaris Software Recall
69330 | 204072 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:15 | 2024 Alaris Software Recall | Completed by BD
69331 | 204073 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-18 13:37 | 2024 Alaris Software Recall
69332 | 204074 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-12 08:57 | 2024 Alaris Software Recall | Completed by BD
69333 | 204075 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 12:55 | 2024 Alaris Software Recall | (TC 08/09/2024 12:55:27 PM, Patel, Rutvik) by BD
69334 | 204076 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 14:45 | 2024 Alaris Software Recall | by BD
69335 | 204077 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-19 20:27 | 2024 Alaris Software Recall
69336 | 204078 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-19 20:31 | 2024 Alaris Software Recall
69337 | 204079 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 14:53 | 2024 Alaris Software Recall
69338 | 204080 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 11:54 | 2024 Alaris Software Recall | Completed by BD
69339 | 204081 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-08-09 12:43 | 2024 Alaris Software Recall | (TC 08/09/2024 12:43:34 PM, Patel, Rutvik) by BD
69340 | 204082 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-11 16:47 | 2024 Alaris Software Recall
69341 | 204083 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-30 13:38 | 2024 Alaris Software Recall
69342 | 204084 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 10:49 | 2024 Alaris Software Recall
69343 | 204085 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-21 18:03 | 2024 Alaris Software Recall
69344 | 204086 | PS | CNP | ENL | 2024-03-01 16:02 | 2024-08-16 15:43 | 2024 Alaris Software Recall
69345 | 204087 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:27 | 2024 Alaris Software Recall | Completed by BD
69346 | 204088 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-12 14:14 | 2024 Alaris Software Recall
69347 | 204089 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-15 09:20 | 2024 Alaris Software Recall | Recall completed
69348 | 204090 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 13:50 | 2024 Alaris Software Recall | (TC 03/06/2024 01:50:03 PM, Huang, Jessica) completed by BD
69349 | 204091 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-28 11:06 | 2024 Alaris Software Recall | Completed
69350 | 204092 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-15 14:24 | 2024 Alaris Software Recall | Completed by BD
69351 | 204093 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-19 21:57 | 2024 Alaris Software Recall
69352 | 204094 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-13 09:49 | 2024 Alaris Software Recall
69353 | 204095 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-19 11:28 | 2024 Alaris Software Recall | Completed by BD
69354 | 204096 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-04 13:03 | 2024 Alaris Software Recall | (TC 04/04/2024 01:02:37 PM, Rhuland, Courtney) Recall completed by BD
69355 | 204097 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 11:09 | 2024 Alaris Software Recall | Completed
69356 | 204098 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-12 09:08 | 2024 Alaris Software Recall | Completed by BD
69357 | 204099 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-27 16:22 | 2024 Alaris Software Recall
69358 | 204100 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-14 14:24 | 2024 Alaris Software Recall | Recall completed
69359 | 204101 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-18 15:12 | 2024 Alaris Software Recall
69360 | 204102 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-04-04 13:01 | 2024 Alaris Software Recall | (TC 04/04/2024 01:00:16 PM, Rhuland, Courtney) Recall completed by BD
69361 | 204103 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-05-15 12:48 | 2024 Alaris Software Recall | (TC 05/15/2024 12:47:11 PM, Baghirzade, Nara) Software upgrade completed.
69362 | 204104 | PS | CMPLT | CMPLT | 2024-03-01 16:02 | 2024-03-22 13:11 | 2024 Alaris Software Recall
69363 | 204105 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-26 10:19 | 2024 Alaris Software Recall
69364 | 204106 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-16 15:07 | 2024 Alaris Software Recall | Completed by BD.
69365 | 204107 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 14:02 | 2024 Alaris Software Recall | (TC 03/06/2024 02:02:15 PM, Huang, Jessica) completed by BD
69366 | 204108 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-20 10:38 | 2024 Alaris Software Recall
69367 | 204109 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 13:23 | 2024 Alaris Software Recall | (TC 03/04/2024 01:23:33 PM, Huang, Jessica) Completed by BD
69368 | 204110 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 13:20 | 2024 Alaris Software Recall
69369 | 204111 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-12 09:05 | 2024 Alaris Software Recall | Completed by BD
69370 | 204112 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 20:50 | 2024 Alaris Software Recall
69371 | 204113 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 09:45 | 2024 Alaris Software Recall
69372 | 204114 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-12 09:01 | 2024 Alaris Software Recall | Completed by BD
69373 | 204115 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-09 13:01 | 2024 Alaris Software Recall | by BD
69374 | 204116 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 10:49 | 2024 Alaris Software Recall
69375 | 204117 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 13:09 | 2024 Alaris Software Recall | (TC 03/13/2024 01:08:27 PM, Baghirzade, Nara) Software recall completed by BD.
69376 | 204118 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 09:45 | 2024 Alaris Software Recall | (TC 03/12/2024 09:44:16 AM, Rendon Lopez , Elizabeth) Recall completed
69377 | 204119 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 20:04 | 2024 Alaris Software Recall
69378 | 204120 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 15:00 | 2024 Alaris Software Recall
69379 | 204121 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-05-15 12:44 | 2024 Alaris Software Recall | (TC 05/15/2024 12:43:43 PM, Baghirzade, Nara) Completed software recall.
69380 | 204122 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:32 | 2024 Alaris Software Recall | Completed
69381 | 204123 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 13:47 | 2024 Alaris Software Recall
69382 | 204124 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 20:49 | 2024 Alaris Software Recall
69383 | 204125 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:55 | 2024 Alaris Software Recall | Completed
69384 | 204126 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-07-17 10:36 | 2024 Alaris Software Recall | (TC 07/17/2024 10:35:46 AM, Rendon Lopez , Elizabeth) Recall completed
69385 | 204127 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 20:33 | 2024 Alaris Software Recall
69386 | 204128 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 11:35 | 2024 Alaris Software Recall
69387 | 204129 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 13:39 | 2024 Alaris Software Recall | Recall completed
69388 | 204130 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:41 | 2024 Alaris Software Recall | Completed
69389 | 204131 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:39 | 2024 Alaris Software Recall | Completed
69390 | 204132 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 14:50 | 2024 Alaris Software Recall
69391 | 204133 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:21 | 2024 Alaris Software Recall | Completed
69392 | 204134 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-18 13:46 | 2024 Alaris Software Recall
69393 | 204135 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-24 09:38 | 2024 Alaris Software Recall
69394 | 204136 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 14:29 | 2024 Alaris Software Recall | (TC 04/03/2024 02:29:02 PM, Rhuland, Courtney) Recall completed by BD
69395 | 204137 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 20:33 | 2024 Alaris Software Recall
69396 | 204138 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-22 13:36 | 2024 Alaris Software Recall | Recall completed by BD
69397 | 204139 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 14:38 | 2024 Alaris Software Recall | (TC 03/08/2024 02:38:35 PM, Huang, Jessica) completed by BD
69398 | 204140 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 10:24 | 2024 Alaris Software Recall | Recall completed
69399 | 204141 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 14:40 | 2024 Alaris Software Recall
69400 | 204142 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 18:26 | 2024 Alaris Software Recall
69401 | 204143 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-16 15:09 | 2024 Alaris Software Recall | Completed by BD.
69402 | 204144 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 14:47 | 2024 Alaris Software Recall | Completed by BD
69403 | 204145 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 14:31 | 2024 Alaris Software Recall
69404 | 204146 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 10:49 | 2024 Alaris Software Recall
69405 | 204147 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 15:58 | 2024 Alaris Software Recall
69406 | 204148 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 09:49 | 2024 Alaris Software Recall | Recall completed
69407 | 204149 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 13:43 | 2024 Alaris Software Recall
69408 | 204150 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 14:39 | 2024 Alaris Software Recall | (TC 04/03/2024 02:39:00 PM, Rhuland, Courtney) Recall completed by BD
69409 | 204151 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 11:38 | 2024 Alaris Software Recall
69410 | 204152 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:45 | 2024 Alaris Software Recall | Completed
69411 | 204153 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 13:31 | 2024 Alaris Software Recall
69412 | 204154 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-18 13:43 | 2024 Alaris Software Recall
69413 | 204155 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:43 | 2024 Alaris Software Recall | Completed
69414 | 204156 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-11 16:42 | 2024 Alaris Software Recall
69415 | 204157 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:35 | 2024 Alaris Software Recall | Completed
69416 | 204158 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 09:51 | 2024 Alaris Software Recall | (TC 03/12/2024 09:51:19 AM, Rendon Lopez , Elizabeth) Recall completed
69417 | 204159 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 20:42 | 2024 Alaris Software Recall
69418 | 204160 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-05-16 11:35 | 2024 Alaris Software Recall | (TC 05/16/2024 11:35:09 AM, Baghirzade, Nara) Software upgrade completed.
69419 | 204161 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-12 08:58 | 2024 Alaris Software Recall | Completed by BD
69420 | 204162 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 13:45 | 2024 Alaris Software Recall
69421 | 204163 | PS | CNP | ENL | 2024-03-01 16:03 | 2024-08-16 15:43 | 2024 Alaris Software Recall
69422 | 204164 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 13:41 | 2024 Alaris Software Recall
69423 | 204165 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 10:15 | 2024 Alaris Software Recall
69424 | 204166 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-15 11:52 | 2024 Alaris Software Recall | Completed by BD
69425 | 204167 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 15:51 | 2024 Alaris Software Recall | Recall completed
69426 | 204168 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 11:16 | 2024 Alaris Software Recall | Completed
69427 | 204169 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-06-06 10:09 | 2024 Alaris Software Recall | (TC 06/06/2024 10:08:48 AM, Rendon Lopez , Elizabeth) Software upgrade completed
69428 | 204170 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 20:43 | 2024 Alaris Software Recall
69429 | 204171 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 13:36 | 2024 Alaris Software Recall | Recall completed
69430 | 204172 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 15:45 | 2024 Alaris Software Recall | Recall completed
69431 | 204173 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 20:57 | 2024 Alaris Software Recall
69432 | 204174 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 11:02 | 2024 Alaris Software Recall | Completed
69433 | 204175 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 14:34 | 2024 Alaris Software Recall | (TC 03/08/2024 02:33:45 PM, Huang, Jessica) completed by BD
69434 | 204176 | PS | CNP | ENL | 2024-03-01 16:03 | 2024-08-16 15:44 | 2024 Alaris Software Recall
69435 | 204177 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 08:37 | 2024 Alaris Software Recall
69436 | 204178 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-15 14:22 | 2024 Alaris Software Recall | Completed by BD
69437 | 204179 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-28 10:58 | 2024 Alaris Software Recall | Completed
69438 | 204180 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 11:44 | 2024 Alaris Software Recall
69439 | 204181 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 13:33 | 2024 Alaris Software Recall
69440 | 204182 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 14:25 | 2024 Alaris Software Recall
69441 | 204183 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 13:43 | 2024 Alaris Software Recall
69442 | 204184 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-17 10:16 | 2024 Alaris Software Recall | Completed by BD
69443 | 204185 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-26 16:59 | 2024 Alaris Software Recall
69444 | 204186 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 10:20 | 2024 Alaris Software Recall | (TC 03/04/2024 10:21:24 AM, Huang, Jessica) BD completed
69445 | 204187 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 14:24 | 2024 Alaris Software Recall
69446 | 204188 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:17 | 2024 Alaris Software Recall | Completed
69447 | 204189 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-09 12:50 | 2024 Alaris Software Recall | (TC 08/09/2024 12:49:35 PM, Patel, Rutvik) by BD
69448 | 204190 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 09:53 | 2024 Alaris Software Recall | (TC 03/12/2024 09:53:02 AM, Rendon Lopez , Elizabeth) Recall completed
69449 | 204191 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-28 09:55 | 2024 Alaris Software Recall | Completed
69450 | 204192 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-16 15:12 | 2024 Alaris Software Recall | Completed by BD.
69451 | 204193 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-30 13:36 | 2024 Alaris Software Recall
69452 | 204194 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 11:12 | 2024 Alaris Software Recall
69453 | 204195 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-07-03 13:42 | 2024 Alaris Software Recall
69454 | 204196 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 10:54 | 2024 Alaris Software Recall
69455 | 204197 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 10:52 | 2024 Alaris Software Recall
69456 | 204198 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 18:09 | 2024 Alaris Software Recall
69457 | 204199 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-18 13:27 | 2024 Alaris Software Recall
69458 | 204200 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 10:59 | 2024 Alaris Software Recall
69459 | 204201 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-26 09:02 | 2024 Alaris Software Recall
69460 | 204202 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-18 13:31 | 2024 Alaris Software Recall
69461 | 204203 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-09 11:42 | 2024 Alaris Software Recall | (TC 08/09/2024 11:41:39 AM, Patel, Rutvik) by BD
69462 | 204204 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 11:02 | 2024 Alaris Software Recall | Completed
69463 | 204205 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 10:52 | 2024 Alaris Software Recall | Completed
69464 | 204206 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 18:31 | 2024 Alaris Software Recall
69465 | 204207 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:09 | 2024 Alaris Software Recall | Completed
69466 | 204208 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-15 11:40 | 2024 Alaris Software Recall | Completed by BD
69467 | 204209 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 10:05 | 2024 Alaris Software Recall
69468 | 204210 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 14:47 | 2024 Alaris Software Recall | Recall completed
69469 | 204211 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-23 10:25 | 2024 Alaris Software Recall
69470 | 204212 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 11:03 | 2024 Alaris Software Recall | Completed
69471 | 204213 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 10:56 | 2024 Alaris Software Recall
69472 | 204214 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-12 08:43 | 2024 Alaris Software Recall | Completed by BD
69473 | 204215 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-24 09:53 | 2024 Alaris Software Recall
69474 | 204216 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 14:29 | 2024 Alaris Software Recall
69475 | 204217 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-12 08:42 | 2024 Alaris Software Recall | Completed by BD
69476 | 204218 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-09 12:29 | 2024 Alaris Software Recall | (TC 08/09/2024 12:29:25 PM, Patel, Rutvik) by BD
69477 | 204219 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 10:39 | 2024 Alaris Software Recall
69478 | 204220 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-15 11:41 | 2024 Alaris Software Recall | Completed by BD
69479 | 204221 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 11:25 | 2024 Alaris Software Recall
69480 | 204222 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-05-16 11:33 | 2024 Alaris Software Recall | (TC 05/16/2024 11:32:57 AM, Baghirzade, Nara) Software upgrade completed.
69481 | 204223 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-15 11:12 | 2024 Alaris Software Recall | (TC 08/15/2024 11:11:46 AM, Rendon Lopez , Elizabeth) Recall completed.
69482 | 204224 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-11 16:45 | 2024 Alaris Software Recall
69483 | 204225 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 16:16 | 2024 Alaris Software Recall
69484 | 204226 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-16 15:13 | 2024 Alaris Software Recall | Completed by BD.
69485 | 204227 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-28 09:54 | 2024 Alaris Software Recall | Completed
69486 | 204228 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-16 15:16 | 2024 Alaris Software Recall | Completed by BD.
69487 | 204229 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-05-16 11:30 | 2024 Alaris Software Recall | (TC 05/16/2024 11:30:00 AM, Baghirzade, Nara) Software upgrade completed.
69488 | 204230 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-23 10:29 | 2024 Alaris Software Recall
69489 | 204231 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 11:15 | 2024 Alaris Software Recall
69490 | 204232 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 14:37 | 2024 Alaris Software Recall
69491 | 204233 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-23 10:22 | 2024 Alaris Software Recall
69492 | 204234 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 18:29 | 2024 Alaris Software Recall
69493 | 204235 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:25 | 2024 Alaris Software Recall | Completed
69495 | 204236 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-10 09:16 | 2024 Alaris Software Recall | Completed
69496 | 204237 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 14:56 | 2024 Alaris Software Recall | Recall completed
69497 | 204238 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-26 09:50 | 2024 Alaris Software Recall
69498 | 204239 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-10 09:13 | 2024 Alaris Software Recall | Completed
69499 | 204240 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-28 10:16 | 2024 Alaris Software Recall | Completed
69500 | 204241 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 10:59 | 2024 Alaris Software Recall | Recall completed
69501 | 204242 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-04 12:46 | 2024 Alaris Software Recall | (TC 04/04/2024 12:46:10 PM, Rhuland, Courtney) Recall completed by BD
69502 | 204243 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-20 10:03 | 2024 Alaris Software Recall
69503 | 204244 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:42 | 2024 Alaris Software Recall | Completed
69504 | 204245 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 09:34 | 2024 Alaris Software Recall
69505 | 204246 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 14:03 | 2024 Alaris Software Recall
69506 | 204247 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 13:59 | 2024 Alaris Software Recall
69507 | 204248 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 10:37 | 2024 Alaris Software Recall | (TC 03/04/2024 10:36:48 AM, Huang, Jessica) BD completed
69508 | 204249 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-05-02 14:15 | 2024 Alaris Software Recall | (TC 05/02/2024 02:14:30 PM, Baghirzade, Nara) Recall completed by BD.
69509 | 204250 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 09:32 | 2024 Alaris Software Recall
69510 | 204251 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 21:49 | 2024 Alaris Software Recall
69511 | 204252 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-21 10:52 | 2024 Alaris Software Recall
69512 | 204253 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-18 13:36 | 2024 Alaris Software Recall
69513 | 204254 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-30 13:58 | 2024 Alaris Software Recall
69514 | 204255 | PS | CNP | ENL | 2024-03-01 16:03 | 2024-08-16 15:44 | 2024 Alaris Software Recall
69515 | 204256 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-26 09:33 | 2024 Alaris Software Recall
69516 | 204257 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-04 12:44 | 2024 Alaris Software Recall | (TC 04/04/2024 12:43:52 PM, Rhuland, Courtney) Recall completed by BD
69517 | 204258 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 11:46 | 2024 Alaris Software Recall
69518 | 204259 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-25 13:17 | 2024 Alaris Software Recall
69519 | 204260 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-05-08 14:47 | 2024 Alaris Software Recall
69520 | 204261 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-12 10:43 | 2024 Alaris Software Recall
69521 | 204262 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-26 09:20 | 2024 Alaris Software Recall
69522 | 204263 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:30 | 2024 Alaris Software Recall | Completed
69523 | 204264 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 10:30 | 2024 Alaris Software Recall | (TC 03/04/2024 10:30:32 AM, Huang, Jessica) BD completed
69524 | 204265 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 16:23 | 2024 Alaris Software Recall
69525 | 204266 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 14:02 | 2024 Alaris Software Recall | Completed by BD
69526 | 204267 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 11:22 | 2024 Alaris Software Recall | Completed
69527 | 204268 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 16:08 | 2024 Alaris Software Recall
69528 | 204269 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-12 08:47 | 2024 Alaris Software Recall | Completed by BD
69529 | 204270 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-09 15:59 | 2024 Alaris Software Recall | Completed
69530 | 204271 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 13:34 | 2024 Alaris Software Recall | Recall completed
69531 | 204272 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-14 13:51 | 2024 Alaris Software Recall | Recall completed
69532 | 204273 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 14:05 | 2024 Alaris Software Recall
69533 | 204274 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-19 21:25 | 2024 Alaris Software Recall
69534 | 204275 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-28 10:19 | 2024 Alaris Software Recall | Completed
69535 | 204276 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-17 10:15 | 2024 Alaris Software Recall | Completed by BD
69536 | 204277 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-18 14:03 | 2024 Alaris Software Recall
69537 | 204278 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-12 09:07 | 2024 Alaris Software Recall | Completed by BD
69538 | 204279 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 13:53 | 2024 Alaris Software Recall
69539 | 204280 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-08-16 15:18 | 2024 Alaris Software Recall | Completed by BD.
69540 | 204281 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 09:37 | 2024 Alaris Software Recall
69541 | 204282 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-15 14:13 | 2024 Alaris Software Recall | Completed by BD
69542 | 204283 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 13:31 | 2024 Alaris Software Recall
69543 | 204284 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:39 | 2024 Alaris Software Recall | Completed
69544 | 204285 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-15 14:17 | 2024 Alaris Software Recall | Completed by BD
69545 | 204286 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:37 | 2024 Alaris Software Recall | Completed
69546 | 204287 | PS | CNP | ENL | 2024-03-01 16:03 | 2024-08-16 15:44 | 2024 Alaris Software Recall
69547 | 204288 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:36 | 2024 Alaris Software Recall | Completed
69548 | 204289 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-19 11:26 | 2024 Alaris Software Recall | Completed by BD
69549 | 204290 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-28 10:23 | 2024 Alaris Software Recall | Completed
69550 | 204291 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 10:59 | 2024 Alaris Software Recall
69551 | 204292 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-04 15:10 | 2024 Alaris Software Recall | Recall completed by BD
69552 | 204293 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 13:54 | 2024 Alaris Software Recall
69553 | 204294 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-13 14:35 | 2024 Alaris Software Recall | Completed by BD
69554 | 204295 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 11:14 | 2024 Alaris Software Recall | Completed
69555 | 204296 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-03 11:07 | 2024 Alaris Software Recall | Completed
69556 | 204297 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-27 10:14 | 2024 Alaris Software Recall | Completed
69557 | 204298 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-05-15 14:57 | 2024 Alaris Software Recall | (TC 05/15/2024 02:56:53 PM, Rendon Lopez , Elizabeth) Recall completed
69558 | 204299 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-04-02 14:17 | 2024 Alaris Software Recall
69559 | 204300 | PS | CMPLT | CMPLT | 2024-03-01 16:03 | 2024-03-15 10:06 | 2024 Alaris Software Recall | Recall completed
69560 | 204301 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:44 | 2024 Alaris Software Recall
69561 | 204302 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-30 13:37 | 2024 Alaris Software Recall
69562 | 204303 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-21 20:28 | 2024 Alaris Software Recall
69563 | 204304 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-25 11:30 | 2024 Alaris Software Recall
69564 | 204305 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 20:07 | 2024 Alaris Software Recall
69565 | 204306 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 10:16 | 2024 Alaris Software Recall | Completed
69566 | 204307 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-15 13:27 | 2024 Alaris Software Recall
69567 | 204308 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-15 13:31 | 2024 Alaris Software Recall
69568 | 204309 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-23 10:30 | 2024 Alaris Software Recall | Recall completed by BD
69569 | 204310 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-25 11:27 | 2024 Alaris Software Recall
69570 | 204311 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-03 11:12 | 2024 Alaris Software Recall | Completed
69571 | 204312 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-09 15:29 | 2024 Alaris Software Recall | Completed
69572 | 204313 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 20:09 | 2024 Alaris Software Recall
69573 | 204314 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 15:15 | 2024 Alaris Software Recall | Recall completed
69574 | 204315 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 11:46 | 2024 Alaris Software Recall | Completed by BD
69575 | 204316 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:45 | 2024 Alaris Software Recall
69576 | 204317 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 20:14 | 2024 Alaris Software Recall
69577 | 204318 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-12 08:45 | 2024 Alaris Software Recall | Completed by BD
69578 | 204319 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-19 11:24 | 2024 Alaris Software Recall | Completed by BD
69579 | 204320 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-25 11:20 | 2024 Alaris Software Recall
69580 | 204321 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 10:12 | 2024 Alaris Software Recall | Completed
69581 | 204322 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 07:20 | 2024 Alaris Software Recall | (TC 03/14/2024 07:19:59 AM, Rhuland, Courtney) Recall completed by BD
69582 | 204323 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 20:01 | 2024 Alaris Software Recall
69583 | 204324 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-26 09:18 | 2024 Alaris Software Recall
69584 | 204325 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 10:23 | 2024 Alaris Software Recall | (TC 03/06/2024 02:00:00 PM, Huang, Jessica) completed by BD
69585 | 204326 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-08-16 15:19 | 2024 Alaris Software Recall | Completed by BD.
69586 | 204327 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-24 09:37 | 2024 Alaris Software Recall
69587 | 204328 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 14:19 | 2024 Alaris Software Recall
69588 | 204329 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 21:38 | 2024 Alaris Software Recall
69589 | 204330 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-12 09:43 | 2024 Alaris Software Recall
69590 | 204331 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-12 10:44 | 2024 Alaris Software Recall | (TC 03/12/2024 10:43:41 AM, Rendon Lopez , Elizabeth) Recall completed
69591 | 204332 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-24 09:43 | 2024 Alaris Software Recall
69592 | 204333 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-12 09:44 | 2024 Alaris Software Recall
69593 | 204334 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-18 13:58 | 2024 Alaris Software Recall
69594 | 204335 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-12 10:37 | 2024 Alaris Software Recall
69595 | 204336 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 14:10 | 2024 Alaris Software Recall
69596 | 204337 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-11 18:46 | 2024 Alaris Software Recall
69597 | 204338 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 14:01 | 2024 Alaris Software Recall
69598 | 204339 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-28 10:11 | 2024 Alaris Software Recall | Completed
69599 | 204340 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 10:27 | 2024 Alaris Software Recall | Completed
69600 | 204341 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-21 17:57 | 2024 Alaris Software Recall
69601 | 204342 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-28 10:17 | 2024 Alaris Software Recall | Completed
69602 | 204343 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-15 09:30 | 2024 Alaris Software Recall | Recall completed
69603 | 204344 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-24 09:34 | 2024 Alaris Software Recall | Recall completed by BD
69604 | 204345 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-07-22 11:54 | 2024 Alaris Software Recall | (TC 07/22/2024 11:53:51 AM, Rendon Lopez , Elizabeth) Recall completed
69605 | 204346 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 16:28 | 2024 Alaris Software Recall
69606 | 204347 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-13 14:33 | 2024 Alaris Software Recall | Completed by BD
69607 | 204348 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-12 08:53 | 2024 Alaris Software Recall | Completed by BD
69608 | 204349 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:45 | 2024 Alaris Software Recall
69609 | 204350 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 15:13 | 2024 Alaris Software Recall | Recall completed
69610 | 204351 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-25 11:17 | 2024 Alaris Software Recall
69611 | 204352 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 11:45 | 2024 Alaris Software Recall | Completed by BD
69612 | 204353 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 14:05 | 2024 Alaris Software Recall | (TC 03/06/2024 02:05:08 PM, Huang, Jessica) completed by BD
69613 | 204354 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-15 09:54 | 2024 Alaris Software Recall | Recall completed
69614 | 204355 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-12 10:00 | 2024 Alaris Software Recall
69615 | 204356 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 11:50 | 2024 Alaris Software Recall | Completed by BD
69616 | 204357 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-12 09:50 | 2024 Alaris Software Recall | (TC 03/12/2024 09:49:29 AM, Rendon Lopez , Elizabeth) Recall completed
69617 | 204358 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 16:32 | 2024 Alaris Software Recall
69618 | 204359 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-17 09:58 | 2024 Alaris Software Recall | Completed by BD
69619 | 204360 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-07-03 13:48 | 2024 Alaris Software Recall
69620 | 204361 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-22 10:46 | 2024 Alaris Software Recall
69621 | 204362 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-17 10:00 | 2024 Alaris Software Recall | Completed by BD
69622 | 204363 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-08-16 15:21 | 2024 Alaris Software Recall | Completed by BD.
69623 | 204364 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-21 21:03 | 2024 Alaris Software Recall
69624 | 204365 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-24 09:47 | 2024 Alaris Software Recall
69625 | 204366 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-12 10:41 | 2024 Alaris Software Recall
69626 | 204367 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:45 | 2024 Alaris Software Recall
69627 | 204368 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-26 09:52 | 2024 Alaris Software Recall
69628 | 204369 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-04 12:49 | 2024 Alaris Software Recall | (TC 04/04/2024 12:48:50 PM, Rhuland, Courtney) Recall completed by BD
69629 | 204370 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-15 13:30 | 2024 Alaris Software Recall | Recall completed
69630 | 204371 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 11:42 | 2024 Alaris Software Recall | Completed by BD
69631 | 204372 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:46 | 2024 Alaris Software Recall
69632 | 204373 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-21 20:56 | 2024 Alaris Software Recall
69633 | 204374 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 14:14 | 2024 Alaris Software Recall | Completed by BD
69634 | 204375 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-12 10:57 | 2024 Alaris Software Recall
69635 | 204376 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-28 11:13 | 2024 Alaris Software Recall | Completed
69636 | 204377 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-28 10:55 | 2024 Alaris Software Recall | Completed
69637 | 204378 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-07-30 12:08 | 2024 Alaris Software Recall | (TC 07/30/2024 12:08:06 PM, Baghirzade, Nara) Software upgrade completed.
69638 | 204379 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-30 13:34 | 2024 Alaris Software Recall
69639 | 204380 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 14:35 | 2024 Alaris Software Recall
69640 | 204381 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 14:31 | 2024 Alaris Software Recall
69641 | 204382 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-10 09:22 | 2024 Alaris Software Recall | Completed
69642 | 204383 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-28 11:11 | 2024 Alaris Software Recall | Completed
69643 | 204384 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 13:47 | 2024 Alaris Software Recall
69644 | 204385 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 14:37 | 2024 Alaris Software Recall
69645 | 204386 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:46 | 2024 Alaris Software Recall
69646 | 204387 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 16:03 | 2024 Alaris Software Recall
69647 | 204388 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-04 14:57 | 2024 Alaris Software Recall | Recall completed by BD
69648 | 204389 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 14:59 | 2024 Alaris Software Recall
69649 | 204390 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-28 10:21 | 2024 Alaris Software Recall | Completed
69650 | 204391 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 07:50 | 2024 Alaris Software Recall | Completed by BD
69651 | 204392 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-12 09:15 | 2024 Alaris Software Recall | Completed by BD
69652 | 204393 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 11:12 | 2024 Alaris Software Recall | Completed
69653 | 204394 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-10 09:06 | 2024 Alaris Software Recall | Completed
69654 | 204395 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 15:53 | 2024 Alaris Software Recall
69655 | 204396 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 21:13 | 2024 Alaris Software Recall
69656 | 204397 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 13:57 | 2024 Alaris Software Recall | Recall completed
69657 | 204398 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:46 | 2024 Alaris Software Recall
69658 | 204399 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-21 18:01 | 2024 Alaris Software Recall
69659 | 204400 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 11:39 | 2024 Alaris Software Recall | Completed by BD
69660 | 204401 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-15 14:33 | 2024 Alaris Software Recall
69661 | 204402 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-04 12:58 | 2024 Alaris Software Recall | (TC 04/04/2024 12:57:41 PM, Rhuland, Courtney) Recall completed by BD
69662 | 204403 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 14:25 | 2024 Alaris Software Recall | Completed by BD
69663 | 204404 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-04 12:50 | 2024 Alaris Software Recall | (TC 04/04/2024 12:50:10 PM, Rhuland, Courtney) Recall completed by BD
69664 | 204405 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 13:18 | 2024 Alaris Software Recall
69665 | 204406 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-21 18:08 | 2024 Alaris Software Recall
69666 | 204407 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 15:49 | 2024 Alaris Software Recall
69667 | 204408 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:47 | 2024 Alaris Software Recall
69668 | 204409 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-08-09 14:56 | 2024 Alaris Software Recall | By BD
69669 | 204410 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-17 10:52 | 2024 Alaris Software Recall | Completed by BD
69670 | 204411 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 11:17 | 2024 Alaris Software Recall | Completed
69671 | 204412 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-02 14:43 | 2024 Alaris Software Recall
69672 | 204413 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-08-16 15:23 | 2024 Alaris Software Recall | Completed by BD.
69673 | 204414 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-15 13:47 | 2024 Alaris Software Recall | Recall completed
69674 | 204415 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-08-16 15:24 | 2024 Alaris Software Recall | Completed by BD.
69675 | 204416 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-14 14:05 | 2024 Alaris Software Recall | Recall completed
69676 | 204417 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 10:59 | 2024 Alaris Software Recall | Completed
69677 | 204418 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-08-16 15:26 | 2024 Alaris Software Recall | Completed by BD.
69678 | 204419 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-19 21:19 | 2024 Alaris Software Recall
69679 | 204420 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-04-15 14:20 | 2024 Alaris Software Recall | Completed by BD
69680 | 204421 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 11:15 | 2024 Alaris Software Recall | Completed
69681 | 204422 | PS | CNP | ENL | 2024-03-01 16:04 | 2024-08-16 15:47 | 2024 Alaris Software Recall
69682 | 204423 | PS | CMPLT | CMPLT | 2024-03-01 16:04 | 2024-03-27 11:04 | 2024 Alaris Software Recall | Completed
| 204424 | RDS | CMPLT | CMPLT | 2024-03-04 08:11 | 2024-03-04 15:22 | Weekly check for ED
| 204425 | RDS | CMPLT | CMPLT | 2024-03-04 08:12 | 2024-03-22 15:10 | Weekly check for ICU
| 204426 | RDS | CMPLT | CMPLT | 2024-03-04 08:17 | 2024-03-04 08:17 | weekly check for 14S and 17M | weekly check for 14S and 17M
1054310 | 204427 | CM | CMPLT | CMPLT | 2024-03-04 08:30 | 2024-03-04 09:01 | Reading 1.5C lower than other devices | (TC 03/04/2024 08:59:57 AM, Canjimavadivel, Kamallesh ) Passed calibration after 3rd attempt
| 204428 | RDS | CMPLT | CMPLT | 2024-03-04 09:28 | 2024-03-04 09:30 | Weekly checks in CCU | (TC 03/04/2024 09:28:25 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed. Room 12 had a broken O2 flowmeter that was causing an O2 leak. Removed the broken one and replaced it with a working flowmeter.
1055573 | 204429 | CM | CMPLT | CMPLT | 2024-03-04 09:45 | 2024-03-04 09:45 | replacement unit from masimo | updated SN on TMS. put Sinai config and wireless setup. Replaced temp probe. tested. functional.
1055598 | 204430 | CM | CMPLT | CMPLT | 2024-03-04 09:48 | 2024-03-04 09:48 | Masimo replacement unit | updated SN on TMS. put Sinai config. tested. functional.
66713 | 204431 | CM | CMPLT | CMPLT | 2024-03-04 10:30 | 2024-03-04 10:31 | Broken scale plate | (TC 03/04/2024 10:30:43 AM, Rhuland, Courtney) Replaced broken plexiglass plate Calibrated RH Functional check - OK
1050474 | 204432 | CM | CMPLT | CMPLT | 2024-03-04 11:30 | 2024-03-04 11:31 | Paper door jammed | (TC 03/04/2024 11:29:37 AM, Rhuland, Courtney) Printer paper tray would not open Appeared to be jammed in the back Pulled tray out and reseated tray Issue resolved
| 204433 | PA | CMPLT | CMPLT | 2024-03-04 12:12 | 2024-04-08 09:14 | PIC+ patient interface cables for BIS | (TC 03/04/2024 12:17:08 PM, Scero, Tomlyn) placed order on ereq for new cables (TC 04/08/2024 09:13:32 AM, Scero, Tomlyn) Product received
1055580 | 204434 | CM | CMPLT | CMPLT | 2024-03-04 13:38 | 2024-03-04 13:38 | replacement unit | updated TMS with new serial#. put Sinai config. tested. Functional.
| 204435 | CM | CMPLT | CMPLT | 2024-03-04 13:40 | 2024-03-04 13:40 | biomed meeting | biomed meeting
1047852 | 204436 | PM | CMPLT | CMPLT | 2024-03-04 13:49 | 2024-03-04 14:14 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 03/04/2024 02:14:04 PM, Rhuland, Courtney) PM complete All tests passed Refer to procedure for details
1061210 | 204437 | CM | CMPLT | CMPLT | 2024-03-04 14:21 | 2024-03-04 14:22 | Monitor 3 not working | (TC 03/04/2024 02:21:32 PM, Reid, Emmerson ) Monitor 3 was not working, went up with Rutvik. It was on the wrong input, working now.
1055733 | 204438 | PM | CMPLT | CMPLT | 2024-03-04 14:34 | 2024-03-07 09:47 | PM MAJOR - Annual PM for: Microstream Extension | (TC 03/07/2024 09:45:21 AM, Scero, Tomlyn) PM complete. Calibration passed.
63073 | 204439 | CM | CMPLT | CMPLT | 2024-03-04 15:07 | 2024-03-04 15:14 | Created by Failed PM WO # 188911 : PM-MAJOR - Annual PM For: Monitors, Physiologic, Patient Transport | (TC 03/04/2024 03:12:10 PM, Chukri, Samir) Repaired and open the monitor check the display cable and calibrated the touch screen. reinstalled into the transport incubator. | (TC 03/04/2024 03:06:01 PM, Chukri, Samir) the screen in mirror setting
1070144 | 204440 | CM | CMPLT | CMPLT | 2024-03-05 09:15 | 2024-03-04 15:20 | CO2 Pressure/Hose issue | Call re: not working. Went to check. Was informed that CO2 tank emptied out itself. Turned unit on. Found low pressure indication. Removed regulator. Found green gasket missing due to which CO2 leaked overnight and CO2 tank emptied out. Replaced CO2 tank. Anita recommended to reroute the CO2 hose inside the back panel. Did that. She requested again to find how much CO2 remains when low pressure indication is on. Contacted sales rep. Could not get a definite answer. Again was requested to remove the unit and replaced it with olumpus. Removed unit and place it on shelf as requested.
1053883 | 204441 | CM | CMPLT | CMPLT | 2024-03-05 09:22 | 2024-03-04 13:25 | Installed on Emegency Cart | Reuested by Anita to install the unit back on emerg cart in Endo. Removed Steris unit and replaced this on emerge cart. When connected to CO2 tank, found regulator was leaking. Made multiple adjustment but did not work. Tried different tank. Issue persisted. Found O ring in CO2 regulator was worn out. Replaced it. Ensured secured connection. Checked functionality. Tested ok.
| 204442 | CM | CMPLT | CMPLT | 2024-03-05 09:33 | 2024-03-04 14:33 | Biomed Weekly Huddle | Biomed Weekly Huddle
| 204443 | RDS | CMPLT | CMPLT | 2024-03-05 09:36 | 2024-03-05 09:39 | Weekly Checks For PACU | (TC 03/05/2024 09:38:32 AM, Reid, Emmerson ) Rounds complete
| 204444 | RDS | CMPLT | CMPLT | 2024-03-05 09:39 | 2024-03-05 09:43 | Weekly Checks For OR | (TC 03/05/2024 09:41:36 AM, Reid, Emmerson ) Rounds complete, Also repaired a neumatic tool connector for OR2
1054379 | 204445 | CM | CMPLT | CMPLT | 2024-03-05 10:19 | 2024-03-05 10:46 | Having an issue sending data | (TC 03/05/2024 10:19:35 AM, Reid, Emmerson ) Checked the network settings and everything was correct, also tested connection and it passed. Sending the unit back.
1055579 | 204446 | CM | CMPLT | CMPLT | 2024-03-05 10:37 | 2024-03-05 10:37 | replacement unit from Masimo | updated SN on TMS. put Sinai config and wireless setup. Tested. functional.
| 204447 | CM | CMPLT | CMPLT | 2024-03-05 10:38 | 2024-03-05 10:38 | meeting | meeting
| 204448 | CM | CMPLT | CMPLT | 2024-03-05 10:39 | 2024-03-05 10:39 | worked with RT | worked with RT
1063791 | 204449 | CM | CMPLT | CMPLT | 2024-03-05 11:02 | 2024-05-14 08:37 | Fan not working, overheating | (TC 03/05/2024 11:02:16 AM, Reid, Emmerson ) Tested and found that after ~5mins the light shut down because it was overheating due to the fan not working. Reached out to the vendor for a warranty repair. (TC 03/14/2024 09:10:03 AM, Reid, Emmerson ) Another headlight was sent down. I was able to swap battery packs and make one working unit out of the two. Sending this unit out with KJV (TC 05/09/2024 01:45:04 PM, Reid, Emmerson ) Put in request for NO CHARGE PO under SC 1003227968 (TC 05/14/2024 08:34:49 AM, Reid, Emmerson ) Came back under PO 7500206395.
| 204450 | CM | CMPLT | CMPLT | 2024-03-05 11:23 | 2024-04-17 12:49 | Pump parts for stock | (TC 03/05/2024 11:31:57 AM, Rendon Lopez , Elizabeth) Ordering parts for repairs during recall (TC 03/12/2024 02:10:50 PM, Rendon Lopez , Elizabeth) Received front cases and right IUI (TC 04/17/2024 12:46:14 PM, Rendon Lopez , Elizabeth) Received all parts
1055953 | 204451 | CM | CMPLT | CMPLT | 2024-03-05 12:54 | 2024-03-05 12:54 | replacement unit from Masimo | updated SN on TMS. put Sinai config and wireless setup. Replaced temp probe and SpO2 sensor. tested. functional.
| 204452 | CM | CMPLT | CMPLT | 2024-03-05 15:01 | 2024-03-05 15:18 | Power Bar Purchase | Purchase of Power bar for Endo and Biomed SC #1003183286
| 204453 | CM | CMPLT | CMPLT | 2024-03-06 09:21 | 2024-04-04 10:16 | Philips Order - NBP hose and inlet | (TC 04/04/2024 10:12:39 AM, Samanian, Kamran) All parts received PO# 7202376342
64225 | 204454 | CM | CMPLT | CMPLT | 2024-03-06 09:53 | 2024-03-06 09:53 | SW Upgrade |  | SW Upgraded to N.00.05 by Philips FSE
64226 | 204455 | CM | CMPLT | CMPLT | 2024-03-06 09:53 | 2024-03-06 09:53 | SW Upgrade |  | SW Upgraded to N.00.05 by Philips FSE
64227 | 204456 | CM | CMPLT | CMPLT | 2024-03-06 09:53 | 2024-03-06 09:53 | SW Upgrade |  | SW Upgraded to N.00.05 by Philips FSE
64228 | 204457 | CM | CMPLT | CMPLT | 2024-03-06 09:53 | 2024-03-06 09:53 | SW Upgrade |  | SW Upgraded to N.00.05 by Philips FSE
64229 | 204458 | CM | CMPLT | CMPLT | 2024-03-06 09:53 | 2024-03-06 09:53 | SW Upgrade |  | SW Upgraded to N.00.05 by Philips FSE
64223 | 204459 | CM | CMPLT | CMPLT | 2024-03-06 09:53 | 2024-03-06 09:53 | SW Upgrade |  | SW Upgraded to N.00.05 by Philips FSE
1057522 | 204460 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1057535 | 204461 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1057004 | 204462 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1057532 | 204463 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1057528 | 204464 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1057536 | 204465 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1057006 | 204466 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1057534 | 204467 | CM | CMPLT | CMPLT | 2024-03-06 10:09 | 2024-03-06 10:09 | SW Upgraded by Philips FSE |  | Philips FSE upgraded SW to N.00.06 and pushed CCU configuration
1061203 | 204468 | CM | CMPLT | CMPLT | 2024-03-06 11:03 | 2024-03-07 10:56 | Speakers not working | (TC 03/06/2024 11:02:47 AM, Reid, Emmerson ) Speakers were not working. Checked the audio settings and cables, both looked good. Reset VMC and working now.
69494 | 204469 | PM | CMPLT | PMFAIL | 2024-03-06 13:07 | 2024-03-14 15:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69496 | 204470 | PM | CMPLT | PMFAIL | 2024-03-06 13:09 | 2024-03-14 14:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1053348 | 204471 | CM | CMPLT | CMPLT | 2024-03-06 13:16 | 2024-03-19 10:39 | MSH BI 1272 H3 GANTRY ROTATING COLLISION WITHPT BIOPSY CHAIR | (TC 03/19/2024 10:38:23 AM, Whittington, Scott) MSH BI 1272 H3 GANTRY ROTATING COLLISION WITHPT BIOPSY CHAIR FOLLOWUP | MSH BI 1272 H3 GANTRY ROTATING COLLISION WITHPT BIOPSY CHAIR
66061 | 204472 | PM | CMPLT | PMFAIL | 2024-03-06 13:37 | 2024-03-14 14:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1057522 | 204473 | PM | CMPLT | CMPLT | 2024-03-06 14:11 | 2024-03-06 14:13 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/06/2024 02:11:52 PM, Scero, Tomlyn) PM complete. Battery replaced.
1057535 | 204474 | PM | CMPLT | CMPLT | 2024-03-06 14:16 | 2024-03-06 14:17 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/06/2024 02:16:02 PM, Scero, Tomlyn) PM complete. Battery replaced.
1057004 | 204475 | PM | CMPLT | CMPLT | 2024-03-06 14:29 | 2024-03-06 14:30 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 03/06/2024 02:28:42 PM, Scero, Tomlyn) PM complete. Battery replaced.
1057532 | 204476 | PM | CMPLT | CMPLT | 2024-03-06 14:36 | 2024-03-06 14:42 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/06/2024 02:40:34 PM, Scero, Tomlyn) PM complete and battery replaced
1057528 | 204477 | PM | CMPLT | CMPLT | 2024-03-06 14:55 | 2024-03-06 14:56 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/06/2024 02:55:28 PM, Scero, Tomlyn) Pm complete and battery replaced
| 204478 | CM | CMPLT | CMPLT | 2024-03-06 15:34 | 2024-03-06 15:36 | Endo Cart Power Bar install | Found Endoscopy utilizing residential power bar to charge their endo emerg video cart. Spoke with Samir and acquired a new power bar. Installed it in Endoscopy Hallway on the wooden wall panel using double sided tape. Tested ok.
1053982 | 204479 | CM | CMPLT | CMPLT | 2024-03-06 15:37 | 2024-03-06 15:39 | Monitor/Endo cart testing in OR | Request from Anita to test the Endo emerge cart in OR. Worked with ER to schedule the testing. Brought the endo cart to OR; Tested monitor functionality. Tested ok. Delivered back to Endoscopy.
| 204480 | CM | CMPLT | CMPLT | 2024-03-06 15:40 | 2024-03-06 15:41 | Meeting | Meeting with AC and SC Discussion wirth Anita Chin on endo ER cart Discussion on IV pump remediation
1056911 | 204481 | CM | CMPLT | CMPLT | 2024-03-06 15:42 | 2024-03-06 15:43 | Recording Issue | Call from OR re: not able to access previous Microscope video recordings. Went to check. Waited for patient move out. Assisted nurse in moving microscope out of the OR. Checked fucntionlity. Browsed through menu and accessed previous video recording. Taught nurse and physician how to access saved recordings. Tested ok.
1053348 | 204482 | CM | CMPLT | CMPLT | 2024-03-06 20:19 | 2024-03-06 20:22 | MSH BI 1272 H3 GANTRY ROTATED AND STRUCK CHAIR, SYSTEM WOULD NOT MOVE | (TC 03/06/2024 08:18:51 PM, Whittington, Scott) MSH BI 1272 H3 GANTRY ROTATED AND STRUCK CHAIR, SYSTEM WOULD NOT MOVE FUSE F8 12 AMP MDL BLOWN NO LOCAL STOCK RYAN LAROUQUE CAME TO SITE WITH LAST OF HIS STOCK INSTALLED OK ASKED HOLOGIC FOR SPARE FUSE STOCK CONCERNED THAT ONLY WAY TO STOP A CRUSH BY GANTRY IS BLOWING FUSE (TC 03/06/2024 08:22:18 PM, Whittington, Scott) MSH BI 1272 H3 GANTRY ROTATED AND STRUCK CHAIR, SYSTEM WOULD NOT MOVE FUSE F8 12 AMP MDL BLOWN NO LOCAL STOCK RYAN LAROUQUE CAME TO SITE WITH LAST OF HIS STOCK INSTALLED OK ASKED HOLOGIC FOR SPARE FUSE STOCK CONCERNED THAT ONLY WAY TO STOP A CRUSH BY GANTRY IS BLOWING FUSE (TC 03/07/2024 02:20:00 PM, Whittington, Scott) MSH BI 1272 H3 GANTRY ROTATED AND STRUCK CHAIR, SYSTEM WOULD NOT MOVE FUSE F8 12 AMP MDL BLOWN NO LOCAL STOCK RYAN LAROUQUE CAME TO SITE WITH LAST OF HIS STOCK INSTALLED OK ASKED HOLOGIC FOR SPARE FUSE STOCK CONCERNED THAT ONLY WAY TO STOP A CRUSH BY GANTRY IS BLOWING FUSE FOLLOWUP (TC 03/06/2024 08:18:51 PM, Whittington, Scott) MSH BI 1272 H3 GANTRY ROTATED AND STRUCK CHAIR, SYSTEM WOULD NOT MOVE FUSE F8 12 AMP MDL BLOWN NO LOCAL STOCK RYAN LAROUQUE CAME TO SITE WITH LAST OF HIS STOCK INSTALLED OK ASKED HOLOGIC FOR SPARE FUSE STOCK CONCERNED THAT ONLY WAY TO STOP A CRUSH BY GANTRY IS BLOWING FUSE | MSH BI 1272 H3 GANTRY ROTATED AND STRUCK CHAIR, SYSTEM WOULD NOT MOVE
| 204483 | CM | CMPLT | CMPLT | 2024-03-07 07:42 | 2024-03-06 07:41 | meeting | meeting
| 204484 | CM | CMPLT | CMPLT | 2024-03-07 07:43 | 2024-03-06 07:43 | worked with RT | worked with RT
1057536 | 204485 | PM | CMPLT | CMPLT | 2024-03-07 09:09 | 2024-03-07 09:11 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/07/2024 09:09:41 AM, Scero, Tomlyn) PM complete and battery replaced.
1054523 | 204486 | CM | CMPLT | CMPLT | 2024-03-07 09:15 | 2024-03-07 09:30 | Not working | (TC 03/07/2024 09:25:46 AM, Canjimavadivel, Kamallesh ) Testing with snap connectors showed intermittent signal. Tested without and working normal. Replaced snap connectors. Also replaced keyboard cover as it was ripped.
1057006 | 204487 | PM | CMPLT | CMPLT | 2024-03-07 09:22 | 2024-03-07 09:34 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 03/07/2024 09:32:38 AM, Scero, Tomlyn) PM complete and battery replaced.
1057534 | 204488 | PM | CMPLT | CMPLT | 2024-03-07 09:36 | 2024-03-07 09:38 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/07/2024 09:36:56 AM, Scero, Tomlyn) PM complete and battery replaced
| 204489 | CM | CMPLT | CMPLT | 2024-03-07 10:04 | 2024-03-07 10:07 | Hug a Vac Repair | Call re: not working. Went to check and brought it down to biomed. Consulted emmerson to understand the functionality. Arranged for the tubing and suction. Checked functionlity in SC's office. Tested ok. Tried to 3 more time to ensure there is no leakage. Tested ok. Returned to OR.
1055734 | 204490 | PM | CMPLT | CMPLT | 2024-03-07 10:16 | 2024-03-07 10:17 | PM ACTIVE-MICROSTREAM EXTENSION | (TC 03/07/2024 10:16:27 AM, Scero, Tomlyn) PM complete. Calibration done
1064418 | 204491 | CM | CMPLT | CMPLT | 2024-03-07 10:18 | 2024-04-30 08:47 | Port Doors falling off | (TC 03/07/2024 10:18:46 AM, Ferreira, Qwynn) replaced two doors on the MRI will need one more as all were tearing off. (TC 04/30/2024 08:46:19 AM, Ferreira, Qwynn) replaced third door
1055735 | 204492 | PM | CMPLT | CMPLT | 2024-03-07 10:20 | 2024-03-07 10:21 | PM ACTIVE-MICROSTREAM EXTENSION | (TC 03/07/2024 10:20:07 AM, Scero, Tomlyn) Pm complete and Calibration done
1055736 | 204493 | PM | CMPLT | CMPLT | 2024-03-07 10:30 | 2024-03-07 10:31 | PM ACTIVE-MICROSTREAM EXTENSION | (TC 03/07/2024 10:30:35 AM, Scero, Tomlyn) PM complete and calibration done
1048530 | 204494 | CM | CMPLT | CMPLT | 2024-03-07 10:33 | 2024-03-07 10:34 | User reported channel error | (TC 03/07/2024 10:33:08 AM, Baghirzade, Nara) Did not find any errors connecting to the PCU. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1048531 | 204495 | CM | CMPLT | CMPLT | 2024-03-07 10:34 | 2024-03-07 10:34 | User reported channel error | (TC 03/07/2024 10:33:56 AM, Baghirzade, Nara) Did not find any errors connecting to the PCU. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65877 | 204496 | CM | CMPLT | CMPLT | 2024-03-07 10:34 | 2024-03-07 10:35 | User reported channel error | (TC 03/07/2024 10:34:28 AM, Baghirzade, Nara) Did not find any errors connecting to the PCU. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65951 | 204497 | CM | CMPLT | CMPLT | 2024-03-07 10:35 | 2024-03-07 10:38 | Broken rear and front case, left handle, and defective pressure sensor | (TC 03/07/2024 10:35:11 AM, Baghirzade, Nara) Replaced: -front case -rear case -left handle -mylar gasket -pressure sensor disc Performed full calibration for: -barrel clamp -plunger force -plunger position -pressure sensor disc Check-in tests passed and completed. IUI connectors and device cleaned. Reprinted and replaced asset tag and SN barcode on front case. Device is functional.
1057283 | 204498 | CM | CMPLT | CMPLT | 2024-03-07 10:39 | 2024-03-07 10:40 | No MSH configuration found on the device | (TC 03/07/2024 10:39:20 AM, Baghirzade, Nara) Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config.
1056436 | 204499 | CM | CMPLT | CMPLT | 2024-03-07 10:40 | 2024-03-07 10:41 | User reported that SpO2 was not working | (TC 03/07/2024 10:40:16 AM, Baghirzade, Nara) Red LNC-10 cable was defective; replaced cable. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
63872 | 204500 | CM | CMPLT | CMPLT | 2024-03-07 10:42 | 2024-03-07 10:43 | Tilt mechanism was crooked, not locking in | (TC 03/07/2024 10:42:31 AM, Baghirzade, Nara) Re-seated tilt mechanism; latching/locking is functional. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
| 204501 | MEP | CMPLT | CMPLT | 2024-03-07 10:57 | 2024-04-15 10:48 | BD Software Remediation Project | (TC 03/07/2024 10:57:50 AM, Reid, Emmerson ) Helping to clean IUI connectors and get pumps ready for remediation. Had many sent back and had to replace/ repair them. Met with Adeel and team to discuss how to handle this situation, and we came up with a revised pan to replace IUI infront of BD with every round of pumps delivered. (TC 03/07/2024 02:16:00 PM, Reid, Emmerson ) Gathered two PCs and LVPs for PACU ahead of their switch tomorrow. Communicated with them that I dropped them off at comm station 2 (TC 03/18/2024 11:15:30 AM, Reid, Emmerson ) Did a pump run and repaired two pumps, both latches. (TC 03/19/2024 10:38:24 AM, Reid, Emmerson ) Did a run and repaired some pumps (TC 04/08/2024 10:43:31 AM, Reid, Emmerson ) Did a run, replaced IUIs and brought down some pumps for repair. (TC 04/09/2024 02:57:43 PM, Reid, Emmerson ) Cleaned up the pump remediation space on 14S
66693 | 204502 | CM | CMPLT | CMPLT | 2024-03-07 12:18 | 2024-03-07 12:19 | Broken porthole door | (TC 03/07/2024 12:18:26 PM, Baghirzade, Nara) Replaced porthole door. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.4%RH Incubator is fully functional.
65992 | 204503 | CM | CMPLT | CMPLT | 2024-03-07 13:08 | 2024-03-07 13:09 | User reported channel error | (TC 03/07/2024 01:08:29 PM, Baghirzade, Nara) No issues connecting to PCU; no channel error observed. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
| 204504 | PA | ACTIV | ISSUE | 2024-03-07 13:47 |  | Ordering parts for stock | (TC 03/07/2024 01:52:43 PM, Baghirzade, Nara) SC 1003185256 (TC 03/19/2024 02:28:42 PM, Baghirzade, Nara) Received 3/4 ordered items. Did not receive power cord straps. PO 7202373773 (TC 03/25/2024 10:25:24 AM, Baghirzade, Nara) Received 15/25 power cord straps. PO 7202373773 (TC 06/03/2024 07:16:03 AM, Baghirzade, Nara) Received 2 power cord straps, total 17/25 power cord straps received from the order. PO 7202373773
65919 | 204505 | CM | CMPLT | CMPLT | 2024-03-07 14:14 | 2024-03-08 14:08 | Installing plunger assembly/calibration/functional check | (TC 03/07/2024 02:14:10 PM, Baghirzade, Nara) Installed new plunger assembly. Successfully completed full calibration of: -barrel clamp -plunger force -plunger position -pressure sensor disc During check-in tests, syringe failed keypad test; none of the keys were being registered. Will inspect front case and display board. (TC 03/08/2024 02:07:38 PM, Baghirzade, Nara) Replaced front case. Keypad is functional. Refer to WO 204518 Gave pump to BD for software upgrade.
65994 | 204506 | CM | CMPLT | CMPLT | 2024-03-07 14:34 | 2024-03-07 14:38 | Broken rear case, left handle, and latch | (TC 03/07/2024 02:34:06 PM, Baghirzade, Nara) Replaced rear case, left handle, and latch.
1054774 | 204507 | CM | CMPLT | CMPLT | 2024-03-07 14:44 | 2024-03-07 14:45 | Defective temperature probe | (TC 03/07/2024 02:43:56 PM, Baghirzade, Nara) Replaced temperature probe and well. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1070213 | 204508 | ACC | CMPLT | CMPLT | 2024-03-07 14:56 | 2024-03-11 11:17 | MSH 5 US PROBE PVT781 SN# FSD2422078 ACC | (TC 03/11/2024 11:17:17 AM, Whittington, Scott) MSH 5 US PROBE PVT781 SN# FSD2422078 ACC | MSH 5 US PROBE PVT781 SN# FSD2422078 ACC
| 204509 | MEP | CMPLT | CMPLT | 2024-03-07 15:22 | 2024-03-14 11:26 | BD Alaris Remediation | Worked on BD remediation. Inspected and cleaned IV pumps. Replaced IUI connectors and other parts and delivered IV pumps to BD.
| 204510 | CM | CMPLT | CMPLT | 2024-03-07 16:17 | 2024-03-07 16:24 | unable to calibrate gas for CPET | (TC 03/07/2024 04:17:32 PM, Cortez, John) "reference gas out of date" Reached out to Antony from Novus. He wasn't familiar with error Tried manually changing date under certificates, this worked Original date didn't even match date on gas tank. Antony mentioned certificate ID isnt relevant/important
1053381 | 204511 | CM | CMPLT | CMPLT | 2024-03-08 08:22 | 2024-03-07 08:21 | expiratory cassette membrane | replaced expiratory cassette membrane. tested. functional. tested 2 x expiratory cassettes with error. passed tests. informed Rex.
66141 | 204512 | CM | CMPLT | CMPLT | 2024-03-08 08:25 | 2024-03-07 08:25 | broken arm holder | replaced holder. tested. functional. tested 2 x heater cables with error. did not pass test. removed from service. informed Rex.
| 204513 | CM | CMPLT | CMPLT | 2024-03-08 08:28 | 2024-03-07 08:27 | meeting | meeting
66692 | 204514 | PA | CMPLT | CMPLT | 2024-03-08 10:18 | 2024-03-19 14:39 | Ordering hood | (TC 03/08/2024 10:20:54 AM, Baghirzade, Nara) SC 1003185885 (TC 03/19/2024 02:38:17 PM, Baghirzade, Nara) PO 7202375524
68580 | 204515 | CM | CMPLT | CMPLT | 2024-03-08 13:59 | 2024-03-08 14:03 | Broken rear case and left handle | (TC 03/08/2024 01:59:22 PM, Baghirzade, Nara) Replaced rear case. Replaced left handle.
68597 | 204516 | CM | CMPLT | CMPLT | 2024-03-08 14:03 | 2024-03-08 14:03 | Broken rear case and left handle | (TC 03/08/2024 02:03:06 PM, Baghirzade, Nara) Replaced rear case. Replaced left handle.
68600 | 204517 | CM | CMPLT | CMPLT | 2024-03-08 14:04 | 2024-03-08 14:05 | Broken front case | (TC 03/08/2024 02:03:43 PM, Baghirzade, Nara) Replaced front case. Replaced mylar gasket.
65919 | 204518 | CM | CMPLT | CMPLT | 2024-03-08 14:06 | 2024-03-08 14:07 | Broken front case | (TC 03/08/2024 02:05:39 PM, Baghirzade, Nara) Replaced front case. Replaced mylar gasket. Gave pump back to BD for software upgrade.
| 204519 | CM | CMPLT | CMPLT | 2024-03-08 14:10 | 2024-03-08 14:10 | Working with Kamran to configure and connect INVOS and Philips monitor with ICM+ software | (TC 03/08/2024 02:09:49 PM, Baghirzade, Nara) Working with Kamran to configure and connect INVOS and Philips monitor with ICM+ software for Dr. Deshpande's study.
| 204520 | RDS | CMPLT | CMPLT | 2024-03-08 14:13 | 2024-03-08 14:15 | Weekly Rounds: 11 North | (TC 03/08/2024 02:13:17 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 204521 | RDS | CMPLT | CMPLT | 2024-03-08 14:15 | 2024-03-08 14:16 | Weekly Rounds: 11 South | (TC 03/08/2024 02:15:20 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 204522 | RDS | CMPLT | CMPLT | 2024-03-08 14:17 | 2024-03-08 14:19 | Weekly Rounds: NICU & RESUS | (TC 03/08/2024 02:16:43 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
1063797 | 204523 | CM | CMPLT | CMPLT | 2024-03-08 14:38 | 2024-03-08 14:38 | Intermittent power. | (TC 03/08/2024 02:37:42 PM, Green, Paul) From: Green, Paul <Paul.Green@uhn.ca> Sent: March 8, 2024 10:32 AM To: Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca>; Paz, Ramon <Ramon.Paz@uhn.ca>; Gushie, John - Sinai Health System <John.Gushie@sinaihealth.ca> Subject: Re: [External] RE: Outlet / Vital Moniter Slave unit Hi Jason, We checked out the monitor this morning and power seems to be ok, cords seem ok etc. If the MRT’s experience this again, check red light on the receptacle as John mentioned. This will show if the power to the plug is on or off. Let us know. Regards, Paul Paul Green, CET Senior Technologist Medical Engineering, Imaging Service Group – UHN, MSH & WCH Toronto General Hospital, TGH-1-NU160 200 Elizabeth Street, Toronto ON, M5G 2C4 Imaging Service Support: 437-688-9325 Cell: (416) 919-7245 Email: paul.green@uhn.ca ________________________________________ From: Gushie, John <John.Gushie@sinaihealth.ca> Sent: Friday, March 8, 2024 7:00:00 AM To: Green, Paul <Paul.Green@uhn.ca>; Fisico, Jason - Sinai Health System <Jason.Fisico@sinaihealth.ca> Subject: Re: [External] RE: Outlet / Vital Moniter Slave unit Correct, the unit is plugged in beside the control station for the 1.5T on the 5th Floor. John ________________________________________ From: Green, Paul <Paul.Green@uhn.ca> Sent: Thursday, March 7, 2024 4:31 PM To: Gushie, John; Fisico, Jason Subject: Re: [External] RE: Outlet / Vital Moniter Slave unit CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Ok, will have a look tomorrow morning. This is in the 1.5T control room? Paul Green, CET Senior Technologist Medical Engineering, Imaging Service Group – UHN, MSH & WCH Toronto General Hospital, TGH-1-NU160 200 Elizabeth Street, Toronto ON, M5G 2C4 Imaging Service Support: 437-688-9325 Cell: (416) 919-7245 Email: paul.green@uhn.ca ________________________________________ From: Fisico, Jason <Jason.Fisico@sinaihealth.ca> Sent: Thursday, March 7, 2024 1:40:00 PM To: Gushie, John - Sinai Health System <John.Gushie@sinaihealth.ca> Cc: Green, Paul <Paul.Green@uhn.ca> Subject: [External] RE: Outlet / Vital Moniter Slave unit Thanks John. Paul would you be able to take a look at the Iradimed base unit at MSH when you have a chance? It seems to be losing power intermittently, but turns back on if unplugged and plugged back in. Thanks, Jason Fisico M.MedRadSc(MRI), MRT(N)(MR) MRI Supervisor Medical Imaging Research Co-Lead JDMI Practice Council Co-Chair jason.fisico@SinaiHealth.ca T 416-586-4800 x1591 F 416-586-8321 From: Gushie, John Sent: March 07, 2024 1:23 PM To: Fisico, Jason <Jason.Fisico@sinaihealth.ca> Subject: Outlet / Vital Moniter Slave unit I was just looking at the issue with maintenance and he pointed out a couple of things that make us think its the unit, not the outlet. 1) Alarm clock is running off the same plug, but it doesn't lose power when the moniter is off. 2) If outlet breaker went out, it would require maintenance to reset the breaker. He couldn't really trouble shoot the outlet as it's currently working normally. We are guessing it's an issue with the monitor. Moving forward the next time the slave unit loses power I would be curious if the power light on the outlet is off. For your consideration, John
502089 | 204524 | CM | CMPLT | CMPLT | 2024-03-08 14:39 | 2024-03-08 14:43 | persistant error | (TC 03/08/2024 02:39:35 PM, Green, Paul) Customer installed PH collimator and a persistant error on TUB1 LVA. Was able to clear and home the system.
1061384 | 204525 | CM | CMPLT | CMPLT | 2024-03-08 14:52 | 2024-03-08 14:55 | Display Glitching | Call re: not working. Went to check. Found display artifacts on the external Boom monitors. Checked connection. Could not find any issue. Attempted to switch connection port. Issue persisted. Rebooted the system. checked functionlity. Tested ok.
| 204526 | CM | ACTIV | ISSUE | 2024-03-08 14:57 |  | CMMS Committee work | Worked on CMMS committee meeting agenda coming up on Mar 12. (TC 03/12/2024 04:00:48 PM, Patel, Rutvik) Worked on Agenda and TMS Improvements and conducted CMMS committee meeting (TC 03/18/2024 03:58:23 PM, Patel, Rutvik) Researched and worked on TMS improvement list. (TC 03/21/2024 03:17:41 PM, Patel, Rutvik) Compiled minutes of last meeting Worked on Agenda items for next meering Researched and created TMS improvement list (TC 03/25/2024 03:24:50 PM, Patel, Rutvik) CMMS Committee meeting Worked on TMS clean up and improvement (TC 03/28/2024 02:31:55 PM, Patel, Rutvik) CMMS Committee meeting and Worked on TMS improvement (TC 05/09/2024 11:23:10 AM, Patel, Rutvik) Conducted Meeting. Recorded notes and minutes. (TC 05/15/2024 03:47:20 PM, Patel, Rutvik) Conducted Meeting. Recorded notes and minutes. (TC 05/22/2024 03:25:58 PM, Patel, Rutvik) CMMS Committee mneeting (TC 06/03/2024 03:12:07 PM, Patel, Rutvik) CMMS Committee meeting (TC 06/10/2024 03:54:49 PM, Patel, Rutvik) Conducted CMMS committee meeting (TC 06/17/2024 03:27:48 PM, Patel, Rutvik) CMMS Committee Meeting (TC 07/08/2024 02:47:42 PM, Patel, Rutvik) CMMS Committee meeting (TC 07/15/2024 03:29:15 PM, Patel, Rutvik) CMMS Committee meeting (TC 07/22/2024 03:22:28 PM, Patel, Rutvik) CMMS Committee meeting (TC 07/24/2024 08:38:03 AM, Patel, Rutvik) Created and Updated TMS category list for SHS-3 segment (TC 07/24/2024 12:38:00 PM, Patel, Rutvik) Created and Updated TMS category list for SHS-3 segment (TC 07/25/2024 03:56:26 PM, Patel, Rutvik) Started working on subcategories in TMS. (TC 07/29/2024 01:10:51 PM, Patel, Rutvik) worked on updating TMS Subcategory (TC 07/29/2024 08:11:35 AM, Patel, Rutvik) CMMS Committee meeting (TC 07/30/2024 03:24:00 PM, Patel, Rutvik) Updated Subcategory list (TC 07/31/2024 03:43:33 PM, Patel, Rutvik) Completed updating SHS cat/subcat Analyzed Philips asset data.. found missing cat/subcat. Communicated with committee. (TC 08/06/2024 03:48:56 PM, Patel, Rutvik) Data analysis in TMS. (TC 08/12/2024 03:56:58 PM, Patel, Rutvik) CMMS Committee meeting. (TC 08/13/2024 03:42:55 PM, Patel, Rutvik) Analyzed assets and completed cat/subcat list as discussed in previous CMMS meeting. (TC 08/19/2024 03:45:18 PM, Patel, Rutvik) CMMS Committee meeting (TC 08/20/2024 03:46:06 PM, Patel, Rutvik) Data analysis for vendor and model creation Mass edited assets for CMAC laryngoscope as previous assets were under incorrect vendor, cat and subcat. (TC 09/23/2024 09:05:38 AM, Patel, Rutvik) CMMS Committee meeting (TC 09/25/2024 03:27:41 PM, Patel, Rutvik) Created Guidance doc on adding vendor and model in TMS. (TC 09/30/2024 03:49:00 PM, Patel, Rutvik) CMMS Committee meeting (TC 10/07/2024 03:27:00 PM, Patel, Rutvik) Conducted CMMS committee meeting. (TC 10/28/2024 04:07:29 PM, Patel, Rutvik) CMMS committee meeting (TC 11/18/2024 04:04:37 PM, Patel, Rutvik) CMMS Committee meeting Vendor clean up work (TC 11/25/2024 03:41:52 PM, Patel, Rutvik) CMMS Committee meeting and vendor/model cleanup
| 204527 | CM | CMPLT | CMPLT | 2024-03-08 15:12 | 2024-04-04 09:08 | TAAAC Work | Apr 03 - Started developing the plan for Biomed test equipment and medical device training. Apr 04 - Worked on training plan. Contacted MECTA for ECT parts and accessories.
| 204528 | CM | ACTIV | ISSUE | 2024-03-08 15:13 |  | Miscellaneous / Clerical / Departmental Work | (TC 03/08/2024 03:12:57 PM, Patel, Rutvik) Email communication and correspondances Discarded boxes and packing material. (TC 03/12/2024 10:40:42 AM, Patel, Rutvik) Ordered stuff from Uline. (TC 03/13/2024 03:48:22 PM, Patel, Rutvik) Elearn courses Uline item returns inquiry (TC 03/14/2024 03:21:58 PM, Patel, Rutvik) Received ULine shipment. Unload and Unboxed everything. Arranged common material into cabinets. Assembled stool. (TC 03/15/2024 02:36:46 PM, Patel, Rutvik) - Received and Unboxed desk drawer. Took measurements under the table. Marked and drilled holes. And mounted drawer on the work bench. - Unboxed and arranged bins on parts shelf. - CMBES Registration - Cut and discarded boxes and material (TC 03/21/2024 03:16:09 PM, Patel, Rutvik) G&T account Setup (TC 03/22/2024 03:16:30 PM, Patel, Rutvik) Email communcations and correspondances Desk cleaning and organization (TC 03/26/2024 03:50:54 PM, Patel, Rutvik) Email communications and correspondances Endo steris equipment inventory and sorting. (TC 04/02/2024 03:39:57 PM, Patel, Rutvik) Email communicatuion and correspondances (TC 04/03/2024 03:40:22 PM, Patel, Rutvik) - PO Sorting - Relocated boxes and Organized Philips parts on shelve - Distributed delivered boxes to staff to make space in the delivery area - Desk organization (TC 04/04/2024 03:47:16 PM, Patel, Rutvik) Boxes/Garbage disposal. Email communication and correspondances (TC 04/05/2024 03:32:45 PM, Patel, Rutvik) - PO Sorting. Removed and organized label printer Tapes on shelve - Distributed delivered boxes to staff to make space in the delivery area -Email communcation and correspondances (TC 04/08/2024 03:11:58 PM, Patel, Rutvik) Email communication and correspondances Assisted KS with OR equipment images (TC 04/11/2024 04:07:22 PM, Patel, Rutvik) Again, Assisted KS with OR equipment images. (TC 04/12/2024 03:35:41 PM, Patel, Rutvik) Email communication and correspondances Discarded boxes and packing material. Assisted KS with OR equipment images. Desk clean up and organization. Explored vendors for Microscope display screen protectors (TC 04/16/2024 04:10:59 PM, Patel, Rutvik) Parts and inventoy organization Repaired and repalced banana clips for ESU testing Watched you tube webinar on FT10 PM. (TC 04/17/2024 03:45:20 PM, Patel, Rutvik) Helped out John C with moving bed from storage to ICU; ICU to 18th floor and mattress from 18th floor to ICU. (TC 04/25/2024 03:53:10 PM, Patel, Rutvik) Email communication and correspondances Discarded boxes and packing material Looked into ECG machine arrangements for 12 N (TC 04/29/2024 03:45:06 PM, Patel, Rutvik) Signed up and started Covidien FT10 online self paced training (TC 04/30/2024 02:56:57 PM, Patel, Rutvik) Foud Paper shredder bin full. Disposed, cleaned and replaced it. (TC 05/01/2024 02:58:31 PM, Patel, Rutvik) Email communcations and correspondances Cut and Disposed card board boxes Went to 522 Uni ave to pick up benefit coverage card. (TC 05/03/2024 03:48:52 PM, Patel, Rutvik) Email communcations and correspondances Cut and Disposed card board boxes Sorted Organized masimo inventory On call documentation review for my first time on call. (TC 05/07/2024 02:31:41 PM, Patel, Rutvik) Email communications and correspondances 12 N move preparation Help desk issue with date and time (TC 05/30/2024 03:26:19 PM, Patel, Rutvik) Worked to record stationary devices in OR. Completed OR 17,18,19. (TC 05/31/2024 03:34:17 PM, Patel, Rutvik) Email communication and correspondances. Cleaned and organized desk and inventory (TC 06/07/2024 03:20:47 PM, Patel, Rutvik) Email communications and correspondances Microwave/lunch area clean up PArts organization. (TC 06/12/2024 04:01:28 PM, Patel, Rutvik) Email communication and correspondances Desk organization Parts sorting (TC 06/10/2024 04:02:21 PM, Patel, Rutvik) Email communication and correspondances (TC 06/14/2024 04:03:22 PM, Patel, Rutvik) Baxter training Registration. (TC 06/17/2024 03:37:53 PM, Patel, Rutvik) Email communications and correspondances OR Inventory OR Issue with Laser set up with Microscope (TC 06/18/2024 03:44:56 PM, Patel, Rutvik) Email communications and correspondances Battery and boxes disposal (TC 06/19/2024 03:38:22 PM, Patel, Rutvik) Email communication and correspondances (TC 06/20/2024 04:31:42 PM, Patel, Rutvik) Email communication and correspondaces (TC 06/24/2024 01:52:03 PM, Patel, Rutvik) Email comminications and correspondances Sorted and organized cable bins on the shelves. (TC 06/25/2024 04:02:21 PM, Patel, Rutvik) Parts and Inventory organization Garbage and electronic waste disposal (TC 06/27/2024 04:01:28 PM, Patel, Rutvik) VR glass testing Desk clean up and organization Email communication and correspondances (TC 06/28/2024 02:08:35 PM, Patel, Rutvik) OR Inventory OR 10,11,12 Updated TMS record. (TC 07/10/2024 06:11:59 PM, Patel, Rutvik) Email communications and correspondances Parts sorting and organization Desk area clean up (TC 07/16/2024 03:24:42 PM, Patel, Rutvik) Dealing with Code grey due to hospitla wide flooding caused by heavy rain (TC 07/19/2024 02:49:19 PM, Patel, Rutvik) Email communications and correspondances Worked on compiling biomed training quotes Desk clean up. (TC 08/08/2024 03:50:17 PM, Patel, Rutvik) Email communiation and correspondances Naptune issue with ER OR 11 Wall monitor issue Desk clean up. (TC 10/29/2024 11:24:54 AM, Patel, Rutvik) Assisted Qwynn with OR 3D printing issue. (TC 11/06/2024 03:13:29 PM, Patel, Rutvik) IS Service Desk NARF request for Increasing outlook storage space and SAP Concur access (TC 11/19/2024 03:08:12 PM, Patel, Rutvik) Assisted Naz with returning goods that was biught for MDRD under phase 3A.
1059878 | 204529 | CM | CMPLT | CMPLT | 2024-03-11 07:04 | 2024-03-11 07:04 | Rear case broken | (TC 03/08/2024 03:00:00 PM, Baghirzade, Nara) Replaced rear case.
| 204530 | RDS | CMPLT | CMPLT | 2024-03-11 07:47 | 2024-03-11 10:39 | Weekly check for ED
| 204531 | PA | CMPLT | CMPLT | 2024-03-11 07:51 | 2024-03-15 13:17 | ECG parts for stock | (TC 03/11/2024 08:06:15 AM, Canjimavadivel, Kamallesh ) Placed order for parts (TC 03/15/2024 01:16:45 PM, Canjimavadivel, Kamallesh ) Received parts, added to stock shelf and labelled
1054771 | 204532 | CM | CMPLT | CMPLT | 2024-03-11 08:07 | 2024-03-11 08:24 | Rubber feet melting | (TC 03/11/2024 08:22:12 AM, Rhuland, Courtney) Removed rubber feet from bottom of device Cleaned melted rubber off Functoinal check - OK
1050672 | 204533 | CM | CMPLT | CMPLT | 2024-03-11 08:37 | 2024-03-11 08:40 | V5 not working | (TC 03/11/2024 08:37:47 AM, Canjimavadivel, Kamallesh ) Replaced chest lead set V4-V6. Tested ok with simulator
| 204534 | PA | CMPLT | CMPLT | 2024-03-11 09:16 | 2024-03-21 14:35 | TC70 spare parts for stock | (TC 03/11/2024 09:18:57 AM, Scero, Tomlyn) Parts ordered from Philips for Biomed stock (TC 03/20/2024 02:33:39 PM, Scero, Tomlyn) Parts received and placed on parts shelf for use PO 7202375013
32520 | 204535 | CM | CMPLT | CMPLT | 2024-03-11 11:23 | 2024-03-13 09:38 | MSH DR2 TUBE HEAD | (TC 03/12/2024 09:38:00 PM, Whittington, Scott) MSH DR2 TUBE HEAD REPAIR | MSH DR2 TUBE HEAD
| 204536 | RDS | CMPLT | CMPLT | 2024-03-11 11:41 | 2024-03-11 13:00 | Weekly checks in CCU | (TC 03/11/2024 12:59:58 PM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
68593 | 204537 | PM | CMPLT | PMFAIL | 2024-03-11 12:47 | 2024-03-11 12:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed by BD due to cracked front and rear cases
68593 | 204538 | CM | CMPLT | CMPLT | 2024-03-11 12:48 | 2024-03-13 15:05 | Created by Failed PM WO # 204537 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/11/2024 12:48:28 PM, Rhuland, Courtney) Replaced broken front and rear cases Replaced missing screws (TC 03/13/2024 03:04:37 PM, Rhuland, Courtney) Calibrated and PM performed All tests passed Refer to ASM for details | PM failed by BD due to cracked front and rear cases
| 204539 | CM | CMPLT | CMPLT | 2024-03-11 12:50 | 2024-04-26 10:48 | BD Pump upgrade work | (TC 03/11/2024 12:49:57 PM, Rhuland, Courtney) Cleaned pump IUI connectors Brought pumps back up to BD and brought more down to clean (TC 03/11/2024 02:59:00 PM, Rhuland, Courtney) Cleaned and replaced IUI connectors (TC 03/12/2024 10:40:00 AM, Rhuland, Courtney) Cleaning and replacing IUI connectors (TC 03/12/2024 01:52:47 PM, Rhuland, Courtney) Replaced IUI connectors (TC 03/13/2024 02:59:11 PM, Rhuland, Courtney) Replaced IUI connectors (TC 03/14/2024 02:32:10 PM, Rhuland, Courtney) Replaced and cleaned IUI connectors (TC 03/15/2024 02:00:44 PM, Rhuland, Courtney) Cleaning and replacing IUI connectors Other work on pumps (TC 03/21/2024 02:57:38 PM, Rhuland, Courtney) Replaced IUI connectors/repairs on pumps (TC 04/03/2024 01:10:24 PM, Rhuland, Courtney) Replaced IUI connectors (TC 04/03/2024 02:52:21 PM, Rhuland, Courtney) Completed work orders (TC 04/04/2024 01:33:10 PM, Rhuland, Courtney) Closed work orders for PM/recall (TC 04/05/2024 02:53:41 PM, Rhuland, Courtney) Closed PM and recall work orders (TC 04/10/2024 02:54:44 PM, Rhuland, Courtney) Closed PM and recall work orders (TC 04/24/2024 02:56:53 PM, Rhuland, Courtney) Closed PM and recall work orders (TC 04/26/2024 10:48:06 AM, Rhuland, Courtney) Closed remaining work orders for PM/recall
68377 | 204540 | CM | CMPLT | CMPLT | 2024-03-11 13:03 | 2024-03-12 14:46 | Incident Report File 69715 | (TC 03/11/2024 02:22:25 PM, Scero, Tomlyn) DSR sent to Draeger Germany for Debug explanation. Response received from them: The probable cause of the issue was the ventilator problem during the case. In the Debug Log and it seems that the blower had a problem to ventilate because of a too high difference between Pinsp and Pexp pressure sensor. This problem was also present during the tests at 16:12 and 16:21! 2024-02-14T16:12:04.8280;0035;"Service log entry with error class K2, ID 0035.00004 has been created. (Pressure Sensor Difference Out Of Tolerance) / ActiveValve : Pe = 9.574, Pi = 13.318 2024-02-14T16:21:14.7960;0035;"Service log entry with error class K2, ID 0035.00004 has been created. (Pressure Sensor Difference Out Of Tolerance) / ActiveValve : Pe = 11.383, Pi = 13.531 As it can be seen, the Pe sensor drifted but became better and up from 21.02.2024 the tests were OK again. Finally the following reasons are possible: - Faulty inspiratory pressure sensor (3.20). - Faulty expiratory pressure sensor (3.8). - Obstructed filters upstream of sensors. - There is water in the pressure sensor hoses. - Leaks in the pressure sensor hoses. - Measurement disturbed by operator. Because the problem was resolved without any action It is possible that water was the root cause for all the issues.. So It is suggested to do the following: - Perform active valve test (2076) with HIT. (The linearity test is performed here) - Check sensors and replace if necessary. - Check hoses. - Check filters and replace if necessary (TC 03/12/2024 02:24:04 PM, Scero, Tomlyn) Performed active valve test (2076) with HIT. Test passed. Self test passed. Checked sensors, hoses, and filters. All were ok. Just to note, before the device was given to me for testing the flow sensors were replaced. Before throwing them out I did noticed that 1 of them was broken/damaged. (TC 03/12/2024 02:26:25 PM, Scero, Tomlyn) Followed up with the BME Incident report attached.
1056125 | 204541 | PM | CMPLT | CMPLT | 2024-03-11 13:26 | 2024-03-11 13:27 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/11/2024 01:26:12 PM, Canjimavadivel, Kamallesh ) Cleaned hoses and valves. Cleaning cycle complete
1053391 | 204542 | CM | CMPLT | CMPLT | 2024-03-11 14:20 | 2024-03-11 14:21 | QC Checklist | (TC 03/11/2024 02:20:14 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 124 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 126 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a n hour. Device is functional. Returned to 11S.
1059778 | 204543 | FC | CMPLT | CMPLT | 2024-03-11 14:21 | 2024-03-11 14:22 | QC Checklist | (TC 03/11/2024 02:21:17 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
| 204544 | CM | ACTIV | ISSUE | 2024-03-11 15:14 |  | Fixing missing or broken otoscopes | (TC 03/11/2024 03:13:54 PM, Scero, Tomlyn) replaced what I could of the missing otoscopes and then tryin gto find replacements. I may have to order new heads as these are most likely discontinued.
| 204545 | CM | CMPLT | CMPLT | 2024-03-12 08:17 | 2024-03-11 15:20 | Endo Cart 1 Organization | Worked on Endo emerge cart to relocate the power bar from top to middle shelf and arrange cabling as reuqested by Anita Chin. Toured with Adeel and informed him about the chnages I made as well.
68601 | 204546 | PM | CMPLT | PMFAIL | 2024-03-12 08:29 | 2024-03-12 08:29 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 08:29:07 AM, Baghirzade, Nara) Broken rear case, left IUI, front case, and left handle.
68601 | 204547 | CM | CMPLT | CMPLT | 2024-03-12 08:29 | 2024-03-21 13:13 | Created by Failed PM WO # 204546 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 08:29:55 AM, Baghirzade, Nara) Replaced rear case and left handle. Front case and left IUI is on back order. (TC 03/21/2024 01:11:52 PM, Baghirzade, Nara) Replaced front case. Replaced left IUI. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. (TC 03/21/2024 01:11:52 PM, Baghirzade, Nara) Replaced front case. Replaced left IUI. Replaced mylar gasket. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68605 | 204548 | CM | CMPLT | CMPLT | 2024-03-12 08:32 | 2024-03-12 08:36 | Created by Failed PM WO # 200602 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 08:32:33 AM, Baghirzade, Nara) Replaced: -front case -rear case -right IUI -mylar gasket -pressure sensor disc holder (the one on the inside) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/12/2024 08:31:12 AM, Baghirzade, Nara) Broken front case and rear case, defective right IUI
1063811 | 204549 | CM | CMPLT | CMPLT | 2024-03-12 08:51 | 2024-03-22 11:08 | Wrench on remote flashing. Bed won't move up or down. | (TC 03/12/2024 08:51:35 AM, Scero, Tomlyn) I first went upstairs to see if unplugging the unit would work. It didn't change. A case was starting so I had to leave and come back this afternoon. I emailed the Manufacturer IBIOM to get a copy of the service manual. (TC 03/12/2024 02:19:41 PM, Scero, Tomlyn) Troubleshooting over the phone with Umano Medical (IBIOM). Looks as though the foot side column isn't raising past a certain point and is causing the error. Back and forth phone calls with Umano, but they're going to send a service tech to do the repair which will be covered by warranty. I also stopped in Diannes office to update her on the progress. (TC 03/15/2024 01:45:20 PM, Scero, Tomlyn) Paul and Mike from Umano Medical came on site and I directed them to the ECHO lab. Together we worked through the issue and confimed that the foot side column wasn't raising properly. This is actually due to the head side column malfuncitoning. The techs replaced the column and we confirmed that the issue has been resolved. they said they would complete the WO next week and send me a copy of the service report. Emailed Dianne an update. (TC 03/22/2024 11:07:33 AM, Scero, Tomlyn) service report received and attached to WO
66025 | 204550 | PM | CMPLT | CMPLT | 2024-03-12 09:05 | 2024-03-12 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:05:22 AM, Rendon Lopez , Elizabeth) PM completed by BD
66022 | 204551 | PM | CMPLT | CMPLT | 2024-03-12 09:08 | 2024-03-12 09:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:08:04 AM, Rendon Lopez , Elizabeth) PM completed by BD
68690 | 204552 | PM | CMPLT | CMPLT | 2024-03-12 09:10 | 2024-03-12 09:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:10:31 AM, Rendon Lopez , Elizabeth) PM completed by BD
65899 | 204553 | PM | CMPLT | CMPLT | 2024-03-12 09:21 | 2024-03-12 09:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
68931 | 204554 | PM | CMPLT | CMPLT | 2024-03-12 09:21 | 2024-03-12 09:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:23:56 AM, Rendon Lopez , Elizabeth) PM completed by BD
65992 | 204555 | PM | CMPLT | CMPLT | 2024-03-12 09:25 | 2024-03-12 09:25 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
68300 | 204556 | PM | CMPLT | CMPLT | 2024-03-12 09:25 | 2024-03-12 09:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68947 | 204557 | PM | CMPLT | CMPLT | 2024-03-12 09:29 | 2024-03-12 09:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:28:46 AM, Rendon Lopez , Elizabeth) PM completed by BD
69296 | 204558 | PM | CMPLT | CMPLT | 2024-03-12 09:29 | 2024-03-12 09:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68583 | 204559 | PM | CMPLT | CMPLT | 2024-03-12 09:30 | 2024-03-12 09:31 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
69255 | 204560 | PM | CMPLT | CMPLT | 2024-03-12 09:30 | 2024-03-12 09:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68783 | 204561 | PM | CMPLT | CMPLT | 2024-03-12 09:30 | 2024-03-12 09:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:30:41 AM, Rendon Lopez , Elizabeth) PM completed by BD
69199 | 204562 | PM | CMPLT | CMPLT | 2024-03-12 09:32 | 2024-03-12 09:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69509 | 204563 | PM | CMPLT | CMPLT | 2024-03-12 09:33 | 2024-03-12 09:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69504 | 204564 | PM | CMPLT | CMPLT | 2024-03-12 09:34 | 2024-03-12 09:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69193 | 204565 | PM | CMPLT | CMPLT | 2024-03-12 09:39 | 2024-03-12 09:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68581 | 204566 | PM | CMPLT | CMPLT | 2024-03-12 09:39 | 2024-03-12 09:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
69049 | 204567 | PM | CMPLT | CMPLT | 2024-03-12 09:40 | 2024-03-12 09:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68660 | 204568 | PM | CMPLT | CMPLT | 2024-03-12 09:42 | 2024-03-12 09:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69589 | 204569 | PM | CMPLT | CMPLT | 2024-03-12 09:43 | 2024-03-12 09:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69376 | 204570 | PM | CMPLT | CMPLT | 2024-03-12 09:44 | 2024-03-12 09:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:43:54 AM, Rendon Lopez , Elizabeth) PM completed by BD
69592 | 204571 | PM | CMPLT | CMPLT | 2024-03-12 09:45 | 2024-03-12 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69371 | 204572 | PM | CMPLT | CMPLT | 2024-03-12 09:46 | 2024-03-12 09:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69161 | 204573 | PM | CMPLT | CMPLT | 2024-03-12 09:47 | 2024-03-12 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:47:20 AM, Rendon Lopez , Elizabeth) PM completed by BD
69616 | 204574 | PM | CMPLT | CMPLT | 2024-03-12 09:49 | 2024-03-12 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:48:55 AM, Rendon Lopez , Elizabeth) PM completed by BD
69416 | 204575 | PM | CMPLT | CMPLT | 2024-03-12 09:51 | 2024-03-12 09:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:50:53 AM, Rendon Lopez , Elizabeth) PM completed by BD
69448 | 204576 | PM | CMPLT | CMPLT | 2024-03-12 09:52 | 2024-03-12 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:52:35 AM, Rendon Lopez , Elizabeth) PM completed by BD
69286 | 204577 | PM | CMPLT | CMPLT | 2024-03-12 09:53 | 2024-03-12 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69328 | 204578 | PM | CMPLT | CMPLT | 2024-03-12 09:54 | 2024-03-12 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:54:16 AM, Rendon Lopez , Elizabeth) PM completed by BD
69057 | 204579 | PM | CMPLT | CMPLT | 2024-03-12 09:56 | 2024-03-12 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:56:04 AM, Rendon Lopez , Elizabeth) PM completed by BD
68600 | 204580 | PM | CMPLT | CMPLT | 2024-03-12 09:58 | 2024-03-12 09:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
68667 | 204581 | PM | CMPLT | CMPLT | 2024-03-12 09:58 | 2024-03-12 09:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 09:57:57 AM, Rendon Lopez , Elizabeth) PM completed by BD
68602 | 204582 | PM | CMPLT | CMPLT | 2024-03-12 10:01 | 2024-03-12 10:02 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
1060462 | 204583 | CM | CMPLT | CMPLT | 2024-03-12 10:07 | 2024-03-20 17:30 | Scope Guide Cable Broken | Mar 11 - Call re: not working. Went to check. Found Scope guide cable broken. Removed cable and routed and installed short cable. Installed and secured cable protector on the floor due to posibbility cable of being overrun by strechers. Also found arm covers broken. Replaced both arm covers. Mar 12 - Contacted Olympus for part# and pricing. Received. Created SC #1003187918 and sent it for approval. Approval Received. Awaiting Delivery. Mar 20 - Coordinated and scheduled replacement. Removed temporary cable and Replaced and routed new scope guide cable. Calibrated teh scope guide position. Tested ok.
1055166 | 204584 | PM | CMPLT | CMPLT | 2024-03-12 10:15 | 2024-03-12 10:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058243 | 204585 | PM | CMPLT | CMPLT | 2024-03-12 10:23 | 2024-03-12 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058250 | 204586 | PM | CMPLT | CMPLT | 2024-03-12 10:25 | 2024-03-12 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69594 | 204587 | PM | CMPLT | CMPLT | 2024-03-12 10:37 | 2024-03-12 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1070230 | 204588 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 13:52 | Incoming inspection | (TC 03/12/2024 10:54:42 AM, Rhuland, Courtney) Tagged and functional check - OK
1070231 | 204589 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:55 | Incoming inspection | (TC 03/12/2024 10:54:13 AM, Rhuland, Courtney) Tagged and functional check - OK
1070232 | 204590 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:54 | Incoming inspection | (TC 03/12/2024 10:53:45 AM, Rhuland, Courtney) Tagged and functional check - OK
1070233 | 204591 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:54 | Incoming inspection | (TC 03/12/2024 10:53:15 AM, Rhuland, Courtney) Tagged and functional check - OK
1070234 | 204592 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:53 | Incoming inspection | (TC 03/12/2024 10:52:45 AM, Rhuland, Courtney) Tagged and functional check - OK
1070235 | 204593 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:53 | Incoming inspection | (TC 03/12/2024 10:52:16 AM, Rhuland, Courtney) Tagged and functional check - OK
1070236 | 204594 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:52 | Incoming inspection | (TC 03/12/2024 10:51:47 AM, Rhuland, Courtney) Tagged and functional check - OK
1070237 | 204595 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:52 | Incoming inspection | (TC 03/12/2024 10:51:13 AM, Rhuland, Courtney) Tagged and functional check - OK
1070238 | 204596 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:51 | Incoming inspection | (TC 03/12/2024 10:50:33 AM, Rhuland, Courtney) Tagged and functional check - OK
1070239 | 204597 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:50 | Incoming inspection | (TC 03/12/2024 10:49:50 AM, Rhuland, Courtney) Tagged and functional check - OK
1070240 | 204598 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:50 | Incoming inspection | (TC 03/12/2024 10:49:26 AM, Rhuland, Courtney) Tagged and functional check - OK
1070241 | 204599 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:49 | Incoming inspection | (TC 03/12/2024 10:48:59 AM, Rhuland, Courtney) Tagged and functional check - OK
1070242 | 204600 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:48 | Incoming inspection | (TC 03/12/2024 10:48:09 AM, Rhuland, Courtney) Tagged and functional check - OK
1070243 | 204601 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:48 | Incoming inspection | (TC 03/12/2024 10:47:40 AM, Rhuland, Courtney) Tagged and functional check - OK
1070244 | 204602 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-20 10:03 | Incoming inspection | (TC 03/19/2024 09:53:31 AM, Cortez, John) Verified scale readings
1070245 | 204603 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-20 10:08 | Incoming inspection | (TC 03/19/2024 10:06:53 AM, Cortez, John) Verified scale readings
1070246 | 204604 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:47 | Incoming inspection | (TC 03/12/2024 10:46:08 AM, Rhuland, Courtney) Tagged and functional check - OK
1070247 | 204605 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:46 | Incoming inspection | (TC 03/12/2024 10:45:38 AM, Rhuland, Courtney) Tagged and functional check - OK
1070248 | 204606 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:45 | Incoming inspection | (TC 03/12/2024 10:45:00 AM, Rhuland, Courtney) Tagged and functional check - OK
1070249 | 204607 | CM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:45 | Incoming inspection | (TC 03/12/2024 10:42:59 AM, Rhuland, Courtney) Tagged and functional check - OK
69590 | 204608 | PM | CMPLT | CMPLT | 2024-03-12 10:43 | 2024-03-12 10:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 10:43:18 AM, Rendon Lopez , Elizabeth) PM completed by BD
69520 | 204609 | PM | CMPLT | CMPLT | 2024-03-12 10:44 | 2024-03-12 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047047 | 204610 | PM | CMPLT | CMPLT | 2024-03-12 10:45 | 2024-03-12 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 10:45:21 AM, Rendon Lopez , Elizabeth) PM completed by BD
69081 | 204611 | PM | CMPLT | CMPLT | 2024-03-12 10:47 | 2024-03-12 10:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69114 | 204612 | PM | CMPLT | CMPLT | 2024-03-12 10:48 | 2024-03-12 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055179 | 204613 | PM | CMPLT | CMPLT | 2024-03-12 10:48 | 2024-03-12 10:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 10:48:20 AM, Rendon Lopez , Elizabeth) PM completed by BD
69342 | 204614 | PM | CMPLT | CMPLT | 2024-03-12 10:49 | 2024-03-12 10:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69374 | 204615 | PM | CMPLT | PMFAIL | 2024-03-12 10:49 | 2024-03-13 15:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:50:20 AM, Cortez, John) Patient side occlusion failed
65971 | 204616 | PM | CMPLT | CMPLT | 2024-03-12 10:50 | 2024-03-12 10:51 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 10:50:32 AM, Rendon Lopez , Elizabeth) PM completed by BD
65946 | 204617 | PM | CMPLT | CMPLT | 2024-03-12 10:52 | 2024-03-12 10:53 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 10:52:31 AM, Rendon Lopez , Elizabeth) PM completed by BD
69254 | 204618 | PM | CMPLT | CMPLT | 2024-03-12 10:53 | 2024-03-12 10:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:53:33 AM, Cortez, John) Female IUI replaced PM complete
68611 | 204619 | PM | CMPLT | CMPLT | 2024-03-12 10:57 | 2024-03-12 10:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 10:57:00 AM, Rendon Lopez , Elizabeth) PM completed by BD
69634 | 204620 | PM | CMPLT | CMPLT | 2024-03-12 10:58 | 2024-03-12 10:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69221 | 204621 | PM | CMPLT | CMPLT | 2024-03-12 11:01 | 2024-03-12 11:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 11:01:23 AM, Cortez, John) Cleaned male IUI PM complete
1057392 | 204622 | PM | CMPLT | CMPLT | 2024-03-12 11:02 | 2024-03-12 11:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058236 | 204623 | PM | CMPLT | CMPLT | 2024-03-12 11:03 | 2024-03-12 11:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66039 | 204624 | PM | CMPLT | CMPLT | 2024-03-12 11:21 | 2024-03-12 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
66042 | 204625 | PM | CMPLT | CMPLT | 2024-03-12 11:22 | 2024-03-12 11:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69010 | 204626 | PM | CMPLT | CMPLT | 2024-03-12 11:27 | 2024-03-12 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68988 | 204627 | PM | CMPLT | CMPLT | 2024-03-12 11:30 | 2024-03-12 11:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68804 | 204628 | PM | CMPLT | CMPLT | 2024-03-12 11:34 | 2024-03-12 11:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68845 | 204629 | PM | CMPLT | CMPLT | 2024-03-12 11:37 | 2024-03-12 11:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68792 | 204630 | PM | CMPLT | CMPLT | 2024-03-12 11:40 | 2024-03-12 11:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68852 | 204631 | PM | CMPLT | CMPLT | 2024-03-12 11:41 | 2024-03-12 11:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68976 | 204632 | PM | CMPLT | CMPLT | 2024-03-12 11:43 | 2024-03-12 11:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68929 | 204633 | PM | CMPLT | CMPLT | 2024-03-12 11:44 | 2024-03-12 11:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68949 | 204634 | PM | CMPLT | CMPLT | 2024-03-12 11:46 | 2024-03-12 11:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1069363 | 204635 | CM | CMPLT | CMPLT | 2024-03-12 11:54 | 2024-09-25 09:34 | MSH OPG SPP GE VOLUSON EXPERT 22 CONNECTION TO HSC | (TC 03/12/2024 11:56:07 AM, Whittington, Scott) MSH OPG SPP GE VOLUSON EXPERT 22 CONNECTION TO HSC From: Whittington, Scott Sent: March 12, 2024 11:26 AM To: Adnan Al-Jalam <adnan.al-jalam@sickkids.ca>; Wallace, Steven - The Hospital for Sick Children (SickKids) <steve@sickkids.ca>; Raigon Kuriakkose <raigon.kuriakkose@sickkids.ca> Cc: Riaz, Mirza - Sinai Health System <Mirza.Riaz@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca>; Hollingshead, Andrea - Sinai Health System <Andrea.Hollingshead@sinaihealth.ca>; Stratulat, Vasilica - Sinai Health System <Vasilica.Stratulat@sinaihealth.ca>; Ellul, Katie - Sinai Health System <Katie.Ellul@sinaihealth.ca> Subject: RE: [External] RE: MSH SPP Image t4ransfer from MSH to HSC echo lab Hi Adnan Here are the AE titles from our database. The systems that are no longer connecting to HSC may have been moved to different rooms. I am not sure if the network jacks/switch that are used have to be programmed to allow to go through the firewall/ Mirza Riaz from MSH Networks is copied here. Thanks, Scott Asset Room Network Jack LAN MAC STATIC IP AE 1057302 3-525 (D7) 3N-125D c400ad541fd7 172.20.220.38 MSH_US_SPP_VOLF 1069363 3-529 (D4) 3N-126D 0013955af331 172.19.72.5 MSH_US_SPP_VOLI (CAPITAL I) 1063331 3-537 (C3) 3N-D071A c400ad9ee7e4 172.20.220.19 MSH_US_SPP_VOLG 502453 3-539 (C2) 3N-D069B c400ad4b123d 172.20.220.16 MSH_US_SPP_VOLB 1063808 msh 15 procedure C4:00:AD:9D:94:3B 172.20.220.39 MSH_US_SPP_VOLH (TC 03/28/2024 09:37:46 AM, Whittington, Scott) MSH OPG SPP GE VOLUSON EXPERT 22 CONNECTION TO HSC | MSH OPG SPP GE VOLUSON EXPERT 22 CONNECTION TO HSC
68642 | 204636 | PM | CMPLT | CMPLT | 2024-03-12 12:19 | 2024-03-12 12:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 12:19:11 PM, Baghirzade, Nara) Replaced AIL assembly. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1047104 | 204637 | PM | CMPLT | CMPLT | 2024-03-12 12:57 | 2024-03-12 12:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055105 | 204638 | PM | CMPLT | CMPLT | 2024-03-12 13:00 | 2024-03-12 13:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1063331 | 204639 | CM | CMPLT | CMPLT | 2024-03-12 13:09 | 2024-03-12 13:09 | PM by GE | Uploaded report.
1065686 | 204640 | PM | CMPLT | CMPLT | 2024-03-12 13:12 | 2024-03-12 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065698 | 204641 | PM | CMPLT | CMPLT | 2024-03-12 13:15 | 2024-03-12 13:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065710 | 204642 | PM | CMPLT | CMPLT | 2024-03-12 13:18 | 2024-03-12 13:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065700 | 204643 | PM | CMPLT | CMPLT | 2024-03-12 13:22 | 2024-03-12 13:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065716 | 204644 | PM | CMPLT | CMPLT | 2024-03-12 13:23 | 2024-03-12 13:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1063808 | 204645 | CM | CMPLT | CMPLT | 2024-03-12 13:25 | 2024-03-12 14:27 | PM by GE
1065727 | 204646 | PM | CMPLT | CMPLT | 2024-03-12 13:26 | 2024-03-12 13:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065734 | 204647 | PM | CMPLT | CMPLT | 2024-03-12 13:29 | 2024-03-12 13:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68923 | 204648 | CM | CMPLT | CMPLT | 2024-03-12 13:58 | 2024-03-12 14:00 | Damaged power cord | (TC 03/12/2024 01:58:28 PM, Baghirzade, Nara) Replaced power cord.
1070257 | 204649 | ACC | CMPLT | CMPLT | 2024-03-12 14:00 | 2024-03-13 11:39 | Incoming inspection | (TC 03/12/2024 03:10:53 PM, Canjimavadivel, Kamallesh ) Unboxed and added to database Assembled case Upgraded software rev and uploaded config (TC 03/13/2024 11:35:31 AM, Canjimavadivel, Kamallesh ) Upgraded software to latest revision Added network certificates Added connectors and modified with security feature Performed electrical safety Checked into dashboard
65898 | 204650 | PM | CMPLT | CMPLT | 2024-03-12 14:09 | 2024-03-12 14:10 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 02:09:24 PM, Baghirzade, Nara) PM completed by BD.
65922 | 204651 | PM | CMPLT | CMPLT | 2024-03-12 14:13 | 2024-03-12 14:13 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 02:13:06 PM, Baghirzade, Nara) PM completed by BD.
69346 | 204652 | PM | CMPLT | CMPLT | 2024-03-12 14:14 | 2024-03-12 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 02:14:38 PM, Reid, Emmerson ) Did rate calibration, was running low. PM Complete.
69445 | 204653 | PM | CMPLT | CMPLT | 2024-03-12 14:24 | 2024-03-13 08:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 08:05:04 AM, Reid, Emmerson ) Ran rate calibration, pm complete.
68604 | 204654 | CM | CMPLT | CMPLT | 2024-03-12 14:31 | 2024-03-19 13:56 | Broken rear and front case | (TC 03/12/2024 02:31:14 PM, Baghirzade, Nara) Replaced rear case. Broken front case; front case on backorder. (TC 03/19/2024 01:55:43 PM, Baghirzade, Nara) Replaced front case and left IUI. Returned syringe to BD for remediation. (TC 03/19/2024 01:55:43 PM, Baghirzade, Nara) Replaced front case and left IUI. Replaced mylar gasket. Returned syringe to BD for remediation.
| 204655 | CM | CMPLT | CMPLT | 2024-03-12 14:32 | 2024-04-16 13:47 | Replacing IUIs after BD inspection | (TC 03/12/2024 02:32:48 PM, Baghirzade, Nara) Replacing IUIs after BD has inspected (TC 03/18/2024 09:59:14 AM, Baghirzade, Nara) Replaced right IUIs. Left IUIs are on backorder. Cleaned left IUIs for pumps and PCUs; sent the cleaned ones back to BD for upgrade and PM. (TC 03/19/2024 08:35:58 AM, Baghirzade, Nara) Replacing left and right IUIs. (TC 03/18/2024 02:30:00 PM, Baghirzade, Nara) Replacing IUIs (TC 03/20/2024 09:15:02 AM, Baghirzade, Nara) Replaced IUIs on LVPs, syringes, and PCUs. Brought defective pumps/PCUs for repair to biomed. (TC 03/20/2024 11:51:33 AM, Baghirzade, Nara) Replaced IUIs and brought devices to biomed for repair. (TC 03/21/2024 09:29:42 AM, Baghirzade, Nara) Replaced IUIs on PCUs, syringes, and LVPs. Brought pumps/PCUs back up to BD for remediation and PM. Picked up pumps/PCUs that require repairs. (TC 03/22/2024 10:20:30 AM, Baghirzade, Nara) Replacing IUIs and sending pumps back to BD for remediation and PM.
| 204656 | RDS | CMPLT | CMPLT | 2024-03-12 15:36 | 2024-03-12 15:36 | weekly check for 14S and 17M | weekly check for 14S and 17M
| 204657 | CM | CMPLT | CMPLT | 2024-03-12 15:39 | 2024-03-12 15:39 | worked with RT | worked with RT
| 204658 | CM | CMPLT | CMPLT | 2024-03-12 15:40 | 2024-03-12 15:40 | meeting | meeting
1070256 | 204659 | CM | CMPLT | CMPLT | 2024-03-12 15:42 | 2024-03-12 15:42 | Incoming inspection | (TC 03/12/2024 03:41:56 PM, Rendon Lopez , Elizabeth) Incoming inspection, added tester to TMS. Labelled H1200 tester
1051664 | 204660 | CM | CMPLT | CMPLT | 2024-03-12 15:43 | 2024-03-12 15:44 | Broken | (TC 03/11/2024 03:42:45 PM, Rendon Lopez , Elizabeth) Damaged case and screen. Unable to repair. Notified PCM, she advised they don't need a replacement. Retire from service
66517 | 204661 | CM | CMPLT | CMPLT | 2024-03-12 15:47 | 2024-03-12 15:48 | Moved to room 1 | (TC 03/11/2024 03:47:09 PM, Rendon Lopez , Elizabeth) Moved spare monitor to L&D room 1 as replacement for compatible software revision with CL stations. Tested all functions-OK
68029 | 204662 | CM | CMPLT | NFF | 2024-03-12 15:50 | 2024-03-12 15:50 | Will not connect-beeping | (TC 03/11/2024 03:49:57 PM, Rendon Lopez , Elizabeth) Tested, no issues found. Returned to L&D Room 1
66465 | 204663 | CM | CMPLT | CMPLT | 2024-03-12 15:51 | 2024-03-12 15:52 | Battery maintenance recommended alert | (TC 03/11/2024 03:51:50 PM, Rendon Lopez , Elizabeth) Turned off battery status alert. Tested functions-OK
1048552 | 204664 | CM | CMPLT | CMPLT | 2024-03-12 16:03 | 2024-03-20 15:14 | Repair - Battery Error | Mar 12 - Call re: not working. Went to check and brought it down to biomed. Found unit indicating battery error. Attempted to troubleshoot it. Found we do not have any battery in stock. Contacted Zimmer to arrnage for a loaner and requested battery Info. Mar 14 - Was informed that Zimmer does not sell any parts and unit will be required to send in for repair. Spoke with KC and SC and was advised to swap battery with a working unit. Swapped battery and found unit works. Contacted BBM for battery info. Received. Created SC 1003189746 and sent it for approval. Received. Issued PO via MMC. Awaiting delivery. Mar 19 - Followed up. Was informed that battery was delivered on Mar 15. Checked with the staff. No one had received the battery. Went to receiving and searched for it. Found battery with receiving dept. Received and replaced it in the unit. Left unit to charge over night. Mar 20 - Went through the service manual and completed calibration and PM. Tested ok. Returned to OR.
67905 | 204665 | CM | ACTIV | ISSUE | 2024-03-13 08:45 |  | dark screen
1061203 | 204666 | CM | CMPLT | CMPLT | 2024-03-13 08:54 | 2024-03-21 10:32 | bluetooth adapter | (TC 03/14/2024 10:29:02 AM, Huang, Jessica) got the call in OR 2 and couldnt connect to the bluetooth. checked the cables and restarted the VMC. Didnt xif the issue. Called Olympus for tech support. Disconnected the blutooth adapter and reconnected it. It didnt work either. Replaced the bluetooth and working fine.
65971 | 204667 | CM | CMPLT | CMPLT | 2024-03-13 09:01 | 2024-03-13 09:02 | Missing barcode | (TC 03/13/2024 09:01:18 AM, Rhuland, Courtney) Printed barcode for syringe Functional check - OK
68954 | 204668 | PM | CMPLT | CMPLT | 2024-03-13 09:08 | 2024-03-13 09:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/11/2024 09:08:00 AM, Rendon Lopez , Elizabeth) PM completed by BD
68590 | 204669 | CM | CMPLT | CMPLT | 2024-03-13 09:12 | 2024-03-15 14:03 | Created by Failed PM WO # 198758 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 09:12:43 AM, Rhuland, Courtney) PM failed by BD Cleaned left IUI Replaced rear case Replaced left handle (TC 03/15/2024 02:02:43 PM, Rhuland, Courtney) PM completed by BD
69024 | 204670 | PM | CMPLT | CMPLT | 2024-03-13 09:13 | 2024-03-13 09:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:13:32 AM, Rendon Lopez , Elizabeth) PM completed by BD
68779 | 204671 | PM | CMPLT | CMPLT | 2024-03-13 09:15 | 2024-03-13 09:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:14:53 AM, Rendon Lopez , Elizabeth) PM completed by BD
68953 | 204672 | PM | CMPLT | CMPLT | 2024-03-13 09:16 | 2024-03-13 09:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:16:33 AM, Rendon Lopez , Elizabeth) PM completed by BD
68938 | 204673 | PM | CMPLT | CMPLT | 2024-03-13 09:19 | 2024-03-13 09:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:19:00 AM, Rendon Lopez , Elizabeth) PM completed by BD
69022 | 204674 | PM | CMPLT | CMPLT | 2024-03-13 09:22 | 2024-03-13 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:22:04 AM, Rendon Lopez , Elizabeth) PM completed by BD
1047095 | 204675 | PM | CMPLT | CMPLT | 2024-03-13 09:25 | 2024-03-13 09:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:25:24 AM, Rendon Lopez , Elizabeth) PM completed by BD
1050244 | 204676 | PM | CMPLT | CMPLT | 2024-03-13 09:29 | 2024-03-13 09:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:29:38 AM, Rendon Lopez , Elizabeth) PM completed by BD
1055119 | 204677 | PM | CMPLT | CMPLT | 2024-03-13 09:31 | 2024-03-13 09:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:31:28 AM, Rendon Lopez , Elizabeth) PM completed by BD
69540 | 204678 | PM | CMPLT | CMPLT | 2024-03-13 09:37 | 2024-03-13 09:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:37:15 AM, Cortez, John) female IUI replaced, module latch cleaned PM complete
1050619 | 204679 | CM | CMPLT | CMPLT | 2024-03-13 09:37 | 2024-03-15 14:04 | Created by Failed PM WO # 199609 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 09:37:56 AM, Rhuland, Courtney) PM failed by BD Replaced rear case Replaced left IUI Replaced missing screws (TC 03/15/2024 02:04:03 PM, Rhuland, Courtney) PM completed by BD
1055115 | 204680 | PM | CMPLT | CMPLT | 2024-03-13 09:44 | 2024-03-13 09:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:45:35 AM, Rendon Lopez , Elizabeth) PM completed by BD
69352 | 204681 | PM | CMPLT | CMPLT | 2024-03-13 09:49 | 2024-03-13 10:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:49:27 AM, Cortez, John) Module latch cleaned, Female IUI replaced PM complete
69352 | 204682 | PM | CANCL | DUP | 2024-03-13 09:49 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058267 | 204683 | PM | CMPLT | CMPLT | 2024-03-13 09:54 | 2024-03-13 09:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:54:39 AM, Rendon Lopez , Elizabeth) PM completed by BD
1058270 | 204684 | PM | CMPLT | CMPLT | 2024-03-13 09:56 | 2024-03-13 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:55:53 AM, Rendon Lopez , Elizabeth) PM completed by BD
| 204685 | PA | CMPLT | CMPLT | 2024-03-13 09:56 | 2024-03-15 07:10 | Ordering sterile water for altrix cleaning NICU | (TC 03/13/2024 09:59:34 AM, Baghirzade, Nara) SC 1003188929 (TC 03/15/2024 07:09:20 AM, Baghirzade, Nara) Order # 4150273572
68023 | 204686 | CM | CMPLT | CMPLT | 2024-03-13 09:57 | 2024-03-27 12:12 | ordering VN parts | received test lungs PO#7202376160
1049901 | 204687 | CM | CMPLT | CMPLT | 2024-03-13 09:58 | 2024-03-14 10:56 | MSH R61 NICU COLLIMATOR LAMP BUTTON IS INTERMITTENT | (TC 03/13/2024 09:58:28 AM, Whittington, Scott) MSH R61 NICU COLLIMATOR LAMP BUTTON IS INTERMITTENT PN# SP9J0042 HARNESS- COLLIMATOR LIGHT SWITCH (TC 03/13/2024 09:58:28 AM, Whittington, Scott) MSH R61 NICU COLLIMATOR LAMP BUTTON IS INTERMITTENT PN# SP9J0042 HARNESS- COLLIMATOR LIGHT SWITCH SWITCH CONFIRMED FAULTY OBTAINED REPLACEMENT FROM LOCAL STOCK REPLACED AA BATTERIES IN REMOTE SO COLLIMATOR LAMP CAN BE ACTIVATED ORDER PARTS FROM CARESTREAM (TC 03/13/2024 09:58:28 AM, Whittington, Scott) MSH R61 NICU COLLIMATOR LAMP BUTTON IS INTERMITTENT PN# SP9J0042 HARNESS- COLLIMATOR LIGHT SWITCH SWITCH CONFIRMED FAULTY OBTAINED REPLACEMENT FROM LOCAL STOCK REPLACED AA BATTERIES IN REMOTE SO COLLIMATOR LAMP CAN BE ACTIVATED SWITCH REPLACED ORDER PARTS FROM CARESTREAM (TC 03/14/2024 09:58:28 AM, Whittington, Scott) march 13 MSH R61 NICU COLLIMATOR LAMP BUTTON IS INTERMITTENT PN# SP9J0042 HARNESS- COLLIMATOR LIGHT SWITCH SWITCH CONFIRMED FAULTY OBTAINED REPLACEMENT FROM LOCAL STOCK REPLACED AA BATTERIES IN REMOTE SO COLLIMATOR LAMP CAN BE ACTIVATED SWITCH REPLACED ORDER PARTS FROM CARESTREAM | MSH R61 NICU COLLIMATOR LAMP BUTTON IS INTERMITTENT PN# SP9J0042 HARNESS- COLLIMATOR LIGHT SWITCH
1058281 | 204688 | PM | CMPLT | CMPLT | 2024-03-13 09:59 | 2024-03-13 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 09:59:07 AM, Rendon Lopez , Elizabeth) PM completed by BD
32510 | 204689 | CM | CMPLT | CMPLT | 2024-03-13 10:01 | 2024-03-14 10:59 | MSH DR1 OGP ERROR | (TC 03/13/2024 10:00:59 AM, Whittington, Scott) MSH DR1 OGP ERROR (TC 03/14/2024 10:00:59 AM, Whittington, Scott) MARCH 13 MSH DR1 OGP ERROR LATHAN REPORTED AN ERROR IN MOVEMENT ABLE TO CLEAR (TC 03/14/2024 10:00:59 AM, Whittington, Scott) MARCH 13 MSH DR1 OGP ERROR LATHAN REPORTED AN ERROR IN MOVEMENT ABLE TO CLEAR | MSH DR1 OGP ERROR LATHAN REPORTED AN ERROR IN MOVEMENT
1070267 | 204690 | CM | ACTIV | ISSUE | 2024-03-13 10:07 |  | Static IP address added- device installed in DMZ- set up needs to be completed by Philips Tcon
69284 | 204691 | PM | CMPLT | CMPLT | 2024-03-13 10:08 | 2024-03-13 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:07:40 AM, Cortez, John) Female IUI replaced PM complete
69423 | 204692 | PM | CMPLT | CMPLT | 2024-03-13 10:15 | 2024-03-13 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:15:52 AM, Cortez, John) Female IUI replaced PM complete
| 204693 | CM | CMPLT | CMPLT | 2024-03-13 10:18 | 2024-03-21 09:27 | Alaris pump spare parts | (TC 03/13/2024 10:23:25 AM, Rendon Lopez , Elizabeth) Alaris pumps spare parts for repair and stock Received
1070268 | 204694 | ACC | CMPLT | CMPLT | 2024-03-13 10:29 | 2024-03-13 13:33 | Incoming inspection | (TC 03/13/2024 01:32:40 PM, Canjimavadivel, Kamallesh ) Added to database, charged fully
66036 | 204695 | CM | CMPLT | CMPLT | 2024-03-13 10:42 | 2024-03-15 13:25 | Created by Failed PM WO # 201377 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 10:43:56 AM, Rhuland, Courtney) PM failed by BD Replaced rear and front case Replaced left IUI Replaced power cord retainer PM completed by BD
69324 | 204696 | PM | CMPLT | CMPLT | 2024-03-13 10:46 | 2024-03-13 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:46:08 AM, Cortez, John) Female IUI replaced PM complete
69455 | 204697 | PM | CMPLT | CMPLT | 2024-03-13 10:52 | 2024-03-13 10:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:52:12 AM, Cortez, John) Female IUI replaced PM complete
69550 | 204698 | PM | CMPLT | CMPLT | 2024-03-13 11:00 | 2024-03-13 11:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:59:42 AM, Cortez, John) Female IUI replaced PM complete
1059828 | 204699 | CM | CMPLT | CMPLT | 2024-03-13 13:01 | 2024-03-13 13:01 | Cleaning | (TC 03/13/2024 01:00:57 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
1063746 | 204700 | CM | CMPLT | CMPLT | 2024-03-13 13:01 | 2024-03-13 13:02 | Cleaning | (TC 03/13/2024 01:01:33 PM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water.
69375 | 204701 | CM | CMPLT | CMPLT | 2024-03-13 13:07 | 2024-03-13 13:08 | Created by Failed PM WO # 193968 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:07:30 PM, Baghirzade, Nara) Replaced bezel. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/13/2024 01:06:56 PM, Baghirzade, Nara) Failed rate calibration
65969 | 204702 | PM | CMPLT | PMFAIL | 2024-03-13 13:10 | 2024-03-13 13:10 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 01:10:15 PM, Baghirzade, Nara) Broken rear case
65969 | 204703 | CM | CMPLT | CMPLT | 2024-03-13 13:10 | 2024-03-13 13:11 | Created by Failed PM WO # 204702 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 01:10:42 PM, Baghirzade, Nara) Replaced rear case. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/13/2024 01:10:15 PM, Baghirzade, Nara) Broken rear case
69298 | 204704 | CM | CMPLT | CMPLT | 2024-03-13 13:12 | 2024-03-13 13:14 | Created by Failed PM WO # 198199 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:12:58 PM, Baghirzade, Nara) Replaced bezel. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/13/2024 01:11:39 PM, Baghirzade, Nara) Failed rate calibration.
68858 | 204705 | PM | CMPLT | PMFAIL | 2024-03-13 13:16 | 2024-03-13 13:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:16:04 PM, Baghirzade, Nara) Left IUI defective.
68858 | 204706 | CM | CMPLT | CMPLT | 2024-03-13 13:16 | 2024-03-13 13:17 | Created by Failed PM WO # 204705 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:16:34 PM, Baghirzade, Nara) Replaced left IUI. PM passed and completed. Battery was replaced by BD. Electrical Safety - Ground resistance: 144 mohms Leakage current: 71.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. | (TC 03/13/2024 01:16:04 PM, Baghirzade, Nara) Left IUI defective.
66028 | 204707 | CM | CMPLT | CMPLT | 2024-03-13 13:18 | 2024-03-13 13:19 | Missing power cord strap and accessories for the power cord strap | (TC 03/13/2024 01:17:59 PM, Baghirzade, Nara) Replaced power cord strap, strap screw, and power cord strap snap.
68752 | 204708 | PM | CMPLT | PMFAIL | 2024-03-13 13:21 | 2024-03-13 13:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:20:37 PM, Baghirzade, Nara) Right IUI defective
68752 | 204709 | CM | CMPLT | CMPLT | 2024-03-13 13:21 | 2024-03-13 13:22 | Created by Failed PM WO # 204708 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 01:21:07 PM, Baghirzade, Nara) Replaced right IUI. PM passed and completed. Battery was replaced by BD. Electrical Safety - Ground resistance: 142mohms Leakage current: 66.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. | (TC 03/13/2024 01:20:37 PM, Baghirzade, Nara) Right IUI defective
69228 | 204710 | PM | CMPLT | CMPLT | 2024-03-13 13:41 | 2024-03-13 13:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1060030 | 204711 | CM | CMPLT | CMPLT | 2024-03-13 13:42 | 2024-03-14 14:08 | Parts return | (TC 03/13/2024 01:42:50 PM, Baghirzade, Nara) Returned part to Chantal Gourlay at Novus as per her instructions. MS4746 part was incorrect for the bodybox and was not needed. Part returned for credit for cost: $990.00 (TC 03/14/2024 02:07:17 PM, Baghirzade, Nara) Attached credit document to the WO.
69525 | 204712 | PM | CMPLT | CMPLT | 2024-03-13 14:03 | 2024-03-13 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1062325 | 204713 | CM | CMPLT | CMPLT | 2024-03-13 14:05 | 2024-03-19 13:16 | Loud noise when raising the head of the bed | (TC 03/13/2024 02:05:16 PM, Baghirzade, Nara) Loud noise when raising the head portion of the bed. Requested service from Stryker (TC 03/19/2024 01:16:03 PM, Baghirzade, Nara) Attached service report. Bed is back in use.
| 204714 | CM | CMPLT | CMPLT | 2024-03-13 14:05 | 2024-11-04 14:59 | BD Alaris plastics recall part order, for biomed stock | (TC 03/13/2024 02:05:33 PM, Cortez, John) Ordered 20x front cases for syringe + PCU (TC 06/04/2024 05:23:13 PM, Cortez, John) Submitted form to BD, requested nocharge PO TC10013702 Front Case 10 49000459 Rear Case 20 49000218 Front Case 7 10013671 Rear Case 10 TC10007985 Left Handle 4 (TC 07/18/2024 09:59:42 AM, Cortez, John) Followed up with BD, parts haven't shipped yet (TC 10/30/2024 03:10:18 PM, Cortez, John) Followed up with BD (TC 11/04/2024 02:59:04 PM, Cortez, John) Remaining parts arrived
68914 | 204715 | PM | CMPLT | CMPLT | 2024-03-13 14:09 | 2024-03-13 14:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:09:02 PM, Baghirzade, Nara) PM completed by BD.
69027 | 204716 | PM | CMPLT | CMPLT | 2024-03-13 14:10 | 2024-03-13 14:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:10:20 PM, Baghirzade, Nara) Software recall completed by BD.
68864 | 204717 | PM | CMPLT | CMPLT | 2024-03-13 14:13 | 2024-03-13 14:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:13:20 PM, Baghirzade, Nara) PM completed by BD.
68816 | 204718 | PM | CMPLT | CMPLT | 2024-03-13 14:15 | 2024-03-13 14:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:14:41 PM, Baghirzade, Nara) PM completed by BD.
68867 | 204719 | PM | CMPLT | CMPLT | 2024-03-13 14:16 | 2024-03-13 14:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:16:17 PM, Baghirzade, Nara) PM completed by BD.
69089 | 204720 | PM | CMPLT | CMPLT | 2024-03-13 14:18 | 2024-03-13 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
65963 | 204721 | PM | CMPLT | CMPLT | 2024-03-13 14:20 | 2024-03-13 14:20 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 02:19:57 PM, Rhuland, Courtney) PM completed by BD
68923 | 204722 | PM | CMPLT | CMPLT | 2024-03-13 14:23 | 2024-03-13 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:22:54 PM, Baghirzade, Nara) PM completed by BD.
65987 | 204723 | PM | CMPLT | CMPLT | 2024-03-13 14:23 | 2024-03-13 14:26 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 02:23:11 PM, Rhuland, Courtney) PM completed by BD
68705 | 204724 | PM | CMPLT | CMPLT | 2024-03-13 14:24 | 2024-03-13 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:24:33 PM, Baghirzade, Nara) PM completed by BD.
65985 | 204725 | PM | CMPLT | CMPLT | 2024-03-13 14:27 | 2024-03-13 14:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 02:27:06 PM, Rhuland, Courtney) PM completed by BD
69167 | 204726 | PM | CMPLT | CMPLT | 2024-03-13 14:27 | 2024-03-13 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058279 | 204727 | PM | CMPLT | CMPLT | 2024-03-13 14:27 | 2024-03-13 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 02:27:22 PM, Baghirzade, Nara) PM completed by BD.
69553 | 204728 | PM | CMPLT | CMPLT | 2024-03-13 14:35 | 2024-03-13 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
65878 | 204729 | PM | CMPLT | CMPLT | 2024-03-13 14:36 | 2024-03-13 14:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 02:36:12 PM, Rhuland, Courtney) PM completed by BD
1050773 | 204730 | PM | CMPLT | CMPLT | 2024-03-13 14:42 | 2024-03-13 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1050617 | 204731 | PM | CMPLT | CMPLT | 2024-03-13 14:42 | 2024-03-13 14:42 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/13/2024 02:41:57 PM, Rhuland, Courtney) PM completed by BD
69402 | 204732 | PM | CMPLT | CMPLT | 2024-03-13 14:48 | 2024-03-13 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
52117 | 204733 | CM | CMPLT | NFF | 2024-03-13 14:49 | 2024-03-13 14:50 | Loud noise | (TC 03/13/2024 02:48:42 PM, Rhuland, Courtney) Noise was heard by SA but nurse reported no noise RH calibrated Functional check - OK No fault found
69037 | 204734 | PM | CMPLT | CMPLT | 2024-03-13 14:58 | 2024-03-13 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69062 | 204735 | PM | CMPLT | CMPLT | 2024-03-13 15:03 | 2024-03-13 15:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69225 | 204736 | PM | CMPLT | CMPLT | 2024-03-13 15:05 | 2024-03-13 15:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69039 | 204737 | PM | CMPLT | CMPLT | 2024-03-13 15:09 | 2024-03-13 15:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69374 | 204738 | CM | CANCL | DUP | 2024-03-13 15:21 |  | Created by Failed PM WO # 204615 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE |  | (TC 03/07/2024 10:50:20 AM, Cortez, John) Patient side occlusion failed
69374 | 204739 | PM | CMPLT | CMPLT | 2024-03-13 15:22 | 2024-03-13 15:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/13/2024 03:21:52 PM, Cortez, John) Calibrated pressure PM complete
| 204740 | CM | CMPLT | CMPLT | 2024-03-13 15:34 | 2024-03-13 15:34 | meeting | meeting
| 204741 | CM | ACTIV | ISSUE | 2024-03-13 15:50 |  | Meetings/Training | (TC 03/13/2024 03:50:34 PM, Patel, Rutvik) OR Asset Division meeting with OR Team, AA and SC Searched through TMS OR Asset inventory to understand how they are organized (TC 03/15/2024 02:39:27 PM, Patel, Rutvik) Meeting with AC, ER and JH about Stryker Neptune unit. (TC 03/18/2024 03:57:22 PM, Patel, Rutvik) Biomed weekly huddle (TC 03/19/2024 03:17:28 PM, Patel, Rutvik) OR Asset distribution meeting with AA, SC, ER and JH. (TC 03/20/2024 03:21:45 PM, Patel, Rutvik) OR Equipment Repair meeting with AC, SC, ER, JH CMBES Peer review meeting Meeting with Anita to discuss endo cart CO2 tank regulator. (TC 03/22/2024 03:14:42 PM, Patel, Rutvik) CMBES International Outreach Committee meeting (TC 03/25/2024 03:23:04 PM, Patel, Rutvik) Meetign with ER and BN for TMS superuser fuctionality (TC 03/26/2024 03:52:09 PM, Patel, Rutvik) MMC E REQ training webinar (TC 04/03/2024 03:29:14 PM, Patel, Rutvik) TAAAC Meeting with AA. (TC 04/11/2024 04:02:40 PM, Patel, Rutvik) Meeting for OR walkthrough (TC 04/15/2024 04:23:04 PM, Patel, Rutvik) Endo monthly Meeting (TC 04/16/2024 04:07:54 PM, Patel, Rutvik) Biomed Monthyl Meeting 12 N Meeting with NB - 6th floor storage visit (TC 04/17/2024 03:41:54 PM, Patel, Rutvik) Departmental Meeting with MSH Biomeds CMBES Peer Review Meeting (TC 04/22/2024 04:22:13 PM, Patel, Rutvik) Biomed Huddle (TC 04/24/2024 02:25:46 PM, Patel, Rutvik) OR PM Meeting Biomed PM WO Meeting 12N Alaris implementation meeting (TC 04/26/2024 02:48:47 PM, Patel, Rutvik) 12 N walthrough tour CMBES Outreach Committee meeting (TC 04/29/2024 03:42:04 PM, Patel, Rutvik) Monday Huddle 12 N Walk through (TC 05/03/2024 03:47:58 PM, Patel, Rutvik) On call overview with KS. (TC 05/10/2024 01:53:17 PM, Patel, Rutvik) In service training for Olympus Laser. (TC 05/13/2024 03:51:18 PM, Patel, Rutvik) Biomed huddle (TC 05/15/2024 03:47:59 PM, Patel, Rutvik) CMBES Peer review meeting (TC 05/17/2024 01:54:21 PM, Patel, Rutvik) iECG configuration meeting for 12N Conmed Airseal Meeting (TC 05/21/2024 03:25:29 PM, Patel, Rutvik) OR Monthly check up meeting (TC 05/28/2024 10:03:26 AM, Patel, Rutvik) CMBES-CESO Conference. (TC 06/03/2024 03:12:33 PM, Patel, Rutvik) Biomed Huddle (TC 06/12/2024 02:19:01 PM, Patel, Rutvik) Departmental meeting with NE and team. (TC 06/14/2024 03:54:33 PM, Patel, Rutvik) OR accreditation meeting (TC 06/21/2024 11:08:58 AM, Patel, Rutvik) IS Division meeting (TC 06/24/2024 01:51:46 PM, Patel, Rutvik) OR Asset meeting (TC 07/09/2024 02:47:13 PM, Patel, Rutvik) Biomed Huddle (TC 07/10/2024 06:13:18 PM, Patel, Rutvik) CMBES Peer review Meeting Started Welch Allyn VSM training (TC 07/17/2024 03:25:58 PM, Patel, Rutvik) OR PM meeting (TC 07/22/2024 03:46:08 PM, Patel, Rutvik) Biomed Huddle (TC 07/24/2024 08:39:09 AM, Patel, Rutvik) PM Huddle (TC 07/23/2024 01:39:00 PM, Patel, Rutvik) PM Huddle (TC 07/29/2024 08:12:21 AM, Patel, Rutvik) Team meeting with Nima and team (TC 07/30/2024 03:24:33 PM, Patel, Rutvik) FT10 bipolar resection meeting and demo with OR staff. (TC 07/30/2024 03:24:33 PM, Patel, Rutvik) FT10 bipolar resection meeting and demo with OR staff. (TC 08/07/2024 03:29:00 PM, Patel, Rutvik) Finalized OR Asset inventory. OR Asset PM workflow meeting. (TC 08/09/2024 06:41:09 PM, Patel, Rutvik) Started WA CVSM online Training (TC 08/09/2024 06:41:09 PM, Patel, Rutvik) complete two modules of WA CVSM online Training (TC 08/12/2024 03:56:06 PM, Patel, Rutvik) Monday Huddle (TC 08/12/2024 03:56:28 PM, Patel, Rutvik) Completed WE CVSM 6000 training., (TC 08/19/2024 03:44:55 PM, Patel, Rutvik) Biomed Huddle (TC 08/21/2024 11:11:00 AM, Patel, Rutvik) Department meeting with NE and Team (TC 08/22/2024 01:37:52 PM, Patel, Rutvik) Patient monitoring games meeting (TC 09/23/2024 03:04:00 PM, Patel, Rutvik) Biomed huddle OR PM meeting OR 12 operational meeting (TC 09/27/2024 01:38:41 PM, Patel, Rutvik) CMBES Outreach committee meeting. CMBES Webinar (TC 09/30/2024 03:48:39 PM, Patel, Rutvik) Biomed Huddle (TC 10/07/2024 03:26:38 PM, Patel, Rutvik) Biomed Huddle (TC 10/25/2024 01:54:00 PM, Patel, Rutvik) CMBES Outreach committee meeting (TC 10/28/2024 04:07:58 PM, Patel, Rutvik) Stryker Lifenet meeting (TC 10/31/2024 11:15:07 AM, Patel, Rutvik) Meetign with SC about Bed PM and Battery replacement schedules for 12N and 9S. (TC 11/04/2024 10:05:56 AM, Patel, Rutvik) Biomed Huddle and Laser discussion (TC 11/07/2024 02:28:09 PM, Patel, Rutvik) OR 12 Operationalize Meeting (TC 11/07/2024 02:28:34 PM, Patel, Rutvik) MMC E REQ Training (TC 11/18/2024 04:03:45 PM, Patel, Rutvik) OR PM Meeting Endoscopy Meeting OR 12 Operational Meeting Biomed Huddle (TC 11/25/2024 03:42:41 PM, Patel, Rutvik) Biomed huddle (TC 11/27/2024 04:15:39 PM, Patel, Rutvik) Dialysis Machine training (TC 11/27/2024 04:15:58 PM, Patel, Rutvik) CMBES Peer review meeting
1056407 | 204742 | CM | CMPLT | CMPLT | 2024-03-13 16:39 | 2024-03-13 16:40 | "not working" | (TC 03/13/2024 04:39:08 PM, Cortez, John) SEC mode was on, switched it off Everything else tested ok
67106 | 204743 | CM | CMPLT | CMPLT | 2024-03-13 16:50 | 2024-03-13 16:52 | "BP not working" | (TC 03/13/2024 04:50:26 PM, Cortez, John) Cuff is worn, replaced Temp probe damaged, replaced All parameters tested ok
65918 | 204744 | PM | CMPLT | CMPLT | 2024-03-14 07:19 | 2024-03-14 07:19 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 07:18:55 AM, Rhuland, Courtney) PM completed by BD
65976 | 204745 | PM | CMPLT | PMFAIL | 2024-03-14 07:39 | 2024-03-14 07:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed by BD
65976 | 204746 | CM | CMPLT | CMPLT | 2024-03-14 07:39 | 2024-03-14 07:40 | Created by Failed PM WO # 204745 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 07:39:21 AM, Rhuland, Courtney) Rear case replaced Left handle replaced PM completed All tests passed Refer to ASM for details
1050258 | 204747 | PM | CMPLT | CMPLT | 2024-03-14 07:58 | 2024-03-14 07:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058164 | 204748 | PM | CMPLT | CMPLT | 2024-03-14 08:07 | 2024-03-14 08:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058245 | 204749 | PM | CMPLT | CMPLT | 2024-03-14 08:09 | 2024-03-14 08:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
| 204750 | PA | CMPLT | CMPLT | 2024-03-14 08:23 | 2024-03-25 13:38 | Ordering LVP Bezels for stock | (TC 03/14/2024 08:24:44 AM, Baghirzade, Nara) SC 1003189703 (TC 03/25/2024 01:37:29 PM, Baghirzade, Nara) PO 7202376499
66540 | 204751 | CM | CMPLT | CMPLT | 2024-03-14 08:39 | 2024-03-14 08:39 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66541 | 204752 | CM | CMPLT | CMPLT | 2024-03-14 08:41 | 2024-03-14 08:42 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66542 | 204753 | CM | CMPLT | CMPLT | 2024-03-14 08:42 | 2024-03-14 08:43 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66543 | 204754 | CM | CMPLT | CMPLT | 2024-03-14 08:43 | 2024-03-14 08:43 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66544 | 204755 | CM | CMPLT | CMPLT | 2024-03-14 08:44 | 2024-03-14 08:44 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66545 | 204756 | CM | CMPLT | CMPLT | 2024-03-14 08:45 | 2024-03-14 08:45 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66546 | 204757 | CM | CMPLT | CMPLT | 2024-03-14 08:46 | 2024-03-14 08:46 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66574 | 204758 | CM | CMPLT | CMPLT | 2024-03-14 08:49 | 2024-03-14 08:49 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66566 | 204759 | CM | CMPLT | CMPLT | 2024-03-14 08:50 | 2024-03-14 08:50 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66567 | 204760 | CM | CMPLT | CMPLT | 2024-03-14 08:51 | 2024-03-14 08:51 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66568 | 204761 | CM | CMPLT | CMPLT | 2024-03-14 08:51 | 2024-03-14 08:52 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66569 | 204762 | CM | CMPLT | CMPLT | 2024-03-14 08:52 | 2024-03-14 08:52 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66570 | 204763 | CM | CMPLT | CMPLT | 2024-03-14 08:53 | 2024-03-14 08:53 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66571 | 204764 | CM | CMPLT | CMPLT | 2024-03-14 08:54 | 2024-03-14 08:54 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66572 | 204765 | CM | CMPLT | CMPLT | 2024-03-14 08:54 | 2024-03-14 08:54 | SW upgraded to Rev.N by Philips | SW upgraded to Rev.N by Philips
66547 | 204766 | CM | CMPLT | CMPLT | 2024-03-14 09:00 | 2024-03-14 09:00 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66548 | 204767 | CM | CMPLT | CMPLT | 2024-03-14 09:00 | 2024-03-14 09:00 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66550 | 204768 | CM | CMPLT | CMPLT | 2024-03-14 09:04 | 2024-03-14 09:04 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66551 | 204769 | CM | CMPLT | CMPLT | 2024-03-14 09:04 | 2024-03-14 09:05 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66552 | 204770 | CM | CMPLT | CMPLT | 2024-03-14 09:05 | 2024-03-14 09:05 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66554 | 204771 | CM | CMPLT | CMPLT | 2024-03-14 09:05 | 2024-03-14 09:06 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66555 | 204772 | CM | CMPLT | CMPLT | 2024-03-14 09:06 | 2024-03-14 09:06 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66557 | 204773 | CM | CMPLT | CMPLT | 2024-03-14 09:07 | 2024-03-14 09:07 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66559 | 204774 | CM | CMPLT | CMPLT | 2024-03-14 09:08 | 2024-03-14 09:08 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66562 | 204775 | CM | CMPLT | CMPLT | 2024-03-14 09:09 | 2024-03-14 09:09 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66563 | 204776 | CM | CMPLT | CMPLT | 2024-03-14 09:09 | 2024-03-14 09:10 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66565 | 204777 | CM | CMPLT | CMPLT | 2024-03-14 09:10 | 2024-03-14 09:10 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66864 | 204778 | CM | CMPLT | CMPLT | 2024-03-14 09:12 | 2024-03-14 09:12 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66868 | 204779 | CM | CMPLT | CMPLT | 2024-03-14 09:12 | 2024-03-14 09:12 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66869 | 204780 | CM | CMPLT | CMPLT | 2024-03-14 09:13 | 2024-03-14 09:13 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
67817 | 204781 | CM | CMPLT | CMPLT | 2024-03-14 09:14 | 2024-03-14 09:14 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
67821 | 204782 | CM | CMPLT | CMPLT | 2024-03-14 09:15 | 2024-03-14 09:15 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
67823 | 204783 | CM | CMPLT | CMPLT | 2024-03-14 09:15 | 2024-03-14 09:15 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
67825 | 204784 | CM | CMPLT | CMPLT | 2024-03-14 09:16 | 2024-03-14 09:16 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
65744 | 204785 | CM | CMPLT | CMPLT | 2024-03-14 09:20 | 2024-03-14 09:20 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
66870 | 204786 | CM | CMPLT | CMPLT | 2024-03-14 09:21 | 2024-03-14 09:22 | SW upgraded to Rev.M by Philips | SW upgraded to Rev.M by Philips
68896 | 204787 | PM | CMPLT | CMPLT | 2024-03-14 09:35 | 2024-03-14 09:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 09:34:53 AM, Huang, Jessica) replaced the female iui connector and completed PM
68896 | 204788 | PM | CANCL | CANCL | 2024-03-14 09:35 | 2024-03-14 09:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68889 | 204789 | PM | CMPLT | CMPLT | 2024-03-14 09:42 | 2024-03-14 09:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 09:42:05 AM, Huang, Jessica) replaced the female iui connector, completed PM
68851 | 204790 | PM | CMPLT | CMPLT | 2024-03-14 09:47 | 2024-03-14 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 09:52:59 AM, Huang, Jessica) replaced power cord wrap and completed PM
68850 | 204791 | PM | CMPLT | CMPLT | 2024-03-14 09:56 | 2024-03-14 09:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 09:57:34 AM, Huang, Jessica) replaced the female iui connector, completed PM
69444 | 204792 | PM | CMPLT | CMPLT | 2024-03-14 10:18 | 2024-03-14 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 10:19:05 AM, Huang, Jessica) replaced the female iui connector, completed PM
69523 | 204793 | PM | CMPLT | CMPLT | 2024-03-14 10:31 | 2024-03-14 10:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 10:31:54 AM, Huang, Jessica) replaced the female iui connector and completed PM
1050622 | 204794 | PM | CMPLT | PMFAIL | 2024-03-14 10:40 | 2024-03-14 10:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 10:40:13 AM, Rendon Lopez , Elizabeth) PM fails force sensor assembly test.
1050622 | 204795 | CM | CMPLT | CMPLT | 2024-03-14 10:41 | 2024-03-21 15:23 | Created by Failed PM WO # 204794 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 03:22:05 PM, Rendon Lopez , Elizabeth) Replaced plunger head assembly. Calibrated syringe. PM completed, all tests passed | (TC 03/14/2024 10:40:13 AM, Rendon Lopez , Elizabeth) PM fails force sensor assembly test.
| 204796 | CM | ACTIV | ISSUE | 2024-03-14 10:43 |  | MSH OPG SPP GE VOLUSON US SYSTEMS DID NOT AUTOMATICALLY UPDATE TO DAYLIGHT SAVINGS TIME |  | MSH OPG SPP GE VOLUSON US SYSTEMS DID NOT AUTOMATICALLY UPDATE TO DAYLIGHT SAVINGS TIME VASSI MANUALLY UPDATED
69507 | 204797 | PM | CMPLT | CMPLT | 2024-03-14 10:47 | 2024-03-14 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 10:47:43 AM, Huang, Jessica) replaced the male iui connector and completed PM
502088 | 204798 | RQ | CMPLT | CMPLT | 2024-03-14 11:14 | 2024-03-20 11:11 | ct quality check problem | (TC 03/19/2024 01:52:12 PM, Green, Paul) Travelled to site. Inspected CT phantom and holder. Found that the bracket was bent. Brought back to TGH shop and disassembled phantom and bracket. Was able to bend back both bracket afixed to phantom, and PHS interface L bracket. Reassebled and took back to MSH. (TC 03/20/2024 11:10:01 AM, Green, Paul) Checked with customer and CT QC had no problems. | during ct quality check consistent get prompted to hit MOVE button before hitting START
1059938 | 204799 | CM | CMPLT | CMPLT | 2024-03-14 11:34 | 2024-03-14 11:41 | Not connecting to network | (TC 03/14/2024 11:38:12 AM, Canjimavadivel, Kamallesh ) Was not connecting with network. Reloaded network configuration. Verified all inboxes were assigned. Moved to different area of workshop and saw connection. Compared with another device and noticed similar signal strength. Removed and reseated network card. Able to connect and time sync.
1054113 | 204800 | CM | CMPLT | CMPLT | 2024-03-14 11:54 | 2024-03-14 11:55 | Upgraded SW rev to check Fetalink connection | (TC 03/14/2024 11:54:22 AM, Rendon Lopez , Elizabeth) Upgraded SV Rev from L33.61 to L.33.65 to check connectivity to Fetalink. Upgrade and connection were successful
65919 | 204801 | PM | CMPLT | CMPLT | 2024-03-14 13:22 | 2024-03-14 13:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 01:22:27 PM, Baghirzade, Nara) PM completed by BD.
65897 | 204802 | PM | CMPLT | CMPLT | 2024-03-14 13:24 | 2024-03-14 13:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 01:23:56 PM, Baghirzade, Nara) PM completed by BD.
65944 | 204803 | PM | CMPLT | CMPLT | 2024-03-14 13:25 | 2024-03-14 13:26 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 01:25:16 PM, Baghirzade, Nara) PM completed by BD.
69348 | 204804 | PM | CMPLT | CMPLT | 2024-03-14 13:26 | 2024-03-14 13:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 01:26:30 PM, Huang, Jessica) replaced the female iui connector and completed PM
68589 | 204805 | CM | CMPLT | CMPLT | 2024-03-14 13:26 | 2024-03-14 13:27 | Broken rear case, left handle, and defective IUIs | (TC 03/14/2024 01:26:29 PM, Baghirzade, Nara) Replaced rear case, left handle, right and left IUI. Returned to BD for upgrade and PM.
65965 | 204806 | PM | CMPLT | CMPLT | 2024-03-14 13:28 | 2024-03-14 13:28 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 01:27:58 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69530 | 204807 | PM | CMPLT | PMFAIL | 2024-03-14 13:31 | 2024-03-14 13:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69530 | 204808 | CM | CMPLT | CMPLT | 2024-03-14 13:32 | 2024-03-14 13:34 | Created by Failed PM WO # 204807 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 01:33:15 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
68809 | 204809 | PM | CMPLT | PMFAIL | 2024-03-14 13:35 | 2024-03-14 13:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68809 | 204810 | CM | CMPLT | CMPLT | 2024-03-14 13:36 | 2024-03-14 13:36 | Created by Failed PM WO # 204809 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 01:36:06 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69387 | 204811 | PM | CMPLT | PMFAIL | 2024-03-14 13:38 | 2024-03-14 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69387 | 204812 | CM | CMPLT | CMPLT | 2024-03-14 13:38 | 2024-03-14 13:40 | Created by Failed PM WO # 204811 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 01:39:33 PM, Rendon Lopez , Elizabeth) Replaced both IUI. PM completed, all tests passed | PM failed by BD
69201 | 204813 | PM | CMPLT | PMFAIL | 2024-03-14 13:41 | 2024-03-14 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69201 | 204814 | CM | CMPLT | CMPLT | 2024-03-14 13:42 | 2024-03-14 13:43 | Created by Failed PM WO # 204813 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/04/2024 01:42:35 PM, Rendon Lopez , Elizabeth) Replaced left IUI. and front case/keypad PM completed, all tests passed | PM failed by BD
66048 | 204815 | PM | CMPLT | PMFAIL | 2024-03-14 13:46 | 2024-03-14 13:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
66048 | 204816 | CM | CMPLT | CMPLT | 2024-03-14 13:46 | 2024-03-14 13:50 | Created by Failed PM WO # 204815 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 01:46:53 PM, Rendon Lopez , Elizabeth) Replaced rear case and handle. PM completed, all tests passed | PM failed by BD
65877 | 204817 | CM | CMPLT | CMPLT | 2024-03-14 13:46 | 2024-03-14 13:47 | Cracked rear case | (TC 03/14/2024 01:46:05 PM, Rhuland, Courtney) Rear case replaced Left handle replaced Returned to BD for software recall and PM
69531 | 204818 | PM | CMPLT | PMFAIL | 2024-03-14 13:52 | 2024-03-14 13:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69531 | 204819 | CM | CMPLT | CMPLT | 2024-03-14 13:52 | 2024-03-14 13:53 | Created by Failed PM WO # 204818 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 01:52:34 PM, Rendon Lopez , Elizabeth) Replaced both IUI. PM completed, all tests passed | PM failed by BD
1058259 | 204820 | PM | CMPLT | CMPLT | 2024-03-14 13:55 | 2024-03-14 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 01:55:07 PM, Huang, Jessica) replaced the female iui connector and completed PM
66011 | 204821 | CM | CMPLT | CMPLT | 2024-03-14 13:55 | 2024-03-14 13:56 | Created by Failed PM WO # 193878 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 01:55:25 PM, Rendon Lopez , Elizabeth) Replaced right IUI. PM completed, all tests passed | PM failed by BD
69656 | 204822 | CM | CMPLT | CMPLT | 2024-03-14 13:58 | 2024-03-14 13:59 | Created by Failed PM WO # 201436 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 01:58:32 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
65996 | 204823 | PM | CMPLT | CMPLT | 2024-03-14 14:00 | 2024-03-14 14:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065889 | 204824 | CM | CMPLT | CMPLT | 2024-03-14 14:01 | 2024-03-14 14:02 | Created by Failed PM WO # 199932 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 02:01:20 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69675 | 204825 | CM | CMPLT | CMPLT | 2024-03-14 14:04 | 2024-03-14 14:06 | Created by Failed PM WO # 199577 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 02:05:14 PM, Rendon Lopez , Elizabeth) Replaced left IUI. Front case/keypad assembly PM completed, all tests passed | PM failed by BD
69612 | 204826 | PM | CMPLT | CMPLT | 2024-03-14 14:04 | 2024-03-14 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:04:06 PM, Huang, Jessica) replaced the male iui connector and completed PM
66047 | 204827 | PM | CMPLT | CMPLT | 2024-03-14 14:05 | 2024-03-14 14:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69590 | 204828 | CM | CMPLT | CMPLT | 2024-03-14 14:08 | 2024-03-14 14:10 | Damaged front case | (TC 03/05/2024 02:08:21 PM, Rendon Lopez , Elizabeth) Repairs before recall. Replaced front case/Keypad, LCD assy, both IUI.
68618 | 204829 | PM | CMPLT | PMFAIL | 2024-03-14 14:15 | 2024-03-14 14:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68618 | 204830 | CM | CMPLT | CMPLT | 2024-03-14 14:15 | 2024-03-14 14:17 | Created by Failed PM WO # 204829 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:15:19 PM, Rendon Lopez , Elizabeth) Replaced left IUI. RJ45 board PM completed, all tests passed | PM failed by BD
66070 | 204831 | PM | CMPLT | CMPLT | 2024-03-14 14:16 | 2024-03-14 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66018 | 204832 | PM | CMPLT | CMPLT | 2024-03-14 14:19 | 2024-03-14 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66061 | 204833 | CM | CMPLT | CMPLT | 2024-03-14 14:20 | 2024-03-14 14:21 | Created by Failed PM WO # 204472 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 02:20:38 PM, Rendon Lopez , Elizabeth) Replaced front case. PM completed, all tests passed | PM failed by BD
68293 | 204834 | PM | CMPLT | CMPLT | 2024-03-14 14:22 | 2024-03-14 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69358 | 204835 | PM | CMPLT | PMFAIL | 2024-03-14 14:25 | 2024-03-14 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69358 | 204836 | CM | CMPLT | CMPLT | 2024-03-14 14:25 | 2024-03-14 14:27 | Created by Failed PM WO # 204835 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 02:25:58 PM, Rendon Lopez , Elizabeth) Replaced door latch. PM completed, all tests passed | PM failed by BD
66026 | 204837 | PM | CMPLT | CMPLT | 2024-03-14 14:32 | 2024-03-14 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69403 | 204838 | PM | CMPLT | CMPLT | 2024-03-14 14:33 | 2024-03-14 14:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1056583 | 204839 | CM | CMPLT | CMPLT | 2024-03-14 14:35 | 2024-03-14 14:38 | Cabinet battery depleted | (TC 03/14/2024 02:37:22 PM, Rhuland, Courtney) Replaced 9V battery in cabinet Alarm test - OK
65999 | 204840 | PM | CMPLT | CMPLT | 2024-03-14 14:35 | 2024-03-14 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69282 | 204841 | PM | CMPLT | CMPLT | 2024-03-14 14:35 | 2024-03-14 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:35:51 PM, Huang, Jessica) replaced the female iui connector and completed PM
69490 | 204842 | PM | CMPLT | CMPLT | 2024-03-14 14:37 | 2024-03-14 14:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68769 | 204843 | PM | CMPLT | PMFAIL | 2024-03-14 14:38 | 2024-03-14 14:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68769 | 204844 | CM | CMPLT | CMPLT | 2024-03-14 14:38 | 2024-03-14 14:41 | Created by Failed PM WO # 204843 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:38:32 PM, Rendon Lopez , Elizabeth) Replaced RJ45 board, and power cord. PM completed, all tests passed | PM failed by BD
69399 | 204845 | PM | CMPLT | CMPLT | 2024-03-14 14:40 | 2024-03-14 14:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66067 | 204846 | PM | CMPLT | CMPLT | 2024-03-14 14:41 | 2024-03-14 14:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69261 | 204847 | PM | CMPLT | CMPLT | 2024-03-14 14:42 | 2024-03-14 14:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68971 | 204848 | PM | CMPLT | PMFAIL | 2024-03-14 14:45 | 2024-03-14 14:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68971 | 204849 | CM | CMPLT | CMPLT | 2024-03-14 14:45 | 2024-03-14 14:46 | Created by Failed PM WO # 204848 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 02:45:36 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
1048527 | 204850 | PM | CMPLT | CMPLT | 2024-03-14 14:47 | 2024-03-14 14:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 02:47:24 PM, Baghirzade, Nara) PM passed and completed. Replaced rear case. Reprinted a new barcode SN sticker. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69468 | 204851 | PM | CMPLT | PMFAIL | 2024-03-14 14:47 | 2024-03-14 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69468 | 204852 | CM | CMPLT | CMPLT | 2024-03-14 14:48 | 2024-03-14 14:48 | Created by Failed PM WO # 204851 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 02:47:54 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
50409 | 204853 | CM | CMPLT | CMPLT | 2024-03-14 14:49 | 2024-03-14 14:50 | Missing power cord | (TC 03/14/2024 02:49:13 PM, Rhuland, Courtney) Replaced power cord with spare Electrical safety - OK
69088 | 204854 | CM | CMPLT | CMPLT | 2024-03-14 14:49 | 2024-03-14 14:51 | Created by Failed PM WO # 201413 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/05/2024 02:50:52 PM, Rendon Lopez , Elizabeth) Replaced both IUI. PM completed, all tests passed | PM failed by BD
68289 | 204855 | PM | CMPLT | CMPLT | 2024-03-14 14:50 | 2024-03-14 14:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69390 | 204856 | PM | CMPLT | CMPLT | 2024-03-14 14:51 | 2024-03-14 14:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68857 | 204857 | PM | CMPLT | PMFAIL | 2024-03-14 14:54 | 2024-03-14 14:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68857 | 204858 | CM | CMPLT | CMPLT | 2024-03-14 14:54 | 2024-03-14 14:55 | Created by Failed PM WO # 204857 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:54:29 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
68906 | 204859 | PM | CMPLT | CMPLT | 2024-03-14 14:56 | 2024-03-14 14:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69496 | 204860 | CM | CMPLT | CMPLT | 2024-03-14 14:57 | 2024-03-14 14:57 | Created by Failed PM WO # 204470 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 02:56:57 PM, Rendon Lopez , Elizabeth) Replaced left IUI. cleaned module latch PM completed, all tests passed | PM failed by BD
68984 | 204861 | PM | CMPLT | CMPLT | 2024-03-14 15:00 | 2024-03-14 15:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69494 | 204862 | PS | CMPLT | CMPLT | 2024-03-14 15:01 | 2024-03-14 15:01 | 2024 Alaris Software Recall | Recall completed
69494 | 204863 | CM | CMPLT | CMPLT | 2024-03-14 15:02 | 2024-03-14 15:02 | Created by Failed PM WO # 204469 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 03:02:11 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
68903 | 204864 | PM | CMPLT | CMPLT | 2024-03-14 15:02 | 2024-03-14 15:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050733 | 204865 | CM | CMPLT | CMPLT | 2024-03-14 15:03 | 2024-03-14 15:05 | No extracting smoke | (TC 03/06/2024 03:03:23 PM, Rendon Lopez , Elizabeth) Device in OR 4. Found activation cable damaged, replaced from their stock. Tested-Ok
66021 | 204866 | CM | CMPLT | CMPLT | 2024-03-14 15:06 | 2024-03-14 15:10 | Faulty power cord | (TC 03/06/2024 03:08:26 PM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced power cord
69079 | 204867 | PM | CMPLT | CMPLT | 2024-03-14 15:11 | 2024-03-14 15:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66021 | 204868 | PM | CMPLT | CMPLT | 2024-03-14 15:11 | 2024-03-14 15:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM completed by BD
69609 | 204869 | PM | CMPLT | PMFAIL | 2024-03-14 15:13 | 2024-03-14 15:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68811 | 204870 | PM | CMPLT | CMPLT | 2024-03-14 15:13 | 2024-03-14 15:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69609 | 204871 | CM | CMPLT | CMPLT | 2024-03-14 15:14 | 2024-03-14 15:14 | Created by Failed PM WO # 204869 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 03:13:51 PM, Rendon Lopez , Elizabeth) Replaced left IUI. Cleaned module latch PM completed, all tests passed | PM failed by BD
69573 | 204872 | CM | CMPLT | CMPLT | 2024-03-14 15:16 | 2024-03-14 15:16 | Created by Failed PM WO # 201434 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 03:16:08 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
68863 | 204873 | PM | CMPLT | CMPLT | 2024-03-14 15:17 | 2024-03-14 15:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68944 | 204874 | PM | CMPLT | CMPLT | 2024-03-14 15:24 | 2024-03-14 15:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68680 | 204875 | PM | CMPLT | PMFAIL | 2024-03-14 15:35 | 2024-03-14 15:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68680 | 204876 | CM | CMPLT | CMPLT | 2024-03-14 15:35 | 2024-03-14 15:37 | Created by Failed PM WO # 204875 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 03:36:36 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69120 | 204877 | PM | CMPLT | PMFAIL | 2024-03-14 15:38 | 2024-03-14 15:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69120 | 204878 | CM | CMPLT | CMPLT | 2024-03-14 15:39 | 2024-03-14 15:39 | Created by Failed PM WO # 204877 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 03:38:49 PM, Rendon Lopez , Elizabeth) Replaced both IUI. PM completed, all tests passed | PM failed by BD
| 204879 | CM | CMPLT | CMPLT | 2024-03-14 15:42 | 2024-03-14 15:42 | meeting | meeting
| 204880 | CM | CMPLT | CMPLT | 2024-03-14 15:43 | 2024-03-14 15:43 | worked with RT | worked with RT
69430 | 204881 | PM | CMPLT | PMFAIL | 2024-03-14 15:45 | 2024-03-14 15:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69430 | 204882 | CM | CMPLT | CMPLT | 2024-03-14 15:46 | 2024-03-14 15:46 | Created by Failed PM WO # 204881 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 03:46:02 PM, Rendon Lopez , Elizabeth) Replaced both IUI. PM completed, all tests passed | PM failed by BD
1055117 | 204883 | PM | CMPLT | PMFAIL | 2024-03-14 15:48 | 2024-03-14 15:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1055117 | 204884 | CM | CMPLT | CMPLT | 2024-03-14 15:49 | 2024-03-14 15:49 | Created by Failed PM WO # 204883 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 03:48:53 PM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69425 | 204885 | PM | CMPLT | PMFAIL | 2024-03-14 15:52 | 2024-03-14 15:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69425 | 204886 | CM | CMPLT | CMPLT | 2024-03-14 15:52 | 2024-03-14 15:53 | Created by Failed PM WO # 204885 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 03:52:41 PM, Rendon Lopez , Elizabeth) Replaced both IUI. PM completed, all tests passed | PM failed by BD
69654 | 204887 | PM | CMPLT | CMPLT | 2024-03-14 15:56 | 2024-03-14 15:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69278 | 204888 | PM | CMPLT | CMPLT | 2024-03-14 16:20 | 2024-03-14 16:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65936 | 204889 | PA | CMPLT | CMPLT | 2024-03-15 08:22 | 2024-03-25 13:52 | Defective plunger head; ordering plungerhead assembly | (TC 03/15/2024 08:24:43 AM, Baghirzade, Nara) SC 1003190616 (TC 03/25/2024 01:52:01 PM, Baghirzade, Nara) PO 7500203710
69347 | 204890 | PM | CMPLT | PMFAIL | 2024-03-15 09:20 | 2024-03-15 09:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69347 | 204891 | CM | CMPLT | CMPLT | 2024-03-15 09:20 | 2024-03-15 09:21 | Created by Failed PM WO # 204890 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 09:20:42 AM, Rendon Lopez , Elizabeth) Replaced both IUI. Cleaned module latch PM completed, all tests passed | PM failed by BD
68865 | 204892 | PM | CMPLT | PMFAIL | 2024-03-15 09:23 | 2024-03-15 09:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68865 | 204893 | CM | CMPLT | CMPLT | 2024-03-15 09:23 | 2024-03-15 09:24 | Created by Failed PM WO # 204892 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 09:23:13 AM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69194 | 204894 | PM | CMPLT | PMFAIL | 2024-03-15 09:25 | 2024-03-15 09:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69194 | 204895 | CM | CMPLT | CMPLT | 2024-03-15 09:28 | 2024-03-15 09:29 | Created by Failed PM WO # 204894 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 09:28:45 AM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69602 | 204896 | PM | CMPLT | PMFAIL | 2024-03-15 09:30 | 2024-03-15 09:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69602 | 204897 | CM | CMPLT | CMPLT | 2024-03-15 09:31 | 2024-03-15 09:31 | Created by Failed PM WO # 204896 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/06/2024 09:30:47 AM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69058 | 204898 | CM | CMPLT | CMPLT | 2024-03-15 09:33 | 2024-03-15 09:33 | Created by Failed PM WO # 202149 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:33:02 AM, Rendon Lopez , Elizabeth) Replaced both IUI. PM completed, all tests passed | PM failed by BD
68297 | 204899 | PM | CMPLT | PMFAIL | 2024-03-15 09:45 | 2024-03-15 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68297 | 204900 | CM | CMPLT | CMPLT | 2024-03-15 09:45 | 2024-03-15 09:46 | Created by Failed PM WO # 204899 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:45:25 AM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed | PM failed by BD
69019 | 204901 | PM | CMPLT | PMFAIL | 2024-03-15 09:47 | 2024-03-15 09:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1061216 | 204902 | CM | CMPLT | CMPLT | 2024-03-15 09:47 | 2024-03-15 09:49 | Monitors flicker while changing to 4K source | (TC 03/15/2024 09:47:12 AM, Reid, Emmerson ) Checked out the room, and the monitors settled after a few seconds. No problem found.
69019 | 204903 | CM | CMPLT | CMPLT | 2024-03-15 09:47 | 2024-03-15 09:48 | Created by Failed PM WO # 204901 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:47:41 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
1061210 | 204904 | CM | CMPLT | CMPLT | 2024-03-15 09:49 | 2024-03-15 09:54 | Monitors are loose and tapped | (TC 03/15/2024 09:50:09 AM, Reid, Emmerson ) All of the moniors were loose. I tightened them and removed the tape. Working now.
69406 | 204905 | PM | CMPLT | PMFAIL | 2024-03-15 09:49 | 2024-03-15 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69406 | 204906 | CM | CMPLT | CMPLT | 2024-03-15 09:50 | 2024-03-15 09:50 | Created by Failed PM WO # 204905 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:50:06 AM, Rendon Lopez , Elizabeth) Cleaned IUIs and moldule latch PM completed, all tests passed | PM failed by BD
68715 | 204907 | PM | CMPLT | PMFAIL | 2024-03-15 09:52 | 2024-03-15 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68715 | 204908 | CM | CMPLT | CMPLT | 2024-03-15 09:53 | 2024-03-15 09:53 | Created by Failed PM WO # 204907 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:53:01 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69613 | 204909 | PM | CMPLT | PMFAIL | 2024-03-15 09:54 | 2024-03-15 09:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69613 | 204910 | CM | CMPLT | CMPLT | 2024-03-15 09:57 | 2024-03-15 09:57 | Created by Failed PM WO # 204909 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:56:50 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69106 | 204911 | PM | CMPLT | PMFAIL | 2024-03-15 09:58 | 2024-03-15 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69106 | 204912 | CM | CMPLT | CMPLT | 2024-03-15 09:59 | 2024-03-15 10:00 | Created by Failed PM WO # 204911 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 09:59:10 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
1059880 | 204913 | CM | CMPLT | CMPLT | 2024-03-15 10:00 | 2024-03-15 10:01 | Rear case broken | (TC 03/07/2024 10:00:35 AM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced rear case
1059875 | 204914 | CM | CMPLT | CMPLT | 2024-03-15 10:01 | 2024-03-15 10:02 | Rear case broken | (TC 03/07/2024 10:01:18 AM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced rear case
68782 | 204915 | PM | CMPLT | PMFAIL | 2024-03-15 10:04 | 2024-03-15 10:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68782 | 204916 | CM | CMPLT | CMPLT | 2024-03-15 10:04 | 2024-03-15 10:05 | Created by Failed PM WO # 204915 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:04:43 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69559 | 204917 | PM | CMPLT | PMFAIL | 2024-03-15 10:06 | 2024-03-15 10:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69559 | 204918 | CM | CMPLT | CMPLT | 2024-03-15 10:07 | 2024-03-15 10:07 | Created by Failed PM WO # 204917 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:06:49 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69320 | 204919 | CM | CMPLT | CMPLT | 2024-03-15 10:09 | 2024-03-15 10:09 | Created by Failed PM WO # 200647 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:08:49 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
68672 | 204920 | PM | CMPLT | PMFAIL | 2024-03-15 10:11 | 2024-03-15 10:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68672 | 204921 | CM | CMPLT | CMPLT | 2024-03-15 10:11 | 2024-03-15 10:11 | Created by Failed PM WO # 204920 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/07/2024 10:11:08 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
68581 | 204922 | CM | CMPLT | CMPLT | 2024-03-15 10:12 | 2024-03-15 10:13 | Rear case broken | (TC 03/08/2024 10:12:14 AM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced rear case
68602 | 204923 | CM | CMPLT | CMPLT | 2024-03-15 10:13 | 2024-03-15 10:14 | Rear case broken | (TC 03/08/2024 10:13:36 AM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced rear case, front case, and left handle
1058265 | 204924 | PM | CMPLT | PMFAIL | 2024-03-15 10:16 | 2024-03-15 10:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1058265 | 204925 | CM | CMPLT | CMPLT | 2024-03-15 10:17 | 2024-03-15 10:19 | Created by Failed PM WO # 204924 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 10:17:00 AM, Rendon Lopez , Elizabeth) RF wireless card issue. Replaced wireless card, downloaded network configuration and data set. Tested network-OK. PM completed, all tests passed | PM failed by BD
68714 | 204926 | PM | CMPLT | PMFAIL | 2024-03-15 10:20 | 2024-03-15 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68714 | 204927 | CM | CMPLT | CMPLT | 2024-03-15 10:20 | 2024-03-15 10:21 | Created by Failed PM WO # 204926 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 10:20:34 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69311 | 204928 | PM | CMPLT | PMFAIL | 2024-03-15 10:23 | 2024-03-15 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69311 | 204929 | CM | CMPLT | CMPLT | 2024-03-15 10:23 | 2024-03-15 10:24 | Created by Failed PM WO # 204928 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 10:23:34 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69398 | 204930 | PM | CMPLT | PMFAIL | 2024-03-15 10:25 | 2024-03-15 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69398 | 204931 | CM | CMPLT | CMPLT | 2024-03-15 10:25 | 2024-03-15 10:26 | Created by Failed PM WO # 204930 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 10:25:33 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
66014 | 204932 | PM | CMPLT | PMFAIL | 2024-03-15 10:27 | 2024-03-15 10:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
66014 | 204933 | CM | CMPLT | CMPLT | 2024-03-15 10:28 | 2024-03-15 10:28 | Created by Failed PM WO # 204932 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 10:27:51 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69270 | 204934 | PM | CMPLT | PMFAIL | 2024-03-15 10:31 | 2024-03-15 10:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69270 | 204935 | CM | CMPLT | CMPLT | 2024-03-15 10:31 | 2024-03-15 10:32 | Created by Failed PM WO # 204934 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 10:31:41 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
69276 | 204936 | PM | CMPLT | PMFAIL | 2024-03-15 10:33 | 2024-03-15 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69276 | 204937 | CM | CMPLT | CMPLT | 2024-03-15 10:34 | 2024-03-15 10:35 | Created by Failed PM WO # 204936 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 10:33:48 AM, Rendon Lopez , Elizabeth) Damaged ground cable from bezel assy. Replaced bezel, calibrated. PM completed, all tests passed | PM failed by BD
1059873 | 204938 | PM | CMPLT | PMFAIL | 2024-03-15 10:38 | 2024-03-15 10:38 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed by BD
1059873 | 204939 | CM | CMPLT | CMPLT | 2024-03-15 10:38 | 2024-03-15 10:39 | Created by Failed PM WO # 204938 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/12/2024 10:38:41 AM, Rendon Lopez , Elizabeth) Replaced rear case. PM completed, all tests passed | PM failed by BD
68905 | 204940 | PM | CMPLT | PMFAIL | 2024-03-15 10:41 | 2024-03-15 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68905 | 204941 | CM | CMPLT | CMPLT | 2024-03-15 10:41 | 2024-03-15 10:42 | Created by Failed PM WO # 204940 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/12/2024 10:41:29 AM, Rendon Lopez , Elizabeth) Replaced power cord. PM completed, all tests passed | PM failed by BD
1050623 | 204942 | CM | CMPLT | CMPLT | 2024-03-15 10:43 | 2024-03-15 10:44 | Rear case broken | (TC 03/12/2024 10:43:25 AM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced rear case
68596 | 204943 | CM | CMPLT | CMPLT | 2024-03-15 10:44 | 2024-03-15 10:45 | Rear case broken | (TC 03/12/2024 10:44:38 AM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced rear case, left IUI, left handle
68608 | 204944 | CM | CMPLT | CMPLT | 2024-03-15 10:45 | 2024-03-15 10:46 | Rear case broken | (TC 03/12/2024 10:45:42 AM, Rendon Lopez , Elizabeth) Repaired before recall. Replaced rear case, left IUI, left handle
69301 | 204945 | CM | CMPLT | CMPLT | 2024-03-15 10:46 | 2024-03-15 10:48 | Mismatched front case/keypad | (TC 03/12/2024 10:47:10 AM, Rendon Lopez , Elizabeth) Repair before recall. Replaced front case with Carefusion branded.
1065944 | 204946 | PM | CMPLT | PMFAIL | 2024-03-15 10:49 | 2024-03-15 10:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1065944 | 204947 | CM | CMPLT | CMPLT | 2024-03-15 10:49 | 2024-03-21 11:50 | Created by Failed PM WO # 204946 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/14/2024 10:49:37 AM, Rendon Lopez , Elizabeth) Error 242.4030 - Motor control board encoder failure. Replaced AIL board, same issue. Needs controller board. Inquired part pricing, ordered (TC 03/21/2024 11:48:15 AM, Rendon Lopez , Elizabeth) Replaced motor controller board. PM completed, all tests passed | PM failed by BD
68847 | 204948 | CM | CMPLT | CMPLT | 2024-03-15 10:52 | 2024-03-15 10:53 | Damaged power cord | (TC 03/14/2024 10:51:45 AM, Rendon Lopez , Elizabeth) Replaced power cord
65931 | 204949 | PM | CMPLT | PMFAIL | 2024-03-15 10:55 | 2024-03-15 10:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed by BD
65931 | 204950 | CM | CMPLT | CMPLT | 2024-03-15 10:55 | 2024-03-15 10:56 | Created by Failed PM WO # 204949 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/14/2024 10:55:47 AM, Rendon Lopez , Elizabeth) Replaced rear case. PM completed, all tests passed | PM failed by BD
69327 | 204951 | PM | CMPLT | PMFAIL | 2024-03-15 10:57 | 2024-03-15 10:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69327 | 204952 | CM | CMPLT | CMPLT | 2024-03-15 10:58 | 2024-03-15 10:58 | Created by Failed PM WO # 204951 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/14/2024 10:57:52 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. Calibrated rate accuracy PM completed, all tests passed | PM failed by BD
69500 | 204953 | PM | CMPLT | PMFAIL | 2024-03-15 11:00 | 2024-03-15 11:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69500 | 204954 | CM | CMPLT | CMPLT | 2024-03-15 11:00 | 2024-03-15 11:00 | Created by Failed PM WO # 204953 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/14/2024 11:00:11 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
68299 | 204955 | PM | CMPLT | PMFAIL | 2024-03-15 11:02 | 2024-03-15 11:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
68299 | 204956 | CM | CMPLT | CMPLT | 2024-03-15 11:02 | 2024-03-15 11:02 | Created by Failed PM WO # 204955 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/14/2024 11:02:08 AM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
68698 | 204957 | PM | CMPLT | CMPLT | 2024-03-15 11:52 | 2024-03-15 11:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/15/2024 11:52:32 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 80 mohms Leakage current: 68.9 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
63927 | 204958 | CM | CMPLT | CMPLT | 2024-03-15 12:22 | 2024-03-15 13:10 | Battery issue | (TC 03/15/2024 01:08:43 PM, Scero, Tomlyn) Batteries need to be replaced due to age. CHarged the device for now and sent the unit back. Contacted Philips for a quote to order new batteries.
68720 | 204959 | CM | CMPLT | CMPLT | 2024-03-15 13:06 | 2024-03-15 13:20 | Bended prong | (TC 03/15/2024 01:19:21 PM, Rendon Lopez , Elizabeth) Replaced power cord
| 204960 | CM | CMPLT | CMPLT | 2024-03-15 13:20 | 2024-04-02 09:45 | Alaris pump spare parts | Received
69566 | 204961 | PM | CMPLT | CMPLT | 2024-03-15 13:27 | 2024-03-15 13:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69567 | 204962 | PM | CMPLT | CMPLT | 2024-03-15 13:31 | 2024-03-15 13:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69629 | 204963 | PM | CMPLT | PMFAIL | 2024-03-15 13:32 | 2024-03-15 13:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69629 | 204964 | CM | CMPLT | CMPLT | 2024-03-15 13:32 | 2024-04-03 13:33 | Created by Failed PM WO # 204963 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/15/2024 01:44:02 PM, Rendon Lopez , Elizabeth) Failed mechanism assy, rate calibration. Ordered bezel assy (TC 04/02/2024 03:23:59 PM, Rendon Lopez , Elizabeth) Replaced bezel assy. (TC 04/03/2024 01:32:45 PM, Rendon Lopez , Elizabeth) Calibrated rate. PM completed, all tets passed | PM failed by BD
69259 | 204965 | PM | CMPLT | CMPLT | 2024-03-15 13:32 | 2024-03-15 13:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65957 | 204966 | PM | CMPLT | CMPLT | 2024-03-15 13:38 | 2024-03-15 13:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 01:38:27 PM, Baghirzade, Nara) PM completed by BD.
69429 | 204967 | CM | CMPLT | CMPLT | 2024-03-15 13:38 | 2024-03-15 13:40 | Created by Failed PM WO # 201427 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/15/2024 01:39:33 PM, Rendon Lopez , Elizabeth) Failed pressure sensor test. Calibrated rate accuracy. PM completed all tests passed | PM failed by BD
69095 | 204968 | PM | CMPLT | CMPLT | 2024-03-15 13:39 | 2024-03-15 13:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65930 | 204969 | PM | CMPLT | CMPLT | 2024-03-15 13:41 | 2024-03-15 13:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 01:40:41 PM, Baghirzade, Nara) PM completed by BD.
65900 | 204970 | PM | CMPLT | PMFAIL | 2024-03-15 13:42 | 2024-03-15 13:42 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 01:41:49 PM, Baghirzade, Nara) Broken rear case
65900 | 204971 | CM | CMPLT | CMPLT | 2024-03-15 13:42 | 2024-03-15 13:43 | Created by Failed PM WO # 204970 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 01:42:23 PM, Baghirzade, Nara) Replaced rear case. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/15/2024 01:41:49 PM, Baghirzade, Nara) Broken rear case
68664 | 204972 | PM | CMPLT | CMPLT | 2024-03-15 13:47 | 2024-03-15 13:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69673 | 204973 | PM | CMPLT | PMFAIL | 2024-03-15 13:48 | 2024-03-15 13:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69673 | 204974 | CM | CMPLT | CMPLT | 2024-03-15 13:48 | 2024-03-15 13:59 | Created by Failed PM WO # 204973 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/15/2024 01:48:09 PM, Rendon Lopez , Elizabeth) Failed Mechanism assy, failed rate calibration. Calibrated rate accuracy. PM completed, all tests passed | PM failed by BD
68659 | 204975 | PM | CMPLT | CMPLT | 2024-03-15 13:55 | 2024-03-15 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69139 | 204976 | PM | CMPLT | CMPLT | 2024-03-15 14:02 | 2024-03-15 14:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66045 | 204977 | CM | CMPLT | CMPLT | 2024-03-15 14:04 | 2024-03-15 14:05 | Created by Failed PM WO # 201378 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/15/2024 02:04:22 PM, Rendon Lopez , Elizabeth) Cleaned IUIs. PM completed, all tests passed | PM failed by BD
68608 | 204978 | PM | CMPLT | CMPLT | 2024-03-15 14:09 | 2024-03-15 14:09 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 02:08:56 PM, Rhuland, Courtney) PM completed by BD
66008 | 204979 | PM | CMPLT | CMPLT | 2024-03-15 14:09 | 2024-03-15 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/15/2024 02:09:27 PM, Baghirzade, Nara) Replaced power cord. PM passed and completed. Battery was replaced by BD. Electrical Safety - Ground resistance: 120 mohms Leakage current: 79.3 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
65954 | 204980 | PM | CMPLT | CMPLT | 2024-03-15 14:12 | 2024-03-15 14:13 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 02:11:53 PM, Rhuland, Courtney) PM completed by BD
68655 | 204981 | PM | CMPLT | CMPLT | 2024-03-15 14:14 | 2024-03-15 14:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65914 | 204982 | PM | CMPLT | CMPLT | 2024-03-15 14:14 | 2024-03-15 14:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/15/2024 02:14:25 PM, Rhuland, Courtney) PM completed by BD
69433 | 204983 | PM | CMPLT | CMPLT | 2024-03-15 14:22 | 2024-03-15 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:22:45 PM, Huang, Jessica) replaced the female iui connector and completed PM
67103 | 204984 | CM | CMPLT | CMPLT | 2024-03-15 14:26 | 2024-03-15 14:29 | Battery Replacemnet | Found unit in repair area. IInspected and cleaned unit. Found BP Cuff missing. Replaced it. Went through manual to understand the device. Found battery completely depleted and not holding charge. Replaced battery. Checedk functionality and accuracy using NIBP simulator. Tested ok. Returned to 9S.
66049 | 204985 | CM | CMPLT | CMPLT | 2024-03-15 14:28 | 2024-03-15 14:29 | Damaged power cord | (TC 03/15/2024 02:28:28 PM, Baghirzade, Nara) Replaced power cord.
69266 | 204986 | PM | CMPLT | CMPLT | 2024-03-15 14:30 | 2024-03-15 14:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69397 | 204987 | PM | CMPLT | CMPLT | 2024-03-15 14:37 | 2024-03-15 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:37:01 PM, Huang, Jessica) replaced the female iui connector and completed PM
69247 | 204988 | PM | CMPLT | CMPLT | 2024-03-15 14:37 | 2024-03-15 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69133 | 204989 | PM | CMPLT | CMPLT | 2024-03-15 14:40 | 2024-03-15 14:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/08/2024 02:40:03 PM, Huang, Jessica) replaced the female iui connector and completed PM
1057388 | 204990 | PM | CMPLT | CMPLT | 2024-03-15 14:41 | 2024-03-15 14:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1057393 | 204991 | PM | CMPLT | CMPLT | 2024-03-15 14:44 | 2024-03-15 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050266 | 204992 | PM | CMPLT | CMPLT | 2024-03-15 14:53 | 2024-03-15 14:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055157 | 204993 | PM | CMPLT | CMPLT | 2024-03-15 15:00 | 2024-03-15 15:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 204994 | OC | CMPLT | CMPLT | 2024-03-17 20:39 | 2024-03-22 13:46 | On-call week March 15-21 | (TC 03/16/2024 09:15:00 AM, Canjimavadivel, Kamallesh ) Call from NICU around 8:10 AM that some of their remediated pumps in the resus area were giving errors overnight. Reported message "calibration required" on the pumps. Found 2 PCUs and 4 LVPs with issues. I asked them to identify how many pumps had this issue and if they had enough pumps. Charge nurse asked to call back after half hour so they can check. Called back at 9 AM, they validated rest of resus area was okay and they still need to validate rest of main unit. Charge nurse called back at 9:15 and confirmed they had enough pumps to last the weekend. I requested they tag the pumps with issues and to send down to biomed.
1058179 | 204995 | CM | CMPLT | CMPLT | 2024-03-18 07:50 | 2024-03-18 07:51 | Updating drug library | (TC 03/18/2024 07:50:29 AM, Reid, Emmerson ) Updated drug library to APS Epidural 2023 Oct
| 204996 | RDS | CMPLT | CMPLT | 2024-03-18 08:07 | 2024-03-18 08:09 | Weekly Checks For PACU | (TC 03/18/2024 08:08:40 AM, Reid, Emmerson ) Rounds complete
| 204997 | RDS | CMPLT | CMPLT | 2024-03-18 08:24 | 2024-03-20 10:06 | Weekly checks in CCU | (TC 03/20/2024 10:05:24 AM, Scero, Tomlyn) weekly checks in ccu complete. 30 J tests passed.
| 204998 | RDS | CMPLT | CMPLT | 2024-03-18 08:27 | 2024-03-18 09:20 | Weekly check for ED
| 204999 | RDS | CMPLT | CMPLT | 2024-03-18 09:53 | 2024-03-18 09:53 | weekly check for 14S and 17M | weekly check for 14S and 17M
65969 | 205000 | CM | CMPLT | CMPLT | 2024-03-18 10:31 | 2024-03-18 10:32 | Alarms calibration required | (TC 03/18/2024 10:31:46 AM, Rendon Lopez , Elizabeth) Performed calibration and check-in tests. All tests passed.
65965 | 205001 | CM | CMPLT | CMPLT | 2024-03-18 10:33 | 2024-03-18 10:33 | Alarming calibration required | (TC 03/18/2024 10:32:43 AM, Rendon Lopez , Elizabeth) Performed calibration and check-in tests. All tests passed.
1050623 | 205002 | CM | CMPLT | CMPLT | 2024-03-18 11:00 | 2024-03-18 11:05 | Syringe requires full calibration | (TC 03/18/2024 11:05:01 AM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1059135 | 205003 | CM | CMPLT | CMPLT | 2024-03-18 11:10 | 2024-03-18 11:12 | SpO2 working intermittently | (TC 03/18/2024 11:10:23 AM, Reid, Emmerson ) Went up to check, patient was on the monitor. I did see the SpO2 fail to read momentarily a couple times. They said it also happened with another patient. I tried on myself and it was working. Making WO to track if there is more of an issue but it seems like only a sensor problem.
1048527 | 205004 | CM | CMPLT | CMPLT | 2024-03-18 11:30 | 2024-03-18 11:31 | Full calibration and functional check | (TC 03/18/2024 11:30:32 AM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
502088 | 205005 | RQ | CMPLT | CMPLT | 2024-03-18 11:35 | 2024-03-20 11:11 | QT QC phantom holder not level? | (TC 03/20/2024 11:10:56 AM, Green, Paul) duplitcate WO. | QT QC phantom holder not level?
1060032 | 205006 | CM | CMPLT | CMPLT | 2024-03-18 12:22 | 2024-03-21 11:34 | Parts return | (TC 03/18/2024 12:00:00 PM, Baghirzade, Nara) Returned part MS 4746 to Novus as the part was incorrect for our device type. Chantal from Novus to issue a credit for this return. Waybill # 423002 (TC 03/21/2024 11:33:16 AM, Baghirzade, Nara) Received credit from Novus for returning MS4746. Attached credit note to this WO.
1048551 | 205007 | CM | CMPLT | CMPLT | 2024-03-18 13:58 | 2024-04-02 16:53 | Repair - Leak | Mar 18 - "not working". Went to check and brought it down to biomed. Charged the unit. Checked functionality and performed cuff leak test. Found pressure leaking. Checked all cuff and hose o rings. Found several that need replacement. Replaced Red hose from JH. Contacted Zimmer for parts inquiry. Awaiting response. Mar 19 - Info received. Created SC 1003192356 for the purchase. Ordered extra items for future stock. Sent it for approval. Mar 21 - Parts received. Repalced o rings on all cuffs and hoses. Performed leak test. 5 cuff still failed the leak test and will required to be repalced. Contacted zimmer for info. Awaiting response. Mar 22 - Followed up. information Received. Created SC 1003192435 and sent it for approval. Received. Awaiting delivery. Apr 02 - Received. Unboxed everything. Checked all cuffs and performed cuff test. Tested ok. Delivered unit back to OR.
66069 | 205008 | PM | CMPLT | CMPLT | 2024-03-18 14:18 | 2024-03-18 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66027 | 205009 | PM | CMPLT | CMPLT | 2024-03-18 14:21 | 2024-03-18 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69004 | 205010 | PM | CMPLT | CMPLT | 2024-03-18 14:31 | 2024-03-18 14:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69028 | 205011 | PM | CMPLT | CMPLT | 2024-03-18 14:32 | 2024-03-18 14:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68822 | 205012 | PM | CMPLT | CMPLT | 2024-03-18 14:35 | 2024-03-18 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68910 | 205013 | PM | CMPLT | CMPLT | 2024-03-18 14:43 | 2024-03-18 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68625 | 205014 | PM | CMPLT | CMPLT | 2024-03-18 14:49 | 2024-03-18 14:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 205015 | CM | CMPLT | CMPLT | 2024-03-18 14:50 | 2024-03-18 14:50 | meeting | Biomed meeting
| 205016 | CM | CMPLT | CMPLT | 2024-03-18 14:52 | 2024-03-18 14:52 | worked with RT | worked with RT
69207 | 205017 | PM | CMPLT | CMPLT | 2024-03-18 14:53 | 2024-03-18 15:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/18/2024 02:58:29 PM, Reid, Emmerson ) ASM PM Complete.
| 205018 | CM | CMPLT | CMPLT | 2024-03-18 14:54 | 2024-03-18 14:54 | RT meeting | RT meeting
68943 | 205019 | PM | CMPLT | CMPLT | 2024-03-18 15:01 | 2024-03-18 15:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68691 | 205020 | PM | CMPLT | CMPLT | 2024-03-18 15:06 | 2024-03-18 15:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051376 | 205021 | PM | CMPLT | CMPLT | 2024-03-18 15:09 | 2024-03-18 15:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055110 | 205022 | PM | CMPLT | CMPLT | 2024-03-18 15:10 | 2024-03-18 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058277 | 205023 | PM | CMPLT | CMPLT | 2024-03-18 15:12 | 2024-03-18 15:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1053349 | 205024 | CM | CMPLT | CMPLT | 2024-03-18 15:20 | 2024-03-18 15:22 | MSH BI 1274 H2 System crashed during tomo | (TC 03/18/2024 03:19:53 PM, Fan, Tak) System crashed, was stuck on windows log in screen. Restarted system. Last image appears acquired but not quite normal. Job queue errors, last image seems to fail going to manger, appears on Coral, and can be pushed to workstation from coral. Ryan looked at system but not sure. Contacted apps, who walked through using "edit" to reprocess the image. Requires clearing out the job and error spoolers. Resaving after going into edit reprocesses. (TC 03/19/2024 09:28:51 AM, Whittington, Scott) MSH BI 1274 H2 System crashed during tomo FOLLOWUP WITH CLINICAL, INCLUDING IVY APPARENTLY A DOWNTIME ON WEEKEND SO A LOT OF CASES INSPOOLER RYAN L FROM HOLOGIC CAME ON-SITE MARCH 18 TO ASSIST TAK | MSH BI 1274 H2 System crashed during tomo
1061214 | 205025 | CM | CMPLT | CMPLT | 2024-03-18 15:51 | 2024-03-18 15:54 | Repair | Call re: a piece came out from the boom. Went to check. Found joint cover on monitor 3 mounting arm got separated from the joint. Was informed to come back later. Went later. Borrowed step ladder from OR 12 and replaced joint cover on the mounting arm. Checked functionlity. Tested ok.
1056421 | 205026 | CM | CMPLT | CMPLT | 2024-03-18 16:38 | 2024-03-18 16:56 | "BP cuff not inflating" | (TC 03/18/2024 04:37:48 PM, Cortez, John) Replaced cuff, cleaned filter, added radical battery, secured power cord, Root missing screw (bottom, securing Root to rollstand), female threading damaged All parameters tested ok
1054521 | 205027 | CM | CMPLT | CMPLT | 2024-03-19 08:24 | 2024-03-19 08:30 | LL lead broken | (TC 03/19/2024 08:29:26 AM, Canjimavadivel, Kamallesh ) Replaced LL lead. Tested ok with simulator
| 205028 | CM | CMPLT | CMPLT | 2024-03-19 08:48 | 2024-03-18 14:56 | Endo ER Cart CO2 Tank Leak issue | Call re: CO2 tank not working. Went to check and found CO2 tank emptied out. Checked with the staff to see if someone left the insufflator on. Was informed no. Removed the regulator and replaced the CO2 tank. Checked functionality. Tested ok.
| 205029 | CM | CMPLT | CMPLT | 2024-03-19 09:03 | 2024-03-19 09:03 | PM Update/Huddle
65877 | 205030 | PM | CMPLT | CMPLT | 2024-03-19 09:33 | 2024-03-19 10:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/19/2024 09:33:09 AM, Reid, Emmerson ) Ran Calibration, and ASM PM PM Complete.
68579 | 205031 | PM | CMPLT | CMPLT | 2024-03-19 09:37 | 2024-03-19 09:37 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/19/2024 09:36:52 AM, Rhuland, Courtney) Calibration and PM completed All tests passed Refer to ASM for details
1061204 | 205032 | CM | CMPLT | CMPLT | 2024-03-19 09:48 | 2024-03-19 09:48 | Light arm is loose | (TC 03/19/2024 09:47:42 AM, Reid, Emmerson ) The light arm was loose, i tightened it. Working now.
1066107 | 205033 | CM | CMPLT | CMPLT | 2024-03-19 10:09 | 2024-09-25 09:34 | MSH GIM GE VENUE 40 CAN'T BE FOUND | (TC 03/19/2024 10:09:10 AM, Whittington, Scott) From: Whittington, Scott Sent: March 19, 2024 10:03 AM To: Chow, Gabriel - Sinai Health System <Gabriel.Chow@sinaihealth.ca> Cc: Ferreira, Qwynn <Qwynn.Ferreira@sinaihealth.ca> Subject: RE: [External] Ultrasonic Scanner - 1066107 Hi Gabriel The GE VENUE 40 POCUS had a lock down on 10th floor. You cant find it? Thanks, Scott 416-402-2462 SCOTT WHITTINGTON SINAI HEALTH SYSTEM TL MEDICAL IMAGING ENGINEERING DIGITAL RADIOGRAPHY, ULTRASOUND, MAMMOGRAPHY UHN, SHS, WCH 416-586-4800 x2113 OFFICE PHONE/SERVICE LINE 416-402-2462 MOBILE scott.whittington@uhn.ca MOUNT SINAI HOSPITAL 600 UNIVERSITY AVE, TORONTO ON ROOM MSH 5-152 M5G 1X5 From: Ferreira, Qwynn <Qwynn.Ferreira@sinaihealth.ca> Sent: March 19, 2024 9:52 AM To: Chow, Gabriel - Sinai Health System <Gabriel.Chow@sinaihealth.ca> Cc: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] Ultrasonic Scanner - 1066107 Good morning Gabe, Thank you for the call. The item you have with the asset 1066107 belongs to JDMI department (Scott from imaging), and seems to be coming from the 10th floor. It might be coming up with patients that require imaging on the unit, but a porter should be able to bring the device back to the 10th floor. I have also cc’d Scott so he might be able to provide more details. Best, Qwynn Ferreira, B.Eng., EIT Biomedical Engineer, Biomedical Engineering T 416-586-4800 x 6060 C 437-288-4358 qwynn.ferreira@sinaihealth.ca Mount Sinai Hospital, Joseph & Wolf Lebovic Health Complex 139-600 University Avenue, Toronto, ON M5G 1X5 | MSH GIM GE VENUE 40 CAN'T BE FOUND From: Whittington, Scott Sent: March 19, 2024 10:03 AM To: Chow, Gabriel - Sinai Health System <Gabriel.Chow@sinaihealth.ca> Cc: Ferreira, Qwynn <Qwynn.Ferreira@sinaihealth.ca> Subject: RE: [External] Ultrasonic Scanner - 1066107 Hi Gabriel The GE VENUE 40 POCUS had a lock down on 10th floor. You cant find it? Thanks, Scott 416-402-2462 SCOTT WHITTINGTON SINAI HEALTH SYSTEM TL MEDICAL IMAGING ENGINEERING DIGITAL RADIOGRAPHY, ULTRASOUND, MAMMOGRAPHY UHN, SHS, WCH 416-586-4800 x2113 OFFICE PHONE/SERVICE LINE 416-402-2462 MOBILE scott.whittington@uhn.ca MOUNT SINAI HOSPITAL 600 UNIVERSITY AVE, TORONTO ON ROOM MSH 5-152 M5G 1X5 From: Ferreira, Qwynn <Qwynn.Ferreira@sinaihealth.ca> Sent: March 19, 2024 9:52 AM To: Chow, Gabriel - Sinai Health System <Gabriel.Chow@sinaihealth.ca> Cc: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] Ultrasonic Scanner - 1066107 Good morning Gabe, Thank you for the call. The item you have with the asset 1066107 belongs to JDMI department (Scott from imaging), and seems to be coming from the 10th floor. It might be coming up with patients that require imaging on the unit, but a porter should be able to bring the device back to the 10th floor. I have also cc’d Scott so he might be able to provide more details. Best, Qwynn Ferreira, B.Eng., EIT Biomedical Engineer, Biomedical Engineering T 416-586-4800 x 6060 C 437-288-4358 qwynn.ferreira@sinaihealth.ca Mount Sinai Hospital, Joseph & Wolf Lebovic Health Complex 139-600 University Avenue, Toronto, ON M5G 1X5
65932 | 205034 | PM | CMPLT | CMPLT | 2024-03-19 10:40 | 2024-03-19 10:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/19/2024 10:40:09 AM, Rhuland, Courtney) Rear case replaced IUIs replaced Calibration and PM complete All tests passed Refer to ASM for details
65900 | 205035 | CM | CMPLT | CMPLT | 2024-03-19 11:25 | 2024-03-19 11:27 | Full calibration required | (TC 03/19/2024 11:26:11 AM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65931 | 205036 | CM | CMPLT | CMPLT | 2024-03-19 12:11 | 2024-03-19 12:12 | Full calibration required | (TC 03/19/2024 12:12:04 PM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1048529 | 205037 | CM | CMPLT | CMPLT | 2024-03-19 12:14 | 2024-03-19 12:15 | Broken rear case | (TC 03/19/2024 12:14:18 PM, Baghirzade, Nara) Replaced rear case.
65982 | 205038 | CM | CMPLT | CMPLT | 2024-03-19 12:14 | 2024-03-19 12:15 | Cracked rear case | (TC 03/19/2024 12:14:28 PM, Rhuland, Courtney) Replaced rear case Replaced IUI
| 205039 | CM | CMPLT | CMPLT | 2024-03-19 13:04 | 2024-03-19 13:04 | meeting | meeting
| 205040 | CM | ACTIV | ISSUE | 2024-03-19 13:04 |  | TMS and OR Assets Pt.3 meeting | (TC 03/19/2024 01:04:19 PM, Chukri, Samir) Meeting with O.R team (TC 04/24/2024 01:55:36 PM, Chukri, Samir) OR Monthly Meeting (TC 08/07/2024 07:47:32 AM, Chukri, Samir) meeting
65894 | 205041 | CM | CMPLT | CMPLT | 2024-03-19 13:12 | 2024-03-19 13:13 | Cracked front case | (TC 03/19/2024 01:12:18 PM, Rhuland, Courtney) Replaced front case Replaced internal housing Replaced left handle Replaced mylar gasket Replaced barcode Returned to BD for recall & PM
65976 | 205042 | CM | CMPLT | CMPLT | 2024-03-19 13:14 | 2024-03-19 13:14 | Full calibration required | (TC 03/19/2024 01:13:50 PM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1059873 | 205043 | CM | CMPLT | CMPLT | 2024-03-19 14:04 | 2024-03-19 14:04 | Full calibration required | (TC 03/19/2024 02:04:03 PM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1063000 | 205044 | PM | CMPLT | CMPLT | 2024-03-19 14:10 | 2024-03-19 14:10 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/19/2024 02:10:02 PM, Cortez, John) Completed cleaning cycle
65896 | 205045 | CM | CMPLT | CMPLT | 2024-03-19 14:16 | 2024-03-20 11:08 | Broken rear case, latch and left handle | (TC 03/19/2024 02:16:16 PM, Baghirzade, Nara) Replaced rear case and latch. Left handle on backorder. (TC 03/20/2024 11:07:18 AM, Baghirzade, Nara) Replaced front case, left handle, and mylar gasket.
65924 | 205046 | CM | CMPLT | CMPLT | 2024-03-19 14:51 | 2024-03-19 14:53 | Cracked front case | (TC 03/19/2024 02:51:25 PM, Rhuland, Courtney) Replaced front case Replaced rear case Replaced left handle Replaced mylar gasket Reprinted barcode Replaced IUI Returned to BD for recall and PM
69648 | 205047 | PM | CMPLT | CMPLT | 2024-03-19 14:59 | 2024-03-19 15:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69043 | 205048 | PM | CMPLT | CMPLT | 2024-03-19 15:04 | 2024-03-19 15:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69067 | 205049 | PM | CMPLT | CMPLT | 2024-03-19 15:06 | 2024-03-19 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055942 | 205050 | CM | CMPLT | CMPLT | 2024-03-19 15:55 | 2024-03-19 15:55 | Blood pressure not working | replaced adult cuff. installed battery on Radical 7. tested BP, temp and SpO2. Functional.
| 205051 | CM | CMPLT | CMPLT | 2024-03-19 15:57 | 2024-03-19 15:57 | worked with RT | worked with RT
1070482 | 205052 | CM | CMPLT | CMPLT | 2024-03-20 07:27 | 2024-03-20 07:29 | Incoming inspection | (TC 03/20/2024 07:28:59 AM, Rhuland, Courtney) Asset tagged and functional check - OK
65973 | 205053 | CM | CMPLT | CMPLT | 2024-03-20 09:00 | 2024-03-20 09:01 | Cracked rear case | (TC 03/20/2024 09:00:43 AM, Rhuland, Courtney) Rear case replaced IUI replaced Returned to BD for recall and PM
65903 | 205054 | PM | CMPLT | CMPLT | 2024-03-20 09:18 | 2024-03-20 09:18 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 09:17:56 AM, Baghirzade, Nara) PM completed by BD.
65923 | 205055 | PM | CMPLT | CMPLT | 2024-03-20 09:20 | 2024-03-20 09:20 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 09:19:50 AM, Baghirzade, Nara) PM completed by BD.
65912 | 205056 | PM | CMPLT | CMPLT | 2024-03-20 09:21 | 2024-03-20 09:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 09:21:02 AM, Baghirzade, Nara) PM completed by BD.
68596 | 205057 | PM | CMPLT | CMPLT | 2024-03-20 09:22 | 2024-03-20 09:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 09:22:30 AM, Baghirzade, Nara) PM completed by BD.
65882 | 205058 | PM | CMPLT | CMPLT | 2024-03-20 09:24 | 2024-03-20 09:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 09:24:08 AM, Baghirzade, Nara) PM completed by BD.
65984 | 205059 | PM | CMPLT | CMPLT | 2024-03-20 09:25 | 2024-03-20 09:26 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 09:25:22 AM, Baghirzade, Nara) PM completed by BD.
65911 | 205060 | PM | CMPLT | CMPLT | 2024-03-20 09:27 | 2024-03-20 09:35 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 09:34:27 AM, Baghirzade, Nara) PM completed by BD.
1070483 | 205061 | CM | CMPLT | CMPLT | 2024-03-20 10:02 | 2024-03-20 10:14 | Loaner device | (TC 03/20/2024 10:13:42 AM, Baghirzade, Nara) Assembled loaner on stand and swapped the probes. Dropped off loaner device with Sharmaine. | Unit owned bladder scanner was sent to the vendor for calibration
69502 | 205062 | PM | CMPLT | CMPLT | 2024-03-20 10:04 | 2024-03-20 10:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65881 | 205063 | CM | CMPLT | CMPLT | 2024-03-20 10:14 | 2024-03-20 10:19 | Broken case | (TC 03/20/2024 10:14:45 AM, Rhuland, Courtney) Replaced front and rear case Replaced mylar gasket Replaced IUI connector Reprinted barcode Returned to BD for recall and PM
69677 | 205064 | PM | CMPLT | CMPLT | 2024-03-20 10:29 | 2024-03-20 10:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69245 | 205065 | PS | CMPLT | CMPLT | 2024-03-20 10:34 | 2024-03-20 10:35 | 2024 Alaris Software Recall
69245 | 205066 | PM | CMPLT | CMPLT | 2024-03-20 10:35 | 2024-03-20 10:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69366 | 205067 | PM | CMPLT | CMPLT | 2024-03-20 10:38 | 2024-03-20 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055167 | 205068 | PM | CMPLT | CMPLT | 2024-03-20 10:40 | 2024-03-20 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055171 | 205069 | PM | CMPLT | CMPLT | 2024-03-20 10:45 | 2024-03-20 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055158 | 205070 | PM | CMPLT | CMPLT | 2024-03-20 10:48 | 2024-03-20 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65917 | 205071 | CM | CMPLT | CMPLT | 2024-03-20 11:05 | 2024-03-20 11:07 | Broken rear case, front case, and defective left IUI | (TC 03/20/2024 11:05:22 AM, Baghirzade, Nara) Replaced: -front case -rear case -left IUI -mylar gasket -pressure disc flag
1049218 | 205072 | CM | CMPLT | CMPLT | 2024-03-20 13:36 | 2024-03-20 13:41 | Manual restart required | (TC 03/20/2024 01:39:27 PM, Scero, Tomlyn) email received from philips focal point saying High GDI Resource Usage. Manual restart suggested. I went up to CCU and restarted the central station with no issues.
68370 | 205073 | CM | CMPLT | CMPLT | 2024-03-20 13:37 | 2024-03-20 14:54 | N2O not being flushed out of circuit | (TC 03/19/2024 02:00:54 PM, Scero, Tomlyn) Email received from Peter regarding device still showing N2O by the gas analyzer 2 hrs after N2O was turned off. A picture was sent. See attached. I went to OR 19 and downloaded the logs (as per Dregers request) and it was sent to the escalation desk. (TC 03/20/2024 02:51:42 PM, Scero, Tomlyn) After Draeger analyzed the DSR they sent back their reccomendations. Hi Tomlyn; Please see the reply from the escalation desk: That's totally normal, because 20 minutes after the end of the nitrous oxide dose, 10% of the initial amount of nitrous oxide is still in the blood. (see attached link and N2O curve over time in the middle) https://pubs.asahq.org/anesthesiology/article/114/3/596/11058/Nitrous-Oxide-Diffusion-and-the-Second-Gas-Effect As you can see in the flight log, they stopped at 12:30 and switched to carrier gas Air at 12:35. At 12:50 the PGA measured still expiratory 3-4% N2O and they switched into Auto Mode. That means only uptake of O2 and Sevo was delivered in closed system with circle flow and for this reason the insp N2O could rise a little up to 5-6% insp and exp. At about 13:05 N2O value stopped to rise, but the user did not start a flush phase (which would help to wash out rest gas) and for this reason the N2O concentration was present until they 15:27 when they increased the O2 setting and switched into FG Mode. Finally no device failure;-) This response was sent to Dr Teresi.
1063026 | 205074 | CM | ACTIV | ISSUE | 2024-03-20 13:48 |  | Viewpoint SW upgrade testing. | (TC 03/20/2024 01:48:25 PM, Fan, Tak) Met with Dianne. Jose will setup test viewpoint server with GE. Testing will be on newest GE vivid e95.
1055745 | 205075 | PM | CMPLT | PMFAIL | 2024-03-20 13:51 | 2024-03-20 13:52 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 01:51:42 PM, Baghirzade, Nara) Broken rear case and syringe requires full calibration.
1055745 | 205076 | CM | CMPLT | CMPLT | 2024-03-20 13:52 | 2024-03-20 14:00 | Created by Failed PM WO # 205075 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 01:59:37 PM, Baghirzade, Nara) Replaced rear case. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/20/2024 01:51:42 PM, Baghirzade, Nara) Broken rear case and syringe requires full calibration.
68598 | 205077 | PM | CMPLT | CMPLT | 2024-03-20 13:57 | 2024-03-20 14:29 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 02:28:06 PM, Reid, Emmerson ) Replaced the front case, ran calibration and did PM. PM Complete
68594 | 205078 | CM | CMPLT | CMPLT | 2024-03-20 14:25 | 2024-03-20 14:32 | Created by Failed PM WO # 202102 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe |  | (TC 03/20/2024 02:24:19 PM, Baghirzade, Nara) Broken rear case and requires full calibration.
68594 | 205079 | PM | CMPLT | CMPLT | 2024-03-20 14:25 | 2024-03-20 14:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/20/2024 02:31:34 PM, Baghirzade, Nara) Replaced rear case. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1055920 | 205080 | CM | CMPLT | CMPLT | 2024-03-20 14:35 | 2024-03-20 14:38 | Replaced handle and all monitoring accessories | (TC 03/20/2024 02:35:13 PM, Baghirzade, Nara) Replaced: -handle -NIBP hose -Temperature probe and well -SpO2 Red LNC-10 cable -SpO2 finger sensor Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
62696 | 205081 | CM | CMPLT | CMPLT | 2024-03-20 15:20 | 2024-03-20 15:24 | Not picking up any signals | (TC 03/20/2024 03:20:47 PM, Canjimavadivel, Kamallesh ) RL and LA leads were intermittently working. Replaced both and tested with simulator.
| 205082 | CM | CNP | ENL | 2024-03-20 15:24 | 2024-04-24 13:26 | Endo cart CO2 regulator Issue | Mar 19 - Call from Anita re: delivery of steris co2 regulator with two dial indicators. Went to check. Anita requested to install it on the tank. Checked and found co2 regulato came with steris connection and was not compatible with olympus hose. Requested help from ER to remove the olumpus hose from the CO2 yolk. Consulted regarding how to make appropriate connection to CO2 tank using Olymous hose. Was advised to change the regulator adaptor fitting. Repalced adaptor and connected the tank to UCR. Checked functionality. Tested ok. Mar 20 - call and meeting with Anita. Was requested to source a CO2 pressure guage with one dial. Contacted a couple of vendors. Awaiting response. Mar 26 - Followed up but could not find any canadian vendor that sells CO2 yoke assembly with pressure gauge. Will have to explore US vendors. Mar 27 - Found US vendor WT Farley that makes CO2 yoke with pressuregauge. Reached out and inquired about canadian distributor. Was givben contact for West care Medical in BC. Reached out to West care for purchase info. Awaiting response. Apr 02 - Followed up. Awaiting response. Apr 12 - Followed up. Was informed that they are awaiting response from manufacturer. Apr 24 - Followed up multiple times but no response. WO Closed.
| 205083 | CM | CMPLT | CMPLT | 2024-03-21 07:16 | 2024-05-09 14:35 | Breast Pump incoming inspections | (TC 03/20/2024 03:00:00 PM, Rhuland, Courtney) Tagged and moved breast pumps to storage (TC 04/04/2024 02:40:40 PM, Rhuland, Courtney) Tagged devices and assembled rolling stands (TC 04/15/2024 09:59:37 AM, Rhuland, Courtney) Assembled rolling stands and mounted breast pumps (TC 04/15/2024 10:11:34 AM, Baghirzade, Nara) Assembled rolling stands and mounted breast pumps (TC 04/17/2024 10:56:09 AM, Rhuland, Courtney) Tagged and stored breast pumps in rental cases (TC 04/17/2024 10:56:39 AM, Baghirzade, Nara) Tagged and stored breast pumps in rental cases (TC 04/25/2024 02:12:18 PM, Baghirzade, Nara) Tagged and stored breast pumps in rental cases (TC 04/30/2024 02:48:14 PM, Rhuland, Courtney) Tagged and stored breast pumps in rental cases (TC 05/08/2024 09:55:04 AM, Rhuland, Courtney) Tagged and stored breast pumps in rental cases
68610 | 205084 | CM | CMPLT | CMPLT | 2024-03-21 08:09 | 2024-03-21 08:10 | Broken front case | (TC 03/21/2024 08:09:04 AM, Rhuland, Courtney) Replaced front case Replaced mylar gasket Replaced IUI Returned to BD for recall and PM
1055743 | 205085 | PM | CMPLT | CMPLT | 2024-03-21 08:19 | 2024-03-21 08:19 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 08:18:46 AM, Rhuland, Courtney) PM completed by BD
1055746 | 205086 | PM | CMPLT | CMPLT | 2024-03-21 08:26 | 2024-03-21 08:26 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 08:25:43 AM, Rhuland, Courtney) PM completed by BD
65888 | 205087 | PM | CMPLT | CMPLT | 2024-03-21 08:32 | 2024-03-21 08:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 08:32:31 AM, Rhuland, Courtney) PM completed by BD
| 205088 | CM | CMPLT | CMPLT | 2024-03-21 08:45 | 2024-04-04 09:54 | Ordering parts for syringes
| 205089 | ACC | CMPLT | CMPLT | 2024-03-21 09:11 | 2024-09-25 09:35 | MSH BI HOLOGIC GAID AI SYSTEM DEMO | (TC 03/20/2024 12:10:00 PM, Whittington, Scott) MSH BI HOLOGIC GAID AI SYSTEM DEMO MEETING | MSH BI HOLOGIC GAID AI SYSTEM DEMO
66050 | 205090 | PM | CMPLT | CMPLT | 2024-03-21 09:32 | 2024-03-21 09:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68935 | 205091 | CM | CMPLT | CMPLT | 2024-03-21 09:41 | 2024-03-21 09:42 | Missing power cord strap and relevant accessories | (TC 03/21/2024 09:40:37 AM, Baghirzade, Nara) Replaced power cord strap, power cord strap snap, and screw.
66033 | 205092 | PM | CMPLT | CMPLT | 2024-03-21 09:41 | 2024-03-21 09:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68886 | 205093 | CM | CMPLT | CMPLT | 2024-03-21 09:42 | 2024-03-21 09:43 | Defective power cord | (TC 03/21/2024 09:42:20 AM, Baghirzade, Nara) Replaced power cord
68291 | 205094 | PM | CMPLT | CMPLT | 2024-03-21 09:44 | 2024-03-21 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68733 | 205095 | PM | CMPLT | CMPLT | 2024-03-21 09:47 | 2024-03-21 09:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68993 | 205096 | PM | CMPLT | CMPLT | 2024-03-21 09:49 | 2024-03-21 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68735 | 205097 | PM | CMPLT | CMPLT | 2024-03-21 09:52 | 2024-03-21 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059778 | 205098 | FC | CMPLT | CMPLT | 2024-03-21 09:54 | 2024-03-21 09:55 | QC checklist | (TC 03/21/2024 09:54:17 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
68791 | 205099 | PM | CMPLT | CMPLT | 2024-03-21 09:55 | 2024-03-21 09:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68823 | 205100 | PM | CMPLT | CMPLT | 2024-03-21 10:02 | 2024-03-21 10:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68708 | 205101 | PM | CMPLT | CMPLT | 2024-03-21 10:05 | 2024-03-21 10:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68807 | 205102 | PM | CMPLT | CMPLT | 2024-03-21 10:07 | 2024-03-21 10:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68978 | 205103 | PM | CMPLT | CMPLT | 2024-03-21 10:08 | 2024-03-21 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68844 | 205104 | PM | CMPLT | CMPLT | 2024-03-21 10:13 | 2024-03-21 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68758 | 205105 | PM | CMPLT | CMPLT | 2024-03-21 10:21 | 2024-03-21 10:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68883 | 205106 | PM | CMPLT | CMPLT | 2024-03-21 10:24 | 2024-03-21 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68880 | 205107 | PM | CMPLT | CMPLT | 2024-03-21 10:26 | 2024-03-21 10:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050250 | 205108 | PM | CMPLT | CMPLT | 2024-03-21 10:28 | 2024-03-21 10:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055113 | 205109 | PM | CMPLT | CMPLT | 2024-03-21 10:31 | 2024-03-21 10:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 205110 | CM | CMPLT | CMPLT | 2024-03-21 10:33 | 2024-03-25 11:07 | Sonomed Escalon VuMAX remote troubleshooting with vendor | (TC 03/21/2024 10:33:30 AM, Baghirzade, Nara) Had a scheduled remote troubleshooting session with Erik & Andrea from Sonomed. Connected Courtney's bluetooth network to ultrasound however Sonomed has issues with their teamviewer software and were unable to meet the appointment. Andrea to update us when they are able to remote in. (TC 03/25/2024 11:04:31 AM, Rhuland, Courtney) Connected to network via bluetooth for remote troubleshooting session Tech was able to perform tests and physician was satisfied with results
1061218 | 205111 | CM | CMPLT | CMPLT | 2024-03-21 10:33 | 2024-03-21 10:33 | ipad was not working | (TC 02/15/2024 10:33:53 AM, Huang, Jessica) ipad frozen display and showing passcode window. removed it from the wall and disconnected the power cord. Tried to power cycle it but it didn't turn off. Reported it to Maureen and informed them to switch the room. passcode window disappeared and the ipad started working. put it back to the room and informed Maureen
1058269 | 205112 | PM | CMPLT | CMPLT | 2024-03-21 10:34 | 2024-03-21 10:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69477 | 205113 | PM | CMPLT | CMPLT | 2024-03-21 10:39 | 2024-03-21 10:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1061204 | 205114 | CM | CMPLT | CMPLT | 2024-03-21 10:40 | 2024-03-21 10:40 | bluetooth not working | (TC 02/13/2024 10:40:29 AM, Huang, Jessica) checked the bluetooth in the room and connected it for the surgeon.
1061202 | 205115 | CM | CMPLT | CMPLT | 2024-03-21 10:45 | 2024-03-21 10:45 | replaced the dvi connector in OR 1
69308 | 205116 | PM | CMPLT | CMPLT | 2024-03-21 10:46 | 2024-03-21 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68646 | 205117 | PM | CMPLT | CMPLT | 2024-03-21 10:50 | 2024-03-21 10:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69511 | 205118 | PM | CMPLT | CMPLT | 2024-03-21 10:52 | 2024-03-21 10:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69454 | 205119 | PM | CMPLT | CMPLT | 2024-03-21 10:54 | 2024-03-21 10:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69094 | 205120 | PM | CMPLT | CMPLT | 2024-03-21 10:55 | 2024-03-21 10:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69471 | 205121 | PM | CMPLT | CMPLT | 2024-03-21 10:57 | 2024-03-21 10:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69091 | 205122 | PM | CMPLT | CMPLT | 2024-03-21 10:58 | 2024-03-21 10:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69458 | 205123 | PM | CMPLT | CMPLT | 2024-03-21 10:59 | 2024-03-21 10:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66040 | 205124 | PM | CMPLT | PMFAIL | 2024-03-21 11:27 | 2024-03-21 11:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
66040 | 205125 | CM | CMPLT | CMPLT | 2024-03-21 11:28 | 2024-03-21 11:38 | Created by Failed PM WO # 205124 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/21/2024 11:28:41 AM, Rendon Lopez , Elizabeth) Replaced front case, rear case and power cord. PM completed, all tests passed | PM failed by BD
65881 | 205126 | PM | CMPLT | CMPLT | 2024-03-21 12:45 | 2024-03-21 12:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 12:45:28 PM, Baghirzade, Nara) PM completed by BD.
68604 | 205127 | PM | CMPLT | CMPLT | 2024-03-21 12:47 | 2024-03-21 12:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 12:47:04 PM, Baghirzade, Nara) PM completed by BD.
| 205128 | CM | CMPLT | CMPLT | 2024-03-21 12:49 | 2024-03-21 12:54 | Vacation coverage form and emails | (TC 03/21/2024 12:52:48 PM, Baghirzade, Nara) Completed vacation coverage form to send to Adeel. Sent emails to 11N, 11S, NICU, and PFT/CPET. Followed up with Jessica, Courtney, and John regarding the coverage for the respective units.
1048522 | 205129 | CM | CMPLT | CMPLT | 2024-03-21 12:51 | 2024-03-21 12:56 | Broken cases | (TC 03/21/2024 12:54:52 PM, Rhuland, Courtney) Rear case replaced Front case replaced Mylar gasket replaced Reseated spring in plungerhead assembly Returned to BD for recall and PM
| 205130 | PS | ACTIV | ISSUE | 2024-03-21 13:40 |  | CADD-Solis OneRecall Notice | (TC 03/21/2024 01:40:26 PM, Reid, Emmerson ) Read the OneRecall Alert and reached out to the vendor to see about what our next steps are to be. (TC 03/22/2024 08:40:06 AM, Reid, Emmerson ) Did a pump run, elevators were very busy and I was stuck waiting. (TC 05/01/2024 01:48:59 PM, Reid, Emmerson ) Reached out to the vendor again regarding questions about the recall update and how it will affect our drug libraries. (TC 05/09/2024 02:52:00 PM, Reid, Emmerson ) Finally got my answer back from ICU medical and it looks like we should have no issues. I tested upgrading one pump to 4.3 and I was able to send the library. I sent over a list of all affected SNs and they are going to send back a workbook for us to fill out. (TC 05/10/2024 03:11:14 PM, Reid, Emmerson ) emailed between vendor and downloaded the workbook they sent (TC 06/10/2024 03:04:06 PM, Reid, Emmerson ) Downloaded the PharmGuard, and Cadd-solis download programs on John, Courtney, and Elizabeth's compluters so they can update their CADD pumps to the latest version (V4.3). Also moved the workbook onto the T-drive and shared the location with the team.
68607 | 205131 | PM | CMPLT | PMFAIL | 2024-03-21 13:58 | 2024-03-21 13:59 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 01:58:27 PM, Baghirzade, Nara) Broken rear and front case. Syringe requires full calibration.
68607 | 205132 | CM | CMPLT | CMPLT | 2024-03-21 13:59 | 2024-03-21 14:19 | Created by Failed PM WO # 205131 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/21/2024 02:18:16 PM, Baghirzade, Nara) Replaced: -rear case -front case -mylar gasket Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/21/2024 01:58:27 PM, Baghirzade, Nara) Broken rear and front case. Syringe requires full calibration.
1057550 | 205133 | PM | CMPLT | CMPLT | 2024-03-21 14:35 | 2024-03-21 14:38 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 03/21/2024 02:37:16 PM, Scero, Tomlyn) PM tests complete after repair was done. Battery replaced. Borrowed from ED stock.
1047284 | 205134 | CM | CMPLT | CMPLT | 2024-03-21 15:36 | 2024-05-01 08:17 | MSH XG PROBE PLT-1204 BX CRYSTAL REPORTED ARTIFACT | (TC 03/21/2024 03:36:00 PM, Whittington, Scott) MSH XG PROBE PLT-1204 BX CRYSTAL REPORTED ARTIFACT INITIAL EVIDENCE OF ARTIFACT, CRYSTAL FIRING PROBE LOCK WAS LOOSE & 90DEG OFF REPAIRED , PLACE IN RM 3 FOR USE | MSH XG PROBE PLT-1204 BX CRYSTAL REPORTED ARTIFACT
68599 | 205135 | CM | CMPLT | CMPLT | 2024-03-21 15:52 | 2024-04-02 11:38 | Created by Failed PM WO # 202103 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 03:03:04 PM, Rendon Lopez , Elizabeth) Replaced front case, rear case. Calibrated. PM completed, all tests passed. Waiting on left handle (TC 04/02/2024 11:37:54 AM, Rendon Lopez , Elizabeth) Replaced left handle-OK | PM failed by BD
65960 | 205136 | PM | CMPLT | PMFAIL | 2024-03-21 15:55 | 2024-03-21 15:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed by BD
65960 | 205137 | CM | CMPLT | CMPLT | 2024-03-21 15:55 | 2024-04-02 11:38 | Created by Failed PM WO # 205136 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 02:21:41 PM, Rendon Lopez , Elizabeth) Replaced front case, rear case. Calibrated. PM completed, all tests passed. Waiting on left handle (TC 04/02/2024 11:37:20 AM, Rendon Lopez , Elizabeth) Replaced left handle-OK | PM failed by BD
69556 | 205138 | CM | CMPLT | CMPLT | 2024-03-22 07:45 | 2024-03-22 07:47 | Broken latch | (TC 03/22/2024 07:46:57 AM, Rhuland, Courtney) Replaced door latch spring and screw Replaced IUI Returned to BD for recall and PM
1050035 | 205139 | CM | CMPLT | CMPLT | 2024-03-22 08:41 | 2024-03-22 08:42 | Repair before recall | (TC 03/21/2024 08:41:41 AM, Rendon Lopez , Elizabeth) Replaced left IUI, sent to BD for recall
69349 | 205140 | CM | CMPLT | CMPLT | 2024-03-22 08:43 | 2024-03-22 08:44 | Repair before recall | (TC 03/21/2024 08:43:07 AM, Rendon Lopez , Elizabeth) Replaced left and right IUI. Gave to BD for recall
65947 | 205141 | CM | CMPLT | CMPLT | 2024-03-22 08:55 | 2024-03-22 08:56 | Cracked cases | (TC 03/22/2024 08:55:32 AM, Rhuland, Courtney) Replaced front and rear cases Replaced IUI Replaced mylar gasket Reprinted barcode Returned to BD for recall and PM
65889 | 205142 | CM | CMPLT | CMPLT | 2024-03-22 09:34 | 2024-03-22 09:35 | Broken case | (TC 03/22/2024 09:34:35 AM, Rhuland, Courtney) Replaced rear case Replaced module latch
65616 | 205143 | CM | CMPLT | CMPLT | 2024-03-22 09:51 | 2024-03-22 09:52 | Bipolar Leads broken | (TC 03/22/2024 09:51:02 AM, Reid, Emmerson ) Was able to replace the bipolar leads with the last replacement. Tested device to make sure it was working.
66004 | 205144 | CM | CMPLT | CMPLT | 2024-03-22 09:58 | 2024-04-10 13:49 | Broken cases | (TC 04/10/2024 01:49:03 PM, Rhuland, Courtney) When replacing rear case, logic board was found to be damaged Replaced logic board Flashed PCU with all necessary softwares Flashed software was not the most recently updated software To be looked into further
68542 | 205145 | CM | CMPLT | CMPLT | 2024-03-22 10:02 | 2024-04-04 14:53 | Unable to see CO2 from AGM on philips monitors | (TC 03/22/2024 10:03:15 AM, Rendon Lopez , Elizabeth) Replaced EC10 module, cable, no light on module, not receiving information. Moved EC5 from Medibus 2 to 1, check configs, not resolved. Inquired infomration about new EC5 (TC 04/04/2024 02:51:44 PM, Rendon Lopez , Elizabeth) Removed adapter from AGM 68545, and replaced on this device. Tested-working. Placed AGM in OR 3
| 205146 | CM | CMPLT | CMPLT | 2024-03-22 10:03 | 2024-04-04 09:54 | Ordering parts for PCU - stock
68626 | 205147 | CM | CMPLT | CMPLT | 2024-03-22 10:22 | 2024-03-22 10:24 | Replacing power cord and power cord strap | (TC 03/22/2024 10:21:40 AM, Baghirzade, Nara) Replaced power cord. Replaced power cord strap, strap snap and screw.
69014 | 205148 | CM | CMPLT | CMPLT | 2024-03-22 10:24 | 2024-03-22 10:25 | Missing power cord strap | (TC 03/22/2024 10:24:07 AM, Baghirzade, Nara) Replaced power cord strap, strap snap and screw.
68951 | 205149 | CM | CMPLT | CMPLT | 2024-03-22 10:26 | 2024-03-22 10:27 | Missing power cord strap | (TC 03/22/2024 10:25:37 AM, Baghirzade, Nara) Replaced power cord strap, strap snap and screw.
68974 | 205150 | CM | CMPLT | CMPLT | 2024-03-22 10:27 | 2024-03-22 10:27 | Damaged power cord | (TC 03/22/2024 10:27:08 AM, Baghirzade, Nara) Replaced power cord
68990 | 205151 | CM | CMPLT | CMPLT | 2024-03-22 10:28 | 2024-03-22 10:28 | Damaged power cord | (TC 03/22/2024 10:27:39 AM, Baghirzade, Nara) Replaced power cord.
68772 | 205152 | CM | CMPLT | CMPLT | 2024-03-22 10:28 | 2024-03-22 10:28 | Damaged power cord | (TC 03/22/2024 10:28:10 AM, Baghirzade, Nara) Replaced power cord.
65940 | 205153 | PM | CMPLT | CMPLT | 2024-03-22 10:30 | 2024-03-22 10:30 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 10:29:47 AM, Baghirzade, Nara) PM completed by BD.
68610 | 205154 | PM | CMPLT | CMPLT | 2024-03-22 10:31 | 2024-03-22 10:31 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 10:30:58 AM, Baghirzade, Nara) PM completed by BD.
69620 | 205155 | PM | CMPLT | CMPLT | 2024-03-22 10:48 | 2024-03-22 11:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2024 11:06:14 AM, Reid, Emmerson ) Replaced front cover. Completed calibration, and PM. PM Complete.
68592 | 205156 | CM | CMPLT | CMPLT | 2024-03-22 10:57 | 2024-03-22 11:53 | Broken rear case | (TC 03/22/2024 10:57:06 AM, Baghirzade, Nara) Replaced rear case. Replaced both IUIs. Printed new asset sticker.
1065114 | 205157 | CM | CMPLT | CMPLT | 2024-03-22 11:21 | 2024-03-22 11:22 | Printer not working | (TC 03/22/2024 11:20:49 AM, Rhuland, Courtney) Printer would not clear cancelled doc Restarted computer and doc cleared Test print - OK
1048525 | 205158 | CM | CMPLT | CMPLT | 2024-03-22 11:59 | 2024-03-22 12:00 | Created by Failed PM WO # 199593 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 11:59:55 AM, Baghirzade, Nara) Replaced rear case. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/22/2024 11:58:48 AM, Baghirzade, Nara) Broken rear case and syringe requires full calibration
65942 | 205159 | PM | CMPLT | PMFAIL | 2024-03-22 12:02 | 2024-03-22 12:04 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 12:02:48 PM, Baghirzade, Nara) Broken front case and wrong asset number attached. Asset number does not match the serial number on TMS.
65942 | 205160 | CM | CMPLT | CMPLT | 2024-03-22 12:04 | 2024-03-22 12:43 | Created by Failed PM WO # 205159 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/22/2024 12:42:34 PM, Baghirzade, Nara) Replaced mylar gasket and pressure disc flag. Replaced front case. Reprinted new asset sticker, previous asset number sticker was incorrect as it did not match the serial number. (previous asset sticker had # 65924 and the correct asset # is 65942). Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/22/2024 12:02:48 PM, Baghirzade, Nara) Broken front case and wrong asset number attached. Asset number does not match the serial number on TMS.
68595 | 205161 | CM | CMPLT | CMPLT | 2024-03-22 13:04 | 2024-03-25 13:58 | Broken case | (TC 03/22/2024 01:04:49 PM, Rhuland, Courtney) Replaced internal frame Replaced top disk holder Replaced rear case Replaced both IUI connectors Replaced module latch Waiting on left handles (TC 03/25/2024 01:57:14 PM, Rhuland, Courtney) Handle replaced Returned to BD for recall and PM
1048529 | 205162 | CM | CMPLT | CMPLT | 2024-03-22 13:44 | 2024-03-22 13:45 | Broken internal frame | (TC 03/22/2024 01:43:49 PM, Rhuland, Courtney) Replaced internal frame PM to be completed in PM workorder
68640 | 205163 | PM | CMPLT | CMPLT | 2024-03-22 14:08 | 2024-03-22 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/22/2024 02:23:00 PM, Reid, Emmerson ) Ran calibration, and then ASM PM again. PM Complete.
1057021 | 205164 | CM | CMPLT | CMPLT | 2024-03-22 15:04 | 2024-03-22 15:07 | Not turning on | (TC 03/22/2024 03:07:20 PM, Canjimavadivel, Kamallesh ) Rechargeable battery was dead. Replaced with regular batteries. Need to order rechargeable.
1057018 | 205165 | CM | CMPLT | CMPLT | 2024-03-22 15:05 | 2024-03-22 15:05 | Not turning on | (TC 03/22/2024 03:05:01 PM, Canjimavadivel, Kamallesh ) Rechargeable battery was dead. Replaced with regular batteries. Need to order rechargeable.
1054725 | 205166 | CM | CMPLT | CMPLT | 2024-03-22 15:59 | 2024-03-22 15:59 | MSH 5 IR EPIQ US TRACKBALL CLEANING | (TC 03/22/2024 03:59:13 PM, Whittington, Scott) MSH 5 IR EPIQ US TRACKBALL CLEANING | MSH 5 IR EPIQ US TRACKBALL CLEANING
1054537 | 205167 | PM | CMPLT | CMPLT | 2024-03-22 16:06 | 2024-03-22 16:08 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/22/2024 04:06:32 PM, Cortez, John) There seems to be an intermittent issue with port 3 not being detected. Will look into replacing sensor. Completed cleaning cycle
66049 | 205168 | CM | CMPLT | CMPLT | 2024-03-25 07:52 | 2024-03-25 07:53 | Broken front case | (TC 03/25/2024 07:52:25 AM, Rhuland, Courtney) Replaced front case
| 205169 | PA | CMPLT | CMPLT | 2024-03-25 08:16 | 2024-04-15 10:07 | Battery modules for servo-u/n | (TC 03/25/2024 08:21:38 AM, Canjimavadivel, Kamallesh ) Placed order for battery modules
65981 | 205170 | CM | CMPLT | CMPLT | 2024-03-25 08:30 | 2024-03-25 08:31 | Channel error | (TC 03/25/2024 08:30:20 AM, Baghirzade, Nara) Downloaded error logs; error was pertaining to the channel being used with a PCU that has not had a software upgrade. Incompatible softwares between the syringe and PCU. Connected the syringe to an upgrade PCU and no errors observed.
65937 | 205171 | CM | CMPLT | CMPLT | 2024-03-25 08:32 | 2024-03-25 08:32 | Channel error | (TC 03/25/2024 08:31:41 AM, Baghirzade, Nara) Downloaded error logs; error was pertaining to the channel being used with a PCU that has not had a software upgrade. Incompatible softwares between the syringe and PCU. Connected the syringe to an upgrade PCU and no errors observed.
68904 | 205172 | CM | CMPLT | CMPLT | 2024-03-25 08:33 | 2024-03-25 08:33 | Broken power cord and both IUIs defective | (TC 03/25/2024 08:32:44 AM, Baghirzade, Nara) Replaced power cord. Replaced both left and right IUIs. Returned PCU to BD for remediation and PM.
| 205173 | RDS | CMPLT | CMPLT | 2024-03-25 09:01 | 2024-03-26 11:47 | Weekly checks in CCU | (TC 03/25/2024 02:21:13 PM, Scero, Tomlyn) Weekly checks complete. 30 J tests passed.
66034 | 205174 | PM | CMPLT | CMPLT | 2024-03-25 09:41 | 2024-03-25 09:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66028 | 205175 | PM | CMPLT | CMPLT | 2024-03-25 09:47 | 2024-03-25 09:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65935 | 205176 | PM | CMPLT | CMPLT | 2024-03-25 09:49 | 2024-03-25 09:49 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 09:48:44 AM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68292 | 205177 | PM | CMPLT | CMPLT | 2024-03-25 09:50 | 2024-03-25 09:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68875 | 205178 | PM | CMPLT | CMPLT | 2024-03-25 09:53 | 2024-03-25 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68750 | 205179 | PM | CMPLT | CMPLT | 2024-03-25 09:56 | 2024-03-25 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68962 | 205180 | PM | CMPLT | CMPLT | 2024-03-25 09:59 | 2024-03-25 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1053349 | 205181 | CM | CMPLT | CMPLT | 2024-03-25 09:59 | 2024-05-01 08:16 | MSH BI H2 ROOM 1274 HOLOGIC COMPUTER UPGRADE MAC ADDRESS CHANGING! confirming service for April 25 and 26 for the upgrade in Room 1274 (H2). | (TC 05/01/2024 08:14:29 AM, Whittington, Scott) MSH BI H2 ROOM 1274 HOLOGIC COMPUTER UPGRADE MAC ADDRESS CHANGING! confirming service for April 25 and 26 for the upgrade in Room 1274 (H2). COMPLETED | MSH BI H2 ROOM 1274 HOLOGIC COMPUTER UPGRADE MAC ADDRESS CHANGING! confirming service for April 25 and 26 for the upgrade in Room 1274 (H2).
68797 | 205182 | PM | CMPLT | CMPLT | 2024-03-25 10:02 | 2024-03-25 10:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68833 | 205183 | PM | CMPLT | CMPLT | 2024-03-25 10:11 | 2024-03-25 10:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058263 | 205184 | CM | CMPLT | CMPLT | 2024-03-25 10:18 | 2024-03-25 10:20 | Network communication error 600.6870 | (TC 03/25/2024 10:18:38 AM, Rendon Lopez , Elizabeth) Replaced wireless assembly, pushed network configuration. Connecting, no errors. Returned to BD for upgrade
68922 | 205185 | PM | CMPLT | CMPLT | 2024-03-25 10:21 | 2024-03-25 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68922 | 205186 | PM | CANCL | DUP | 2024-03-25 10:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68717 | 205187 | PM | CMPLT | CMPLT | 2024-03-25 10:37 | 2024-03-25 10:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68717 | 205188 | PM | CANCL | DUP | 2024-03-25 10:37 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68747 | 205189 | PM | CMPLT | CMPLT | 2024-03-25 10:41 | 2024-03-25 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69452 | 205190 | PM | CMPLT | CMPLT | 2024-03-25 11:12 | 2024-03-25 11:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66000 | 205191 | CM | CMPLT | CMPLT | 2024-03-25 11:16 | 2024-03-25 11:17 | Created by Failed PM WO # 201375 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/25/2024 11:16:15 AM, Rendon Lopez , Elizabeth) Damaged power cord, replaced. PM completed, all tests passed | PM failed by BD
69092 | 205192 | PM | CMPLT | CMPLT | 2024-03-25 11:22 | 2024-03-25 11:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69479 | 205193 | PM | CMPLT | CMPLT | 2024-03-25 11:25 | 2024-03-25 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69170 | 205194 | PM | CMPLT | PMFAIL | 2024-03-25 11:28 | 2024-03-25 11:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69170 | 205195 | CM | CMPLT | CMPLT | 2024-03-25 11:28 | 2024-03-25 11:29 | Created by Failed PM WO # 205194 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/25/2024 11:28:34 AM, Rendon Lopez , Elizabeth) Keypad damaged. Replaced front case/keypad. PM completed, all tests done | PM failed by BD
69173 | 205196 | PM | CMPLT | CMPLT | 2024-03-25 11:29 | 2024-03-25 11:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69563 | 205197 | PM | CMPLT | CMPLT | 2024-03-25 11:31 | 2024-03-25 11:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69188 | 205198 | PM | CMPLT | CMPLT | 2024-03-25 11:33 | 2024-03-25 11:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69409 | 205199 | PM | CMPLT | CMPLT | 2024-03-25 11:38 | 2024-03-25 11:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69108 | 205200 | PM | CMPLT | CMPLT | 2024-03-25 11:40 | 2024-03-25 11:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69396 | 205201 | CM | CMPLT | CMPLT | 2024-03-25 11:47 | 2024-04-11 13:13 | Created by Failed PM WO # 200650 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/25/2024 11:47:11 AM, Rendon Lopez , Elizabeth) Replaced rear case. Fails rate accuracy, needs new bezel. (TC 04/02/2024 02:34:01 PM, Rendon Lopez , Elizabeth) Replaced Bezel assy. Platen/ Spring hinge Assembly damaged, inquired pricing info (TC 04/11/2024 11:12:53 AM, Rendon Lopez , Elizabeth) Replaced platen assy. Calibrated. PM completed, all tests passed | PM failed by BD
69047 | 205202 | PM | CMPLT | CMPLT | 2024-03-25 11:49 | 2024-03-25 11:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68651 | 205203 | PM | CMPLT | CMPLT | 2024-03-25 12:02 | 2024-03-25 12:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65917 | 205204 | PM | CMPLT | CMPLT | 2024-03-25 13:13 | 2024-03-25 13:14 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:13:28 PM, Baghirzade, Nara) PM completed by BD.
65955 | 205205 | PM | CMPLT | CMPLT | 2024-03-25 13:15 | 2024-03-25 13:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:14:59 PM, Baghirzade, Nara) PM completed by BD.
65986 | 205206 | PM | CMPLT | CMPLT | 2024-03-25 13:18 | 2024-03-25 13:18 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:18:04 PM, Baghirzade, Nara) PM completed by BD.
68592 | 205207 | PM | CMPLT | CMPLT | 2024-03-25 13:20 | 2024-03-25 13:20 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:19:40 PM, Baghirzade, Nara) PM completed by BD.
1050618 | 205208 | PM | CMPLT | CMPLT | 2024-03-25 13:21 | 2024-03-25 13:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:20:46 PM, Baghirzade, Nara) PM completed by BD.
63437 | 205209 | CM | CMPLT | CMPLT | 2024-03-25 13:27 | 2024-03-25 13:28 | Pump not charging | (TC 03/25/2024 01:27:20 PM, Baghirzade, Nara) Replaced battery. Ran feed for a couple of hours, no issues or errors observed. Rotor is functional. Occlusion alarm is functional.
1050614 | 205210 | PM | CMPLT | PMFAIL | 2024-03-25 13:45 | 2024-03-25 13:45 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:45:13 PM, Baghirzade, Nara) Broken rear case
1050614 | 205211 | CM | CMPLT | CMPLT | 2024-03-25 13:45 | 2024-03-25 13:46 | Created by Failed PM WO # 205210 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:45:43 PM, Baghirzade, Nara) Replaced rear case. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 03/25/2024 01:45:13 PM, Baghirzade, Nara) Broken rear case
65907 | 205212 | PM | CMPLT | PMFAIL | 2024-03-25 13:48 | 2024-03-25 13:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:47:57 PM, Baghirzade, Nara) Pressure disc holder broken.
65907 | 205213 | PM | CANCL | CANCL | 2024-03-25 13:48 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65907 | 205214 | PM | CMPLT | CMPLT | 2024-03-25 13:48 | 2024-03-25 13:50 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/25/2024 01:49:38 PM, Baghirzade, Nara) Replaced pressure sensor disc holder. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. (TC 03/25/2024 01:49:38 PM, Baghirzade, Nara) Replaced pressure sensor disc holder. Mylar gasket replaced. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65907 | 205215 | CM | CMPLT | CMPLT | 2024-03-25 13:48 | 2024-03-25 13:50 | Created by Failed PM WO # 205212 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe |  | (TC 03/25/2024 01:47:57 PM, Baghirzade, Nara) Pressure disc holder broken.
| 205216 | CM | CMPLT | CMPLT | 2024-03-25 14:03 | 2024-03-26 13:56 | Infusion testing for NICU | (TC 03/25/2024 02:02:58 PM, Rhuland, Courtney) Performed infusion testing for NICU as requested by Janet Testing 1.2u filter with infusion rates to determine best location for the filter Using lipids and filter determined location is on the syringe side as that is what is recommended by BD and occlusion location can be determined.
69292 | 205217 | PM | CMPLT | CMPLT | 2024-03-25 15:07 | 2024-03-25 15:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | BD Remediation
69177 | 205218 | PM | CMPLT | CMPLT | 2024-03-25 15:10 | 2024-03-25 15:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | BD Remediation
| 205219 | CM | CMPLT | CMPLT | 2024-03-25 15:21 | 2024-03-25 15:22 | Suction Issue - OR 13 | Call re: Suction not working. Went to check. Waited for surgeons to allow into the room. Checekd and found suction filter full of fluid. Removed it. Borrowed a filter from ER. Replaced filter. Tested ok.
| 205220 | CM | CMPLT | CMPLT | 2024-03-25 15:34 | 2024-03-25 15:43 | IVF - Lightsource issue | Call from IVF clini re: lightsource (S/N: 7667256) not working. Went to 250 Dundas to check. Waited for the staff provide access to room. Checked lightsource functionality. Could not find any issue. Requersted waited for staff to bring a scope. Checked functionality again. Could not find any issue. Determined it was a use error. Advised physician on proper usage. Tested ok. Returned to service.
| 205221 | CM | CMPLT | CMPLT | 2024-03-25 16:20 | 2024-04-26 13:35 | MSH OPG 3 CEOU RM 11 IU22 NOT STARTING RELIABLY | (TC 03/25/2024 04:20:38 PM, Dolinar, John) MSH OPG 3 CEOU RM 11 IU22 NOT STARTING RELIABLY | MSH OPG 3 CEOU RM 11 IU22 NOT STARTING RELIABLY
| 205222 | RDS | CMPLT | CMPLT | 2024-03-26 08:30 | 2024-03-26 08:32 | Weekly Checks For PACU | (TC 03/26/2024 08:31:02 AM, Reid, Emmerson ) All checks good, also found a seet of unremediated pumps in PACU and OR thaT i brought down
1061202 | 205223 | CM | CMPLT | NFF | 2024-03-26 08:32 | 2024-03-26 08:33 | Monitor not working with Microscope | (TC 03/26/2024 08:32:18 AM, Reid, Emmerson ) Went up to check and it was working.
65627 | 205224 | CM | CMPLT | CMPLT | 2024-03-26 09:12 | 2024-03-26 10:40 | Waveforms not all showing | (TC 03/25/2024 09:11:42 AM, Canjimavadivel, Kamallesh ) Amanda found issue when trying to use in skills lab. Verified config on monitor (TC 03/26/2024 10:32:12 AM, Canjimavadivel, Kamallesh ) Compared the options with another monitor in the unit, they were not the same. Asked Kamran to help with adding options to the monitor. Tested ok, informed Amanda
66711 | 205225 | CM | CMPLT | CMPLT | 2024-03-26 10:25 | 2024-03-26 10:25 | Broken doors | (TC 03/26/2024 10:24:52 AM, Rhuland, Courtney) Repaired broken porthole doors Calibrated RH Functional check - OK
65953 | 205226 | CM | CMPLT | CMPLT | 2024-03-26 10:44 | 2024-03-26 10:55 | Broken cases | (TC 03/26/2024 10:43:59 AM, Rhuland, Courtney) Replaced front case Replaced mylar gasket Replaced internal frame Replaced rear case Replaced left handle Replaced both IUIs Returned to BD for recall and PM
1063746 | 205227 | PM | CMPLT | CMPLT | 2024-03-26 11:03 | 2024-03-26 11:04 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/26/2024 11:03:10 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
68935 | 205228 | PM | CMPLT | CMPLT | 2024-03-26 11:16 | 2024-03-26 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM completed
68956 | 205229 | PM | CMPLT | CMPLT | 2024-03-26 11:54 | 2024-03-26 11:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
68878 | 205230 | PM | CMPLT | CMPLT | 2024-03-26 11:56 | 2024-03-26 11:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
69030 | 205231 | PM | CMPLT | CMPLT | 2024-03-26 11:58 | 2024-03-26 11:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
69151 | 205232 | PM | CMPLT | CMPLT | 2024-03-26 12:01 | 2024-03-26 12:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68982 | 205233 | PM | CMPLT | CMPLT | 2024-03-26 12:01 | 2024-03-26 12:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
1069509 | 205234 | CM | CMPLT | CMPLT | 2024-03-26 12:02 | 2024-03-26 12:03 | MSH 5 GEN US PRISM PROBE PLI1205 BX REPORTED ARTIFACT IN ROOM 8 | (TC 03/26/2024 12:02:26 PM, Whittington, Scott) MSH 5 GEN US PRISM PROBE PLI1205 BX REPORTED ARTIFACT IN ROOM 8 TESTED IN RM 6 OK RETESTED IN RM 8 OK PROBABLY JUST MOMENTARY PIN CONNECTION PROBLEM | MSH 5 GEN US PRISM PROBE PLI1205 BX REPORTED ARTIFACT IN ROOM 8
68796 | 205235 | PM | CMPLT | CMPLT | 2024-03-26 12:03 | 2024-03-26 12:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
68711 | 205236 | PM | CMPLT | CMPLT | 2024-03-26 12:05 | 2024-03-26 12:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
69035 | 205237 | PS | CMPLT | CMPLT | 2024-03-26 12:07 | 2024-03-26 12:07 | 2024 Alaris Software Recall
69035 | 205238 | PM | CMPLT | CMPLT | 2024-03-26 12:07 | 2024-03-26 12:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68622 | 205239 | PM | CMPLT | CMPLT | 2024-03-26 12:08 | 2024-03-26 12:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
32809 | 205240 | CM | CMPLT | CMPLT | 2024-03-26 12:08 | 2024-04-12 09:16 | MSH DR CR CASSETTES FOR LLI REQ BAR CODE LABELS | (TC 03/26/2024 12:08:06 PM, Whittington, Scott) MSH DR CR CASSETTES FOR LLI REQ BAR CODE LABELS (TC 03/26/2024 12:28:51 PM, Dolinar, John) MSH DR CR CASSETTES FOR LLI REQ BAR CODE LABELS MAKE LABELS WITH BROTHER PRINTER | MSH DR CR CASSETTES FOR LLI REQ BAR CODE LABELS
69265 | 205241 | PM | CMPLT | CMPLT | 2024-03-26 12:13 | 2024-03-26 12:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68946 | 205242 | PM | CMPLT | CMPLT | 2024-03-26 12:13 | 2024-03-26 12:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
1050249 | 205243 | PM | CMPLT | CMPLT | 2024-03-26 12:16 | 2024-03-26 12:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
1047003 | 205244 | PM | CMPLT | CMPLT | 2024-03-26 12:17 | 2024-03-26 12:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051377 | 205245 | PM | CMPLT | CMPLT | 2024-03-26 12:19 | 2024-03-26 12:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM Completed
1047004 | 205246 | PM | CMPLT | CMPLT | 2024-03-26 12:23 | 2024-03-26 12:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1048500 | 205247 | CM | CMPLT | CMPLT | 2024-03-26 12:26 | 2024-04-26 13:34 | MSH US EPIQ WIRELESS NOT STAYING CONNECT TO MSHMED WIFI | (TC 03/26/2024 01:30:59 PM, Whittington, Scott) MSH US EPIQ WIRELESS NOT STAYING CONNECT TO MSHMED WIFI KRISTA REPORTED WIRED IS INFREQUENTLY USED DISABLE WIRED, LEAVING ONLY WIRELESS ENABELD | MSH US EPIQ WIRELESS NOT STAYING CONNECT TO MSHMED WIFI From: Calabrese, Angela <Angela.Calabrese@sinaihealth.ca> Sent: March 26, 2024 12:04 PM To: Ultrasound Support <Ultrasound.Support@uhn.ca> Cc: Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca> Subject: [External] Wireless connectivity issues with Philips Epiq unit at SHS We have been experiencing some intermittent issues with the WIFI connection when using the Philips Epiq at SHS – Level 5. This unit is used portably, so WIFI is used regularly. The techs have to repeatedly reconnect to WIFI. This is not an urgent issue, but I would appreciate if someone could take a look at the unit to see if anything can be done to correct the issue. Thanks, Angela Calabrese
1047006 | 205248 | PM | CMPLT | CMPLT | 2024-03-26 12:28 | 2024-03-26 12:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047007 | 205249 | PM | CMPLT | CMPLT | 2024-03-26 12:33 | 2024-03-26 12:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047010 | 205250 | PM | CMPLT | CMPLT | 2024-03-26 12:40 | 2024-03-26 12:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047011 | 205251 | PM | CMPLT | CMPLT | 2024-03-26 12:47 | 2024-03-26 12:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68951 | 205252 | PM | CMPLT | CMPLT | 2024-03-26 12:56 | 2024-03-26 12:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68626 | 205253 | PM | CMPLT | CMPLT | 2024-03-26 12:59 | 2024-03-26 12:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1047014 | 205254 | PM | CMPLT | CMPLT | 2024-03-26 13:00 | 2024-03-26 13:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68701 | 205255 | PM | CMPLT | CMPLT | 2024-03-26 13:01 | 2024-03-26 13:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68990 | 205256 | PM | CMPLT | CMPLT | 2024-03-26 13:03 | 2024-03-26 13:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1047015 | 205257 | PM | CMPLT | CMPLT | 2024-03-26 13:07 | 2024-03-26 13:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68940 | 205258 | PM | CMPLT | CMPLT | 2024-03-26 13:08 | 2024-03-26 13:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68952 | 205259 | PM | CMPLT | CMPLT | 2024-03-26 13:11 | 2024-03-26 13:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68679 | 205260 | PM | CMPLT | CMPLT | 2024-03-26 13:13 | 2024-03-26 13:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68916 | 205261 | PM | CMPLT | CMPLT | 2024-03-26 13:15 | 2024-03-26 13:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68974 | 205262 | PM | CMPLT | CMPLT | 2024-03-26 13:17 | 2024-03-26 13:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1047049 | 205263 | PM | CMPLT | CMPLT | 2024-03-26 13:24 | 2024-03-26 13:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047050 | 205264 | PM | CMPLT | CMPLT | 2024-03-26 13:30 | 2024-03-26 13:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055161 | 205265 | PM | CMPLT | CMPLT | 2024-03-26 13:32 | 2024-03-26 13:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1054035 | 205266 | CM | CMPLT | CMPLT | 2024-03-26 13:33 | 2024-03-26 13:34 | "Loss of U/S data message" | (TC 03/25/2024 01:33:19 PM, Cortez, John) Confirmed that it was a probe issue, replaced with spare
1065812 | 205267 | PM | CMPLT | CMPLT | 2024-03-26 13:35 | 2024-03-26 13:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1054413 | 205268 | CM | CMPLT | CMPLT | 2024-03-26 13:36 | 2024-03-26 13:36 | Battery dead, probe holder low | (TC 03/26/2024 01:35:49 PM, Cortez, John) Readjusted stand, charged battery overnight
1065910 | 205269 | PM | CMPLT | CMPLT | 2024-03-26 13:37 | 2024-03-26 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065970 | 205270 | PM | CMPLT | CMPLT | 2024-03-26 13:40 | 2024-03-26 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065989 | 205271 | PM | CMPLT | CMPLT | 2024-03-26 13:43 | 2024-03-26 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066005 | 205272 | PM | CMPLT | CMPLT | 2024-03-26 13:47 | 2024-03-26 13:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066114 | 205273 | FC | CMPLT | CMPLT | 2024-03-26 13:54 | 2024-03-26 13:56 | QC checklist | (TC 03/26/2024 01:54:34 PM, Rhuland, Courtney) QC checklist complete
| 205274 | CM | CMPLT | CMPLT | 2024-03-26 13:57 | 2024-04-10 14:29 | PCU IP address issue | (TC 03/26/2024 01:56:58 PM, Rhuland, Courtney) Issue coming from BD saying there are 2 PCUs with the same serial number and 2 IP addresses Searched for the device that is in TMS with the specified serial number (66001) could not find Asked IT to tell us where the IP address is and searched that area for the PCU and could not find (TC 03/27/2024 02:58:18 PM, Rhuland, Courtney) Continued search for PCU Found asset and serial number from MAC address on TMS Gave info to BD when performing recall (TC 03/28/2024 12:42:53 PM, Rhuland, Courtney) Looked for PCU again Checked which unit PCUs were returned to during the first week Found one PCU with IP address ending in 232 Brought it to biomed and disabled network and will keep in biomed until IP address ending in 235 is found (TC 04/03/2024 01:07:54 PM, Rhuland, Courtney) Second PCU was upgradeed by BD - listed as the same asset and serial number Searched BD workspace and endo, 10S, oncology Will wait for device to be turned on again to search (TC 04/10/2024 02:28:36 PM, Rhuland, Courtney) Other PCU was returned from BD Serial numbers have been updated
1054962 | 205275 | CM | CMPLT | CMPLT | 2024-03-26 14:04 | 2024-03-26 14:07 | "No bp hose" | (TC 03/26/2024 02:04:18 PM, Cortez, John) Added hose, cuff Cleaned filter, added rad battery All parameters tested ok
68805 | 205276 | PM | CMPLT | CMPLT | 2024-03-26 14:08 | 2024-03-26 14:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68772 | 205277 | PM | CMPLT | CMPLT | 2024-03-26 14:14 | 2024-03-26 14:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68921 | 205278 | PM | CMPLT | CMPLT | 2024-03-26 14:16 | 2024-03-26 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65880 | 205279 | PM | CMPLT | CMPLT | 2024-03-26 14:19 | 2024-03-26 14:19 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/26/2024 02:18:39 PM, Rhuland, Courtney) PM completed by BD
68997 | 205280 | PM | CMPLT | CMPLT | 2024-03-26 14:19 | 2024-03-26 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68834 | 205281 | PM | CMPLT | CMPLT | 2024-03-26 14:21 | 2024-03-26 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
65989 | 205282 | PM | CMPLT | CMPLT | 2024-03-26 14:22 | 2024-03-26 14:23 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/26/2024 02:21:41 PM, Rhuland, Courtney) PM completed by BD
68720 | 205283 | PM | CMPLT | CMPLT | 2024-03-26 14:26 | 2024-03-26 14:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68968 | 205284 | PM | CMPLT | CMPLT | 2024-03-26 14:28 | 2024-03-26 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68870 | 205285 | PM | CMPLT | CMPLT | 2024-03-26 14:29 | 2024-03-26 14:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205286 | CM | CMPLT | CMPLT | 2024-03-26 14:36 | 2024-04-04 09:55 | Bulb for light in amnio room | (TC 03/26/2024 02:40:41 PM, Rhuland, Courtney) Went over to OPG to look at the light Light not working, looks very old Ordered replacement bulbs to see if this is the issue
68729 | 205287 | PM | CMPLT | CMPLT | 2024-03-26 14:36 | 2024-03-26 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68932 | 205288 | PM | CMPLT | CMPLT | 2024-03-26 14:39 | 2024-03-26 14:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205289 | CM | ACTIV | ISSUE | 2024-03-26 14:40 |  | Philips monitor Configuration changes on Zeus AGM | (TC 03/26/2024 02:44:29 PM, Scero, Tomlyn) Monitor config changes on device that was down for PM (TC 08/19/2024 11:05:56 AM, Scero, Tomlyn) 3 Zeus configs complete in the main OR (TC 08/29/2024 02:57:34 PM, Scero, Tomlyn) monitor config changes (TC 09/05/2024 02:15:28 PM, Scero, Tomlyn) changed monitor config on 1 of the L&D zeus that was downstairs for repair (TC 09/10/2024 02:24:57 PM, Scero, Tomlyn) Started monitor changes in L&D (TC 09/20/2024 11:01:29 AM, Scero, Tomlyn) L&D monitor changes. OR 1 and 5 (TC 09/24/2024 11:56:12 AM, Scero, Tomlyn) Finished with OR 3 and 4 in L&D (TC 09/25/2024 02:43:06 PM, Scero, Tomlyn) changed OR 14 (TC 10/07/2024 01:48:56 PM, Scero, Tomlyn) Found 2 more AGMs to change config.
1050035 | 205290 | PM | CMPLT | CMPLT | 2024-03-26 14:42 | 2024-03-26 14:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1050240 | 205291 | PM | CMPLT | CMPLT | 2024-03-26 14:47 | 2024-03-26 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205292 | RDS | CMPLT | CMPLT | 2024-03-26 14:48 | 2024-03-26 14:48 | weekly check for 14S and 17M | weekly check for 14S and 17M
| 205293 | CM | CMPLT | CMPLT | 2024-03-26 14:52 | 2024-03-26 14:52 | worked with RT | worked with RT
| 205294 | CM | CMPLT | CMPLT | 2024-03-26 15:18 | 2024-03-26 15:18 | worked on TMS | worked on TMS
1062002 | 205295 | CM | CMPLT | CMPLT | 2024-03-26 16:11 | 2024-03-26 16:12 | "bed won't move up/down" | (TC 03/26/2024 04:10:50 PM, Cortez, John) Power cycled bed, cleared error log. Issue seems to resolve
69025 | 205296 | PM | CMPLT | CMPLT | 2024-03-26 16:47 | 2024-03-26 16:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/26/2024 04:47:53 PM, Cortez, John) Added RJ45 connector cover PM complete
69025 | 205297 | PM | CANCL | DUP | 2024-03-26 16:47 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1061957 | 205298 | CM | CMPLT | CMPLT | 2024-03-27 09:23 | 2024-03-27 09:24 | "no bed functions" | (TC 03/27/2024 09:23:06 AM, Cortez, John) error 3009-3-1-1306 Power cycling seems to resolve issue, unable to fully test with patient in bed. Bed alarm works
1061203 | 205299 | CM | CMPLT | CMPLT | 2024-03-27 09:27 | 2024-03-27 14:41 | PACS PC not on monitors | (TC 03/27/2024 02:39:45 PM, Reid, Emmerson ) They had the monitors on extend and wanted douplicate. I made the change and showed them what to do next time.
1059778 | 205300 | FC | CMPLT | CMPLT | 2024-03-27 09:28 | 2024-03-27 09:28 | QC checklist | (TC 03/27/2024 09:27:51 AM, Rhuland, Courtney) QC checklist completed Returned to 11S
1059828 | 205301 | PM | CMPLT | CMPLT | 2024-03-27 09:37 | 2024-03-27 09:38 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/27/2024 09:37:17 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
69125 | 205302 | PM | CMPLT | CMPLT | 2024-03-27 10:08 | 2024-03-27 10:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69465 | 205303 | PM | CMPLT | CMPLT | 2024-03-27 10:09 | 2024-03-27 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69306 | 205304 | PM | CMPLT | CMPLT | 2024-03-27 10:11 | 2024-03-27 10:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69580 | 205305 | PM | CMPLT | CMPLT | 2024-03-27 10:13 | 2024-03-27 10:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69565 | 205306 | PM | CMPLT | CMPLT | 2024-03-27 10:16 | 2024-03-27 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69446 | 205307 | PM | CMPLT | CMPLT | 2024-03-27 10:18 | 2024-03-27 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69317 | 205308 | PM | CMPLT | CMPLT | 2024-03-27 10:19 | 2024-03-27 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69118 | 205309 | PM | CMPLT | CMPLT | 2024-03-27 10:20 | 2024-03-27 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69391 | 205310 | PM | CMPLT | CMPLT | 2024-03-27 10:22 | 2024-03-27 10:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1056125 | 205311 | PM | CMPLT | CMPLT | 2024-03-27 10:23 | 2024-03-27 10:24 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 03/27/2024 10:23:04 AM, Canjimavadivel, Kamallesh ) Cleaned hoses and nozzles Cleaning cycle complete
69172 | 205312 | PM | CMPLT | CMPLT | 2024-03-27 10:23 | 2024-03-27 10:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69257 | 205313 | PM | CMPLT | CMPLT | 2024-03-27 10:25 | 2024-03-27 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205314 | RDS | CMPLT | CMPLT | 2024-03-27 10:25 | 2024-04-02 08:45 | Weekly check for ED
| 205315 | RDS | CMPLT | CMPLT | 2024-03-27 10:27 | 2024-03-27 10:36 | Weekly check for ED
69599 | 205316 | PM | CANCL | CANCL | 2024-03-27 10:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68649 | 205317 | PM | CMPLT | CMPLT | 2024-03-27 10:30 | 2024-03-27 10:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69293 | 205318 | PM | CMPLT | CMPLT | 2024-03-27 10:33 | 2024-03-27 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69547 | 205319 | PM | CMPLT | CMPLT | 2024-03-27 10:36 | 2024-03-27 10:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69545 | 205320 | PM | CMPLT | CMPLT | 2024-03-27 10:38 | 2024-03-27 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69543 | 205321 | PM | CMPLT | CMPLT | 2024-03-27 10:40 | 2024-03-27 10:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69227 | 205322 | PM | CMPLT | CMPLT | 2024-03-27 10:41 | 2024-03-27 10:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
66249 | 205323 | CM | CMPLT | CMPLT | 2024-03-27 10:42 | 2024-03-27 11:27 | Functional Check | Unit was sent to biomed without any issue description. Found battery low. Repalced battery. Went through manual and checked operational settings. Tested ok. Returned to OR.
69503 | 205324 | PM | CMPLT | CMPLT | 2024-03-27 10:43 | 2024-03-27 10:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69068 | 205325 | PM | CMPLT | CMPLT | 2024-03-27 10:44 | 2024-03-27 10:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69205 | 205326 | PM | CMPLT | CMPLT | 2024-03-27 10:48 | 2024-03-27 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68666 | 205327 | PM | CMPLT | CMPLT | 2024-03-27 10:50 | 2024-03-27 10:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69054 | 205328 | PM | CMPLT | CMPLT | 2024-03-27 10:52 | 2024-03-27 10:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69051 | 205329 | PM | CMPLT | CMPLT | 2024-03-27 10:54 | 2024-03-27 10:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69240 | 205330 | PM | CMPLT | CMPLT | 2024-03-27 10:57 | 2024-03-27 10:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69676 | 205331 | PM | CMPLT | CMPLT | 2024-03-27 11:00 | 2024-03-27 11:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69432 | 205332 | PM | CMPLT | CMPLT | 2024-03-27 11:02 | 2024-03-27 11:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69682 | 205333 | PM | CMPLT | CMPLT | 2024-03-27 11:05 | 2024-03-27 11:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68669 | 205334 | PM | CMPLT | CMPLT | 2024-03-27 11:07 | 2024-03-27 11:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69355 | 205335 | PM | CMPLT | CMPLT | 2024-03-27 11:09 | 2024-03-27 11:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69232 | 205336 | PM | CMPLT | CMPLT | 2024-03-27 11:11 | 2024-03-27 11:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69652 | 205337 | PM | CMPLT | CMPLT | 2024-03-27 11:12 | 2024-03-27 11:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69290 | 205338 | PM | CMPLT | CMPLT | 2024-03-27 11:14 | 2024-03-27 11:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69680 | 205339 | PM | CMPLT | CMPLT | 2024-03-27 11:16 | 2024-03-27 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69670 | 205340 | PM | CMPLT | CMPLT | 2024-03-27 11:17 | 2024-03-27 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1047048 | 205341 | PM | CMPLT | CMPLT | 2024-03-27 11:21 | 2024-03-27 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1050263 | 205342 | PM | CMPLT | CMPLT | 2024-03-27 11:22 | 2024-03-27 11:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1050267 | 205343 | PM | CMPLT | CMPLT | 2024-03-27 11:25 | 2024-03-27 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065877 | 205344 | PM | CMPLT | CMPLT | 2024-03-27 11:27 | 2024-03-27 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065992 | 205345 | PM | CMPLT | CMPLT | 2024-03-27 11:28 | 2024-03-27 11:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205346 | CM | CMPLT | CMPLT | 2024-03-27 12:23 | 2024-03-27 12:23 | meeting | meeting
68369 | 205347 | CM | CMPLT | CMPLT | 2024-03-27 13:18 | 2024-04-25 14:16 | Device stopped ventilating during case | (TC 03/27/2024 01:18:25 PM, Scero, Tomlyn) Faheem called to say they were pulling the device from OR3 because the ventilator stopped working during the case and they had to manually bag the patient. I brought the device downstairs to investigate. DSR downloaded and will send to Draeger for review. Self test passed ok. Note: that the breathing system (16) was extrememly wet and had to be dried out, both sensors and internally. No water traps were present, along with the patient tubing and sample line. CO2 absorber was 1/2 purple. (TC 04/25/2024 02:13:33 PM, Scero, Tomlyn) PM tests were passed. All that could be determined from the DSR was that the in and ex pressure sensors drifted from each other causing the auto flow to shut off, volume control ventilation was also unavailable due to alll th ewater in the breathing system.
66046 | 205348 | PM | CMPLT | CMPLT | 2024-03-27 13:50 | 2024-03-27 13:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68290 | 205349 | PM | CMPLT | CMPLT | 2024-03-27 13:52 | 2024-03-27 13:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68671 | 205350 | PM | CMPLT | CMPLT | 2024-03-27 13:57 | 2024-03-27 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68904 | 205351 | PM | CMPLT | CMPLT | 2024-03-27 14:06 | 2024-03-27 14:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68795 | 205352 | PM | CMPLT | CMPLT | 2024-03-27 14:08 | 2024-03-27 14:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69015 | 205353 | PM | CMPLT | CMPLT | 2024-03-27 14:11 | 2024-03-27 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68689 | 205354 | PM | CMPLT | CMPLT | 2024-03-27 14:18 | 2024-03-27 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68891 | 205355 | PM | CMPLT | CMPLT | 2024-03-27 14:25 | 2024-03-27 14:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69002 | 205356 | PM | CMPLT | CMPLT | 2024-03-27 14:27 | 2024-03-27 14:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68738 | 205357 | PM | CMPLT | CMPLT | 2024-03-27 14:34 | 2024-03-27 14:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1056665 | 205358 | CM | CMPLT | CMPLT | 2024-03-27 14:43 | 2024-08-23 13:19 | Provido not coming up on monitors | (TC 03/27/2024 02:46:03 PM, Reid, Emmerson ) DVI cable had one of the pins damaged. I replaced it and it is working now but the connector connot be screwed in so I will need to fix this. (TC 03/28/2024 08:44:39 AM, Reid, Emmerson ) opened up the unit and found that the DVI connector is attached to the board, reached out to vendor with pictures. (TC 05/16/2024 10:38:31 AM, Reid, Emmerson ) Did a bunch of troubleshooting with Jessica. Found the Scaler was not working. Jessica went to replace. (TC 05/16/2024 10:38:31 AM, Reid, Emmerson ) Did a bunch of troubleshooting with Jessica. Found the Scaler was not working. Jessica went to replace. (TC 08/23/2024 01:18:41 PM, Reid, Emmerson ) Worked with Rutvik on this. He was able to use needle nose pliers to remove the broken screws from the board. We then replaced them from an unused connector and secured the DVI cable in place. Working now.
1051375 | 205359 | PM | CMPLT | CMPLT | 2024-03-27 14:43 | 2024-03-27 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058262 | 205360 | PM | CMPLT | CMPLT | 2024-03-27 14:45 | 2024-03-27 14:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058272 | 205361 | PM | CMPLT | CMPLT | 2024-03-27 14:48 | 2024-03-27 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058258 | 205362 | CM | CMPLT | CMPLT | 2024-03-27 14:56 | 2024-03-27 14:58 | "wireless card malfunction" | (TC 03/27/2024 02:55:43 PM, Cortez, John) Replaced wireless card, BD to PM/remediate
65921 | 205363 | CM | CMPLT | CMPLT | 2024-03-27 14:56 | 2024-03-27 14:57 | Broken front case | (TC 03/27/2024 02:56:11 PM, Rhuland, Courtney) Replaced front case Replaced mylar gasket Replaced module latch
| 205364 | CM | CMPLT | CMPLT | 2024-03-27 15:50 | 2024-03-27 15:54 | PACS Gear Issue | Issue re: not being able to access Hyland images for the patient. Went to check. Rebooted the unit. Checked worklist and found the patient was still on the worklist. Checekd sent prodcedure status. No pending or error. Checked other room. Spoke with Olga to confirm the issue. Found the physician recorded images of a patient under a different patient's name. Explained issue to Olga and Anita. Connected with Clinical informatics and explanied the issue. Noted down both patient's name and provided to clinical analyst. They will connect with Anita and physician to resolve this issue.
65893 | 205365 | CM | CMPLT | CMPLT | 2024-03-27 15:54 | 2024-03-27 15:57 | "cracked rear case, L IUI" | (TC 03/27/2024 03:54:37 PM, Cortez, John) Replaced rear case, and L IUI BD to remediate/PM pump
65938 | 205366 | CM | CMPLT | CMPLT | 2024-03-27 16:06 | 2024-03-27 16:10 | "broken handle" | (TC 03/27/2024 04:05:46 PM, Cortez, John) Replaced left handle, BD to remediate/PM
68700 | 205367 | CM | CMPLT | CMPLT | 2024-03-27 16:12 | 2024-03-27 16:13 | "missing power cord strap" | (TC 03/27/2024 04:12:36 PM, Cortez, John) Added strap BD to PM/remediate
69357 | 205368 | PM | CMPLT | CMPLT | 2024-03-27 16:23 | 2024-03-27 16:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69262 | 205369 | PM | CMPLT | CMPLT | 2024-03-27 16:26 | 2024-03-27 16:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 205370 | CM | CMPLT | CMPLT | 2024-03-27 16:30 | 2024-03-27 16:30 | worked on TMS | worked on TMS
1066223 | 205371 | CM | CMPLT | CMPLT | 2024-03-27 16:37 | 2024-03-27 16:37 | alarming | tested. functional.
1061384 | 205372 | CM | CMPLT | CMPLT | 2024-03-28 08:47 | 2024-03-28 08:48 | Not showing up of the monitors | (TC 03/28/2024 08:46:46 AM, Reid, Emmerson ) The DVI cable had a bent pin, I was able to straighten it and it is working now.
1061202 | 205373 | CM | CMPLT | CMPLT | 2024-03-28 08:48 | 2024-03-28 08:49 | Wanted surigcal light camera attached | (TC 03/28/2024 08:48:05 AM, Reid, Emmerson ) Attached camera and it connected.
1047846 | 205374 | PM | CMPLT | CMPLT | 2024-03-28 09:17 | 2024-03-28 14:15 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 03/28/2024 02:14:34 PM, Rhuland, Courtney) PM complete All tests passed Refer to procedure for details
69226 | 205375 | PM | CMPLT | CMPLT | 2024-03-28 09:48 | 2024-03-28 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69130 | 205376 | PM | CMPLT | CMPLT | 2024-03-28 09:53 | 2024-03-28 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69485 | 205377 | PM | CMPLT | CMPLT | 2024-03-28 09:54 | 2024-03-28 09:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69449 | 205378 | PM | CMPLT | CMPLT | 2024-03-28 09:56 | 2024-03-28 09:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69087 | 205379 | PM | CMPLT | CMPLT | 2024-03-28 09:57 | 2024-03-28 09:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69315 | 205380 | PM | CMPLT | CMPLT | 2024-03-28 10:02 | 2024-03-28 10:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69598 | 205381 | PM | CMPLT | CMPLT | 2024-03-28 10:12 | 2024-03-28 10:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69044 | 205382 | PM | CMPLT | CMPLT | 2024-03-28 10:14 | 2024-03-28 10:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69499 | 205383 | PM | CMPLT | CMPLT | 2024-03-28 10:16 | 2024-03-28 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69601 | 205384 | PM | CMPLT | CMPLT | 2024-03-28 10:18 | 2024-03-28 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69534 | 205385 | PM | CMPLT | CMPLT | 2024-03-28 10:19 | 2024-03-28 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68901 | 205386 | CM | CMPLT | CMPLT | 2024-03-28 10:20 | 2024-04-10 14:31 | Missing screws | (TC 03/28/2024 12:41:33 PM, Rhuland, Courtney) Found PCU in MBU and it had passed PM by BD Both screws holding front case under the handle were missing Brought to BD to investigate how this was missed
69649 | 205387 | PM | CMPLT | CMPLT | 2024-03-28 10:21 | 2024-03-28 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69549 | 205388 | PM | CMPLT | CMPLT | 2024-03-28 10:23 | 2024-03-28 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69112 | 205389 | PM | CMPLT | CMPLT | 2024-03-28 10:25 | 2024-03-28 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69032 | 205390 | PM | CMPLT | CMPLT | 2024-03-28 10:27 | 2024-03-28 10:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69192 | 205391 | PM | CMPLT | CMPLT | 2024-03-28 10:33 | 2024-03-28 10:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68643 | 205392 | PM | CMPLT | CMPLT | 2024-03-28 10:42 | 2024-03-28 10:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69246 | 205393 | PM | CMPLT | CMPLT | 2024-03-28 10:44 | 2024-03-28 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69636 | 205394 | PM | CMPLT | CMPLT | 2024-03-28 10:56 | 2024-03-28 10:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69437 | 205395 | PM | CMPLT | CMPLT | 2024-03-28 10:58 | 2024-03-28 10:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69291 | 205396 | PM | CMPLT | CMPLT | 2024-03-28 11:08 | 2024-03-28 11:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69236 | 205397 | PM | CMPLT | CMPLT | 2024-03-28 11:10 | 2024-03-28 11:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69642 | 205398 | PM | CMPLT | CMPLT | 2024-03-28 11:12 | 2024-03-28 11:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1061210 | 205399 | CM | CMPLT | CMPLT | 2024-03-28 11:12 | 2024-03-28 11:13 | ESU not working | (TC 03/28/2024 11:12:24 AM, Reid, Emmerson ) ESU was an FT10 with the conmed monopolar foot pedal. I switched it with the correct one. Was not able to test becaause they were in a case and brought in a spare
69635 | 205400 | PM | CMPLT | CMPLT | 2024-03-28 11:13 | 2024-03-28 11:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1047037 | 205401 | PM | CMPLT | CMPLT | 2024-03-28 11:17 | 2024-03-28 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1047046 | 205402 | PM | CMPLT | CMPLT | 2024-03-28 11:19 | 2024-03-28 11:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065767 | 205403 | PM | CMPLT | CMPLT | 2024-03-28 11:24 | 2024-03-28 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065768 | 205404 | PM | CMPLT | CMPLT | 2024-03-28 11:26 | 2024-03-28 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065769 | 205405 | PM | CMPLT | CMPLT | 2024-03-28 11:28 | 2024-03-28 11:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065853 | 205406 | PM | CMPLT | CMPLT | 2024-03-28 11:29 | 2024-03-28 11:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065865 | 205407 | PM | CMPLT | CMPLT | 2024-03-28 11:31 | 2024-03-28 11:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065888 | 205408 | PM | CMPLT | CMPLT | 2024-03-28 11:34 | 2024-03-28 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69484 | 205409 | CM | CMPLT | CMPLT | 2024-03-28 11:37 | 2024-03-28 11:48 | AIL assembly cracked | (TC 03/28/2024 11:36:50 AM, Cortez, John) Replaced AIL assembly, both IUI connectors Completed AIL test, IUI connector test, Door ajar test. BD to remediate/PM
| 205410 | CM | CMPLT | CMPLT | 2024-03-28 12:26 | 2024-03-28 12:30 | Uneven scale | (TC 03/28/2024 12:26:16 PM, Rhuland, Courtney) Readjusted feet of scale so it is level Tested with 5kg and 20 lb weights and all tests measured accurately All buttons function properly
65890 | 205411 | PM | CMPLT | CMPLT | 2024-03-28 12:49 | 2024-03-28 12:49 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 12:49:07 PM, Rhuland, Courtney) PM completed by BD
65909 | 205412 | PM | CMPLT | CMPLT | 2024-03-28 12:51 | 2024-03-28 12:51 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 12:50:54 PM, Rhuland, Courtney) PM completed by BD
65913 | 205413 | PM | CMPLT | CMPLT | 2024-03-28 13:01 | 2024-03-28 13:02 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 01:01:43 PM, Rhuland, Courtney) PM completed by BD
65970 | 205414 | PM | CMPLT | CMPLT | 2024-03-28 13:03 | 2024-03-28 13:04 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 01:03:29 PM, Rhuland, Courtney) PM completed by BD
65924 | 205415 | PM | CMPLT | CMPLT | 2024-03-28 13:07 | 2024-03-28 13:07 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 01:07:01 PM, Rhuland, Courtney) PM completed by BD
65947 | 205416 | PM | CMPLT | CMPLT | 2024-03-28 13:09 | 2024-03-28 13:09 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 01:08:40 PM, Rhuland, Courtney) PM completed by BD
1048522 | 205417 | PM | CMPLT | CMPLT | 2024-03-28 13:11 | 2024-03-28 13:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 03/28/2024 01:11:01 PM, Rhuland, Courtney) PM completed by BD
67123 | 205418 | CM | CMPLT | CMPLT | 2024-03-28 14:26 | 2024-03-28 14:05 | Repair | Call re: "Not working". Went to check and brought it down to biomed. Cleaned unit. Founf BP hose cuff connector broken. Removed and replaced it. Battery depleted and last replaced in 2018. Also, two screws on battery cover missing. Replaced battery and battery cover screws. FOund SPO2 probe cable damaged. Searched in the office inventory and replaced it. Checked functionality. Tested ok. Delivered back to 9S.
| 205419 | CM | CMPLT | CMPLT | 2024-03-28 14:40 | 2024-03-28 14:40 | worked on TMS | worked on TMS
1055544 | 205420 | CM | CMPLT | CMPLT | 2024-03-28 14:44 | 2024-05-21 15:40 | Repair | Mar 28 - "not working". Cleaned unit. Found unit battery was completely depleted. Charged unit for sometime. Checked battery diagnostics and found only 75 % battery health remaining. Unit needs new battery. Contacted Masimo. Awaiting Response. Apr 03 - Contacted Masimo again 3 times and waited for over 45 minutes on the phone. Could not reach cust. Service. Will follow up again tomorrow. Apr 05 - Followed up. Info received. Was advised by Masimo tech support to upgrade the SW as well. Reached out to sales rep to confirm if it will impact the Masimo safety net gateway. Awaiting response. Apr 08 - Info Received. Created SC 1003205200. Awaiting delivery. Apr 10 - Moved unit to MCR until the part is delivered. May 21 - Accessories received. Followed procedure and Repalced battery. Replced accessories. Charged it. Checked functionality. Tested ok.
26181 | 205421 | PM | CANCL | CANCL | 2024-04-01 03:23 |  | PM-MAJOR - Annual PM For: LASERS, SURGICAL, CARBON DIOXIDE | No vendor support. No PM or CM available.
50057 | 205422 | PM | CMPLT | CMPLT | 2024-04-01 03:23 | 2024-05-13 14:41 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 05/13/2024 02:33:21 PM, Ferreira, Qwynn) PM Passed Replaced two latches on the main door Replaced Battery Cleaned the inside and fan Worked with Samir to adjust the temp as the value on the tester was lower than the screen temp. Adjust R3 CCW. Replaced battery in X2 Functional Test Passed Returned to the 18th floor. (TC 05/13/2024 02:41:06 PM, Chukri, Samir) Clibrated the temp to mach the setting temp.
52181 | 205423 | PM | CANCL | CANCL | 2024-04-01 03:23 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | Device to be retired and replaced
62618 | 205424 | PM | CMPLT | CMPLT | 2024-04-01 03:23 | 2024-11-07 13:35 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/07/2024 01:34:49 PM, Rhuland, Courtney) Side door latch and side door hinge repaired Heater doors would not close fully when canopy was lowered - resulting in an alarm Opened canopy top and greased keeper plates for both doors Heater doors closed poperly Replaced: 9V battery Calibrated: Mains & VREF voltage Scale Functional check - OK Set 35C, 75% humidity Read 35C, 75.6% humidity Electrical safety - OK Ground resistance - 186 mOhms Leakage current - 75.4 uA
62985 | 205425 | PM | CMPLT | CMPLT | 2024-04-01 03:23 | 2024-06-26 12:56 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/26/2024 12:56:19 PM, Rhuland, Courtney) PM completed Irradiance test - OK Electrical safety - OK
63011 | 205426 | PM | CMPLT | CMPLT | 2024-04-01 03:23 | 2024-07-05 09:03 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/05/2024 09:02:41 AM, Rhuland, Courtney) PM complete Irradiance - 55.5 uW/cm^2/nm Hours - 5666 Leakage current - 14.1 uA Ground resistance - 57 mOhms
63023 | 205427 | PM | CMPLT | CMPLT | 2024-04-01 03:23 | 2024-07-29 14:21 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/29/2024 02:21:17 PM, Rhuland, Courtney) PM completed Power cord replaced Cord retainer replaced Average light intensity - 54 µW•cm-2•nm-1 Leakage current - 9.2 uA Ground resistance - 86 mOhms Hours - 8530
63072 | 205428 | PM | CANCL | CANCL | 2024-04-01 03:23 | 2024-06-27 15:53 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport
63159 | 205429 | PM | ACTIV | ISSUE | 2024-04-01 03:23 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
65421 | 205430 | PM | ACTIV | ISSUE | 2024-04-01 03:23 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
65887 | 205431 | PM | CMPLT | CMPLT | 2024-04-01 03:23 | 2024-04-10 14:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:34:11 PM, Rhuland, Courtney) PM completed by BD
66002 | 205432 | PM | CMPLT | CMPLT | 2024-04-01 03:23 | 2024-04-11 14:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
66005 | 205433 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-01 08:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66017 | 205434 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-03 14:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:02:38 PM, Reid, Emmerson ) PM Complete.
66019 | 205435 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-01 08:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66052 | 205436 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-02 09:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM completed and passed. by BD after SW recall.
66058 | 205437 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-23 10:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/23/2024 10:01:35 AM, Rhuland, Courtney) PM completed by BD
66456 | 205438 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-11 14:14 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced air filter. Run battery check. 25 minutes and 220 minutes. Tested. Functional. SW2.51 Leakage: 83uA and 93 ohm.
66469 | 205439 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-09 15:18 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced air filter. Run battery check. 25 minutes and 200 minutes. Tested. Functional. SW2.51 Leakage: 87uA and 94 ohm.
66709 | 205440 | PM | CANCL | CANCL | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
66715 | 205441 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
66816 | 205442 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-09-13 11:42 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/13/2024 11:41:33 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
66836 | 205443 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-10-16 08:45 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 10/16/2024 08:45:13 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
67758 | 205444 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-04 11:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/04/2024 11:15:06 AM, Reid, Emmerson ) Need to run the volume test 3 times before it passed. PM complete, see procedure for details.
67914 | 205445 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-04 10:08 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/07/2024 01:21:03 PM, Baghirzade, Nara) 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. **Set RH to 75%, measured on the analyzer 81% Set temperature to 35C, measured 35C on all temperature points. Calibrated the RH sensor twice however it failed calibration. Took a RH sensor from a retired omnibed; also failed 2 calibrations. Waiting for new RH sensors to come in. Will repeat the PM. (TC 06/04/2024 10:07:37 AM, Baghirzade, Nara) Device will be retired; new incubators and omnibeds received.
67927 | 205446 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-01 13:22 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/01/2024 01:20:39 PM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Porthole seals on west door replaced. Grommet replaced. Electrical safety results - Leakage current = 14.9 uA Power cord resistance = 185 mohms (TC 08/01/2024 01:21:31 PM, DeGuzman, Leila) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Porthole seals on west door replaced. Grommet replaced. Electrical safety results - Leakage current = 14.9 uA Power cord resistance = 185 mohms
67928 | 205447 | PM | CANCL | CANCL | 2024-04-01 03:24 | 2024-06-04 10:13 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
67935 | 205448 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-02 13:28 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 05/02/2024 01:21:21 PM, Rhuland, Courtney) PM completed Calibrated: REF & line voltage Scale RH Replaced rechargeable 9V battery Repaired tilt platform assembly Flip door was on backwards; fixed All tests passed Functional check - OK Set: 35C, 75% humidity Read: 35C, 75.2% humidity Electrical safety - OK Leakage current - 86.4uA Resistance - 96mOhms
67936 | 205449 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-10 10:48 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/10/2024 10:47:59 AM, Rhuland, Courtney) Replaced broken porthole door PM completed Calibrated: MAINS & LV voltage Scale Functional check - OK Set 35C & 75% humidity Read 35C & 75.5% humidity Electrical safety - OK Leakage current - 87.7 uA Ground resistance - 53 mOhms
68298 | 205450 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-03 10:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68368 | 205451 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68526 | 205452 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Bladder Scanner
68527 | 205453 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Bladder Scanner
68555 | 205454 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-13 11:10 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 05/22/2024 09:30:16 AM, Rendon Lopez , Elizabeth) Removed unit from OR1 and brought down for PM (TC 06/11/2024 11:10:23 AM, Rendon Lopez , Elizabeth) 1 Year PM parts replaced. PM tests started (TC 06/12/2024 11:10:23 AM, Rendon Lopez , Elizabeth) Working with student, PM tests in progress (TC 06/13/2024 11:10:23 AM, Rendon Lopez , Elizabeth) PM tests completed, all tests passed. Attached results sheet. Returned to L&D as spare.
68580 | 205455 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-16 12:50 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD.
68588 | 205456 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-11-21 09:02 | PM-MAJOR-Annual PM for Servo-Ventilator | (TC 11/18/2024 02:50:50 PM, Canjimavadivel, Kamallesh ) PM performed by vendor, see attached service report. Waiting for batteries. (TC 11/21/2024 09:02:20 AM, Canjimavadivel, Kamallesh ) Received batteries and installed
68614 | 205457 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-18 12:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68636 | 205458 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-16 13:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68652 | 205459 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-03 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:46:18 PM, Rhuland, Courtney) PM completed by BD
68712 | 205460 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-11 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68761 | 205461 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-03 14:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:09:25 PM, Rhuland, Courtney) PM completed by BD
68768 | 205462 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-05 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:42:19 PM, Rhuland, Courtney) PM completed by BD
68780 | 205463 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-02 11:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68790 | 205464 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-16 13:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68838 | 205465 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-16 13:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68841 | 205466 | PM | CANCL | CANCL | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68890 | 205467 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-02 11:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68899 | 205468 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-16 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68901 | 205469 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-03 09:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69018 | 205470 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-05 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:41:02 PM, Rhuland, Courtney) Recall completed by BD
69021 | 205471 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-20 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69055 | 205472 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-02 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69090 | 205473 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-30 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69121 | 205474 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-19 21:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69131 | 205475 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69132 | 205476 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-23 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69150 | 205477 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-12 08:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69182 | 205478 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-24 09:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/24/2024 09:49:04 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69222 | 205479 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-19 21:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69314 | 205480 | PM | CMPLT | PMFAIL | 2024-04-01 03:24 | 2024-05-23 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/23/2024 02:33:08 PM, Rendon Lopez , Elizabeth) Door latch not staying up. and bezel ground wire damaged
69330 | 205481 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-15 14:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69377 | 205482 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-19 20:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69393 | 205483 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-24 09:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69420 | 205484 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-02 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69441 | 205485 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-19 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69461 | 205486 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-09 11:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 11:42:16 AM, Patel, Rutvik) by BD
69463 | 205487 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-03 10:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69508 | 205488 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-02 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/02/2024 02:15:27 PM, Baghirzade, Nara) Replaced right IUI. Replaced both pressure sensors. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69541 | 205489 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-15 14:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69571 | 205490 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-09 15:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69638 | 205491 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-30 13:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
32540 | 205492 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-13 16:06 | PM - Monthly | PM Major done Mechanical done - LCSR brushes cleaned Slice count - 37352 TOS phantom a(-994.6, 10.1) b(331.6, 13.7) c(131.6, 14.0) d(95.9, 14.9) e(-106.8, 13.3) f(-0.6, 12.6) -System passed all tests
32560 | 205493 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-07 15:19 | PM - Monthly | PM completed Mechanical PM done sc-1419621 TOS phantom a(-990.1,10.0) b(332.9,13.1) c(133.2,13.1) d(96.4,12.8) e(-106.1,12.4) f(0.3,13.7) -passed all phantom tests
32580 | 205494 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-08 15:01 | PM - Monthly | PM completed Mechanical PM done sc-234568 TOS phantom a(-993.9,10.6) b(333.8,15.4) c(132.8,14.4) d(96.5,13.9) e(-107.5,11.9) f(0.6,14.9) -passed all phantom tests
502088 | 205495 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-24 13:33 | PM - Semi-Annual | (TC 05/24/2024 01:33:15 PM, Green, Paul) PM2 completed with 12 month saftey.
502089 | 205496 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-16 13:07 | PM - Semi-Annual | (TC 05/16/2024 01:06:05 PM, Green, Paul) Performed PM2 as per procedure. Also 12 month saftey was done.
502231 | 205497 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-17 16:54 | PM - Semi-Annual | Harp
502231 | 205498 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-19 14:52 | HARP - Semi Annual | HARP
1046439 | 205499 | PM | ACTIV | READY | 2024-04-01 03:24 |  | Annual PM
1046670 | 205500 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-ELECTRICAL MECHANISM INSPECTION AND OPERATION TEST |  | DEPARTMENT CONTACT THE VENDOR TO PROVIDE PM AND MAINTENANCE
1047848 | 205501 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-04 13:12 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/04/2024 01:11:56 PM, Rhuland, Courtney) Motor rate error present when running occlusion test Replaced worm gear and rebuilt Error cleared PM completed all tests passed Refer to procedure for details
1047854 | 205502 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-08 13:18 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/08/2024 01:18:01 PM, Rhuland, Courtney) Replaced worn clutch PM completed All tests passed Refer to procedure for details
1047855 | 205503 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-07 09:26 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/07/2024 09:25:53 AM, Rhuland, Courtney) Replaced battery PM passed Refer to procedure for details
1047856 | 205504 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-25 14:40 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/19/2024 02:58:12 PM, Rhuland, Courtney) Replaced front case Replaced worn clutch Reseated worm gear and seal (TC 06/25/2024 02:39:40 PM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1047857 | 205505 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-09 11:13 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 07/09/2024 11:12:51 AM, Rhuland, Courtney) Replaced worn clutch Replaced plunger seal PM completed All tests passed Refer to procedure for details
1047858 | 205506 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-07 08:39 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/07/2024 08:38:45 AM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1047863 | 205507 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-05 13:30 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/05/2024 01:29:55 PM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1048734 | 205508 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Ventilators
1049027 | 205509 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-07 11:21 | Planned Maintenance | (TC 05/07/2024 11:20:45 AM, Reid, Emmerson ) PM Complete, see procedure for details.
1049040 | 205510 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Planned Maintenance
1049041 | 205511 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Planned Maintenance
1049052 | 205512 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Planned Maintenance
1049054 | 205513 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Planned Maintenance
1049065 | 205514 | PM | CMPLT | PMFAIL | 2024-04-01 03:24 | 2024-05-24 10:58 | Planned Maintenance | (TC 05/24/2024 10:57:45 AM, Rendon Lopez , Elizabeth) Found device in L&D OR with damaged hose and missing one stand caster. Contacted Stryker to receive replacement unit
1049069 | 205515 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Planned Maintenance
1049074 | 205516 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-24 11:00 | Planned Maintenance | (TC 05/24/2024 10:59:47 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1049076 | 205517 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Planned Maintenance
1053010 | 205518 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-03 14:54 | PM - Annual | (TC 06/03/2024 02:53:37 PM, Fan, Tak) Gnd R = 174 mohm Max A = 0.3 / 163.4 / 314 uA ECG = 2.6 uA ISO = 10.3 uA M5s = 2.9 uA 4Vc = 3.1 uA
1053347 | 205519 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-13 08:28 | MSH BI 1270 H1 HARP MAY 8 TRGT SEMI-ANNUAL | (TC 05/08/2024 04:09:50 PM, Dolinar, John) MSH BI 1270 H1 HARP MAY 8 TRGT SEMI-ANNUAL (TC 05/08/2024 04:10:07 PM, Chau, Ken) MSH BI 1270 H1 HARP MAY 8 TRGT SEMI-ANNUAL (TC 05/08/2024 04:10:35 PM, Chau, Ken) MSH BI 1270 H1 HARP MAY 8 TRGT SEMI-ANNUAL | MSH BI 1270 H1 HARP MAY 8 TRGT SEMI-ANNUAL
1053347 | 205520 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-13 08:28 | MSH BI 1270 -PM SEMI-ANNUAL
1054039 | 205521 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - ANNUAL PM For: Warmers
1054460 | 205522 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-04 14:53 | Lubrication cycle | Completed with PM
1054462 | 205523 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-11 10:35 | Lubrication cycle | Completed with PM
1054468 | 205524 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-15 14:49 | Lubrication cycle
1055094 | 205525 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-03 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:18:56 PM, Rhuland, Courtney) PM completed by BD
1055102 | 205526 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-13 08:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055108 | 205527 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-13 08:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055156 | 205528 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-16 10:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055159 | 205529 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-15 14:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055169 | 205530 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-08-16 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055170 | 205531 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055174 | 205532 | PM | CMPLT | PMFAIL | 2024-04-01 03:24 | 2024-04-04 15:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1055175 | 205533 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-10 10:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/10/2024 10:22:01 AM, DeGuzman, Leila) Cleaned device. After the software updated, performed preventive maintenance. All tests passed. Return to service.
1055176 | 205534 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-15 09:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055772 | 205535 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-18 10:12 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/18/2024 10:11:53 AM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1056055 | 205536 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-10 23:29 | HARP - Semi Annual | HARP was done.
1056055 | 205537 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-03 23:31 | PM - Semi Annual | PM was done.
1057391 | 205538 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-03-11 19:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1057835 | 205539 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057843 | 205540 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057859 | 205541 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
1057860 | 205542 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
1058153 | 205543 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-04 09:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/04/2024 09:36:06 AM, Reid, Emmerson ) PM complete, see procedure for details.
1058170 | 205544 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058171 | 205545 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-06-04 10:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/04/2024 10:32:24 AM, Reid, Emmerson ) PM Complete, see procedure for details.
1058273 | 205546 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-17 10:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1059878 | 205547 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-05 10:36 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/05/2024 10:35:06 AM, Rhuland, Courtney) Calibration and check-in completed All tests passed Refer to ASM for details
1060518 | 205548 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-24 21:52 | HARP - Semi Annual | HARP
1060518 | 205549 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-23 21:54 | PM - Semi Annual | PM was done.
1062926 | 205550 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Planned Maintenance
1063026 | 205551 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-22 14:27 | PM - Annual | (TC 05/22/2024 02:27:00 PM, Fan, Tak) PM completed Gnd R = 101 mOhm MAx a = 0.2 / 173.4 / 340 uA ECG = 2.1 uA ISO = 8 uA M5S-D (380871YP4) = 0.5 / 2.4 uA 4VcD (402792YP6) = 0.5 / 3.0 uA
1063797 | 205552 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Annual PM
1064167 | 205553 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-05-23 10:27 | PM MAJOR | (TC 05/23/2024 10:27:05 AM, Scero, Tomlyn) Device retrieved from OR and PM was performed by Gentinge.
1065736 | 205554 | PM | ACTIV | READY | 2024-04-01 03:24 |  | Annual PM | (TC 05/01/2024 01:10:38 PM, Baghirzade, Nara) Coordinating the service visit from Clarion tech to fix the plate assembly. Called Clarion to clarify the quote. Received quote and sent it to Dr. Yan and Henna; they have informed me that they don't want to pay for travel and labour. Asked Henna to respond to the email to the vendor stating this as the tech has been scheduled to visit the clinic on May 1st. for the repair. Waiting on response.
1066063 | 205555 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Annual PM - For: LIFT
1066064 | 205556 | PM | ACTIV | ISSUE | 2024-04-01 03:24 |  | Annual PM - For: LIFT
1066756 | 205557 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-17 07:48 | Lubrication cycle | Completed with PM
1066757 | 205558 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-11 14:05 | Lubrication cycle | Completed with PM
1067164 | 205559 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-04-18 13:42 | Yearly Calibration | Refer to WO 206283
1070151 | 205560 | PM | CMPLT | CMPLT | 2024-04-01 03:24 | 2024-07-15 13:40 | Quaterly PM for: 3D Pritner | (TC 07/15/2024 01:39:47 PM, Ferreira, Qwynn) PM Complete Lead screws were lubricated filters were inspected Printer internals were cleaned/debris removed No broken or loose cables Jogged the controls for the extruder and build plate/platform - OK Functional Check- OK
| 205561 | RDS | CMPLT | CMPLT | 2024-04-02 07:23 | 2024-04-09 10:46 | Weekly check for ICU
| 205562 | OC | CMPLT | CMPLT | 2024-04-02 08:06 | 2024-04-05 07:53 | Emmerson Reid On-Call March 29 - April 8 2024 | (TC 04/01/2024 06:30:00 PM, Reid, Emmerson ) 14S called about an issue with ECG carts. They had tried 3 different carts on one patient and none of them worked. I asked them to change electrodes 3 different times and they said they did and that there was interference in the room. They also mentioned the carts were working in other areas. I came in and used my simulator on the ECGs and they were working. I checked the electrodes in the drawer and they were expired. I threw them out and let them know and asked them to try again with electrodes that weren't expired. It was working. (TC 04/01/2024 06:30:00 PM, Reid, Emmerson ) 14S called about an issue with ECG carts. They had tried 3 different carts on one patient and none of them worked. I asked them to change electrodes 3 different times and they said they did and that there was interference in the room. They also mentioned the carts were working in other areas. I came in and used my simulator on the ECGs and they were working. I checked the electrodes in the drawer and they were expired. I threw them out and let them know and asked them to try again with electrodes that weren't expired. It was working.
1051734 | 205563 | CM | CMPLT | CMPLT | 2024-04-02 08:40 | 2024-04-02 08:40 | PM Scheduled | (TC 03/28/2024 08:39:43 AM, Ferreira, Qwynn) pm completed Battery replaced last year - OK
501069 | 205564 | CM | CMPLT | CMPLT | 2024-04-02 09:34 | 2024-04-03 14:11 | Left cursor not working. | (TC 04/03/2024 02:10:34 PM, Chau, Ken) Replaced kidney button
68941 | 205565 | PM | CMPLT | CMPLT | 2024-04-02 09:39 | 2024-04-02 09:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68941 | 205566 | PM | CANCL | DUP | 2024-04-02 09:39 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68972 | 205567 | PM | CMPLT | CMPLT | 2024-04-02 09:44 | 2024-04-02 09:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
50517 | 205568 | CM | CMPLT | CMPLT | 2024-04-02 09:56 | 2024-04-04 09:49 | Pedal broken, not working | (TC 04/02/2024 09:59:15 AM, Rendon Lopez , Elizabeth) Brought down to Biomed, confirmed issue. Inquired updated pricing (TC 04/04/2024 09:48:26 AM, Rendon Lopez , Elizabeth) Received replacement foot pedal. Tested, all functions-OK
68789 | 205569 | PM | CMPLT | CMPLT | 2024-04-02 10:10 | 2024-04-02 10:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1054418 | 205570 | CM | CMPLT | CMPLT | 2024-04-02 10:18 | 2024-05-15 10:59 | Not working | (TC 04/02/2024 10:17:56 AM, Rendon Lopez , Elizabeth) Confirmed issue. TOCO Eq malfunction. Kept to send out for repair under warranty. (TC 05/15/2024 10:58:24 AM, Rendon Lopez , Elizabeth) Repaired, incoming inspection-OK
66514 | 205571 | CM | CMPLT | CMPLT | 2024-04-02 10:22 | 2024-04-02 11:36 | Recorder not printing | (TC 04/02/2024 10:22:19 AM, Rendon Lopez , Elizabeth) Confirmed issue, removed from room. Replaced recorder with spare one. Returned monitor to room 17.
1054899 | 205572 | CM | CMPLT | CMPLT | 2024-04-02 10:23 | 2024-04-02 10:24 | Not saving patient data | (TC 04/02/2024 10:22:54 AM, Reid, Emmerson ) Went up to check the monitor. It was setup properly. Get them to run again on patient and it worked. (TC 04/02/2024 10:22:54 AM, Reid, Emmerson ) Went up to check the monitor. It was setup properly. Get them to run again on patient and it worked.
68960 | 205573 | PM | CMPLT | CMPLT | 2024-04-02 10:23 | 2024-04-02 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1057706 | 205574 | CM | CMPLT | CMPLT | 2024-04-02 10:25 | 2024-04-04 09:52 | Broken MSL connector and missing cover | (TC 04/02/2024 01:58:20 PM, Canjimavadivel, Kamallesh ) Contacted Philips for part numbers and quote. Placed order for bezel and MSL board (TC 04/04/2024 09:50:19 AM, Canjimavadivel, Kamallesh ) Swapped out damaged components
| 205575 | RDS | CMPLT | CMPLT | 2024-04-02 10:25 | 2024-04-02 10:27 | Weekly Checks For PACU | (TC 04/02/2024 10:26:29 AM, Reid, Emmerson ) Rounds complete (TC 04/02/2024 10:26:29 AM, Reid, Emmerson ) Rounds complete
| 205576 | RDS | CMPLT | CMPLT | 2024-04-02 10:27 | 2024-04-02 10:30 | Weekly Checks For OR | (TC 04/02/2024 10:29:20 AM, Reid, Emmerson ) rounds complete (TC 04/02/2024 10:29:20 AM, Reid, Emmerson ) rounds complete
69007 | 205577 | PM | CMPLT | CMPLT | 2024-04-02 10:29 | 2024-04-02 10:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68699 | 205578 | PM | CMPLT | CMPLT | 2024-04-02 10:44 | 2024-04-02 10:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68702 | 205579 | PM | CMPLT | CMPLT | 2024-04-02 10:52 | 2024-04-02 10:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68766 | 205580 | PM | CMPLT | CMPLT | 2024-04-02 11:02 | 2024-04-02 11:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 205581 | RDS | CMPLT | CMPLT | 2024-04-02 11:03 | 2024-04-02 11:03 | Weekly check for 14S and 17M | Weekly check for 14S and 17M
68785 | 205582 | PM | CMPLT | CMPLT | 2024-04-02 11:04 | 2024-04-02 11:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68721 | 205583 | PM | CMPLT | CMPLT | 2024-04-02 11:06 | 2024-04-02 11:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066462 | 205584 | FC | CMPLT | CMPLT | 2024-04-02 11:07 | 2024-04-02 11:07 | QC check | passed QC check
67969 | 205585 | CM | CMPLT | CMPLT | 2024-04-02 11:08 | 2024-04-02 11:08 | Alarming | Turn on main power switch. Checked. Functional.
63872 | 205586 | CM | CMPLT | CMPLT | 2024-04-02 11:10 | 2024-04-02 11:10 | alarming, not turning on. | Replaced broken fan rubber. Tested. functional.
68762 | 205587 | PM | CMPLT | CMPLT | 2024-04-02 11:16 | 2024-04-02 11:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68760 | 205588 | PM | CMPLT | CMPLT | 2024-04-02 11:25 | 2024-04-02 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68881 | 205589 | PM | CMPLT | CMPLT | 2024-04-02 11:45 | 2024-04-02 11:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047096 | 205590 | PM | CMPLT | CMPLT | 2024-04-02 11:47 | 2024-04-02 11:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055106 | 205591 | PM | CMPLT | CMPLT | 2024-04-02 11:57 | 2024-04-02 11:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055109 | 205592 | PM | CMPLT | CMPLT | 2024-04-02 11:59 | 2024-04-02 11:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058278 | 205593 | PM | CMPLT | CMPLT | 2024-04-02 12:01 | 2024-04-02 12:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064642 | 205594 | PM | CMPLT | CMPLT | 2024-04-02 12:04 | 2024-04-02 12:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064643 | 205595 | PM | CMPLT | CMPLT | 2024-04-02 12:06 | 2024-04-02 12:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064644 | 205596 | PM | CMPLT | CMPLT | 2024-04-02 12:08 | 2024-04-02 12:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064641 | 205597 | PM | CMPLT | CMPLT | 2024-04-02 12:10 | 2024-04-02 12:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065705 | 205598 | PM | CMPLT | CMPLT | 2024-04-02 12:12 | 2024-04-02 12:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065720 | 205599 | PM | CMPLT | CMPLT | 2024-04-02 12:13 | 2024-04-02 12:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69302 | 205600 | CM | CMPLT | CMPLT | 2024-04-02 13:09 | 2024-04-02 13:13 | Front display damaged | (TC 04/02/2024 01:08:53 PM, Cortez, John) Front display/LVP replaced, IUIs replaced Relevant tests for replaced parts passed BD to remediate/PM
1064974 | 205601 | PA | CMPLT | CMPLT | 2024-04-02 13:11 | 2024-04-16 11:49 | Parts Order - Resins and Tanks | (TC 04/02/2024 02:21:45 PM, Chukri, Samir) Material for 3D printer lab
1066629 | 205602 | CM | CMPLT | CMPLT | 2024-04-02 13:30 | 2024-04-02 13:30 | MSH NICU ZONARE Z2 ACTIVATE MULTI FRAME ARCHIVING | (TC 04/02/2024 01:30:16 PM, Whittington, Scott) MSH NICU ZONARE Z2 ACTIVATE MULTI FRAME ARCHIVING (TC 04/02/2024 01:30:27 PM, Dolinar, John) MSH NICU ZONARE Z2 ACTIVATE MULTI FRAME ARCHIVING | MSH NICU ZONARE Z2 ACTIVATE MULTI FRAME ARCHIVING
69411 | 205603 | PM | CMPLT | CMPLT | 2024-04-02 13:32 | 2024-04-02 13:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1053348 | 205604 | CM | CMPLT | CMPLT | 2024-04-02 13:39 | 2024-04-23 14:08 | MSH BI 1272 H3 CONTROL PANEL LEFT HAND GANTRY, "RETURN HOME " CONTROL INTERMITTENT | (TC 04/02/2024 01:39:17 PM, Whittington, Scott) MSH BI 1272 H3 CONTROL PANEL LEFT HAND GANTRY, "RETURN HOME " CONTROL INTERMITTENT (TC 04/02/2024 01:39:28 PM, Dolinar, John) MSH BI 1272 H3 CONTROL PANEL LEFT HAND GANTRY, "RETURN HOME " CONTROL INTERMITTENT (TC 04/19/2024 01:36:05 PM, Whittington, Scott) MSH BI 1272 H3 CONTROL PANEL LEFT HAND GANTRY, "RETURN HOME " CONTROL INTERMITTENT SN# 3DM160700414 ROSE REF# 01486351 PARTS DELIVERED , TO BE INSTALLED NEXT WEEK (TC 04/19/2024 01:36:50 PM, Dolinar, John) MSH BI 1272 H3 CONTROL PANEL LEFT HAND GANTRY, "RETURN HOME " CONTROL INTERMITTENT SN# 3DM160700414 ROSE REF# 01486351 PARTS DELIVERED , TO BE INSTALLED NEXT WEEK (TC 04/23/2024 02:08:08 PM, Whittington, Scott) MSH BI 1272 H3 CONTROL PANEL LEFT HAND GANTRY, "RETURN HOME " CONTROL INTERMITTENT SN# 3DM160700414 ROSE REF# 01486351 COMPLETED | MSH BI 1272 H3 CONTROL PANEL LEFT HAND GANTRY, "RETURN HOME " CONTROL INTERMITTENT SN# 3DM160700414 ROSE REF# 01486351
69251 | 205605 | PM | CMPLT | CMPLT | 2024-04-02 13:40 | 2024-04-02 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69422 | 205606 | PM | CMPLT | CMPLT | 2024-04-02 13:42 | 2024-04-02 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69407 | 205607 | PM | CMPLT | CMPLT | 2024-04-02 13:44 | 2024-04-02 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69381 | 205608 | PM | CMPLT | CMPLT | 2024-04-02 13:47 | 2024-04-02 13:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69189 | 205609 | PM | CMPLT | CMPLT | 2024-04-02 13:49 | 2024-04-02 13:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69215 | 205610 | PM | CMPLT | CMPLT | 2024-04-02 13:51 | 2024-04-02 13:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69538 | 205611 | PM | CMPLT | CMPLT | 2024-04-02 13:53 | 2024-04-02 13:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69046 | 205612 | PM | CMPLT | CMPLT | 2024-04-02 13:55 | 2024-04-02 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69080 | 205613 | PM | CMPLT | CMPLT | 2024-04-02 13:57 | 2024-04-02 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69506 | 205614 | PM | CMPLT | CMPLT | 2024-04-02 13:59 | 2024-04-02 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69597 | 205615 | PM | CMPLT | CMPLT | 2024-04-02 14:02 | 2024-04-02 14:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69505 | 205616 | PM | CMPLT | CMPLT | 2024-04-02 14:04 | 2024-04-02 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69532 | 205617 | PM | CMPLT | CMPLT | 2024-04-02 14:06 | 2024-04-02 14:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69071 | 205618 | PM | CMPLT | CMPLT | 2024-04-02 14:08 | 2024-04-02 14:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69132 | 205619 | CM | CMPLT | CMPLT | 2024-04-02 14:10 | 2024-04-02 14:11 | Broken door latch | (TC 04/02/2024 02:10:01 PM, Cortez, John) Replaced broken door latch, replace L IUI BD to remediate/PM
69595 | 205620 | PM | CMPLT | CMPLT | 2024-04-02 14:11 | 2024-04-02 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68665 | 205621 | PM | CMPLT | CMPLT | 2024-04-02 14:14 | 2024-04-02 14:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69558 | 205622 | PM | CMPLT | CMPLT | 2024-04-02 14:17 | 2024-04-02 14:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69587 | 205623 | PM | CMPLT | CMPLT | 2024-04-02 14:20 | 2024-04-02 14:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69084 | 205624 | PM | CMPLT | CMPLT | 2024-04-02 14:25 | 2024-04-02 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69137 | 205625 | PM | CMPLT | CMPLT | 2024-04-02 14:27 | 2024-04-02 14:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69474 | 205626 | PM | CMPLT | CMPLT | 2024-04-02 14:30 | 2024-04-02 14:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69640 | 205627 | PM | CMPLT | CMPLT | 2024-04-02 14:31 | 2024-04-02 14:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69639 | 205628 | PM | CMPLT | CMPLT | 2024-04-02 14:36 | 2024-04-02 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69644 | 205629 | PM | CMPLT | CMPLT | 2024-04-02 14:38 | 2024-04-02 14:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69267 | 205630 | PM | CMPLT | CMPLT | 2024-04-02 14:39 | 2024-04-02 14:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69269 | 205631 | PM | CMPLT | CMPLT | 2024-04-02 14:41 | 2024-04-02 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69671 | 205632 | PM | CMPLT | CMPLT | 2024-04-02 14:43 | 2024-04-02 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1047038 | 205633 | PM | CMPLT | CMPLT | 2024-04-02 14:45 | 2024-04-02 14:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1057389 | 205634 | PM | CMPLT | CMPLT | 2024-04-02 14:46 | 2024-04-02 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055162 | 205635 | PM | CMPLT | CMPLT | 2024-04-02 14:48 | 2024-04-02 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058253 | 205636 | PM | CMPLT | CMPLT | 2024-04-02 14:50 | 2024-04-02 14:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064649 | 205637 | PM | CMPLT | CMPLT | 2024-04-02 14:52 | 2024-04-02 14:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064652 | 205638 | PM | CMPLT | CMPLT | 2024-04-02 14:53 | 2024-04-02 14:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064651 | 205639 | PM | CMPLT | CMPLT | 2024-04-02 14:55 | 2024-04-02 14:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1064650 | 205640 | PM | CMPLT | CMPLT | 2024-04-02 14:57 | 2024-04-02 14:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065065 | 205641 | PM | CMPLT | CMPLT | 2024-04-02 14:58 | 2024-04-02 14:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 205642 | CM | CMPLT | CMPLT | 2024-04-02 14:59 | 2024-04-02 14:59 | suction tree | (TC 02/07/2024 02:59:02 PM, Huang, Jessica) got the call from OR desk and brought the suction tree to Biomed. repalced the regulator and sent it back to OR (TC 02/07/2024 02:59:02 PM, Huang, Jessica) got the call from OR desk and brought the suction tree to Biomed. repalced the regulator and sent it back to OR
1065898 | 205643 | PM | CMPLT | CMPLT | 2024-04-02 15:00 | 2024-04-02 15:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065522 | 205644 | CM | CMPLT | CMPLT | 2024-04-02 15:01 | 2024-04-02 15:01 | not working | (TC 02/13/2024 03:01:13 PM, Huang, Jessica) the console and the probe were sent down without the battery. got the battery from Shermaine. passed self-test
1066038 | 205645 | PM | CMPLT | CMPLT | 2024-04-02 15:02 | 2024-04-02 15:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 205646 | CM | CMPLT | CMPLT | 2024-04-02 15:04 | 2024-04-02 15:04 | suction tree | (TC 02/14/2024 03:03:44 PM, Huang, Jessica) got the call from OR desk and brought the suction tree down to Biomed replaced the regulator and sent it back to OR
1064444 | 205647 | CM | CMPLT | CMPLT | 2024-04-02 15:05 | 2024-04-02 15:05 | not working | (TC 02/16/2024 03:05:23 PM, Huang, Jessica) replaced batteries. tested it. working fine
1054888 | 205648 | CM | CMPLT | CMPLT | 2024-04-02 15:07 | 2024-04-02 15:07 | removed from service
1060974 | 205649 | CM | CMPLT | CMPLT | 2024-04-02 15:09 | 2024-04-02 15:09 | not working | (TC 02/28/2024 03:09:19 PM, Huang, Jessica) got the call in OR 13 and went there to check the ESU. the ligasure display was blank and didnt allow to do anything because it was in the middle of the procedure. infromed Camilla to get the spare device and swapped the device. picked up the device from OR desk and brought it to Biomed and could not recreate the same issue. contacted Medtronic for support was informed it could be system glitch by power cycle the device to fix. sent the device back to OR
1061337 | 205650 | CM | CMPLT | CMPLT | 2024-04-02 15:17 | 2024-04-02 15:17 | removed from service
69650 | 205651 | CM | CMPLT | CMPLT | 2024-04-02 15:25 | 2024-04-03 13:47 | Door broken | (TC 04/02/2024 08:34:42 AM, Rendon Lopez , Elizabeth) Door unhinged from bezel, damaged ground wire. Replaced bezel assy, and front door cover. (TC 04/03/2024 01:39:54 PM, Rendon Lopez , Elizabeth) Module latch damaged, replaced. Calibrated rate. PM completed, all tets passed
1054965 | 205652 | CM | CMPLT | CMPLT | 2024-04-02 15:41 | 2024-04-02 15:41 | not working | (TC 03/25/2024 03:41:18 PM, Huang, Jessica) the device was sent down from mother and baby the power cord didnt connect to the monitor properly so the battery was drained and could not be powered on connected the power cord to the monitor properly and charged it. tested it, working fine
1057279 | 205653 | CM | CMPLT | CMPLT | 2024-04-02 15:43 | 2024-04-02 15:43 | not working | (TC 03/25/2024 03:43:13 PM, Huang, Jessica) the device was sent down from mother and baby the power cord didnt connect to the monitor properly so the battery was drained and could not be powered on connected the power cord to the monitor properly and charged it. tested it, working fine
1057115 | 205654 | CM | CMPLT | CMPLT | 2024-04-02 15:44 | 2024-04-02 15:44 | not working | (TC 03/27/2024 03:44:39 PM, Huang, Jessica) received the call from the ward clerk saying one of the masimo monitors could not be powered on went the MBU to check the monitor and noticed the power cord didnt connect to the monitor properly. infromed the nurse and ward clerk. sent the email to Shermain, Tanya and Kim to infrom the issue because it was the third time that the power cord didnt connect to the monitor and the battery got draned.
1064477 | 205655 | CM | CMPLT | CMPLT | 2024-04-02 15:48 | 2024-04-02 15:48 | docking station fell from the stand | (TC 03/26/2024 03:47:50 PM, Huang, Jessica) received the email from Mira (RT) saying the docking station was falling from the stand checked it. all screws were gone replaced the screws and secured the docking station
1046654 | 205656 | CM | CMPLT | CMPLT | 2024-04-02 16:00 | 2024-04-02 16:00 | not working | (TC 03/26/2024 04:01:00 PM, Huang, Jessica) replaced the power cord on the probe and the battery for the console. failed calibration contacted Roxon and was infromed the probe needed to be sent for repair sent the console and probe, forwarded the repair quote to the PCM for approval got the console and the probe back
68658 | 205657 | CM | CMPLT | CMPLT | 2024-04-02 16:07 | 2024-04-02 16:08 | front display damaged, harness for display also damaged | (TC 04/02/2024 04:06:52 PM, Cortez, John) Replaced bezel, front display, right IUI Check in tests completed BD to remediate/PM
1065903 | 205658 | PM | CMPLT | CMPLT | 2024-04-02 16:08 | 2024-04-02 16:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065918 | 205659 | PM | CMPLT | CMPLT | 2024-04-02 16:20 | 2024-04-02 16:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065951 | 205660 | PM | CMPLT | CMPLT | 2024-04-02 16:22 | 2024-04-02 16:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065976 | 205661 | PM | CMPLT | CMPLT | 2024-04-02 16:23 | 2024-04-02 16:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065996 | 205662 | PM | CMPLT | CMPLT | 2024-04-02 16:25 | 2024-04-02 16:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1066012 | 205663 | PM | CMPLT | CMPLT | 2024-04-02 16:26 | 2024-04-02 16:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1066034 | 205664 | PM | CMPLT | CMPLT | 2024-04-02 16:28 | 2024-04-02 16:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1066047 | 205665 | PM | CMPLT | CMPLT | 2024-04-02 16:29 | 2024-04-02 16:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205666 | CM | CMPLT | CMPLT | 2024-04-02 16:30 | 2024-04-02 16:30 | worked on TMS | worked on TMS
1066000 | 205667 | PM | CMPLT | CMPLT | 2024-04-02 16:33 | 2024-04-02 16:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
66008 | 205668 | CM | CMPLT | CMPLT | 2024-04-03 07:18 | 2024-04-03 07:19 | power cord fallen out | (TC 04/03/2024 07:18:19 AM, Rhuland, Courtney) Reseated power cord in holder Functional check - OK Device is charging
1057304 | 205669 | CM | CMPLT | CMPLT | 2024-04-03 08:07 | 2024-04-08 10:04 | Msmt Ext Malfunction | (TC 04/02/2024 08:08:12 AM, Canjimavadivel, Kamallesh ) MSL connector corroded, swapped out with spare (TC 04/03/2024 01:30:40 PM, Canjimavadivel, Kamallesh ) Opened up X3 and removed board. Contacted Philips for quote. Placed order for MSL board (TC 04/08/2024 09:58:14 AM, Canjimavadivel, Kamallesh ) Replaced MSL board and tested in shop. Swapped back with spare in room.
65548 | 205670 | CM | CMPLT | CMPLT | 2024-04-03 08:09 | 2024-04-08 10:05 | Msmt Ext Malfunction | (TC 04/02/2024 08:09:30 AM, Canjimavadivel, Kamallesh ) MSL connector corroded, swapped out with spare (TC 04/04/2024 02:20:27 PM, Canjimavadivel, Kamallesh ) Opened up rack and found part number, need to order the entire assembly. Requested quote from vendor. Placed order for part. (TC 04/08/2024 09:55:13 AM, Canjimavadivel, Kamallesh ) Replaced connector assembly and tested in shop. Swapped back with spare from room
68034 | 205671 | CM | CMPLT | NFF | 2024-04-03 08:37 | 2024-04-03 08:38 | Temperature keeps increasing | (TC 04/03/2024 08:36:59 AM, Rhuland, Courtney) Ran incubator for an hour Device reached and held temperature and humidity No fault found
| 205672 | RDS | CMPLT | CMPLT | 2024-04-03 10:29 | 2024-04-03 13:42 | weekly checks in CCU | (TC 04/03/2024 01:41:06 PM, Scero, Tomlyn) weekly checks complete. 30 J tests passed. replaced paper in 1 defib
69462 | 205673 | PM | CMPLT | CMPLT | 2024-04-03 11:02 | 2024-04-03 11:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69470 | 205674 | PM | CMPLT | CMPLT | 2024-04-03 11:03 | 2024-04-03 11:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68656 | 205675 | PM | CMPLT | CMPLT | 2024-04-03 11:06 | 2024-04-03 11:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69555 | 205676 | PM | CMPLT | CMPLT | 2024-04-03 11:07 | 2024-04-03 11:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69162 | 205677 | PM | CMPLT | CMPLT | 2024-04-03 11:08 | 2024-04-03 11:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69159 | 205678 | PM | CMPLT | CMPLT | 2024-04-03 11:10 | 2024-04-03 11:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69570 | 205679 | PM | CMPLT | CMPLT | 2024-04-03 11:12 | 2024-04-03 11:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205680 | CM | CMPLT | CMPLT | 2024-04-03 11:15 | 2024-04-03 11:18 | chair not working | (TC 04/03/2024 11:17:12 AM, Rhuland, Courtney) Checked chair and power was cutting out intermittently Grabbed power cord from biomed and swapped cords Functional check - OK
69426 | 205681 | PM | CMPLT | CMPLT | 2024-04-03 11:17 | 2024-04-03 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1061204 | 205682 | CM | CMPLT | CMPLT | 2024-04-03 11:18 | 2024-04-03 11:21 | Light iPad frozen | (TC 04/03/2024 11:17:45 AM, Reid, Emmerson ) The iPad was frozen and asking for a password. I was not able to enter anything with the touch screen or home button. I took the iPad out of the wall mount and rebooted it. When it came back on it was working so i returned it into the wall.
69319 | 205683 | PM | CMPLT | CMPLT | 2024-04-03 11:18 | 2024-04-03 11:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69198 | 205684 | PM | CMPLT | CMPLT | 2024-04-03 11:20 | 2024-04-03 11:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69526 | 205685 | PM | CMPLT | CMPLT | 2024-04-03 11:22 | 2024-04-03 11:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
501116 | 205686 | CM | CMPLT | CMPLT | 2024-04-03 12:39 | 2024-04-04 13:11 | Probe Malfunction - Broken Strain Gauge | (TC 04/03/2024 12:39:45 PM, Chau, Ken) case #0123388390 for probe replacement (TC 04/04/2024 01:10:24 PM, Chau, Ken) OLD PROBE SN#B1VKGB NEW PROBE SN#F0BBZ1 TESTED OKAY | case #0123388390 for probe replacement
501066 | 205687 | CM | CMPLT | CMPLT | 2024-04-03 12:41 | 2024-05-02 08:20 | No Downtime Available | (TC 04/03/2024 12:40:41 PM, Chau, Ken) network issue and jdmipacs down -downtime also was not showing, deleted and readded, tested okay after
68800 | 205688 | PM | CMPLT | CMPLT | 2024-04-03 13:36 | 2024-04-03 13:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:35:41 PM, Rhuland, Courtney) PM completed by BD
68986 | 205689 | PM | CMPLT | CMPLT | 2024-04-03 13:38 | 2024-04-03 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:37:42 PM, Rhuland, Courtney) PM completed by BD
1055155 | 205690 | CM | CMPLT | CMPLT | 2024-04-03 13:40 | 2024-04-03 14:00 | door latch spring loose | (TC 04/03/2024 01:40:05 PM, Cortez, John) Replaced door latch (TC 04/03/2024 02:00:05 PM, Cortez, John) BD to remediate/PM
68824 | 205691 | PM | CMPLT | CMPLT | 2024-04-03 13:43 | 2024-04-03 13:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:43:47 PM, Rhuland, Courtney) PM completed by BD
68840 | 205692 | PM | CMPLT | CMPLT | 2024-04-03 13:46 | 2024-04-03 13:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:45:52 PM, Rhuland, Courtney) PM completed by BD
68926 | 205693 | PM | CMPLT | CMPLT | 2024-04-03 13:56 | 2024-04-03 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 01:57:00 PM, Rhuland, Courtney) PM completed by BD
68871 | 205694 | PM | CMPLT | CMPLT | 2024-04-03 14:03 | 2024-04-03 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:03:40 PM, Rhuland, Courtney) PM completed by BD
68876 | 205695 | CM | CMPLT | CMPLT | 2024-04-03 14:17 | 2024-04-03 14:18 | "battery not charging" | (TC 04/03/2024 02:17:08 PM, Cortez, John) Replaced power cord BD to remediate/PM
1048472 | 205696 | CM | CMPLT | CMPLT | 2024-04-03 14:35 | 2024-04-03 14:37 | Small det no connection. | (TC 04/03/2024 02:35:05 PM, Fan, Tak) Checked batteries, one was at 0 charge. Det OK with charged battery. 0% battery eventually started charging. Replaced battery contacts.
69209 | 205697 | PM | CMPLT | CMPLT | 2024-04-03 14:43 | 2024-04-03 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:42:46 PM, Rhuland, Courtney) PM completed by BD
69134 | 205698 | PM | CMPLT | CMPLT | 2024-04-03 14:45 | 2024-04-03 14:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:45:19 PM, Rhuland, Courtney) PM completed by BD
1069354 | 205699 | CM | CMPLT | CMPLT | 2024-04-03 14:47 | 2024-04-03 14:48 | Removed from service | (TC 04/03/2024 02:47:31 PM, Rendon Lopez , Elizabeth) Received our device back from repair. Removed from L&D, shipped back to Dornier
69223 | 205700 | PM | CMPLT | CMPLT | 2024-04-03 14:48 | 2024-04-03 14:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/03/2024 02:49:47 PM, Rhuland, Courtney) PM completed by BD
1051268 | 205701 | CM | CMPLT | CMPLT | 2024-04-03 15:15 | 2024-04-03 15:15 | Physical damage | (TC 04/03/2024 03:14:41 PM, Rendon Lopez , Elizabeth) Replaced bottom cover from retired transducer. Changed SN in TMS Tested-OK
1051282 | 205702 | CM | CMPLT | CMPLT | 2024-04-03 15:16 | 2024-04-03 15:16 | Physical damage | (TC 04/03/2024 03:15:50 PM, Rendon Lopez , Elizabeth) Replaced bottom cover from retired transducer. Changed SN in TMS Tested-OK
501621 | 205703 | CM | CMPLT | CMPLT | 2024-04-03 15:18 | 2024-07-26 14:17 | Hard to install both syringes A and B | (TC 04/03/2024 03:17:45 PM, Green, Paul) Talked to customer about the problem. Waited for patient study to complete. In the interm, got the replacement injector ready to install. Inspected injector and found it hard, sometimes to install the syringes. Will look into syringe interface A and B replacment parts. Emailed Bayer for a quote. (TC 07/26/2024 02:16:41 PM, Green, Paul) Installed A and B syringe interfaces. Tested, ok.
1055647 | 205704 | CM | CMPLT | CMPLT | 2024-04-03 15:18 | 2024-05-14 16:48 | Controls not working + pendant broken | (TC 04/03/2024 03:20:32 PM, Canjimavadivel, Kamallesh ) Tested all controls and pendant, no response. No error codes present (TC 05/02/2024 04:08:17 PM, Cortez, John) Moved broken mattress from other bed to this one. Noted errors (1041,3100) Requested service from Hillrom (TC 05/10/2024 04:17:49 PM, Cortez, John) Vince completed repair, but error 1041 remains. Reached out to Vince, and he suggested to charge overnight and see if error persists. If it does, batteries may need to be replaced again (TC 05/14/2024 04:48:27 PM, Cortez, John) Error remains, will look to order batteries
| 205705 | CM | CMPLT | CMPLT | 2024-04-03 15:22 | 2024-04-03 15:22 | MSH OPG SPP CODE GREY JDMIPACS CORAL DOWNTIME WORKED WITH ANNE TO MAKE SURE DOWNTIME-MSH WAS INSTALLED ON SYSTEMS, INSTRUCTED ON USE | (TC 04/03/2024 03:22:02 PM, Whittington, Scott) MSH OPG SPP CODE GREY JDMIPACS CORAL DOWNTIME WORKED WITH ANNE TO MAKE SURE DOWNTIME-MSH WAS INSTALLED ON SYSTEMS, INSTRUCTED ON USE | MSH OPG SPP CODE GREY JDMIPACS CORAL DOWNTIME WORKED WITH ANNE TO MAKE SURE DOWNTIME-MSH WAS INSTALLED ON SYSTEMS, INSTRUCTED ON USE
1054019 | 205706 | FC | CMPLT | NFF | 2024-04-03 15:56 | 2024-04-03 16:00 | "Unsure of accuracy" | (TC 04/03/2024 03:56:09 PM, Cortez, John) Natalia from 10N reached out saying her staff have been getting inaccurate readings with this scanner. Asked her to send it down. Caliscan test passed. Accuracy test with phantom passed (120-129 mL) Might be user error. Reached back to Natalia for next steps.
| 205707 | CM | CMPLT | CMPLT | 2024-04-03 16:07 | 2024-04-03 16:07 | worked on TMS | worked on TMS
| 205708 | CM | CMPLT | CMPLT | 2024-04-03 16:08 | 2024-04-03 16:08 | worked with RT | worked with RT
| 205709 | CM | CMPLT | CMPLT | 2024-04-03 16:09 | 2024-04-03 16:09 | meeting | meeting
| 205710 | CM | CMPLT | CMPLT | 2024-04-03 16:10 | 2024-05-01 16:58 | misc biomed - April 2024 | (TC 04/02/2024 04:09:58 PM, Cortez, John) Replaced Alaris IUIs (TC 04/03/2024 04:10:43 PM, Cortez, John) Replaced Alaris IUIs (TC 04/03/2024 04:18:09 PM, Cortez, John) Resolved Masimo client down issue on 11S with Jessica (TC 04/04/2024 04:18:09 PM, Cortez, John) Resolved Masimo client down issue on 11S with Jessica (TC 04/04/2024 04:19:38 PM, Cortez, John) Swapped Alaris IUIs (TC 04/05/2024 04:20:12 PM, Cortez, John) Created instructions for how to power cycle procuity beds, for 10N and 10S (TC 04/05/2024 04:21:10 PM, Cortez, John) Swapped IUIs (TC 04/05/2024 04:21:49 PM, Cortez, John) Swapped spare ICU beds with broken ones in ICU, with Kam (TC 04/05/2024 04:22:25 PM, Cortez, John) on-call check (TC 04/05/2024 04:22:47 PM, Cortez, John) Responded to CPET PFT call about a exercise bike not turning on. Power cord was loose. (TC 04/05/2024 04:23:50 PM, Cortez, John) Went with Rutvik to PACU for cal regarding discreptancy in vitals via Philips and manual readings. Monitor tested ok with sim. It was probably an electrode issue (TC 04/09/2024 05:32:06 PM, Cortez, John) Cleared 14N BD space (TC 04/10/2024 05:32:31 PM, Cortez, John) removed IV poles from 14S (TC 04/10/2024 05:32:56 PM, Cortez, John) Took LED covers from broken Alaris LVP front displays (TC 04/18/2024 01:12:31 PM, Cortez, John) Helped Rutvik inspect Endo light sources for leaks (TC 04/18/2024 01:13:12 PM, Cortez, John) Checked 11 flr closet for AP issue. POE power cord was loose at back (TC 04/17/2024 01:14:07 PM, Cortez, John) Swapped ICU beds around, swapped mattress and moved bariatric to 18th floor (TC 04/17/2024 01:15:34 PM, Cortez, John) systems meeting (TC 04/18/2024 01:16:16 PM, Cortez, John) Plugged Masimo/Philips in workshop to emergency power (TC 04/19/2024 01:17:26 PM, Cortez, John) Oncall daily check (TC 04/19/2024 10:36:14 AM, Cortez, John) Resolved Masimo client down issues (TC 04/26/2024 10:01:44 PM, Cortez, John) Close WOs/adjust PM scheds for 210 LVPs/PCUs
1056430 | 205711 | CM | CMPLT | CMPLT | 2024-04-03 16:12 | 2024-04-03 16:14 | "O2 reading unavailable" | (TC 04/03/2024 04:12:14 PM, Cortez, John) Received "replace probe" alert after testing for a bit. Replaced SPO2 sensor with spare. Added cuff Cleaned filter Added Rad battery All parameters tested ok
63344 | 205712 | CM | CMPLT | CMPLT | 2024-04-03 16:15 | 2024-04-22 10:36 | Prints are faint | (TC 04/02/2024 04:15:30 PM, Cortez, John) Confirmed prints are faint on lower side of paper. Wiping printhead with alcohol doesn't resolve issue (TC 04/03/2024 04:16:54 PM, Cortez, John) Disassembled unit to inspect printer assembly. There's a small crack in area of housing where printer head is fastened, but nothing else was unusual Swapped printer assembly with spare and prints are still faint (TC 04/04/2024 04:25:48 PM, Cortez, John) Tried grabbing print head from another spare and it was also faint. Found a slight crack in the housing near where print head is fastened. May try swapping housings (TC 04/04/2024 04:25:48 PM, Cortez, John) Tried grabbing print head from another spare and it was also faint. Issue might be related to slight crack in the housing near where print head is fastened (TC 04/09/2024 05:33:48 PM, Cortez, John) Swapped out board and installed in spare housing (TC 04/10/2024 10:31:32 AM, Cortez, John) Seems to be a paper issue. Loaded paper in cart, and new/packaged 3rd party paper in basket both resulted in faint prints. Installed spare Philips supplied paper and prints were better.
1066114 | 205713 | CM | CMPLT | CMPLT | 2024-04-04 07:22 | 2024-04-04 07:34 | QC Check | (TC 04/04/2024 07:33:26 AM, Rhuland, Courtney) QC checklist complete Returned to unit
1056667 | 205714 | CM | CMPLT | CMPLT | 2024-04-04 08:10 | 2024-04-04 08:12 | Arm was not moving | (TC 04/04/2024 08:10:08 AM, Reid, Emmerson ) Arm was stuck and the locking pin was also stuck. I moved the arm around to try and losen the pin, eventually moving the arm down losened it and I was able to unlock the arm. working now.
69137 | 205715 | CM | CMPLT | CMPLT | 2024-04-04 08:22 | 2024-04-04 08:22 | replaced the male iui connector
1057389 | 205716 | CM | CMPLT | CMPLT | 2024-04-04 08:22 | 2024-04-04 08:22 | ErinC@AMTVantage.com
65883 | 205717 | CM | CMPLT | CMPLT | 2024-04-04 08:24 | 2024-04-04 08:24 | replaced the male iui connector
68841 | 205718 | CM | CMPLT | CMPLT | 2024-04-04 08:24 | 2024-04-04 08:24 | replaced the male iui connector
68888 | 205719 | CM | CMPLT | CMPLT | 2024-04-04 08:26 | 2024-04-04 08:26 | replaced the female iui connector and power cord
1065705 | 205720 | CM | CMPLT | CMPLT | 2024-04-04 08:27 | 2024-04-04 08:27 | replaced the power cord
68897 | 205721 | CM | CMPLT | CMPLT | 2024-04-04 08:28 | 2024-04-04 08:28 | replaced the power cord
69269 | 205722 | CM | CMPLT | CMPLT | 2024-04-04 08:29 | 2024-04-04 08:29 | replaced the both sides iui conncector
69084 | 205723 | CM | CMPLT | CMPLT | 2024-04-04 08:30 | 2024-04-04 08:30 | replaced the both sides iui conncector
69381 | 205724 | CM | CMPLT | CMPLT | 2024-04-04 08:31 | 2024-04-04 08:31 | replaced the both sides iui conncector
68656 | 205725 | CM | CMPLT | CMPLT | 2024-04-04 08:32 | 2024-04-04 08:32 | replaced the both sides iui conncector
69501 | 205726 | CM | CMPLT | CMPLT | 2024-04-04 08:32 | 2024-04-04 08:32 | replaced the both sides iui conncector
69554 | 205727 | CM | CMPLT | CMPLT | 2024-04-04 08:33 | 2024-04-04 08:33 | replaced the both sides iui conncector
69319 | 205728 | CM | CMPLT | CMPLT | 2024-04-04 08:34 | 2024-04-04 08:34 | replaced the both sides iui conncector
68789 | 205729 | CM | CMPLT | CMPLT | 2024-04-04 08:34 | 2024-04-04 08:34 | replaced the both sides iui conncector
69360 | 205730 | CM | CMPLT | CMPLT | 2024-04-04 08:35 | 2024-04-04 08:35 | replaced the both sides iui conncector
69426 | 205731 | CM | CMPLT | CMPLT | 2024-04-04 08:37 | 2024-04-04 08:37 | replaced the female iui connector
69527 | 205732 | CM | CMPLT | CMPLT | 2024-04-04 08:38 | 2024-04-04 08:38 | replaced the female iui connector
68761 | 205733 | CM | CMPLT | CMPLT | 2024-04-04 08:38 | 2024-04-04 08:38 | replaced the female iui connector
69628 | 205734 | CM | CMPLT | CMPLT | 2024-04-04 08:39 | 2024-04-04 08:39 | replaced the female iui connector
68926 | 205735 | CM | CMPLT | CMPLT | 2024-04-04 08:40 | 2024-04-04 08:40 | replaced the female iui connector
69223 | 205736 | CM | CMPLT | CMPLT | 2024-04-04 08:41 | 2024-04-04 08:41 | replaced the female iui connector
68827 | 205737 | CM | CMPLT | CMPLT | 2024-04-04 08:41 | 2024-04-04 08:41 | replaced the female iui connector
69162 | 205738 | CM | CMPLT | CMPLT | 2024-04-04 08:42 | 2024-04-04 08:42 | replaced the female iui connector
69354 | 205739 | CM | CMPLT | CMPLT | 2024-04-04 08:43 | 2024-04-04 08:43 | replaced the female iui connector
1055121 | 205740 | CM | CMPLT | CMPLT | 2024-04-04 08:44 | 2024-04-04 08:44 | replaced the female iui connector
68759 | 205741 | CM | CMPLT | CMPLT | 2024-04-04 08:45 | 2024-04-04 08:45 | replaced the female iui connector
69046 | 205742 | CM | CMPLT | CMPLT | 2024-04-04 08:45 | 2024-04-04 08:45 | replaced the female iui connector
69663 | 205743 | CM | CMPLT | CMPLT | 2024-04-04 08:47 | 2024-04-04 08:47 | replaced the female iui connector
69516 | 205744 | CM | CMPLT | CMPLT | 2024-04-04 08:48 | 2024-04-04 08:48 | replaced the female iui connector
69595 | 205745 | CM | CMPLT | CMPLT | 2024-04-04 08:48 | 2024-04-04 08:48 | replaced the female iui connector
69538 | 205746 | CM | CMPLT | CMPLT | 2024-04-04 08:49 | 2024-04-04 08:49 | replaced the female iui connector
69055 | 205747 | CM | CMPLT | CMPLT | 2024-04-04 08:51 | 2024-04-04 08:51 | replaced the female iui connector
69526 | 205748 | CM | CMPLT | CMPLT | 2024-04-04 08:51 | 2024-04-04 08:51 | replaced the female iui connector
69644 | 205749 | CM | CMPLT | CMPLT | 2024-04-04 08:53 | 2024-04-04 08:53 | replaced the female iui connector
69136 | 205750 | CM | CMPLT | CMPLT | 2024-04-04 08:55 | 2024-04-04 08:55 | replaced the female iui connector
69215 | 205751 | CM | CMPLT | CMPLT | 2024-04-04 08:59 | 2024-04-04 08:59 | replaced the female iui connector
69420 | 205752 | CM | CMPLT | CMPLT | 2024-04-04 09:03 | 2024-04-04 09:03 | replaced the female iui connector
68704 | 205753 | CM | CMPLT | CMPLT | 2024-04-04 09:12 | 2024-04-04 09:12 | replaced the female iui connector
1047096 | 205754 | CM | CMPLT | CMPLT | 2024-04-04 09:13 | 2024-04-04 09:13 | replaced the female iui connector
68960 | 205755 | CM | CMPLT | CMPLT | 2024-04-04 09:15 | 2024-04-04 09:15 | replaced the female iui connector
68785 | 205756 | CM | CMPLT | CMPLT | 2024-04-04 09:17 | 2024-04-04 09:17 | replaced the female iui connector
68665 | 205757 | CM | CMPLT | CMPLT | 2024-04-04 09:18 | 2024-04-04 09:18 | replaced the female iui connector
65925 | 205758 | CM | CMPLT | CMPLT | 2024-04-04 09:48 | 2024-04-04 09:48 | replaced the female iui connector
69405 | 205759 | CM | CMPLT | CMPLT | 2024-04-04 09:49 | 2024-04-04 09:49 | replaced the female iui connector
69231 | 205760 | CM | CMPLT | CMPLT | 2024-04-04 09:52 | 2024-04-04 09:52 | replaced the female iui connector
69605 | 205761 | CM | CMPLT | CMPLT | 2024-04-04 09:53 | 2024-04-04 09:53 | replaced the female iui connector
69524 | 205762 | CM | CMPLT | CMPLT | 2024-04-04 09:54 | 2024-04-04 09:54 | replaced the female iui connector
68907 | 205763 | CM | CMPLT | CMPLT | 2024-04-04 09:55 | 2024-04-04 09:55 | replaced the female iui connector
68690 | 205764 | CM | CMPLT | CMPLT | 2024-04-04 09:56 | 2024-04-04 09:56 | replaced the female iui connector
69198 | 205765 | CM | CMPLT | CMPLT | 2024-04-04 10:00 | 2024-04-04 10:00 | replaced the male iui connector
68702 | 205766 | CM | CMPLT | CMPLT | 2024-04-04 10:01 | 2024-04-04 10:01 | replaced the male iui connector
69329 | 205767 | CM | CMPLT | CMPLT | 2024-04-04 10:08 | 2024-04-04 10:15 | Light house broken | (TC 04/04/2024 10:14:31 AM, Rendon Lopez , Elizabeth) Replaced front case, returned to BD for recall
69164 | 205768 | CM | CMPLT | CMPLT | 2024-04-04 10:24 | 2024-04-04 10:28 | Mismatched with new model keypad | (TC 04/04/2024 10:27:54 AM, Rendon Lopez , Elizabeth) Replaced front case. Returned to BD for recall
69053 | 205769 | CM | CMPLT | CMPLT | 2024-04-04 10:29 | 2024-04-04 10:34 | Cracked light house | (TC 04/04/2024 10:29:10 AM, Rendon Lopez , Elizabeth) Replaced front case. Returned to BD for recall
1066115 | 205770 | CM | CMPLT | CMPLT | 2024-04-04 10:49 | 2024-04-04 10:49 | QC check | passed QC check
| 205771 | CM | CMPLT | CMPLT | 2024-04-04 11:46 | 2024-11-22 07:58 | Pump spare parts | (TC 04/04/2024 11:56:32 AM, Rendon Lopez , Elizabeth) Ordering spare parts (TC 04/30/2024 09:23:17 AM, Rendon Lopez , Elizabeth) All items received except cover doors. Received 2 cover door, waiting for reamining 8. (TC 11/22/2024 07:57:46 AM, Baghirzade, Nara) Received 8 cover doors. Full order completed. PO 7500204781
69516 | 205772 | PM | CMPLT | CMPLT | 2024-04-04 12:44 | 2024-04-04 12:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 12:44:36 PM, Rhuland, Courtney) PM completed by BD
69501 | 205773 | PM | CMPLT | CMPLT | 2024-04-04 12:47 | 2024-04-04 12:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 12:47:07 PM, Rhuland, Courtney) PM completed by BD
69570 | 205774 | CM | CMPLT | CMPLT | 2024-04-04 12:55 | 2024-04-04 12:55 | replaced the male iui connector
69661 | 205775 | PM | CMPLT | CMPLT | 2024-04-04 12:56 | 2024-04-04 12:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 12:57:16 PM, Rhuland, Courtney) PM completed by BD
69639 | 205776 | CM | CMPLT | CMPLT | 2024-04-04 12:56 | 2024-04-04 12:56 | replaced the male iui connector
68802 | 205777 | CM | CMPLT | CMPLT | 2024-04-04 12:58 | 2024-04-04 12:58 | replaced the male iui connector
69671 | 205778 | CM | CMPLT | CMPLT | 2024-04-04 12:59 | 2024-04-04 12:59 | replaced both iui connectors
69555 | 205779 | CM | CMPLT | CMPLT | 2024-04-04 13:00 | 2024-04-04 13:00 | replaced both iui connectors
69360 | 205780 | PM | CMPLT | CMPLT | 2024-04-04 13:01 | 2024-04-04 13:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 01:00:49 PM, Rhuland, Courtney) PM completed by BD
68298 | 205781 | CM | CMPLT | CMPLT | 2024-04-04 13:02 | 2024-04-04 13:02 | replaced both iui connectors
69407 | 205782 | CM | CMPLT | CMPLT | 2024-04-04 13:03 | 2024-04-04 13:03 | replaced both iui connectors
69174 | 205783 | CM | CMPLT | CMPLT | 2024-04-04 13:03 | 2024-04-04 13:03 | replaced both iui connectors
69539 | 205784 | CM | CMPLT | CMPLT | 2024-04-04 13:04 | 2024-04-04 13:04 | replaced both iui connectors
69080 | 205785 | CM | CMPLT | CMPLT | 2024-04-04 13:05 | 2024-04-04 13:05 | replaced both iui connectors
69505 | 205786 | CM | CMPLT | CMPLT | 2024-04-04 13:05 | 2024-04-04 13:05 | replaced both iui connectors
69071 | 205787 | CM | CMPLT | CMPLT | 2024-04-04 13:06 | 2024-04-04 13:06 | replaced both iui connectors
69422 | 205788 | CM | CMPLT | CMPLT | 2024-04-04 13:07 | 2024-04-04 13:07 | replaced both iui connectors
1047038 | 205789 | CM | CMPLT | CMPLT | 2024-04-04 13:08 | 2024-04-04 13:08 | replaced both iui connectors
69007 | 205790 | CM | CMPLT | CMPLT | 2024-04-04 13:08 | 2024-04-04 13:08 | replaced both iui connectors
69251 | 205791 | CM | CMPLT | CMPLT | 2024-04-04 13:09 | 2024-04-04 13:09 | replaced both iui connectors
69164 | 205792 | CM | CMPLT | CMPLT | 2024-04-04 13:11 | 2024-04-04 13:11 | replaced both iui connectors
501068 | 205793 | CM | CMPLT | CMPLT | 2024-04-04 13:11 | 2024-04-04 13:12 | MSH RM#2 AND #4 - Trackball issues | (TC 04/04/2024 01:11:26 PM, Chau, Ken) Cleaned system and tested okay
69436 | 205794 | CM | CMPLT | CMPLT | 2024-04-04 13:12 | 2024-04-04 13:12 | replaced the female iui connector
68766 | 205795 | CM | CMPLT | CMPLT | 2024-04-04 13:13 | 2024-04-04 13:13 | replaced the male iui connector
69639 | 205796 | CM | CMPLT | CMPLT | 2024-04-04 13:14 | 2024-04-04 13:14 | replaced the female iui connector
68881 | 205797 | CM | CMPLT | CMPLT | 2024-04-04 13:15 | 2024-04-04 13:15 | replaced the female iui connector
69587 | 205798 | CM | CMPLT | CMPLT | 2024-04-04 13:16 | 2024-04-04 13:16 | replaced the female iui connector
69474 | 205799 | CM | CMPLT | CMPLT | 2024-04-04 13:18 | 2024-04-04 13:18 | replaced the female iui connector
69506 | 205800 | CM | CMPLT | CMPLT | 2024-04-04 13:18 | 2024-04-04 13:18 | replaced the female iui connector
69267 | 205801 | CM | CMPLT | CMPLT | 2024-04-04 13:19 | 2024-04-04 13:19 | replaced the female iui connector
69411 | 205802 | CM | CMPLT | CMPLT | 2024-04-04 13:20 | 2024-04-04 13:20 | replaced the female iui connector
69532 | 205803 | CM | CMPLT | CMPLT | 2024-04-04 13:22 | 2024-04-04 13:22 | replaced the female iui connector
69558 | 205804 | CM | CMPLT | CMPLT | 2024-04-04 13:23 | 2024-04-04 13:23 | replaced the female iui connector
68699 | 205805 | CM | CMPLT | CMPLT | 2024-04-04 13:24 | 2024-04-04 13:24 | replaced the female iui connector
68824 | 205806 | CM | CMPLT | CMPLT | 2024-04-04 13:25 | 2024-04-04 13:25 | replaced the female iui connector
1051031 | 205807 | CM | CMPLT | CMPLT | 2024-04-04 13:30 | 2024-04-04 13:34 | System slowing down, user requesting maintenance | (TC 04/04/2024 01:30:10 PM, Fan, Tak) Faith reported the touch screen being unresponsive sometimes, requiring a restart to restore functionality. Julia (newer fellow) said touch screen issues have not been a problem lately. Checked system, seems OK. Cloned drive. Replaced drive. Left old inside above BEP. Checked system again, OK.
68721 | 205808 | CM | CMPLT | CMPLT | 2024-04-04 13:30 | 2024-04-04 13:30 | replaced the female iui connector
68780 | 205809 | CM | CMPLT | CMPLT | 2024-04-04 13:31 | 2024-04-04 13:31 | replaced the female iui connector
68754 | 205810 | CM | CMPLT | CMPLT | 2024-04-04 13:32 | 2024-04-04 13:32 | replaced the female iui connector
68967 | 205811 | PM | CMPLT | CMPLT | 2024-04-04 13:38 | 2024-04-04 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 03/18/2024 01:38:42 PM, Huang, Jessica) replaced the power cord and completed PM
1059778 | 205812 | FC | CMPLT | CMPLT | 2024-04-04 13:41 | 2024-04-04 13:43 | QC Checklist | (TC 04/04/2024 01:41:02 PM, Rhuland, Courtney) QC checklist complete Returned to 11S
1050247 | 205813 | CM | CMPLT | CMPLT | 2024-04-04 13:59 | 2024-04-04 13:59 | replaced the female iui connector
1055106 | 205814 | CM | CMPLT | CMPLT | 2024-04-04 14:00 | 2024-04-04 14:00 | replaced the power cord strap
66377 | 205815 | CM | CMPLT | CMPLT | 2024-04-04 14:01 | 2024-04-04 14:03 | repurposed | (TC 04/04/2024 02:01:25 PM, Scero, Tomlyn) Tested unit and charged the battery. Changed the settings to monitor spot check. Gave to the pacemaker clinic temporarily until I can find a suitable 74/7500 welch allyn.
68731 | 205816 | CM | CMPLT | CMPLT | 2024-04-04 14:02 | 2024-04-04 14:02 | replaced the power cord
| 205817 | CM | CMPLT | CMPLT | 2024-04-04 14:47 | 2024-04-04 14:47 | helping Courtney to unbox and tagged new incoming breast pumps and moved them to the storage room
69647 | 205818 | PM | CMPLT | PMFAIL | 2024-04-04 14:57 | 2024-04-04 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
69647 | 205819 | CM | CMPLT | CMPLT | 2024-04-04 14:58 | 2024-04-04 15:03 | Created by Failed PM WO # 205818 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 03:01:49 PM, Rendon Lopez , Elizabeth) Bezel assembly wire damaged. Replaced bezel assembly, calibrated pump. PM completed, all tests passed | PM failed by BD
120009 | 205820 | CM | CMPLT | CMPLT | 2024-04-04 15:01 | 2024-04-04 15:01 | not working | (TC 03/19/2024 03:03:44 PM, Huang, Jessica) replaced the BP cuff and the SPO2 sensor
| 205821 | CM | CMPLT | CMPLT | 2024-04-04 15:08 | 2024-04-04 15:08 | ordering the connector for the laerdal
1048046 | 205822 | CM | CMPLT | CMPLT | 2024-04-04 15:10 | 2024-04-04 15:10 | not working | (TC 03/26/2024 03:10:49 PM, Huang, Jessica) Julie sent the device and saying it was not alarming when it was disconnected from the patient. charged the battery, could not recreat the same issue, the device seemed functional informed Julie and showed her the device she agreed to monitor the device
69551 | 205823 | CM | CMPLT | CMPLT | 2024-04-04 15:11 | 2024-04-04 15:20 | Created by Failed PM WO # 198798 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 03:18:33 PM, Rendon Lopez , Elizabeth) Door latch Sagging. Replaced front case, found bezel ground wire broken. Replaced bezel assy. Calibrated rate. PM completed, all tests passed | PM failed by BD
1046467 | 205824 | CM | CMPLT | CMPLT | 2024-04-04 15:14 | 2024-04-04 15:14 | not working | (TC 03/26/2024 03:14:07 PM, Huang, Jessica) replaced the blood pressure cuff and SPO2
1046521 | 205825 | CM | CMPLT | CMPLT | 2024-04-04 15:15 | 2024-04-04 15:15 | not working | (TC 03/26/2024 03:15:07 PM, Huang, Jessica) replaced the blood pressure cuff and the SPO2 sensor
1067958 | 205826 | CM | CNP | ENL | 2024-04-04 15:17 | 2024-04-12 15:31 | Mounting Issue | Apr 04 - Call re: need to mount this monitor on the wall. I advised to connect with building services but was insisted to visit. Travelled to IVF. Found confirmed that I cannot mount the monitor on the wall as first, the panel needs to be mounted on the wall by building services. Also, upon unboxing and checking the mounting arm, I found that the mount arm is not compatible with CVSM 6000. Advised staff to connect with the sales rep to return the incorrect mounting arm and order the correct mounting arm. Will address the issue when the correct mounting arm is available. Apr 12 - Did not receive any update on this issue from IVF. WO Closed. Will create a new when they get back to us.
1055174 | 205827 | CM | CMPLT | CMPLT | 2024-04-04 15:22 | 2024-04-04 15:47 | Created by Failed PM WO # 205532 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 03:22:38 PM, Rendon Lopez , Elizabeth) Door latch sagging. Replaced front door. Fails rate accuracy during PM, calibrated rate. PM completed, all tests passed | PM failed by BD
541052 | 205828 | CM | CMPLT | CMPLT | 2024-04-04 15:25 | 2024-04-04 15:25 | removed from service
1060187 | 205829 | CM | CMPLT | CMPLT | 2024-04-04 15:27 | 2024-04-04 15:27 | replaced batteries
66242 | 205830 | CM | CMPLT | CMPLT | 2024-04-04 15:28 | 2024-05-02 14:57 | SPO2 not working | Apr 04 - SPO2 not working. Went to check and brought it down to biomed. Cleaned unit. Foudn caster wheel jammed. Cleaned it. SPO2 ext cable and probe missing. Searched in the office and asked staff but found none in the office. Contacted Masimo for part# and pricing. Awaiting response. Apr 10 - Moved unit to MCR until the part is delivered. May 02 - PArts received. Brought the monitor back to biomed. Replaced SPO2 cable and sensor. Checked functionlity. Tested ok.
1069748 | 205831 | CM | CMPLT | CMPLT | 2024-04-04 15:28 | 2024-04-04 15:28 | incoming device for demo
541271 | 205832 | CM | CMPLT | CMPLT | 2024-04-04 15:30 | 2024-04-04 15:30 | removed from service
541021 | 205833 | CM | CMPLT | CMPLT | 2024-04-04 15:32 | 2024-04-04 15:32 | not working | (TC 01/23/2024 03:32:16 PM, Huang, Jessica) replaced the SPO2 sensor and the temperature probe
1069643 | 205834 | CM | CMPLT | CMPLT | 2024-04-04 15:43 | 2024-04-04 15:43 | new incoming | (TC 01/10/2024 03:42:43 PM, Huang, Jessica) unboxed the device, tested and tagged it, sent it to MBU 56moh 95.7 uA
1069644 | 205835 | CM | CMPLT | CMPLT | 2024-04-04 15:45 | 2024-04-04 15:45 | new incoming | (TC 01/10/2024 03:44:49 PM, Huang, Jessica) unboxed the device, tested and tagged it, sent it to MBU 56moh 95.7 uA
1046461 | 205836 | CM | CMPLT | CMPLT | 2024-04-04 15:47 | 2024-04-04 15:47 | not working | (TC 01/16/2024 03:46:53 PM, Huang, Jessica) replaced the bp cuff and the SPO2 sensor
1046462 | 205837 | CM | CMPLT | CMPLT | 2024-04-04 15:48 | 2024-04-04 15:48 | not working | (TC 01/16/2024 03:48:03 PM, Huang, Jessica) replaced the battery and the SPO2 sensor
541061 | 205838 | CM | CMPLT | CMPLT | 2024-04-04 15:49 | 2024-04-04 15:49 | not working | (TC 01/16/2024 03:49:24 PM, Huang, Jessica) replaced the battery and the SPO2 sensor
541313 | 205839 | CM | CMPLT | CMPLT | 2024-04-04 15:50 | 2024-04-04 15:50 | not working | (TC 01/16/2024 03:50:15 PM, Huang, Jessica) replaced the battery and the SPO2 sensor
1069841 | 205840 | CM | CMPLT | CMPLT | 2024-04-04 15:53 | 2024-04-04 15:53 | new incoming | (TC 01/31/2024 03:52:46 PM, Huang, Jessica) unboxed the device and the cart, installed the device on the cart. tested and tagged it, contacted Chris from Medtronic for software upgraded, 56 moh; 98.7 uA took the cart to OR and informed the OR team
1069842 | 205841 | CM | CMPLT | CMPLT | 2024-04-04 15:55 | 2024-04-04 15:55 | new incoming | (TC 01/31/2024 03:54:44 PM, Huang, Jessica) unboxed the device and the cart, contacted Chris from Medtronic for software upgraded installed the device in OR 15. tested and tagged it, 56 moh; 98.7 uA
1047042 | 205842 | PM | CMPLT | PMFAIL | 2024-04-04 15:57 | 2024-04-04 15:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM failed by BD
1047042 | 205843 | CM | CMPLT | CMPLT | 2024-04-04 15:58 | 2024-04-04 16:00 | Created by Failed PM WO # 205842 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/04/2024 03:58:17 PM, Rendon Lopez , Elizabeth) Failed patient and bottle pressure sensor tests. Replaced both pressure sensors. Calibrated rate. PM completed, all tests passed | PM failed by BD
68302 | 205844 | CM | CMPLT | CMPLT | 2024-04-05 07:37 | 2024-04-05 07:38 | Missing latch | (TC 04/05/2024 07:37:39 AM, Rhuland, Courtney) Replaced missing latch module Returned to BD for recall and PM
1061203 | 205845 | CM | CMPLT | CMPLT | 2024-04-05 08:28 | 2024-04-05 08:28 | DVI connector on the boom not working | (TC 02/13/2024 08:28:15 AM, Huang, Jessica) got the call in OR 2 saying the stryker device on the boom was not able to show the image on the monitors. disconnected the DVI cable and reconnected it, could not fix the issue. replaced the DVI cable and it was the same. went back to the room to replace the DVI connector on the boom and tested it, working fine informed the OR team
1061202 | 205846 | CM | CMPLT | CMPLT | 2024-04-05 08:32 | 2024-04-05 08:32 | DVI connector on the boom not working | (TC 02/13/2024 08:32:21 AM, Huang, Jessica) got the call in OR 1 and saying the device was not showing the image on the monitors. tried the different DVI cable and it didnt fix the issue went back to the room and replaced the DVI connector on the foot boom tested it, working fine
1061208 | 205847 | CM | CMPLT | CMPLT | 2024-04-05 08:35 | 2024-04-05 08:35 | DVI connector not working | (TC 02/13/2024 08:34:48 AM, Huang, Jessica) connected the device to the foot boom and did not display the image on the monitors, tried the different DVI cable and it didnt fix the issue went back to the room and replaced the DVI connector on the foot boom tested it, working fine
1061208 | 205848 | CM | CMPLT | CMPLT | 2024-04-05 08:36 | 2024-04-05 08:36 | ipad not working | (TC 03/22/2024 08:36:19 AM, Huang, Jessica) got the call in OR 7 saying the ipad was not working went to the room and reset it, working fine showed the nurse how to get to the lights setting
1061208 | 205849 | CM | CMPLT | CMPLT | 2024-04-05 08:38 | 2024-04-05 08:38 | surgical light 3 drafting | (TC 03/22/2024 08:38:16 AM, Huang, Jessica) got the call in OR 7 saying the light 3 was drafting tighted the screw
1061216 | 205850 | CM | CMPLT | CMPLT | 2024-04-05 08:42 | 2024-04-05 08:42 | images delayed on the monitors | (TC 03/22/2024 08:42:32 AM, Huang, Jessica) got the call in OR 15, the surgeon complained the monitors delayed the image, stayed in the room and could not see the issue; however, the clinical team was complaining the image quality. informed the nurse to get the spare tower on the hallway. helped to set up the tower and stayed in the room for awhile
1061219 | 205851 | CM | CMPLT | CMPLT | 2024-04-05 08:48 | 2024-04-05 08:48 | the gas machine did not connect to the right port | (TC 02/14/2024 08:48:24 AM, Huang, Jessica) got the call from Maureen and saying the user could not connect use open the chat on the gas machine, went to OR 18 and connected the cat cable to right port
1061204 | 205852 | CM | CMPLT | CMPLT | 2024-04-05 08:56 | 2024-04-05 08:56 | olympus touch panel froze | (TC 01/25/2024 08:56:21 AM, Huang, Jessica) got the call in OR 3 and the olympus touch panel froze, restarted the VMC a couple times and it didnt fix the issue. contacted Derrick for support, disconnected the cables from the VMC and reconnected them, swapped the VMC between OR 12 and OR 3, changed the settings, could not get it worked informed Maureen the room could not be used for next day (TC 01/26/2024 08:59:20 AM, Huang, Jessica) Derrick came onsite and helped to contact TVOne Tech Support rebulit the proflie on the CORIOMatrix
| 205853 | CM | CMPLT | CMPLT | 2024-04-05 09:09 | 2024-04-12 09:51 | 2 Dispensing Guns not working | (TC 04/05/2024 09:10:39 AM, Chukri, Samir) Repaired one gun give it to Sheena to try it and let me know if it is okay to do the second one. (TC 04/10/2024 11:01:26 AM, Chukri, Samir) repaired the second dispensing gun .
1058179 | 205854 | CM | CMPLT | CMPLT | 2024-04-05 09:12 | 2024-04-05 09:13 | Dose cord not working | (TC 04/05/2024 09:12:13 AM, Reid, Emmerson ) Dose cord was not working. I replaced it with another one because it was broken for some reason.
1061203 | 205855 | CM | CMPLT | CMPLT | 2024-04-05 09:27 | 2024-04-05 09:27 | showed the user how to use the recorder on the olympus touch panel | (TC 02/23/2024 09:27:36 AM, Huang, Jessica) got the call from the OR desk and went to OR 2 the nurse forgot how to start the new recording on the olympus touch panel; showed him the steps got the call again from the same user to show him again the steps
1061217 | 205856 | CM | CMPLT | CMPLT | 2024-04-05 09:35 | 2024-04-05 09:35 | audio cable broken | (TC 01/24/2024 09:34:52 AM, Huang, Jessica) received the email from HP and saying the audio cable in OR 16 was broken; went to the room and checked it contacted olympus for part number and quote was informed it was regulator audio cable and olympus did not provide it googled the part and informed Adeel to purchase a few of them from grand and toy (TC 02/01/2024 09:39:00 AM, Huang, Jessica) replaced the audio cable in OR 16 and infromed OR team
1061211 | 205857 | CM | CMPLT | CMPLT | 2024-04-05 09:52 | 2024-04-05 09:52 | Dr. Taymaa May questions on the stryker tower | (TC 01/23/2024 09:51:52 AM, Huang, Jessica) got the call from OR 10 and wen to the room. Dr. Taymma May was asking if the sryker tower has been upgraded recently. Apparently TGH has the same tower and it was upgraded and she wanted to know it sinai could get the same. contacted the sale rep Mike Kneeland from stryker and received the emal from him fowarded it to the OR team
1061211 | 205858 | CM | CMPLT | CMPLT | 2024-04-05 09:57 | 2024-04-05 09:57 | Olympu 4K and 200 system connecting on the boom 0 | (TC 01/31/2024 09:57:18 AM, Huang, Jessica) got the call from Rose in OR 10 because the video system cables were messed up went to the room and reconnected them (TC 02/08/2024 09:59:24 AM, Huang, Jessica) got the call from Rose again and she wanted to label the conectors on the boom labelled the connectors
| 205859 | CM | CMPLT | CMPLT | 2024-04-05 10:01 | 2024-04-05 10:02 | Replacing bulb in amnio room | (TC 04/05/2024 10:01:14 AM, Rhuland, Courtney) Brought new bulb over to OPG and replaced bulb Light is functional Informed contact that light was fixed
1051500 | 205860 | CM | CMPLT | CMPLT | 2024-04-05 11:04 | 2024-04-05 11:16 | Broken cord | (TC 04/05/2024 11:14:49 AM, Canjimavadivel, Kamallesh ) Replaced cord, tested ok
1057867 | 205861 | PM | CMPLT | CMPLT | 2024-04-05 13:29 | 2024-04-05 14:24 | Annual PM | (TC 04/05/2024 02:11:35 PM, Reid, Emmerson ) Needed to calibrate the set temperature. Completed PM, see procedure.
| 205862 | MEP | CMPLT | CMPLT | 2024-04-05 13:40 | 2024-05-02 08:21 | TGH GW Room | (TC 04/05/2024 01:40:04 PM, Chau, Ken) IR required GW room for cases -removed med eng equipment from room and organized for IR
| 205863 | CON | CMPLT | CMPLT | 2024-04-05 13:42 | 2024-04-05 13:43 | MSH OPG CEOU - Samsung Incoming | (TC 04/05/2024 01:42:07 PM, Chau, Ken) Received and setup DICOM, tested unit (TC 04/08/2024 11:23:41 AM, Whittington, Scott) MSH OPG CEOU - Samsung Incoming SAMSUNG HERA W10 FOLLOWUP WITH BROCK (TC 04/15/2024 11:27:09 AM, Whittington, Scott) MSH OPG CEOU - Samsung Incoming SYSTE3M REMOVED FROM SITE | MSH OPG CEOU - Samsung Incoming
1070663 | 205864 | CM | CMPLT | CMPLT | 2024-04-05 14:04 | 2024-04-05 14:05 | Incoming inspection | (TC 04/05/2024 02:04:13 PM, Rhuland, Courtney) Helped the Laborie sales rep to unbox and assemble chair Functional check - OK
65921 | 205865 | PM | CMPLT | CMPLT | 2024-04-05 14:06 | 2024-04-05 14:07 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/05/2024 02:06:21 PM, Rhuland, Courtney) Calibration and check in completed All tests passed Refer to ASM for details
68740 | 205866 | PM | CMPLT | CMPLT | 2024-04-05 14:22 | 2024-04-05 15:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 03:15:02 PM, Huang, Jessica) replaced the power cord and completed PM 85 moh 72.3 uA
1069320 | 205867 | CM | CMPLT | CMPLT | 2024-04-05 14:29 | 2024-04-05 14:32 | Mounting monitor to rolling stand | (TC 04/05/2024 02:30:45 PM, Rhuland, Courtney) Mounted monitor onto spare rollling stand as requested by Dr. Poorva Cable tied power cord to arm of stand
68722 | 205868 | PM | CMPLT | CMPLT | 2024-04-05 14:47 | 2024-04-05 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/05/2024 02:47:31 PM, Rhuland, Courtney) PM completed by BD
| 205869 | CM | CMPLT | CMPLT | 2024-04-05 14:54 | 2024-04-05 14:59 | Suction Regulator Issue in Broncospopy Room | Call re" Suction not working" Went to check. Was informed that the suction is in the broncoscopy room on 5th floor. Waited for staff to take me there as a key is needed for access. Checked suction tree with two regulators. Found Metal outlet connector that connects to the canister via threads was loose. Removed it, cleaned old teflon tape, put new teflon tape and and tightened it. Found black o ring in suction cannister connector was worn off and broken. Advised staff to replace the cannister as well. Checked functionality. Tested ok. Returned to Service.
| 205870 | CM | CMPLT | CMPLT | 2024-04-05 15:00 | 2024-04-05 15:02 | YelloFin Stirrup Issue | Call re: not working. S/N : 12223000176. Went to check and brought it down to biomed. Unit cannot be checked without access to OR. Contacted OR desk and reserved OR for functional Check. Went to OR and checked functionality. Could not find any issue. Unit is functioning as intended. Advised OR on proper usage. Returned to service.
| 205871 | CM | CMPLT | CMPLT | 2024-04-05 15:04 | 2024-04-12 15:34 | YelloFin Stirrup Issue | Apr 05 - Call re: not working. S/N : 12223000176. Went to check and brought it down to biomed. Unit cannot be checked without access to OR. Contacted OR desk and reserved OR for functional Check. Went to OR and checked functionality. Found the latch/handle has lost its tension and gas cylinder unable to move back and forth. Contacted tech support and requested repair quote. Awaiting. Apr 08 - Followed up. Repair estimate received. Communicated with OR manager for approval. Awaiting. Apr 11 - Followed up. Was informed to check and deterning the age of the equipment. Unit is not in TMS. Consulted OR staff and JH. Based on S/N unit was found to be old. Was advised bt Reeba Philip to retire the equipment if it is too old. Again, contacted vendor to find out the purhcase date. Awaiting response. Apr 12 - Vendor could not find the purchase date either. Discussed with JH and determined that it is very old. Kept leg pads for future use and disposed unit into electronic waste.
1063783 | 205872 | CM | CMPLT | CMPLT | 2024-04-05 15:19 | 2024-04-05 15:19 | not working | (TC 04/05/2024 03:19:10 PM, Huang, Jessica) picked the light from OR desk and sent it for repair received it back and sent it back to OR
| 205873 | CM | CMPLT | CMPLT | 2024-04-05 15:28 | 2024-04-05 15:28 | brought two Yellowfins stirrups to OR for inspection in OR 19
| 205874 | CM | CMPLT | CMPLT | 2024-04-05 15:29 | 2024-04-05 15:29 | helped Courtney to tag new incoming breast pumps | (TC 04/04/2024 03:29:35 PM, Huang, Jessica) helped Courtney to unbox the new incoming breast pumps, tagged them, assembled them on the stands and moved them to 17th floor
1057870 | 205875 | CM | CMPLT | CMPLT | 2024-04-05 15:31 | 2024-04-05 15:31 | Repair | Call re: not working. Went to check and brought it down to biomed. Found power cord missing. Replaced power cord. Checked functinality and temp accuracy . Tested ok. Delivered back to OR.
1062127 | 205876 | CM | CMPLT | CMPLT | 2024-04-05 16:29 | 2024-04-05 16:30 | "no function" | (TC 04/05/2024 04:29:17 PM, Cortez, John) Power cycling bed seems to resolve issue
1061446 | 205877 | CM | CMPLT | NFF | 2024-04-05 16:30 | 2024-04-05 16:31 | "bed alarm doesn't work" | (TC 04/04/2024 04:30:06 PM, Cortez, John) NFF, bed alarm works normally. Power cycled bed anyway
1055909 | 205878 | CM | CMPLT | CMPLT | 2024-04-05 16:32 | 2024-04-05 16:34 | "BP needs calibrating | (TC 04/04/2024 04:31:57 PM, Cortez, John) Added cuff and hose added filter All parameters tested ok
| 205879 | CM | CMPLT | CMPLT | 2024-04-08 07:18 | 2024-04-11 07:43 | Sterile water for NICU altrix | (TC 04/08/2024 07:17:36 AM, Rhuland, Courtney) Ordered water for NICU PO 4150274893
1055643 | 205880 | CM | CMPLT | CMPLT | 2024-04-08 07:44 | 2024-05-08 15:36 | Bed malfunctioning | (TC 04/05/2024 03:15: PM, Canjimavadivel, Kamallesh ) Bed was in ICU with note saying malfunction. Swapped out with bed from 18th floor (TC 05/02/2024 04:10:39 PM, Cortez, John) Ran a functional test and didn't run into any issues. Will hold off on requesting service
| 205881 | RDS | CMPLT | CMPLT | 2024-04-08 07:54 | 2024-04-08 08:11 | Weekly Checks For OR | (TC 04/08/2024 08:10:42 AM, Reid, Emmerson ) Rounds complete
| 205882 | RDS | CMPLT | CMPLT | 2024-04-08 08:11 | 2024-04-08 08:13 | Weekly Checks For PACU | (TC 04/08/2024 08:12:31 AM, Reid, Emmerson ) Rounds complete
1057941 | 205883 | CM | CMPLT | CMPLT | 2024-04-08 08:17 | 2024-04-08 08:18 | Brakes not engaging | (TC 04/08/2024 08:17:08 AM, Reid, Emmerson ) Table brakes were not engaging and an error was coming up. I reset the table, and engaged the brakes. It is working now.
1046295 | 205884 | CM | CMPLT | CMPLT | 2024-04-08 08:19 | 2024-04-08 08:20 | Smoke Evac not synced | (TC 04/08/2024 08:18:55 AM, Reid, Emmerson ) None of the cables were connected. I setup the cables and foot pedals and tested. Working now.
1066475 | 205885 | CM | CMPLT | CMPLT | 2024-04-08 08:21 | 2024-04-29 14:52 | Dirty Filter | (TC 04/08/2024 08:21:34 AM, Reid, Emmerson ) Pulled the unit because the filter needs to be changed and unit needs investigation. (TC 04/08/2024 10:05:17 AM, Reid, Emmerson ) Created safer report 72561 and communicated info to involved parties. (TC 04/08/2024 03:14:25 PM, Patel, Rutvik) Brought down unit from OR to Biomed. Worked with ER to remove outer covers and started invenstigation on the unit. (TC 04/09/2024 09:06:56 AM, Reid, Emmerson ) Meeting to duscuss this event (TC 04/10/2024 09:03:02 AM, Reid, Emmerson ) Brought the unit back up to OR (TC 04/10/2024 03:05:49 PM, Reid, Emmerson ) Went to bring both units down with Rutvik to MCR storage room as per Maureen's request to have them out of the OR. Ran into Maureen, Tania, and Reeba and discuessed the issue with them. We also ran a cleaning cycle on both units. (TC 04/29/2024 02:51:49 PM, Reid, Emmerson ) Submitted the PO request and confirmed the location of the last working neptune.
1066476 | 205886 | CM | CMPLT | CMPLT | 2024-04-08 08:22 | 2024-04-16 11:28 | Dirty Filter | (TC 04/08/2024 08:22:29 AM, Reid, Emmerson ) Pulled the unit because the filter needs to be changed and unit needs investigation. (TC 04/08/2024 10:04:06 AM, Reid, Emmerson ) Created safer report 72561 and communicated info to involved parties. (TC 04/09/2024 02:58:35 PM, Reid, Emmerson ) Spoke with Styker and biomed team about the issue (TC 04/16/2024 11:02:19 AM, Reid, Emmerson ) Meeting to go over breifing note before meeting to go over briefing note. (TC 04/16/2024 11:28:07 AM, Reid, Emmerson ) Meeting to discuss briefing note.
1061333 | 205887 | FC | CMPLT | CMPLT | 2024-04-08 08:51 | 2024-04-08 08:51 | completed QC test | (TC 04/05/2024 08:51:05 AM, Huang, Jessica) Sarabeth sent the email and requested to complete the QC test completed it and informed Sarabeth
| 205888 | RDS | CMPLT | CMPLT | 2024-04-08 09:14 | 2024-04-08 14:04 | Weekly checks in CCU | (TC 04/08/2024 02:03:13 PM, Scero, Tomlyn) weekly checks complete.
68570 | 205889 | CM | CMPLT | CMPLT | 2024-04-08 09:44 | 2024-04-08 09:45 | set up monitor with camera system | (TC 04/08/2024 09:44:18 AM, Rhuland, Courtney) Set up converter to be compatible with camera system Functional check - OK (TC 05/02/2024 09:05:10 AM, Rhuland, Courtney) Fixed connections again - set up second monitor (TC 05/10/2024 10:57:03 AM, Rhuland, Courtney) Monitors were moved and reconfigured setup guide created for users
69350 | 205890 | CM | CMPLT | CMPLT | 2024-04-08 09:54 | 2024-04-08 09:54 | replaced the female iui connector
69597 | 205891 | CM | CMPLT | CMPLT | 2024-04-08 09:55 | 2024-04-08 09:55 | replaced the female iui connector
69470 | 205892 | CM | CMPLT | CMPLT | 2024-04-08 09:56 | 2024-04-08 09:56 | replaced the female iui connector
68762 | 205893 | CM | CMPLT | CMPLT | 2024-04-08 09:57 | 2024-04-08 09:57 | replaced the female iui connector
68840 | 205894 | CM | CMPLT | CMPLT | 2024-04-08 09:59 | 2024-04-08 09:59 | replaced the female iui connector
68786 | 205895 | CM | CMPLT | CMPLT | 2024-04-08 10:00 | 2024-04-08 10:00 | replaced the female iui connector
66692 | 205896 | PM | CMPLT | CMPLT | 2024-04-08 10:01 | 2024-04-08 10:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 04/08/2024 10:00:40 AM, Rhuland, Courtney) Hood replaced PM completed Rechargeable 9v battery replaced Calibrated: Main and line voltage Scale Functional check - OK Set 35C, 75% humiditiy Read 35C, 76% humidity Electrical safety - OK 87.5uA, 58 mOhms
69463 | 205897 | CM | CMPLT | CMPLT | 2024-04-08 10:01 | 2024-04-08 10:01 | replaced the female iui connector
| 205898 | CM | CMPLT | CMPLT | 2024-04-08 10:02 | 2024-05-06 15:19 | Biomed Stock Parts Ordering | Apr 08 - Contacted vendor and ordered Masimo SPO2 extension cable and SPO2 sensor for WA VSM 6000 for biomed stock. Created SC 1003205200 . Awaiting delivery. Apr 26 - Searched and created SC 1003218636 for biomed soldering parts. May 06 - All parts received. Sorted and organized in appropriate shelves.
| 205899 | RDS | CMPLT | CMPLT | 2024-04-08 10:25 | 2024-04-08 10:27 | Weekly check for ED
1055748 | 205900 | CM | ACTIV | ISSUE | 2024-04-08 10:37 |  | replaced the rear case | (TC 04/03/2024 10:37:43 AM, Huang, Jessica) the pump was sent by BD team while they were upgrading the software replaced the rear case and sent back to BD for software upgraded and PM (TC 04/03/2024 10:37:43 AM, Huang, Jessica) the pump was sent by BD team while they were upgrading the software replaced the rear case and sent back to BD for software upgraded and PM
66054 | 205901 | CM | CMPLT | CMPLT | 2024-04-08 10:39 | 2024-04-08 10:39 | replaced the power cord | (TC 04/03/2024 10:39:26 AM, Huang, Jessica) the pump was sent by BD while they were upgrading the software replaced the power cord and sent back to BD for software upgraded and PM
65906 | 205902 | CM | CMPLT | CMPLT | 2024-04-08 10:42 | 2024-04-08 10:42 | replaced the rear case | (TC 04/03/2024 10:41:47 AM, Huang, Jessica) the pump was sent by BD while they were upgrading the software replaced the rear case and sent back to BD for software upgraded and PM
| 205903 | CM | CMPLT | CMPLT | 2024-04-08 11:18 | 2024-04-08 11:18 | went to 11th floor to check the masimo central monitors
1048524 | 205904 | CM | CMPLT | CMPLT | 2024-04-08 11:29 | 2024-04-08 11:30 | Broken front case | (TC 04/08/2024 11:29:27 AM, Rhuland, Courtney) Replaced front case Replaced mylar gasket Returned to BD for recall and PM
1056125 | 205905 | PM | CMPLT | CMPLT | 2024-04-08 11:33 | 2024-04-08 11:37 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/08/2024 11:34:05 AM, Canjimavadivel, Kamallesh ) Cleaned hoses and nozzles. Cleaning cycle complete
1064895 | 205906 | CM | CMPLT | CMPLT | 2024-04-08 11:46 | 2024-04-08 11:48 | Static noise | (TC 04/08/2024 11:47:25 AM, Rendon Lopez , Elizabeth) Confirmed issue, previously repaired. Retired from service.
1055665 | 205907 | CM | CMPLT | CMPLT | 2024-04-08 11:49 | 2024-04-08 11:49 | unable to hear FHR | (TC 04/08/2024 11:48:45 AM, Rendon Lopez , Elizabeth) Confirmed issue, previously repaired. Retired from service.
1057499 | 205908 | CM | CMPLT | NFF | 2024-04-08 12:08 | 2024-04-08 12:08 | Not picking up | (TC 04/08/2024 12:07:47 PM, Rendon Lopez , Elizabeth) Tested, NFF. Cleaned back membrane.
1070670 | 205909 | ACC | CMPLT | CMPLT | 2024-04-08 13:41 | 2024-04-08 13:43 | Incoming Inspection | (TC 04/08/2024 01:42:26 PM, Reid, Emmerson ) Incoming complete. For a single day demo April 9, 2024
67970 | 205910 | CM | CMPLT | CMPLT | 2024-04-08 14:31 | 2024-04-09 08:48 | Fan always on high error | (TC 04/08/2024 02:30:46 PM, Rhuland, Courtney) Error when running incubator Replaced control board and error still present (TC 04/09/2024 08:46:52 AM, Rhuland, Courtney) Replaced relay board and heat sink thermistor and error cleared PM to be completed in PM work order
| 205911 | CM | CMPLT | CMPLT | 2024-04-08 15:06 | 2024-04-08 15:10 | Power bar issue | Call from Endoscopy re: power bar on ER cart is not working. Went to check. Found power bar plug pronge bent. Spokw with the staff and they confirmed that the power bar was struck by a cart. corrected the pronge position. Secured power bar cable on the wall using tie wrap and cable holders. Tested ok.
1056911 | 205912 | CM | CMPLT | CMPLT | 2024-04-09 08:28 | 2024-07-12 10:05 | Provido Cracked Screen | (TC 04/09/2024 08:37:22 AM, Reid, Emmerson ) Covered crack with tape and spoke with Steve (TC 04/10/2024 09:02:29 AM, Reid, Emmerson ) Took pictures of the monitor's model number and sent to Steve (TC 04/12/2024 09:40:10 AM, Reid, Emmerson ) Spoke with the vendor and communicated with Reeba about how we should proceed. (TC 04/12/2024 01:48:13 PM, Reid, Emmerson ) Removed the monitor and let the OR team know. (TC 07/12/2024 10:05:00 AM, Reid, Emmerson ) Working with Rutvik. Replaced the monitor. Needed to change the video cable to a display cable. Also Vacuumed the inside for dust. Working now.
1055644 | 205913 | CM | CMPLT | CMPLT | 2024-04-09 08:43 | 2024-04-09 08:44 | Swapping nurse call to 37-pin | (TC 04/09/2024 08:43:28 AM, Canjimavadivel, Kamallesh ) Swapped nurse call cable to 37-pin cable
| 205914 | PA | CMPLT | CMPLT | 2024-04-09 08:45 | 2024-04-12 14:08 | Lifeband order for ED | (TC 04/09/2024 08:49:52 AM, Canjimavadivel, Kamallesh ) Placed order for lifebands (TC 04/12/2024 02:03:02 PM, Canjimavadivel, Kamallesh ) Received order and delivered to ED supplies room.
1061207 | 205915 | CM | CMPLT | CMPLT | 2024-04-09 09:15 | 2024-04-09 09:21 | ESUs were moved on the boom | (TC 04/09/2024 09:14:46 AM, Reid, Emmerson ) Had to switch the position of the ESU and SmokeEvac on the boom. I also synced the two.
| 205916 | RDS | CMPLT | CMPLT | 2024-04-09 09:21 | 2024-04-08 09:21 | Weekly check for 14S and17M | Weekly check for 14S and17M
1061207 | 205917 | CM | CMPLT | CMPLT | 2024-04-09 09:23 | 2024-04-09 09:23 | DVI Connector broken | (TC 04/08/2024 09:22:39 AM, Reid, Emmerson ) Replaced the DVI connector
| 205918 | CM | CMPLT | CMPLT | 2024-04-09 09:32 | 2024-04-08 09:32 | worked with RT and Draeger tech | worked with RT and Draeger tech
| 205919 | CM | CMPLT | CMPLT | 2024-04-09 10:46 | 2024-04-29 14:01 | EC5 for AGMs | (TC 04/09/2024 10:52:32 AM, Rendon Lopez , Elizabeth) 1 EC5 for repair and 2 for spare Received
1051327 | 205920 | CM | CMPLT | CMPLT | 2024-04-09 11:37 | 2024-06-04 14:35 | Set up for pacemaker clinic
1070673 | 205921 | CM | CMPLT | CMPLT | 2024-04-09 13:17 | 2024-04-09 14:13 | Incoming inspection | (TC 04/09/2024 02:12:46 PM, Rhuland, Courtney) Brought to OPG and tagged
1070674 | 205922 | CM | CMPLT | CMPLT | 2024-04-09 13:20 | 2024-04-09 14:14 | Incoming inspection | (TC 04/09/2024 02:13:38 PM, Rhuland, Courtney) Brought to OPG and tagged
1065567 | 205923 | CM | CMPLT | CMPLT | 2024-04-09 13:57 | 2024-04-18 14:17 | Large leak | (TC 04/18/2024 02:13:41 PM, Scero, Tomlyn) Replaced 3 year PM parts and retested. Leak is now <50mmHg and passes the self test.
63187 | 205924 | CM | CMPLT | CMPLT | 2024-04-09 14:41 | 2024-05-31 10:47 | Missing focusing sleeve | (TC 04/09/2024 02:44:36 PM, Rhuland, Courtney) Went and looked at light and focusing sleeve was missing Light was hanging out of device Contacting vendor for replacement price (TC 05/31/2024 10:46:37 AM, Rhuland, Courtney) Received focusing sleeve and installed on broken light Functional check - OK
1055790 | 205925 | CM | CMPLT | CMPLT | 2024-04-09 14:57 | 2024-04-09 14:57 | Not plugged in | (TC 04/09/2024 02:57:07 PM, Reid, Emmerson ) Plugged in, working now.
| 205926 | CM | CMPLT | CMPLT | 2024-04-09 15:22 | 2024-04-09 15:22 | worked on TMS | worked on TMS
69493 | 205927 | PM | CMPLT | CMPLT | 2024-04-09 15:25 | 2024-04-09 15:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69086 | 205928 | PM | CMPLT | CMPLT | 2024-04-09 15:28 | 2024-04-09 15:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69522 | 205929 | PM | CMPLT | CMPLT | 2024-04-09 15:30 | 2024-04-09 15:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1069285 | 205930 | CM | CMPLT | CMPLT | 2024-04-09 15:32 | 2024-04-11 13:13 | MSH US ER G7 3rd (Main) Floor CANON PRISM PROBE TV PVT781 FAULTY CRYSTAL | (TC 04/09/2024 03:32:23 PM, Whittington, Scott) MSH US ER G7 3rd (Main) Floor CANON PRISM PROBE TV PVT781 FAULTY CRYSTAL | MSH US ER G7 3rd (Main) Floor CANON PRISM PROBE TV PVT781 FAULTY CRYSTAL APRIL 11 2024 SN# FSD2422089 REPLACED FSD2375605
1057511 | 205931 | CM | CMPLT | CMPLT | 2024-04-09 15:33 | 2024-04-09 15:35 | Bipolar Foor Switch Issue | Call re: not working. Went to check. Or staff busy. Waited for access. Checekd unit and found bipolar foot switch cable was damaged. Came bakc to biomed and found the bipolar footswitch in biomed alsop had the similar damage. Went back to OR and replaced the bipolar footswitch from a spare unit in OR hall way. returned back with faulty footswitch. Tested ok.
69415 | 205932 | PM | CMPLT | CMPLT | 2024-04-09 15:36 | 2024-04-09 15:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1063649 | 205933 | CM | CMPLT | CMPLT | 2024-04-09 15:37 | 2024-04-12 13:13 | MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRST INSIDE THE DOOR INTERMITTENT | (TC 04/09/2024 03:37:49 PM, Whittington, Scott) MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRT INSIDE THE DOOR INTERMITTENT (TC 04/09/2024 03:38:05 PM, Dolinar, John) MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRT INSIDE THE DOOR INTERMITTENT HDMI CABLEW LOOSE, RESET (TC 04/09/2024 03:37:49 PM, Whittington, Scott) MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRT INSIDE THE DOOR INTERMITTENT HDMI CABLEW LOOSE, RESET (TC 04/12/2024 01:12:19 PM, Whittington, Scott) MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRT INSIDE THE DOOR INTERMITTENT RYAN LAROQUE CONFIRMED NO OTHER ISSUES WITH SYSTEM (TC 04/19/2024 01:39:47 PM, Whittington, Scott) MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRT INSIDE THE DOOR INTERMITTENT APRIL 18 SYSTEM AGAIN HAS ISSUES. DID NOT BOOT UP PROPERLY AFTER POSSIBLE CONSTRUCTION CAUSED POWER OUTAGE (TC 04/19/2024 01:41:00 PM, Dolinar, John) MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRT INSIDE THE DOOR INTERMITTENT APRIL 18 SYSTEM AGAIN HAS ISSUES. DID NOT BOOT UP PROPERLY AFTER POSSIBLE CONSTRUCTION CAUSED POWER OUTAGE (TC 05/03/2024 07:55:04 AM, Whittington, Scott) MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRST INSIDE THE DOOR INTERMITTENT APC UPS DELIVERED AND INSTALLED WITH RYAN L ON MAY 1 WS APPEARS TO BE OPERATING NORMALLY (TC 05/27/2024 03:55:10 PM, Whittington, Scott) MSH BI RAD REWIEW STATION 1 INSIDE DOOR "SYNC APP"ERROR | MSH BI RAD REVIEW ROOM 12th Floor - North Side SECUREVIEW WS FIRT INSIDE THE DOOR INTERMITTENT APC UPS 1500VA Smart-UPS with SmartConnect, Pure Sinewave UPS Battery Backup, Line Interactive, 120V Uninterruptible Power Supply Visit the APC by Schneider Electric Store 4.6 4.6 out of 5 stars 875 ratings Amazon's Choice $759.50$759.50 Size: 1500VA
69300 | 205934 | PM | CMPLT | CMPLT | 2024-04-09 15:38 | 2024-04-09 15:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1057511 | 205935 | CM | CMPLT | CMPLT | 2024-04-09 15:38 | 2024-06-04 08:27 | Bipolar Foot Switch/Caster wheel Repair | Apr 09 - Found Two bipolar foorswitches in biomed that had cable damage. Consulted SC. Arranged for the solder and related accessories. Cut the damaged wire and resoldered the connector for both bipolar footswitch. Tested ok. Found one footswitch with the connector broken. Contacted a couple of vendors to inquire about the availability of the connector. Awaiting confirmation. Apr 10 - Found connectors on newark. Created SC 1003207278 for purchase and sent for approval. Awaiting. Apr 11 - Inquired about a spare bipolar footswitch for biomed. Found caster wheels is not balanced and requre replacement. Reached out to vendor but was advised they do not seel caster wheels. Reached out to a couple of third party vendors to see if caster wheels are available. Apr 12 - Repaired and replaced another bipolar footswith cable connector. Followed up with Medtronic re; parts inquiry. Apr 15 - Explored vendors for caster wheels. Apr 16 - Foudn a vendor after extensive search and reached out. Received confirmation on availability. Contacted MMC for account details. Apr 24 - Followed up. Info received. Created SC 1003217065 for purchase and sent for approval. Awaiting delivery. MAy 02 - Followed up. Was informed the ETA is 3-4 weeks. Awaiting delivery. May 22 - Caster wheels received. Attempted to replace the casters. Found threads were damaged due to caster stem bending previously. Awaiting for SC to advise. June 04 - Consulted SC. Repaired threads on two FT10 cars and replaced casters. Tested ok. Delivered to OR.
69389 | 205936 | PM | CMPLT | CMPLT | 2024-04-09 15:39 | 2024-04-09 15:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69388 | 205937 | PM | CMPLT | CMPLT | 2024-04-09 15:42 | 2024-04-09 15:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69413 | 205938 | PM | CMPLT | CMPLT | 2024-04-09 15:43 | 2024-04-09 15:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69294 | 205939 | PM | CMPLT | CMPLT | 2024-04-09 15:52 | 2024-04-09 15:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69383 | 205940 | PM | CMPLT | CMPLT | 2024-04-09 15:56 | 2024-04-09 15:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69299 | 205941 | PM | CMPLT | CMPLT | 2024-04-09 15:58 | 2024-04-09 15:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69529 | 205942 | PM | CMPLT | CMPLT | 2024-04-09 15:59 | 2024-04-09 15:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
32510 | 205943 | CM | CMPLT | CMPLT | 2024-04-10 07:54 | 2024-04-10 07:57 | MSH DR1 DETECTOR WOULD NOT CALIBRATE | (TC 04/10/2024 07:56:37 AM, Whittington, Scott) MSH DR1 DETECTOR WOULD NOT CALIBRATE (TC 04/10/2024 07:56:00 PM, Whittington, Scott) MSH DR1 DETECTOR WOULD NOT CALIBRATE | MSH DR1 DETECTOR WOULD NOT CALIBRATE
32520 | 205944 | CM | CMPLT | CMPLT | 2024-04-10 07:59 | 2024-04-19 13:32 | MSH DR2 TUBE ALPHA DRIVE | (TC 04/10/2024 08:00:47 AM, Whittington, Scott) MSH DR2 TUBE ALPHA DRIVE (TC 04/12/2024 08:31:00 PM, Whittington, Scott) MSH DR2 TUBE ALPHA DRIVE | MSH DR2 TUBE ALPHA DRIVE
68759 | 205945 | CM | CMPLT | CMPLT | 2024-04-10 09:00 | 2024-04-10 09:01 | Corrected serial number | (TC 04/10/2024 08:59:51 AM, Rhuland, Courtney) Internal serial number did not match external serial number Flashed correct serial number Functional check - OK
68759 | 205946 | PM | CMPLT | CMPLT | 2024-04-10 09:02 | 2024-04-10 09:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/10/2024 09:01:45 AM, Rhuland, Courtney) PM completed by BD
1061214 | 205947 | CM | CMPLT | CMPLT | 2024-04-10 09:02 | 2024-04-10 09:02 | Monitor loose | (TC 04/10/2024 09:01:48 AM, Reid, Emmerson ) Tightened the tilt brakes, working now.
69653 | 205948 | PM | CMPLT | CMPLT | 2024-04-10 09:07 | 2024-04-10 09:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68641 | 205949 | PM | CMPLT | CMPLT | 2024-04-10 09:10 | 2024-04-10 09:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
68638 | 205950 | PM | CMPLT | CMPLT | 2024-04-10 09:12 | 2024-04-10 09:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69498 | 205951 | PM | CMPLT | CMPLT | 2024-04-10 09:13 | 2024-04-10 09:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69100 | 205952 | PM | CMPLT | CMPLT | 2024-04-10 09:15 | 2024-04-10 09:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69495 | 205953 | PM | CMPLT | CMPLT | 2024-04-10 09:17 | 2024-04-10 09:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69117 | 205954 | PM | CMPLT | CMPLT | 2024-04-10 09:18 | 2024-04-10 09:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69235 | 205955 | PM | CMPLT | CMPLT | 2024-04-10 09:19 | 2024-04-10 09:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69641 | 205956 | PM | CMPLT | CMPLT | 2024-04-10 09:22 | 2024-04-10 09:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69280 | 205957 | PM | CMPLT | CMPLT | 2024-04-10 09:34 | 2024-04-10 09:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69277 | 205958 | PM | CMPLT | CMPLT | 2024-04-10 09:36 | 2024-04-10 09:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69283 | 205959 | PM | CMPLT | CMPLT | 2024-04-10 09:38 | 2024-04-10 09:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
69273 | 205960 | PM | CMPLT | CMPLT | 2024-04-10 09:39 | 2024-04-10 09:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1055163 | 205961 | PM | CMPLT | CMPLT | 2024-04-10 10:08 | 2024-04-10 10:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065781 | 205962 | PM | CMPLT | CMPLT | 2024-04-10 10:12 | 2024-04-10 10:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065796 | 205963 | PM | CMPLT | CMPLT | 2024-04-10 10:13 | 2024-04-10 10:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065836 | 205964 | PM | CANCL | CANCL | 2024-04-10 10:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065836 | 205965 | PM | CMPLT | CMPLT | 2024-04-10 10:15 | 2024-04-10 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065871 | 205966 | PM | CMPLT | CMPLT | 2024-04-10 10:17 | 2024-04-10 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065886 | 205967 | PM | CMPLT | CMPLT | 2024-04-10 10:19 | 2024-04-10 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
1065959 | 205968 | PM | CMPLT | CMPLT | 2024-04-10 10:21 | 2024-04-10 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed
| 205969 | CM | CMPLT | CMPLT | 2024-04-10 10:22 | 2024-04-10 10:22 | worked on TMS | worked on TMS
| 205970 | CM | CMPLT | CMPLT | 2024-04-10 10:38 | 2024-04-12 15:29 | Parts Purchase for Biomed | Parts purchase for biomed (TC 04/12/2024 03:28:49 PM, Patel, Rutvik) Parts received. Organized then in the cabinet.
1066114 | 205971 | CM | CMPLT | CMPLT | 2024-04-10 13:27 | 2024-04-10 13:27 | QC check | (TC 04/10/2024 01:27:02 PM, Rhuland, Courtney) QC checklist complete
1053379 | 205972 | CM | CMPLT | CMPLT | 2024-04-10 14:36 | 2024-04-24 11:42 | High pressure alarm. CPAP around 60 cmH2O, apparent water on back of ventilator. | Collected logs and sent them to Getinge. Run device check. Passed. Running unit overnight to see/create error. Tested unit on Monday. Functional. No error. Contacted Getinge. They will investigate the logs. Got reply from Getinge, Alex"After reviewing the logs, it appears that the reported issue could be due to user error. Specifically an issue with the circuit (possible leakage, bad position or kinked tubing). I see many instances of user alarms in the Events Log. Alarms such as “Leakage too high”, “Check tubing”, “CPAP low” repeated consistently throughout the day on April 7th. I do not see any indication of technical errors related to the failure of a specific component. However, I do see that some parts of the pre-use check were failing on April 7th around 21:00 (see screenshot below). If these errors are not reproducible during the pre-use check at this time, then we can assume the pressure transducers are ok. " Passed device check again. Run unit 2 days. No eror. no humidity. Functional. Returned back to service. Informed Rex.
65905 | 205973 | PM | CMPLT | CMPLT | 2024-04-10 14:38 | 2024-04-10 14:38 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:37:46 PM, Rhuland, Courtney) PM completed by BD
65933 | 205974 | PM | CMPLT | CMPLT | 2024-04-10 14:40 | 2024-04-10 14:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:39:48 PM, Rhuland, Courtney) PM completed by BD
65983 | 205975 | PM | CMPLT | CMPLT | 2024-04-10 14:46 | 2024-04-10 14:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:46:25 PM, Rhuland, Courtney) PM completed by BD
65951 | 205976 | PM | CMPLT | CMPLT | 2024-04-10 14:48 | 2024-04-10 14:49 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:48:19 PM, Rhuland, Courtney) PM completed by BD
1050616 | 205977 | PM | CMPLT | CMPLT | 2024-04-10 14:52 | 2024-04-10 14:52 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/10/2024 02:51:52 PM, Rhuland, Courtney) PM completed by BD
| 205978 | CM | CMPLT | CMPLT | 2024-04-10 15:57 | 2024-07-29 08:57 | PM Procedure | Apr 10 - Started working on ATS 4000 PM Procedure July 29 - On halt. WO complete.
| 205979 | OC | CMPLT | CMPLT | 2024-04-10 17:27 | 2024-04-15 13:27 | On-call - April 5-11 | (TC 04/07/2024 05:27:11 PM, Cortez, John) Call from OR saying they had issue with Conmed ESU. Smoke evac wasnt actuating when ESU was fired. They had a spare, so they left device at nurse station. Informed Emmerson (TC 04/07/2024 05:28:43 PM, Cortez, John) Call from OR 6 saying that touch display for monitors was dispaying error. "video matrix not available" Asked nurse to try resetting VMC by powering off unit in cabinet. Error didn't resolve. They were set up in this OR already, so they would prefer not to change rooms. Reached out to Emmerson, who suggested to reset Corio as well as VMC. This ended up working. (TC 04/11/2024 09:34:22 PM, Cortez, John) 11N called saying they couldn't see vitals on Masimo viewstation. Tried remoting in. There was no remote desktop shortcut for 11N viewstation in biomed support server (need to resolve this later) , so I had to remote in manually using IP address for 11N (IPs in masimo spreadsheet) Confirmed issue. 11N domain not showing up. Restarting viewstation doesn't resolve issue. As a short term solution, logged into Masimo profile user profile instead, amd opened safetynet application. Vitals can be seen now. Called 11N nurse station to confirm On Monday, will look into issue (TC 04/11/2024 09:50:14 PM, Cortez, John) 11N called back saying Masimo screen is acting up again. Remoted in and saw that windows logged out automatically. Looked into power settings and saw that power/sleep settings set to "never," but screensaver set to 5 mins. Toggled off screensaver and remotely viewed station for 30 mins to confirm if it remains logged in.
1055812 | 205980 | PA | CMPLT | CMPLT | 2024-04-11 07:53 | 2024-04-24 10:16 | Cables for ICU defib | (TC 04/11/2024 07:58:32 AM, Canjimavadivel, Kamallesh ) Received quote from vendor and placed order for cable and connector (TC 04/24/2024 10:06:16 AM, Canjimavadivel, Kamallesh ) Received and added to stock shelf
| 205981 | CM | CMPLT | CMPLT | 2024-04-11 08:07 | 2024-04-11 08:09 | Rounds - NICU | (TC 04/11/2024 08:08:57 AM, Rhuland, Courtney) NICU rounds completed
| 205982 | CM | CMPLT | CMPLT | 2024-04-11 08:10 | 2024-04-11 08:11 | Rounds - RESUS | (TC 04/11/2024 08:10:38 AM, Rhuland, Courtney) RESUS rounds completed
1047844 | 205983 | PM | CMPLT | CMPLT | 2024-04-11 08:49 | 2024-04-11 09:45 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 04/11/2024 09:44:01 AM, Rhuland, Courtney) PM completed Refer to procedure for details
1057840 | 205984 | CM | CMPLT | CMPLT | 2024-04-11 09:13 | 2024-04-11 09:14 | Giving Error | (TC 04/11/2024 09:13:10 AM, Reid, Emmerson ) Noticed unit had the wrong foot pedal. I changed it and tested, was working fine. Sending back.
1059828 | 205985 | PM | CMPLT | CMPLT | 2024-04-11 09:16 | 2024-04-11 09:17 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/11/2024 09:16:20 AM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1046480 | 205986 | CM | CMPLT | CMPLT | 2024-04-11 09:22 | 2024-04-29 13:56 | Smoke Evac Damaged front panel | (TC 04/11/2024 09:29:13 AM, Reid, Emmerson ) Switch with another unit and contacted vendor for a replacement. (TC 04/15/2024 10:45:36 AM, Reid, Emmerson ) Put in a request for a NO CHARGE PO SC 1003210014 (TC 04/15/2024 01:20:34 PM, Reid, Emmerson ) Placed KJV Shipping request (TC 04/29/2024 01:55:09 PM, Reid, Emmerson ) Unboxed the repaired unit and added a new asset tag. Also attached the service report to WO. Working now.
62696 | 205987 | CM | CMPLT | CMPLT | 2024-04-11 09:30 | 2024-04-11 09:31 | Not picking up signal | (TC 04/11/2024 09:30:30 AM, Canjimavadivel, Kamallesh ) RL pin on PIM was loose. Opened up PIM and resoldered pin onto PCB. Tested ok
1057542 | 205988 | PM | CMPLT | CMPLT | 2024-04-11 09:43 | 2024-04-11 14:20 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 04/11/2024 02:19:13 PM, Scero, Tomlyn) Battery replaced. PM tests complete and passed.
1047840 | 205989 | PM | CMPLT | CMPLT | 2024-04-11 09:51 | 2024-06-05 10:23 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/05/2024 10:22:32 AM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1070710 | 205990 | CM | CMPLT | CMPLT | 2024-04-11 12:09 | 2024-04-11 12:10 | Broken drewer holder complete case | (TC 04/11/2024 12:08:56 PM, Chukri, Samir) replaced a completed drewer case Sara brought it to me. checked working fine.
| 205991 | CM | CMPLT | CMPLT | 2024-04-11 12:15 | 2024-05-01 08:12 | MSH IVF 250 DUNDAS CANON PVT 375 PROBE NEEDED GROMMET TAPED PVT 375 SN# | (TC 04/11/2024 12:14:48 PM, Whittington, Scott) MSH IVF 250 DUNDAS CANON PVT 375 PROBE NEEDED GROMMET TAPED | MSH IVF 250 DUNDAS CANON PVT 375 PROBE NEEDED GROMMET TAPED
1069285 | 205992 | ACC | CMPLT | CMPLT | 2024-04-11 12:39 | 2024-09-25 09:39 | MSH ER US CANON PROBE PVT 781 APRIL 11 2024 SN# FSD2422089 REPLACED FSD2375605 | (TC 04/11/2024 12:40:48 PM, Whittington, Scott) MSH ER US CANON PROBE PVT 781 APRIL 11 2024 SN# FSD2422089 REPLACED FSD2375605 | APRIL 11 2024 SN# FSD2422089 REPLACED FSD2375605 ULT050 AIRTAG
1063746 | 205993 | PM | CMPLT | CMPLT | 2024-04-11 12:52 | 2024-04-11 12:53 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/11/2024 12:52:24 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1047851 | 205994 | CM | CMPLT | CMPLT | 2024-04-11 13:04 | 2024-04-11 13:07 | Error | (TC 04/11/2024 01:04:08 PM, Rhuland, Courtney) Pump motor drive off post alarm present when error logs were checked Ran multiple test infusions and was unable to replicate an error Functional check - OK
| 205995 | CM | CMPLT | CMPLT | 2024-04-11 13:50 | 2024-04-11 13:54 | Exam bed would not go down - beeps when button is pressed | (TC 04/11/2024 01:50:18 PM, Rhuland, Courtney) Exam bed stuck in highest position Drawer at the end of the bed was falling off Reattached bracket and drawer to inside of bed Reseated sensors inside drawer that had shifted - causing bed to think there was something stuck under the drawer Once sensors were reseated, bed moved as expected.
65896 | 205996 | PM | CMPLT | PMFAIL | 2024-04-11 13:58 | 2024-04-11 13:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed by BD
65896 | 205997 | CM | CMPLT | CMPLT | 2024-04-11 13:58 | 2024-04-11 14:11 | Created by Failed PM WO # 205996 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/11/2024 02:08:04 PM, Rendon Lopez , Elizabeth) Flange detect sensor failed. ordered replacement sensor assy. Flange detect membrane getting stucked, replaced front case, mylar gasket, flange detect sensor assy, and IUI board. Calibrated. PM completed, all tests passed | PM failed by BD
| 205998 | CM | CMPLT | CMPLT | 2024-04-11 14:16 | 2024-04-11 14:16 | meeting | meeting
501107 | 205999 | CM | CMPLT | CMPLT | 2024-04-11 14:16 | 2024-04-11 14:17 | Caster brakes not working/Monitor dropping (found during philips PM) | (TC 04/10/2024 02:16:13 PM, Fan, Tak) Replaced both caster wheels. (TC 04/10/2024 11:16:00 AM, Fan, Tak) Located system, asked for availability. (TC 04/12/2024 01:15:34 PM, Fan, Tak) Replaced gas spring. (TC 05/08/2024 02:57:46 PM, Fan, Tak) Replaced operator panel gas spring assembly.
| 206000 | CM | CMPLT | CMPLT | 2024-04-11 14:17 | 2024-04-11 14:17 | worked with RT | worked with RT
67944 | 206001 | CM | CMPLT | CMPLT | 2024-04-11 14:19 | 2024-04-15 10:05 | Broken porthole doors | (TC 04/11/2024 02:19:36 PM, Rhuland, Courtney) Replaced 2 broken porthole doors (TC 04/15/2024 10:04:22 AM, Rhuland, Courtney) Functional check - OK
66056 | 206002 | PM | CMPLT | CMPLT | 2024-04-11 14:20 | 2024-04-11 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/11/2024 02:20:14 PM, Reid, Emmerson ) Replaced front and rear case. Downloaded and cleared logs. Ran ASM PM, all tests passed. PM Complete.
| 206003 | CM | CMPLT | CMPLT | 2024-04-11 14:21 | 2024-05-27 09:41 | X3 coloured handles | (TC 04/11/2024 02:22:33 PM, Scero, Tomlyn) Ordered orange handles for the X3s in CCU. Quotation # Q-00295914 (TC 05/22/2024 08:44:11 AM, Scero, Tomlyn) handles delivered (TC 05/27/2024 09:41:29 AM, Scero, Tomlyn) All handles replaced.
66013 | 206004 | PM | CMPLT | CMPLT | 2024-04-11 14:30 | 2024-04-11 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
66013 | 206005 | PM | CANCL | DUP | 2024-04-11 14:30 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66043 | 206006 | PM | CMPLT | CMPLT | 2024-04-11 14:34 | 2024-04-11 14:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68288 | 206007 | PM | CMPLT | CMPLT | 2024-04-11 14:36 | 2024-04-11 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69009 | 206008 | PM | CMPLT | CMPLT | 2024-04-11 14:38 | 2024-04-11 14:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68696 | 206009 | PM | CMPLT | CMPLT | 2024-04-11 14:41 | 2024-04-11 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69001 | 206010 | PM | CMPLT | CMPLT | 2024-04-11 14:42 | 2024-04-11 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68697 | 206011 | PM | CMPLT | CMPLT | 2024-04-11 14:43 | 2024-04-11 14:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68595 | 206012 | PM | CMPLT | PMFAIL | 2024-04-11 14:44 | 2024-04-11 14:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM failed, bent plunger pins
68595 | 206013 | CM | CMPLT | CMPLT | 2024-04-11 14:44 | 2024-04-11 14:50 | Created by Failed PM WO # 206012 : PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/11/2024 02:49:16 PM, Rendon Lopez , Elizabeth) Replaced plunger head. Calibrated. PM completed, all tests passed | PM failed, bent plunger pins
68694 | 206014 | PM | CMPLT | CMPLT | 2024-04-11 14:45 | 2024-04-11 14:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68915 | 206015 | PM | CMPLT | CMPLT | 2024-04-11 14:47 | 2024-04-11 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68955 | 206016 | PM | CMPLT | CMPLT | 2024-04-11 14:47 | 2024-04-11 14:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68933 | 206017 | PM | CMPLT | CMPLT | 2024-04-11 14:49 | 2024-04-11 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68950 | 206018 | PM | CMPLT | CMPLT | 2024-04-11 14:50 | 2024-04-11 14:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69013 | 206019 | PM | CMPLT | CMPLT | 2024-04-11 14:52 | 2024-04-11 14:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68930 | 206020 | PM | CMPLT | CMPLT | 2024-04-11 14:53 | 2024-04-11 14:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68703 | 206021 | PM | CMPLT | CMPLT | 2024-04-11 14:54 | 2024-04-11 14:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68900 | 206022 | PM | CMPLT | CMPLT | 2024-04-11 14:58 | 2024-04-11 14:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68980 | 206023 | PM | CMPLT | CMPLT | 2024-04-11 14:59 | 2024-04-11 14:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68996 | 206024 | PM | CMPLT | CMPLT | 2024-04-11 15:00 | 2024-04-11 15:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68820 | 206025 | PM | CMPLT | CMPLT | 2024-04-11 15:03 | 2024-04-11 15:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68828 | 206026 | PM | CMPLT | CMPLT | 2024-04-11 15:04 | 2024-04-11 15:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68731 | 206027 | PM | CMPLT | CMPLT | 2024-04-11 15:07 | 2024-04-11 15:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68683 | 206028 | PM | CMPLT | CMPLT | 2024-04-11 15:10 | 2024-04-11 15:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68884 | 206029 | PM | CMPLT | CMPLT | 2024-04-11 15:11 | 2024-04-11 15:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055127 | 206030 | PM | CMPLT | CMPLT | 2024-04-11 15:12 | 2024-04-11 15:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058271 | 206031 | PM | CMPLT | CMPLT | 2024-04-11 15:13 | 2024-04-11 15:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058282 | 206032 | PM | CMPLT | CMPLT | 2024-04-11 15:15 | 2024-04-11 15:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058288 | 206033 | PM | CMPLT | CMPLT | 2024-04-11 15:16 | 2024-04-11 15:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065690 | 206034 | PM | CMPLT | CMPLT | 2024-04-11 15:17 | 2024-04-11 15:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065718 | 206035 | PM | CMPLT | CMPLT | 2024-04-11 15:18 | 2024-04-11 15:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1069391 | 206036 | CM | CMPLT | CMPLT | 2024-04-12 08:07 | 2024-04-24 14:26 | "System Failure" error message | (TC 04/12/2024 10:19:09 AM, Canjimavadivel, Kamallesh ) Brought machine down as well as second machine that had issue. Looked into service manual. Error message requires looking at logs via service mode. Reached out to vendor to come on site and take look at both machines. (TC 04/24/2024 02:25:33 PM, Canjimavadivel, Kamallesh ) Received service report from vendor and attached to WO. Device was returned to unit.
1069389 | 206037 | CM | CMPLT | CMPLT | 2024-04-12 08:09 | 2024-04-24 14:30 | Flow problem warning during treatment | (TC 04/24/2024 02:29:23 PM, Canjimavadivel, Kamallesh ) Received service report from vendor and attached to WO. Device was returned to unit.
69475 | 206038 | PM | CMPLT | CMPLT | 2024-04-12 08:41 | 2024-04-12 08:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69472 | 206039 | PM | CMPLT | CMPLT | 2024-04-12 08:43 | 2024-04-12 08:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69577 | 206040 | PM | CMPLT | CMPLT | 2024-04-12 08:44 | 2024-04-12 08:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69528 | 206041 | PM | CMPLT | CMPLT | 2024-04-12 08:47 | 2024-04-12 08:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69160 | 206042 | PM | CMPLT | CMPLT | 2024-04-12 08:49 | 2024-04-12 08:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69158 | 206043 | PM | CMPLT | CMPLT | 2024-04-12 08:50 | 2024-04-12 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69146 | 206044 | PM | CMPLT | CMPLT | 2024-04-12 08:52 | 2024-04-12 08:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69607 | 206045 | PM | CMPLT | CMPLT | 2024-04-12 08:53 | 2024-04-12 08:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69179 | 206046 | PM | CMPLT | CMPLT | 2024-04-12 08:54 | 2024-04-12 08:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69175 | 206047 | PM | CMPLT | CMPLT | 2024-04-12 08:55 | 2024-04-12 08:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69332 | 206048 | PM | CMPLT | CMPLT | 2024-04-12 08:56 | 2024-04-12 08:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69419 | 206049 | PM | CMPLT | CMPLT | 2024-04-12 08:58 | 2024-04-12 08:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69110 | 206050 | PM | CMPLT | CMPLT | 2024-04-12 08:59 | 2024-04-12 08:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69372 | 206051 | PM | CMPLT | CMPLT | 2024-04-12 09:00 | 2024-04-12 09:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69056 | 206052 | PM | CMPLT | CMPLT | 2024-04-12 09:01 | 2024-04-12 09:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69138 | 206053 | PM | CMPLT | CMPLT | 2024-04-12 09:02 | 2024-04-12 09:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69038 | 206054 | PM | CMPLT | CMPLT | 2024-04-12 09:03 | 2024-04-12 09:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69369 | 206055 | PM | CMPLT | CMPLT | 2024-04-12 09:05 | 2024-04-12 09:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69537 | 206056 | PM | CMPLT | CMPLT | 2024-04-12 09:06 | 2024-04-12 09:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69356 | 206057 | PM | CMPLT | CMPLT | 2024-04-12 09:07 | 2024-04-12 09:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69289 | 206058 | PM | CMPLT | CMPLT | 2024-04-12 09:10 | 2024-04-12 09:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69268 | 206059 | PM | CMPLT | CMPLT | 2024-04-12 09:11 | 2024-04-12 09:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69238 | 206060 | PM | CMPLT | CMPLT | 2024-04-12 09:12 | 2024-04-12 09:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69651 | 206061 | PM | CMPLT | CMPLT | 2024-04-12 09:14 | 2024-04-12 09:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058238 | 206062 | PM | CMPLT | CMPLT | 2024-04-12 09:15 | 2024-04-12 09:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058251 | 206063 | PM | CMPLT | CMPLT | 2024-04-12 09:17 | 2024-04-12 09:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058252 | 206064 | PM | CMPLT | CMPLT | 2024-04-12 09:18 | 2024-04-12 09:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065751 | 206065 | PM | CMPLT | CMPLT | 2024-04-12 09:19 | 2024-04-12 09:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065762 | 206066 | PM | CMPLT | CMPLT | 2024-04-12 09:21 | 2024-04-12 09:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065799 | 206067 | PM | CMPLT | CMPLT | 2024-04-12 09:22 | 2024-04-12 09:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065806 | 206068 | PM | CMPLT | CMPLT | 2024-04-12 09:24 | 2024-04-12 09:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065814 | 206069 | PM | CMPLT | CMPLT | 2024-04-12 09:25 | 2024-04-12 09:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065828 | 206070 | PM | CMPLT | CMPLT | 2024-04-12 09:26 | 2024-04-12 09:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065864 | 206071 | PM | CMPLT | CMPLT | 2024-04-12 09:27 | 2024-04-12 09:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065893 | 206072 | PM | CMPLT | CMPLT | 2024-04-12 10:17 | 2024-04-12 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69614 | 206073 | PM | CMPLT | CMPLT | 2024-04-12 10:22 | 2024-04-12 10:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/12/2024 10:23:01 AM, Reid, Emmerson ) Downloaded and cleared logs. Ran ASM PM, all tests passed. PM Complete.
1065928 | 206074 | PM | CMPLT | CMPLT | 2024-04-12 11:23 | 2024-04-12 11:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065974 | 206075 | PM | CMPLT | CMPLT | 2024-04-12 11:24 | 2024-04-12 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065993 | 206076 | PM | CMPLT | CMPLT | 2024-04-12 11:25 | 2024-04-12 11:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066017 | 206077 | PM | CMPLT | CMPLT | 2024-04-12 11:26 | 2024-04-12 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066029 | 206078 | PM | CMPLT | CMPLT | 2024-04-12 11:28 | 2024-04-12 11:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066052 | 206079 | PM | CMPLT | CMPLT | 2024-04-12 11:28 | 2024-04-12 11:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066054 | 206080 | PM | CMPLT | CMPLT | 2024-04-12 11:29 | 2024-04-12 11:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066057 | 206081 | PM | CMPLT | CMPLT | 2024-04-12 11:31 | 2024-04-12 11:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
66533 | 206082 | CM | CMPLT | CMPLT | 2024-04-12 12:46 | 2024-04-12 12:55 | SAFER report - Fetal monitor heartbeat sound louder than FHR measured | (TC 04/12/2024 12:53:08 PM, Rendon Lopez , Elizabeth) Checked monitor, all functions working properly. Checked both US transducers in room asset 1066320 and 1063184, working properly, and reading accurately with no issues of loud noise or static noises. Notified Gillian, she updated report | Received call from Gillian, Monitor in room 6 had a SAFER reported overnight, stated the heart rate sound was louder and faster than the FHR displayed on fetal monitor screen and Fetalink.
1053349 | 206083 | CM | CMPLT | CMPLT | 2024-04-12 13:03 | 2024-04-19 14:29 | MSH BI 1274 H2 MSH H2 CPU UPGRADE | (TC 04/12/2024 01:10:56 PM, Whittington, Scott) MSH BI 1274 H2 MSH H2 CPU UPGRADE (TC 04/13/2024 01:34:56 PM, Whittington, Scott) MSH BI 1274 H2 MSH H2 CPU UPGRADE CONNECTING TO NETWORK (TC 04/13/2024 01:34:56 PM, Whittington, Scott) MSH BI 1274 H2 MSH H2 CPU UPGRADE CONNECTING TO NETWORK | MSH BI 1274 H2 MSH H2 CPU UPGRADE APRIL 13 2024 CPU REPLACED NEW MAC 3C:EC:EF:C4:64:F5 REPLACES AC-1F-6B-B2-AE-DA
501072 | 206084 | CM | CMPLT | CMPLT | 2024-04-12 13:17 | 2024-04-12 13:17 | Brakes not working, monitor dropping (found during philips PM) | (TC 04/12/2024 01:17:00 PM, Fan, Tak) Replaced gas spring and both casters.
68681 | 206085 | PM | CMPLT | CMPLT | 2024-04-12 13:40 | 2024-04-12 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68992 | 206086 | PM | CMPLT | CMPLT | 2024-04-12 13:42 | 2024-04-12 13:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68784 | 206087 | PM | CMPLT | CMPLT | 2024-04-12 13:44 | 2024-04-12 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69017 | 206088 | PM | CMPLT | CMPLT | 2024-04-12 13:45 | 2024-04-12 13:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1061216 | 206089 | CM | CMPLT | CMPLT | 2024-04-12 13:47 | 2024-04-12 13:48 | Pyxis frozen | (TC 04/12/2024 01:46:40 PM, Reid, Emmerson ) Was frozen, tried calling tech support was put on hold. Eventually was able to reboot and resolve the issue.
69029 | 206090 | PM | CMPLT | CMPLT | 2024-04-12 13:47 | 2024-04-12 13:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68803 | 206091 | PM | CMPLT | CMPLT | 2024-04-12 13:49 | 2024-04-12 13:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68810 | 206092 | PM | CMPLT | CMPLT | 2024-04-12 14:16 | 2024-04-12 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68829 | 206093 | PM | CMPLT | CMPLT | 2024-04-12 14:18 | 2024-04-12 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68774 | 206094 | PM | CMPLT | CMPLT | 2024-04-12 14:20 | 2024-04-12 14:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68773 | 206095 | PM | CMPLT | CMPLT | 2024-04-12 14:21 | 2024-04-12 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68826 | 206096 | PM | CMPLT | CMPLT | 2024-04-12 14:22 | 2024-04-12 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68767 | 206097 | PM | CMPLT | CMPLT | 2024-04-12 14:23 | 2024-04-12 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
| 206098 | CM | CMPLT | CMPLT | 2024-04-12 14:29 | 2024-04-12 14:29 | worked on 14S and NICU equipment | worked on 14S and NICU equipment
68765 | 206099 | PM | CMPLT | CMPLT | 2024-04-12 14:33 | 2024-04-12 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68624 | 206100 | PM | CANCL | DUP | 2024-04-12 14:34 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
541077 | 206101 | CM | CMPLT | CMPLT | 2024-04-15 07:18 | 2024-04-15 07:18 | removed from service
541272 | 206102 | CM | CMPLT | CMPLT | 2024-04-15 07:20 | 2024-04-15 07:20 | not working | (TC 04/08/2024 07:19:52 AM, Huang, Jessica) replaced the SPO2 sensor
541041 | 206103 | CM | CMPLT | CMPLT | 2024-04-15 07:21 | 2024-04-15 07:21 | removed from service
541310 | 206104 | CM | CMPLT | CMPLT | 2024-04-15 07:22 | 2024-04-15 07:22 | not working | (TC 04/15/2024 07:22:28 AM, Huang, Jessica) problem not found working fine (TC 04/08/2024 07:22:28 AM, Huang, Jessica) problem not found working fine
1048640 | 206105 | CM | CMPLT | CMPLT | 2024-04-15 07:24 | 2024-04-15 07:24 | not working | (TC 04/08/2024 07:23:59 AM, Huang, Jessica) problem not found working fine
1048683 | 206106 | CM | CMPLT | CMPLT | 2024-04-15 07:25 | 2024-04-15 07:25 | not working | (TC 04/08/2024 07:24:46 AM, Huang, Jessica) problem not found working fine
1048633 | 206107 | CM | CMPLT | CMPLT | 2024-04-15 07:26 | 2024-04-15 07:26 | not working | (TC 04/08/2024 07:25:53 AM, Huang, Jessica) problem not found working fine
1046652 | 206108 | CM | CMPLT | CMPLT | 2024-04-15 07:31 | 2024-04-15 07:31 | replaced the battery | (TC 04/08/2024 07:30:49 AM, Huang, Jessica) received the email from Lianne saying the device was not working properly. informed her to send the device to CSR the battery was defective and ordered it replaced the battery and charged it informed Natalie and Lianne sent the device back
1055924 | 206109 | CM | CMPLT | CMPLT | 2024-04-15 07:34 | 2024-04-15 07:34 | not working | (TC 04/08/2024 07:35:11 AM, Huang, Jessica) the monitor was sent down from MBU without any accessories and not power cord either the battery was completely drained charged the battery and tested it working fine informed Shermaine and Tanya they agreed to have the monitor back and would install all the accessories by the unit (TC 04/10/2024 07:35:11 AM, Huang, Jessica) the monitor was sent down from MBU without any accessories and not power cord either the battery was completely drained charged the battery and tested it working fine informed Shermaine and Tanya they agreed to have the monitor back and would install all the accessories by the unit
1057283 | 206110 | CM | CMPLT | CMPLT | 2024-04-15 07:39 | 2024-04-15 07:39 | not working | (TC 04/10/2024 07:39:00 AM, Huang, Jessica) the monitor was sent down from MBU without any accessories and not power cord either the battery was completely drained charged the battery and tested it working fine informed Shermaine and Tanya they agreed to have the monitor back and would install all the accessories by the unit
1070586 | 206111 | CM | CMPLT | CMPLT | 2024-04-15 07:44 | 2024-04-15 07:44 | new incoming inspection | (TC 04/09/2024 07:45:43 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
1070587 | 206112 | CM | CMPLT | CMPLT | 2024-04-15 07:50 | 2024-04-15 07:50 | new incoming inspection | (TC 04/09/2024 07:50:28 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
1070588 | 206113 | CM | CMPLT | CMPLT | 2024-04-15 07:51 | 2024-04-15 07:51 | new incoming inspection | (TC 04/09/2024 07:51:10 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
| 206114 | RDS | CMPLT | CMPLT | 2024-04-15 07:52 | 2024-04-15 07:52 | weekly check for 14S and 17M | weekly check for 14S and 17M
1070662 | 206115 | CM | CMPLT | CMPLT | 2024-04-15 07:52 | 2024-04-15 07:52 | new incoming inspection | (TC 04/09/2024 07:52:24 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
1070664 | 206116 | CM | CMPLT | CMPLT | 2024-04-15 07:53 | 2024-04-15 07:53 | new incoming inspection | (TC 04/09/2024 07:53:19 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
1070666 | 206117 | CM | CMPLT | CMPLT | 2024-04-15 07:54 | 2024-04-15 07:54 | new incoming inspection | (TC 04/09/2024 07:54:04 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
1070667 | 206118 | CM | CMPLT | CMPLT | 2024-04-15 07:55 | 2024-04-15 07:55 | new incoming inspection | (TC 04/09/2024 07:54:54 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
| 206119 | RDS | CMPLT | CMPLT | 2024-04-15 08:22 | 2024-04-15 13:11 | Weekly checks in CCU | (TC 04/15/2024 01:09:48 PM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
1065731 | 206120 | PM | CMPLT | CMPLT | 2024-04-15 08:23 | 2024-04-15 08:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065709 | 206121 | PM | CMPLT | CMPLT | 2024-04-15 08:24 | 2024-04-15 08:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065693 | 206122 | PM | CMPLT | CMPLT | 2024-04-15 08:26 | 2024-04-15 08:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1057123 | 206123 | CM | CMPLT | CMPLT | 2024-04-15 08:47 | 2024-04-15 08:47 | missing cord | (TC 04/11/2024 08:47:38 AM, Huang, Jessica) the monitor was sent down from MBU and it was missing the power cord. The battery was completely drained. replaced the power cord and charged it went to MBU and informed Tanya because it was the third monitor missing the power cord in this week
| 206124 | RDS | CMPLT | CMPLT | 2024-04-15 09:10 | 2024-04-15 09:12 | Weekly Checks For PACU | (TC 04/15/2024 09:12:01 AM, Reid, Emmerson ) Rounds complete.
| 206125 | RDS | CMPLT | CMPLT | 2024-04-15 09:49 | 2024-04-15 10:02 | Weekly check for ED
| 206126 | RDS | CMPLT | CMPLT | 2024-04-15 09:50 | 2024-04-15 15:19 | Weekly check for ICU
65936 | 206127 | CM | CMPLT | CMPLT | 2024-04-15 10:33 | 2024-04-15 10:33 | Replaced plungerhead assembly and left handle | (TC 04/15/2024 10:32:59 AM, Baghirzade, Nara) Replaced plungerhead assembly. Replaced left handle.
1059778 | 206128 | FC | CMPLT | CMPLT | 2024-04-15 10:46 | 2024-04-15 10:47 | QC checklist | (TC 04/15/2024 10:46:29 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
67958 | 206129 | CM | CMPLT | CMPLT | 2024-04-15 11:10 | 2024-04-15 11:12 | Broken tilt plate | (TC 04/15/2024 11:10:43 AM, Rhuland, Courtney) Repaired bed tilt plate Replaced missing turn plate with spare Calibrated RH Functional check - OK
1065692 | 206130 | PM | CMPLT | CMPLT | 2024-04-15 11:23 | 2024-04-15 11:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058287 | 206131 | PM | CMPLT | CMPLT | 2024-04-15 11:24 | 2024-04-15 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1051374 | 206132 | PM | CMPLT | CMPLT | 2024-04-15 11:25 | 2024-04-15 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1050253 | 206133 | PM | CMPLT | CMPLT | 2024-04-15 11:28 | 2024-04-15 11:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1050252 | 206134 | PM | CMPLT | CMPLT | 2024-04-15 11:29 | 2024-04-15 11:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68676 | 206135 | PM | CMPLT | CMPLT | 2024-04-15 11:30 | 2024-04-15 11:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68741 | 206136 | PM | CMPLT | CMPLT | 2024-04-15 11:32 | 2024-04-15 11:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68887 | 206137 | PM | CMPLT | CMPLT | 2024-04-15 11:32 | 2024-04-15 11:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68630 | 206138 | PM | CMPLT | CMPLT | 2024-04-15 11:33 | 2024-04-15 11:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68624 | 206139 | PM | CMPLT | CMPLT | 2024-04-15 11:35 | 2024-04-15 11:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69659 | 206140 | PM | CMPLT | CMPLT | 2024-04-15 11:38 | 2024-04-15 11:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69466 | 206141 | PM | CMPLT | CMPLT | 2024-04-15 11:40 | 2024-04-15 11:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69478 | 206142 | PM | CMPLT | CMPLT | 2024-04-15 11:41 | 2024-04-15 11:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68633 | 206143 | PM | CMPLT | CMPLT | 2024-04-15 11:43 | 2024-04-15 11:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69093 | 206144 | PM | CMPLT | CMPLT | 2024-04-15 11:44 | 2024-04-15 11:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69611 | 206145 | PM | CMPLT | CMPLT | 2024-04-15 11:45 | 2024-04-15 11:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69574 | 206146 | PM | CMPLT | CMPLT | 2024-04-15 11:46 | 2024-04-15 11:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69174 | 206147 | PM | CMPLT | CMPLT | 2024-04-15 11:48 | 2024-04-15 11:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69168 | 206148 | PM | CMPLT | CMPLT | 2024-04-15 11:49 | 2024-04-15 11:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69615 | 206149 | PM | CMPLT | CMPLT | 2024-04-15 11:50 | 2024-04-15 11:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69303 | 206150 | PM | CMPLT | CMPLT | 2024-04-15 11:51 | 2024-04-15 11:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69424 | 206151 | PM | CMPLT | CMPLT | 2024-04-15 11:52 | 2024-04-15 11:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69310 | 206152 | PM | CMPLT | CMPLT | 2024-04-15 11:52 | 2024-04-15 11:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69338 | 206153 | PM | CMPLT | CMPLT | 2024-04-15 11:53 | 2024-04-15 11:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69218 | 206154 | PM | CMPLT | CMPLT | 2024-04-15 11:55 | 2024-04-15 11:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1054965 | 206155 | CM | ACTIV | ISSUE | 2024-04-15 12:16 |  | not working | (TC 04/12/2024 12:16:38 PM, Huang, Jessica) received the call from MBU and went to check the monitor informed Tanya that the monitor would need to be sent out for repair conntacted Masimo for the RMA
69216 | 206156 | PM | CMPLT | CMPLT | 2024-04-15 14:11 | 2024-04-15 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69214 | 206157 | PM | CMPLT | CMPLT | 2024-04-15 14:12 | 2024-04-15 14:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69633 | 206158 | PM | CMPLT | CMPLT | 2024-04-15 14:14 | 2024-04-15 14:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69220 | 206159 | PM | CMPLT | CMPLT | 2024-04-15 14:16 | 2024-04-15 14:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
501068 | 206160 | CM | CMPLT | CMPLT | 2024-04-15 14:16 | 2024-04-15 14:17 | PM followup (brakes/power cord) | (TC 04/15/2024 02:16:25 PM, Fan, Tak) Replaced brakes and power cord.
69544 | 206161 | PM | CMPLT | CMPLT | 2024-04-15 14:17 | 2024-04-15 14:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69085 | 206162 | PM | CMPLT | CMPLT | 2024-04-15 14:18 | 2024-04-15 14:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
| 206163 | CON | CMPLT | CMPLT | 2024-04-15 14:18 | 2024-04-15 14:18 | GE lunch demo tomorrow | (TC 04/15/2024 02:17:54 PM, Fan, Tak) Dianne requesting video cable for GE demo. Dropped off cable. (TC 04/16/2024 11:52:00 AM, Fan, Tak) Assisted setup in 19F - 317 conference room.
69052 | 206164 | PM | CMPLT | CMPLT | 2024-04-15 14:19 | 2024-04-15 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69679 | 206165 | PM | CMPLT | CMPLT | 2024-04-15 14:19 | 2024-04-15 14:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69041 | 206166 | PM | CMPLT | CMPLT | 2024-04-15 14:20 | 2024-04-15 14:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69436 | 206167 | PM | CMPLT | CMPLT | 2024-04-15 14:21 | 2024-04-15 14:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69263 | 206168 | PM | CMPLT | CMPLT | 2024-04-15 14:23 | 2024-04-15 14:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69350 | 206169 | PM | CMPLT | CMPLT | 2024-04-15 14:24 | 2024-04-15 14:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69662 | 206170 | PM | CMPLT | CMPLT | 2024-04-15 14:25 | 2024-04-15 14:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69623 | 206171 | CM | CMPLT | CMPLT | 2024-04-15 14:26 | 2024-04-15 14:35 | Bottle side sensor defective | (TC 04/15/2024 02:34:25 PM, Baghirzade, Nara) Downloaded error logs; noted errors pertaining to bottle side pressure sensor sensor. Ran analog sensor test on ASM; bottle side pressure sensor remains at 4.9V and does not drop to 0. Replaced bottle side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
69287 | 206172 | PM | CMPLT | CMPLT | 2024-04-15 14:26 | 2024-04-15 14:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1047041 | 206173 | PM | CMPLT | CMPLT | 2024-04-15 14:28 | 2024-04-15 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1050265 | 206174 | PM | CMPLT | CMPLT | 2024-04-15 14:29 | 2024-04-15 14:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1050273 | 206175 | PM | CMPLT | CMPLT | 2024-04-15 14:30 | 2024-04-15 14:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1057390 | 206176 | PM | CMPLT | CMPLT | 2024-04-15 14:32 | 2024-04-15 14:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065776 | 206177 | PM | CMPLT | CMPLT | 2024-04-15 14:33 | 2024-04-15 14:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065787 | 206178 | PM | CMPLT | CMPLT | 2024-04-15 14:34 | 2024-04-15 14:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065804 | 206179 | PM | CMPLT | CMPLT | 2024-04-15 14:35 | 2024-04-15 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065817 | 206180 | PM | CMPLT | CMPLT | 2024-04-15 14:36 | 2024-04-15 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69375 | 206181 | CM | CMPLT | CMPLT | 2024-04-15 14:37 | 2024-04-16 13:46 | Door does not close properly | (TC 04/15/2024 01:45:43 PM, Baghirzade, Nara) Tightened door latch screw. Check-in tests passed and completed. IUI connectors and device cleaned.
1065819 | 206182 | PM | CMPLT | CMPLT | 2024-04-15 14:37 | 2024-04-15 14:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065822 | 206183 | PM | CMPLT | CMPLT | 2024-04-15 14:39 | 2024-04-15 14:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065834 | 206184 | PM | CMPLT | CMPLT | 2024-04-15 14:40 | 2024-04-15 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
| 206185 | CM | CMPLT | CMPLT | 2024-04-15 14:40 | 2024-04-15 14:40 | daily checks | daily checks
| 206186 | CM | CMPLT | CMPLT | 2024-04-15 14:41 | 2024-04-15 14:41 | meeting | meeting
1065931 | 206187 | PM | CMPLT | CMPLT | 2024-04-15 14:41 | 2024-04-15 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
66739 | 206188 | CM | CMPLT | CMPLT | 2024-04-15 14:48 | 2024-04-15 14:49 | Printer will not turn off | (TC 04/15/2024 02:48:27 PM, Rhuland, Courtney) Went over and removed battery Unit power cycled and is working properly
1065932 | 206189 | PM | CMPLT | CMPLT | 2024-04-15 14:49 | 2024-04-15 14:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065941 | 206190 | PM | CMPLT | CMPLT | 2024-04-15 14:50 | 2024-04-15 14:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
| 206191 | CM | CMPLT | CMPLT | 2024-04-15 14:52 | 2024-04-16 12:27 | Stethoscope battery cables broken | (TC 04/15/2024 02:51:54 PM, Rhuland, Courtney) Began soldering new battery connectors to boards inside stethoscopes (TC 04/16/2024 12:26:24 PM, Rhuland, Courtney) Repaired 3 stethoscopes Soldered battery cables and tested Returned to clinic
67932 | 206192 | CM | CMPLT | CMPLT | 2024-04-15 14:54 | 2024-04-15 14:54 | Broken porthole door | (TC 04/15/2024 02:53:57 PM, Rhuland, Courtney) Repaired 4 porthole doors Calibrated RH Functional check - OK
1054538 | 206193 | CM | CMPLT | CMPLT | 2024-04-15 14:56 | 2024-04-16 13:24 | Keyboard cover ripped | (TC 04/15/2024 03:15:24 PM, Canjimavadivel, Kamallesh ) Replaced keyboard cover Replaced LL, LA, RA leads (TC 04/16/2024 01:20:28 PM, Canjimavadivel, Kamallesh ) Replaced V7,V8 leads. Tested okay with simulator
1067334 | 206194 | CM | CMPLT | CMPLT | 2024-04-15 15:19 | 2024-04-15 15:20 | Not zeroing | (TC 04/15/2024 03:19:16 PM, Canjimavadivel, Kamallesh ) Checked scale and it was zeroing. Verified scale was also balanced.
32540 | 206195 | CM | CMPLT | CMPLT | 2024-04-15 17:47 | 2024-02-23 17:49 | Oil Flow error | Oil flow error Troubleshot the issue and confirmed the issue was due to a faulty Heat Exchanger Order and replaced Heat exchanger. Upon test scan sytem wouldn't prep -Further troubleshooting done and found 2 MHR BP boards to be faulty -Replaced boards and system passed all test scans --Both scan and control rooms were above normal operating temperatures. CT was not being used the whole day for that reason while plant maintenance was working on the AC
69383 | 206196 | CM | CMPLT | CMPLT | 2024-04-16 07:22 | 2024-04-16 10:39 | No error stated; functional check | (TC 04/16/2024 10:38:31 AM, Baghirzade, Nara) Downloaded error log; only found software compatibility errors by connecting the LVP to a PCU with a different software version. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
1066040 | 206197 | PM | CMPLT | CMPLT | 2024-04-16 08:36 | 2024-04-16 08:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066022 | 206198 | PM | CMPLT | CMPLT | 2024-04-16 08:37 | 2024-04-16 08:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
63410 | 206199 | CM | CMPLT | CMPLT | 2024-04-16 08:49 | 2024-04-16 08:50 | Removed from service | (TC 04/15/2024 08:49:04 AM, Rendon Lopez , Elizabeth) Clinical team confirmed device is not used in L&D, removed from service. Retired
63412 | 206200 | CM | CMPLT | CMPLT | 2024-04-16 08:51 | 2024-04-16 08:52 | Removed from service | (TC 04/15/2024 08:51:15 AM, Rendon Lopez , Elizabeth) Clinical team confirmed device is not used in L&D, removed from service. Retired
| 206201 | CM | CMPLT | CMPLT | 2024-04-16 10:03 | 2024-05-01 16:54 | Order for spare Servo-U part - NVRAM memory batteries | (TC 04/30/2024 04:53:15 PM, Cortez, John) Parts arrived
1061204 | 206202 | CM | CMPLT | CMPLT | 2024-04-16 10:05 | 2024-04-16 10:20 | Boom DVI Connectors Damaged | (TC 04/16/2024 10:16:18 AM, Reid, Emmerson ) Changed both DVI connectors on the foot boom. Also straightened the pins on the DVI cable for the surgical nav
69273 | 206203 | CM | CMPLT | CMPLT | 2024-04-16 10:40 | 2024-04-16 10:41 | No error stated on yellow biomed tag; functional check | (TC 04/16/2024 10:40:16 AM, Baghirzade, Nara) Downloaded error log; only found software compatibility errors by connecting the LVP to a PCU with a different software version. During check-in tests, the rate accuracy failed (reading too high). Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
69499 | 206204 | CM | CMPLT | CMPLT | 2024-04-16 10:41 | 2024-04-16 10:41 | No error stated on yellow biomed tag; functional check | (TC 04/16/2024 10:41:09 AM, Baghirzade, Nara) Downloaded error log; only found software compatibility errors by connecting the LVP to a PCU with a different software version. During check-in tests, the rate accuracy failed (reading too high). Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
68950 | 206205 | CM | CMPLT | CMPLT | 2024-04-16 10:42 | 2024-04-16 10:43 | No error stated on biomed yellow tag; functional check | (TC 04/16/2024 10:41:55 AM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 180 mohms Leakage current: 71.8 uA
| 206206 | CM | CMPLT | CMPLT | 2024-04-16 10:43 | 2024-04-16 10:44 | PM huddle | (TC 04/16/2024 10:43:19 AM, Baghirzade, Nara) PM huddle biomed
| 206207 | PA | CMPLT | CMPLT | 2024-04-16 10:44 | 2024-04-16 10:50 | Ordering sterile water for NICU altrix cleaning | (TC 04/16/2024 10:49:36 AM, Baghirzade, Nara) SC 1003210983 (TC 04/17/2024 11:16:54 AM, Baghirzade, Nara) Order # 4150275403
1070737 | 206208 | ACC | CMPLT | CMPLT | 2024-04-16 11:21 | 2024-04-16 11:23 | Incoming inspection | (TC 04/16/2024 11:21:34 AM, Baghirzade, Nara) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1061333 | 206209 | FC | CMPLT | CMPLT | 2024-04-16 11:24 | 2024-04-16 11:24 | QC Checklist | (TC 04/16/2024 11:23:55 AM, Baghirzade, Nara) QC checklist completed. Returned to 11North
1070738 | 206210 | CM | CMPLT | CMPLT | 2024-04-16 11:38 | 2024-04-16 13:16 | Incoming inspection | (TC 04/16/2024 01:16:16 PM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Electrical safety - OK Incubator is functional.
1070739 | 206211 | CM | CMPLT | CMPLT | 2024-04-16 11:46 | 2024-04-16 11:51 | Incoming inspection | (TC 04/16/2024 11:50:18 AM, Baghirzade, Nara) Incoming inspection completed. Adjusted light height and removed adjustable knob. Removed aux outlets cover at the rear of the device. Incubator is fully functional.
69450 | 206212 | CM | CMPLT | CMPLT | 2024-04-16 13:01 | 2024-04-16 13:02 | Defective bottle side pressure sensor | (TC 04/16/2024 01:01:03 PM, Baghirzade, Nara) Ran analog test on ASM; bottle side pressure sensor was staying high and not dropping down to 0V. Replaced bottle side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
1050274 | 206213 | CM | CMPLT | CMPLT | 2024-04-16 13:02 | 2024-04-16 13:05 | O-ring on bezel is broken and both pressure sensors defective | (TC 04/16/2024 01:02:20 PM, Baghirzade, Nara) Immediate error when connecting LVP to PCU; ran analog sensor test and both pressure sensors are defective. Bottle-side pressure sensor stays high (at 5V) and patient side pressure sensor stays at 0V with pressure applied. Replaced bezel. Replaced both pressure sensors.
1070740 | 206214 | CM | CMPLT | CMPLT | 2024-04-16 13:03 | 2024-04-16 13:16 | Incoming inspection | (TC 04/16/2024 01:15:36 PM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Electrical safety - OK Incubator is functional.
1050274 | 206215 | PM | CMPLT | CMPLT | 2024-04-16 13:07 | 2024-04-16 13:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 01:07:25 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
63531 | 206216 | CM | CMPLT | CMPLT | 2024-04-16 13:19 | 2024-04-16 13:19 | Does not switch on | (TC 04/16/2024 01:18:39 PM, Baghirzade, Nara) Replaced battery and probe, did not switch on. Will retire and repurpose a different thermometer in its place.
1051698 | 206217 | CM | CMPLT | CMPLT | 2024-04-16 13:23 | 2024-04-16 13:26 | Repurposed thermometer for NICU | (TC 04/16/2024 01:25:11 PM, Baghirzade, Nara) Repurposed for NICU. Updated location on TMS and changed status to active. Printed new asset sticker. Functional tests passed.
1067816 | 206218 | CM | CMPLT | NFF | 2024-04-16 13:24 | 2024-04-16 13:25 | No issue listed | (TC 04/16/2024 01:23:57 PM, Rhuland, Courtney) Functional check - OK No fault found Returned to NICU
1070741 | 206219 | CM | CMPLT | CMPLT | 2024-04-16 13:29 | 2024-04-16 13:31 | E03 error | (TC 04/16/2024 01:29:01 PM, Rhuland, Courtney) Functional check - OK Not in TMS - created asset Could not replicate error
1059496 | 206220 | CM | CMPLT | CMPLT | 2024-04-16 13:40 | 2024-04-16 13:40 | Reported not working | (TC 04/16/2024 01:39:52 PM, Baghirzade, Nara) Cleaned the probe connector. Functional test passed. Thermometer is functional.
1070742 | 206221 | ACC | CMPLT | CMPLT | 2024-04-16 14:30 | 2024-04-16 14:52 | Incoming inspection | (TC 04/16/2024 02:51:06 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1070743 | 206222 | ACC | CMPLT | CMPLT | 2024-04-16 14:35 | 2024-04-16 14:56 | Incoming inspection | (TC 04/16/2024 02:55:35 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1070744 | 206223 | ACC | CMPLT | CMPLT | 2024-04-16 14:37 | 2024-04-16 14:55 | Incoming inspection | (TC 04/16/2024 02:54:27 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
501070 | 206224 | CM | CMPLT | CMPLT | 2024-04-16 15:53 | 2024-05-08 14:56 | PM followup | (TC 04/16/2024 12:53:00 PM, Fan, Tak) Control panel gas spring assembly replacement. Monitor gas spring replacement.
501063 | 206225 | CM | CMPLT | CMPLT | 2024-04-16 15:55 | 2024-04-16 15:56 | PM followup | (TC 04/16/2024 08:55:00 AM, Fan, Tak) Gas spring replacement on monitor.
1063725 | 206226 | CM | CMPLT | CMPLT | 2024-04-16 18:47 | 2024-04-16 18:48 | MSH IR CRYOABLATION SYSTEM PM BY HUNTER FROM BOSTON SCIENTIFIC , USING MSH SHOP | (TC 04/16/2024 06:47:31 PM, Whittington, Scott) MSH IR CRYOABLATION SYSTEM PM BY HUNTER FROM BOSTON SCIENTIFIC , USING MSH SHOP | MSH IR CRYOABLATION SYSTEM PM BY HUNTER FROM BOSTON SCIENTIFIC , USING MSH SHOP
1059547 | 206227 | CM | CMPLT | CMPLT | 2024-04-17 07:39 | 2024-04-17 07:40 | Was shut down | (TC 04/17/2024 07:39:13 AM, Rhuland, Courtney) Received message that 11O was not reporting Went up and central had been turned off Rebooted device and verified connectivity
66055 | 206228 | PM | CMPLT | CMPLT | 2024-04-17 08:41 | 2024-04-17 08:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69012 | 206229 | PM | CMPLT | CMPLT | 2024-04-17 08:42 | 2024-04-17 08:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68957 | 206230 | PM | CMPLT | CMPLT | 2024-04-17 08:43 | 2024-04-17 08:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68778 | 206231 | PM | CMPLT | CMPLT | 2024-04-17 08:44 | 2024-04-17 08:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68621 | 206232 | PM | CMPLT | CMPLT | 2024-04-17 08:45 | 2024-04-17 08:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68843 | 206233 | PM | CMPLT | CMPLT | 2024-04-17 08:47 | 2024-04-17 08:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68723 | 206234 | PM | CMPLT | CMPLT | 2024-04-17 08:48 | 2024-04-17 08:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68994 | 206235 | PM | CMPLT | CMPLT | 2024-04-17 08:49 | 2024-04-17 08:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1051378 | 206236 | PM | CMPLT | CMPLT | 2024-04-17 08:50 | 2024-04-17 08:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055107 | 206237 | PM | CMPLT | CMPLT | 2024-04-17 08:51 | 2024-04-17 08:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065703 | 206238 | PM | CMPLT | CMPLT | 2024-04-17 08:52 | 2024-04-17 08:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1051030 | 206239 | PM | CMPLT | CMPLT | 2024-04-17 08:54 | 2024-04-17 08:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 08:54:23 AM, Rhuland, Courtney) PM completed by BD
68302 | 206240 | PM | CMPLT | CMPLT | 2024-04-17 08:59 | 2024-04-17 08:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1063129 | 206241 | CM | CMPLT | CMPLT | 2024-04-17 09:52 | 2024-04-17 10:06 | Not Sound, not picking up FHR or movement | (TC 04/17/2024 10:04:45 AM, Rendon Lopez , Elizabeth) Damaged probe head. Replaced probe with spare, new batteries. Tested-OK
69621 | 206242 | PM | CMPLT | CMPLT | 2024-04-17 10:00 | 2024-04-17 10:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69535 | 206243 | PM | CMPLT | CMPLT | 2024-04-17 10:02 | 2024-04-17 10:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1056537 | 206244 | CM | CMPLT | CMPLT | 2024-04-17 10:09 | 2024-04-17 10:10 | Battery replacement | (TC 04/17/2024 10:09:21 AM, Canjimavadivel, Kamallesh ) Replaced battery
69442 | 206245 | PM | CMPLT | CMPLT | 2024-04-17 10:15 | 2024-04-17 10:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065863 | 206246 | PM | CMPLT | CMPLT | 2024-04-17 10:17 | 2024-04-17 10:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065942 | 206247 | PM | CMPLT | CMPLT | 2024-04-17 10:18 | 2024-04-17 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065947 | 206248 | PM | CMPLT | CMPLT | 2024-04-17 10:19 | 2024-04-17 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066003 | 206249 | PM | CMPLT | CMPLT | 2024-04-17 10:19 | 2024-04-17 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
65979 | 206250 | PM | CMPLT | CMPLT | 2024-04-17 10:23 | 2024-04-17 10:24 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
65982 | 206251 | PM | CMPLT | CMPLT | 2024-04-17 10:24 | 2024-04-17 10:25 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
65967 | 206252 | PM | CMPLT | CMPLT | 2024-04-17 10:26 | 2024-04-17 10:26 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
1048530 | 206253 | PM | CMPLT | CMPLT | 2024-04-17 10:27 | 2024-04-17 10:27 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
66060 | 206254 | PM | CMPLT | CMPLT | 2024-04-17 10:46 | 2024-04-17 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68746 | 206255 | PM | CMPLT | CMPLT | 2024-04-17 10:48 | 2024-04-17 10:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68998 | 206256 | PM | CMPLT | CMPLT | 2024-04-17 10:49 | 2024-04-17 10:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065696 | 206257 | PM | CMPLT | CMPLT | 2024-04-17 10:51 | 2024-04-17 10:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69234 | 206258 | PM | CMPLT | CMPLT | 2024-04-17 10:52 | 2024-04-17 10:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
65990 | 206259 | PM | CMPLT | CMPLT | 2024-04-17 10:53 | 2024-04-17 10:54 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
1048524 | 206260 | PM | CMPLT | CMPLT | 2024-04-17 10:54 | 2024-04-17 10:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD
| 206261 | CM | CMPLT | CMPLT | 2024-04-17 11:14 | 2024-04-17 11:14 | Capital meeting | (TC 04/17/2024 11:14:09 AM, Baghirzade, Nara) Capital meeting
1061554 | 206262 | CM | CMPLT | CMPLT | 2024-04-17 11:24 | 2024-04-22 08:31 | Dirty Filter | (TC 04/17/2024 12:53:25 PM, Reid, Emmerson ) Brought the unit down to biomed, took pictures of the filter and submitted safer ID 72937 (TC 04/22/2024 08:30:38 AM, Reid, Emmerson ) Neptune investigation meeting with Andrew and with the vendor
| 206263 | CM | CMPLT | CMPLT | 2024-04-17 11:33 | 2024-05-01 08:07 | MSH MIE SPARE TO PMH 2 SOUTH UCC | (TC 04/17/2024 03:13:14 PM, Whittington, Scott) MSH MIE SPARE TO PMH 2 SOUTH UCC (TC 05/01/2024 08:06:58 AM, Whittington, Scott) MSH MIE SPARE TO PMH 2 SOUTH UCC RETURNED TO MSH MIE APRIL 18 WHEN LOANER ARRIVED FROM BOTHELL | MSH MIE SPARE TO PMH 2 SOUTH UCC
1061333 | 206264 | PA | CMPLT | CMPLT | 2024-04-17 13:04 | 2024-04-25 07:53 | Mounting mechanism broken | (TC 04/17/2024 01:04:16 PM, Baghirzade, Nara) Removed mounting mechanism entirely; plastic ring inside the mechanism is broken. Contacted Ifrah from 3M for parts quote. Updated Sarabeth. (TC 04/19/2024 06:57:23 AM, Baghirzade, Nara) Received parts quote and sent it to Sharmaine for approval. (TC 04/25/2024 07:37:03 AM, Baghirzade, Nara) PO 7500205435
1053391 | 206265 | CM | CMPLT | CMPLT | 2024-04-17 13:33 | 2024-04-17 13:34 | Before/after patient use check | (TC 04/17/2024 01:33:07 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 49 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
65943 | 206266 | PM | CMPLT | CMPLT | 2024-04-17 13:41 | 2024-04-17 13:42 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/17/2024 01:41:21 PM, Baghirzade, Nara) Replaced rear case. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1070748 | 206267 | CM | CMPLT | CMPLT | 2024-04-17 13:45 | 2024-04-23 13:20 | Incoming inspection | (TC 04/23/2024 01:20:12 PM, Rhuland, Courtney) Check-in completed Asset and barcode tagged Refer to ASM for details
1070749 | 206268 | CM | CMPLT | CMPLT | 2024-04-17 13:47 | 2024-04-25 14:52 | Incoming inspection | (TC 04/25/2024 02:52:13 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
1070750 | 206269 | CM | CMPLT | CMPLT | 2024-04-17 13:47 | 2024-04-23 13:19 | Incoming inspection | (TC 04/23/2024 01:17:47 PM, Rhuland, Courtney) Check-in completed Asset and barcode tagged Refer to ASM for details
1070751 | 206270 | CM | CMPLT | CMPLT | 2024-04-17 13:48 | 2024-04-25 14:30 | Incoming inspection | (TC 04/25/2024 02:30:11 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
1070752 | 206271 | CM | CMPLT | CMPLT | 2024-04-17 13:49 | 2024-04-25 14:14 | Incoming inspection | (TC 04/25/2024 02:12:57 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
1070753 | 206272 | CM | CMPLT | CMPLT | 2024-04-17 13:50 | 2024-04-25 14:30 | Incoming inspection | (TC 04/25/2024 02:29:42 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
1070754 | 206273 | CM | CMPLT | CMPLT | 2024-04-17 13:50 | 2024-04-25 14:40 | Incoming inspection | (TC 04/25/2024 02:39:27 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
1070755 | 206274 | CM | CMPLT | CMPLT | 2024-04-17 13:51 | 2024-04-25 14:40 | Incoming inspection | (TC 04/25/2024 02:39:50 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
1070756 | 206275 | CM | CMPLT | CMPLT | 2024-04-17 13:52 | 2024-04-25 14:53 | Incoming inspection | (TC 04/25/2024 02:52:57 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
1070757 | 206276 | CM | CMPLT | CMPLT | 2024-04-17 13:52 | 2024-04-25 14:14 | Incoming inspection | (TC 04/25/2024 02:13:50 PM, Rhuland, Courtney) Check-in complete All tests passed Refer to ASM for details
| 206277 | CM | CMPLT | CMPLT | 2024-04-17 14:40 | 2024-04-25 13:36 | Setting up new pumps for NICU | (TC 04/17/2024 02:40:25 PM, Rhuland, Courtney) Created asset tags and barcodes for syringe pumps Tested barcodes and inputted into spreadsheet (TC 04/25/2024 01:35:28 PM, Rhuland, Courtney) Asset tagged and barcode tagged all pumps Verified barcodes were correct
32510 | 206278 | CM | CMPLT | CMPLT | 2024-04-17 15:09 | 2024-04-18 14:15 | MSH DR1 OGP ISSUES | (TC 04/17/2024 08:09:00 PM, Whittington, Scott) MSH DR1 OGP ISSUES | MSH DR1 OGP ISSUES
1069509 | 206279 | CM | ACTIV | ISSUE | 2024-04-17 15:34 |  | MSH 5 PROBE PLI-1205BX KOROSH KHALILI REPORTED ARTIFACTS, PREVIOUSLY REPORTED IN MARCH, BUT PUT DOWN TO CONNECTION ISSUE |  | MSH 5 PROBE PLI-1205BX KOROSH KHALILI REPORTED ARTIFACTS, PREVIOUSLY REPORTED IN MARCH, BUT PUT DOWN TO CONNECTION ISSUE
1069842 | 206280 | CM | CMPLT | CMPLT | 2024-04-17 15:41 | 2024-04-23 15:54 | Preventive Maintenance - Self Learn | Apr 17 - Brought the unit down to learn PM. Partially attended youtube training for learning. Found unit had SW v4 and SM was available in SW 2.1 due to which there were some defferences in procedures and measurements. Contacted Medtronic for updated SM. Received. Completed partial PM. Apr 22 - Completed rest of the tests. Tested ok. Returned to OR.
1061202 | 206281 | CM | CMPLT | CMPLT | 2024-04-17 15:48 | 2024-04-17 15:49 | Computer Issue | Call re: monitor not working. Went to check. Fund Nursing desk computer was turned off and staff did not realize it. Turned the cimputer on. Tested ok.
1070759 | 206282 | CM | CMPLT | CMPLT | 2024-04-18 10:26 | 2024-04-18 11:06 | Incoming inspection | (TC 04/18/2024 11:06:13 AM, Baghirzade, Nara) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1067164 | 206283 | CM | CMPLT | CMPLT | 2024-04-18 10:27 | 2024-06-19 13:31 | Tester calibration | (TC 04/18/2024 01:41:29 PM, Rendon Lopez , Elizabeth) Requested quote, created e-req (TC 06/19/2024 01:30:56 PM, Rendon Lopez , Elizabeth) Received device back. See attached certificate
1070760 | 206284 | CM | CMPLT | CMPLT | 2024-04-18 10:29 | 2024-04-18 11:06 | Incoming inspection | (TC 04/18/2024 11:05:48 AM, Baghirzade, Nara) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070761 | 206285 | CM | CMPLT | CMPLT | 2024-04-18 10:30 | 2024-04-18 11:06 | Incoming inspection | (TC 04/18/2024 11:05:29 AM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070763 | 206286 | CM | CMPLT | CMPLT | 2024-04-18 10:31 | 2024-04-18 11:05 | Incoming inspection | (TC 04/18/2024 11:05:09 AM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
| 206287 | CM | CMPLT | CMPLT | 2024-04-18 11:33 | 2024-04-18 11:33 | Ophthalmology laser meeting | (TC 04/18/2024 11:32:51 AM, Baghirzade, Nara) Ophthalmology laser meeting with Andrew
1047011 | 206288 | CM | CMPLT | CMPLT | 2024-04-18 12:44 | 2024-04-18 12:45 | Defective bottle side pressure sensor | (TC 04/18/2024 12:43:55 PM, Baghirzade, Nara) Downloaded error log; found fatal errors for the pressure sensor. Ran analog test and bottle side pressure sensor stays high at 4.9V. Replaced pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
68635 | 206289 | CM | CMPLT | CMPLT | 2024-04-18 12:46 | 2024-04-18 12:47 | No error stated on yellow biomed tag | (TC 04/18/2024 12:45:59 PM, Baghirzade, Nara) Downloaded error log; found fatal errors for the pressure sensor. Ran analog test and bottle side pressure sensor stays high at 4.9V. Replaced pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
68718 | 206290 | CM | CMPLT | CMPLT | 2024-04-18 12:48 | 2024-04-18 12:51 | No error listed on yellow biomed tag | (TC 04/18/2024 12:49:38 PM, Baghirzade, Nara) Downloaded error logs; did not find any errors. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 106 mohms Leakage current: 74.6 uA
68677 | 206291 | CM | CMPLT | CMPLT | 2024-04-18 12:52 | 2024-04-18 12:52 | Reported "keeps beeping' | (TC 04/18/2024 12:51:40 PM, Baghirzade, Nara) Downloaded error log; found log only communication errors. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 47 mohms Leakage current: 71.0 uA Switched the PCU in normal mode, connected LVPs; no issues or beeping.
66487 | 206292 | CM | CMPLT | CMPLT | 2024-04-18 12:53 | 2024-04-18 12:55 | Light not working | (TC 04/18/2024 12:53:03 PM, Baghirzade, Nara) Connector was on backwards; put heat shrink over the connector to prevent user from unplugging it from the connector end. Left it to charge for the entire day; light intensity is good. Light if functional. Unplugged the power source and left the light on; light started to lose intensity and dim at the 1 hour mark. Let Samir know and he confirmed that its functional. Recharged the light and returned to NICU.
65879 | 206293 | PM | CMPLT | CMPLT | 2024-04-18 12:56 | 2024-04-18 12:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 04/18/2024 12:56:00 PM, Baghirzade, Nara) Replaced: -front case -rear case -mylar gasket -left handle Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1063105 | 206294 | CM | CMPLT | CMPLT | 2024-04-18 13:35 | 2024-04-18 13:36 | Client overview down | (TC 04/18/2024 01:35:05 PM, Rhuland, Courtney) Client overview down message from biomed overview Went up and monitor had been turned off Turned monitor on and waited for it to boot up Verified connectivity and message was gone from biomed
1066987 | 206295 | PM | CMPLT | CMPLT | 2024-04-18 13:36 | 2024-05-15 10:54 | PM-MAJOR-Annual PM for: Bunnell HFV Tester | (TC 04/18/2024 01:39:41 PM, Rendon Lopez , Elizabeth) Requested qupte for yearly calibration, created req (TC 05/15/2024 10:47:21 AM, Rendon Lopez , Elizabeth) Received device back. See attached report. Cost paid by CC.
67945 | 206296 | CM | CMPLT | CMPLT | 2024-04-18 13:43 | 2024-04-18 13:48 | Probe failure | (TC 04/18/2024 01:46:25 PM, Rhuland, Courtney) Tag stated "probe failure" Ran incubator for 5 hours and checked error logs No error found and could not replicate error Assume it was patient probe error Functional - OK
66918 | 206297 | CM | CMPLT | CMPLT | 2024-04-18 13:45 | 2024-04-29 13:43 | Functional test | (TC 04/29/2024 01:42:27 PM, Baghirzade, Nara) Left radical to charge however it would only charge intermittently. Replaced the battery and completed battery upgrade. Left to charge overnight and it was charging ok. SpO2 functional check passed.
63507 | 206298 | CM | CMPLT | CMPLT | 2024-04-18 13:53 | 2024-04-19 06:57 | 50 FTSW Failure error message | (TC 04/18/2024 03:00:00 PM, Baghirzade, Nara) 50 FTSW Failure appears when laser switches on and does not allow the laser to operate on "ready" mode. Users power cycled the laser at least 3 times Called Cook Medical and tech helped troubleshoot that the footswitch is faulty. Rutvik and I fiddled with the cable until we found the point where the connection is weak and causes the error to display on the laser. We secured the footswitch cable to the back of the device with zipties and cable management tools, and had spoken with the surgeon who operates the laser to fiddle with the cable until the error goes away and then not to move the footswitch once he has placed it down on the floor. This is a temporary solution until the new laser arrives on site. Cannot order any parts as this product is out of support. Maureen confirmed that there are no spare footswitches for this laser.
1061219 | 206299 | CM | CMPLT | CMPLT | 2024-04-18 14:01 | 2024-04-18 14:01 | VMC Issue | Call re: Image not showing up on the monitor. Went to check. Waited for access Foudn cable connection loose. Fixed it. Tested ok.
1064352 | 206300 | CM | CMPLT | CMPLT | 2024-04-18 15:42 | 2024-04-18 15:43 | MSH ICU POD A SPARQ NOT BOOTING ; | (TC 04/18/2024 03:42:54 PM, Whittington, Scott) MSH ICU POD A SPARQ NOT BOOTING ; HARD REBOOT SYSTEM BACK UP STAFF PROBABLY REQUIRE INSTRUCTION ON PROPER SLEEP/WAKEUP PROTOCOL | MSH ICU POD A SPARQ NOT BOOTING ; HARD REBOOT SYSTEM BACK UP STAFF PROBABLY REQUIRE INSTRUCTION ON PROPER SLEEP/WAKEUP PROTOCOL
501173 | 206301 | CM | CMPLT | CMPLT | 2024-04-18 15:53 | 2024-04-19 11:08 | No touch screen display | (TC 04/18/2024 03:52:51 PM, Fan, Tak) Touch screen controls work but no display. Checked video card connections. Not the double video card setup so cannot swap to test. Probably bad display. Placed call to Philips. (TC 04/19/2024 11:08:14 AM, Fan, Tak) Swapped touch screen with spare from inactive system, screen works. Jillian will order parts.
1061333 | 206302 | PA | CMPLT | CMPLT | 2024-04-19 08:21 | 2024-05-02 10:32 | Ordering parts for repair | (TC 04/19/2024 08:24:37 AM, Baghirzade, Nara) SC 1003213490 (TC 05/02/2024 10:31:32 AM, Baghirzade, Nara) PO 7500205435
67242 | 206303 | CM | CMPLT | CMPLT | 2024-04-19 08:28 | 2024-10-02 07:58 | Sending pump for exchange/repair | (TC 04/19/2024 08:34:19 AM, Baghirzade, Nara) SC 1003213543 (TC 04/26/2024 10:54:15 AM, Baghirzade, Nara) Received work order for exchange from Cardinal Health. KJV # 430128 (TC 09/10/2024 10:25:19 AM, Baghirzade, Nara) Sent follow up email regarding status of repair or if a replacement unit will be sent. (TC 10/02/2024 07:58:14 AM, Baghirzade, Nara) Received replacement pump. PO 7500205438
1054696 | 206304 | CM | CMPLT | CMPLT | 2024-04-19 08:36 | 2024-10-02 07:59 | Sending pump for repair/exchange | (TC 04/19/2024 08:39:28 AM, Baghirzade, Nara) SC 1003213562 (TC 04/26/2024 10:55:53 AM, Baghirzade, Nara) Received work order for exchange from Cardinal Health. KJV # 430130 (TC 09/10/2024 10:25:58 AM, Baghirzade, Nara) Sent follow up email regarding status of repair or if a replacement unit will be sent. (TC 10/02/2024 07:58:55 AM, Baghirzade, Nara) Received replacement pump. PO 7500205438
1054693 | 206305 | CM | CMPLT | CMPLT | 2024-04-19 08:41 | 2024-10-02 08:00 | Sending pump for repair/exchange | (TC 04/19/2024 08:42:20 AM, Baghirzade, Nara) SC 1003213566 (TC 04/26/2024 10:56:48 AM, Baghirzade, Nara) Received work order for exchange from Cardinal Health. KJV # 430129 (TC 09/10/2024 10:26:23 AM, Baghirzade, Nara) Sent follow up email regarding status of repair or if a replacement unit will be sent. (TC 10/02/2024 08:00:15 AM, Baghirzade, Nara) Received replacement pump. PO 7500205438
68869 | 206306 | PM | CMPLT | CMPLT | 2024-04-19 09:33 | 2024-05-02 09:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
68684 | 206307 | PM | CMPLT | CMPLT | 2024-04-19 09:34 | 2024-04-19 09:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055838 | 206308 | PM | CMPLT | CMPLT | 2024-04-19 09:44 | 2024-04-19 09:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:04:15 AM, Huang, Jessica) replaced the female iui connector completed PM
1048634 | 206309 | PM | CMPLT | CMPLT | 2024-04-19 09:48 | 2024-04-19 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 03:52:54 PM, Huang, Jessica) replaced he both male and female IUI connectors PM completed
1048642 | 206310 | PM | CMPLT | CMPLT | 2024-04-19 09:59 | 2024-04-19 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
| 206311 | RDS | CMPLT | CMPLT | 2024-04-19 10:16 | 2024-04-19 10:18 | Weekly Rounds: 11 North | (TC 04/19/2024 10:16:02 AM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 206312 | RDS | CMPLT | CMPLT | 2024-04-19 10:18 | 2024-04-19 10:21 | Weekly Rounds: 11 South | (TC 04/19/2024 10:18:11 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 206313 | RDS | CMPLT | CMPLT | 2024-04-19 10:21 | 2024-04-19 10:23 | Weekly Rounds: NICU | (TC 04/19/2024 10:21:32 AM, Baghirzade, Nara) Weekly rounds in NICU completed. See procedure for results.
| 206314 | RDS | CMPLT | CMPLT | 2024-04-19 10:23 | 2024-04-19 10:24 | Weekly Rounds: RESUS | (TC 04/19/2024 10:23:18 AM, Baghirzade, Nara) Weekly rounds in RESUS completed. See procedure for results.
1060000 | 206315 | CM | CMPLT | CMPLT | 2024-04-19 10:29 | 2024-04-19 10:31 | Blank screen | (TC 04/19/2024 10:29:18 AM, Baghirzade, Nara) 11S PSN was displaying error "client view down" on 11N PSN Screen was blank. Rebooted the PC. 11 South PSN is functional.
1055773 | 206316 | CM | CMPLT | CMPLT | 2024-04-19 10:32 | 2024-04-19 10:33 | Shows as unplugged from AC on dashboard | (TC 04/19/2024 10:32:24 AM, Baghirzade, Nara) Defib dashboard showed defib was not plugged in to AC. Nurse disconnected from AC during checks; re-did the 30J test to check it in to the dashboard.
1055772 | 206317 | CM | CMPLT | CMPLT | 2024-04-19 10:33 | 2024-04-19 10:34 | Unplugged from AC on dashboard | (TC 04/19/2024 10:33:33 AM, Baghirzade, Nara) Defib dashboard showed defib was not plugged in to AC. Nurse disconnected from AC during checks; re-did the 30J test to check it in to the dashboard.
501974 | 206318 | RQ | CMPLT | CMPLT | 2024-04-19 10:59 | 2024-04-19 15:00 | Probe & Well QC not passing |  | after power outage last night, constancy deviation over 20%
| 206319 | CM | CMPLT | CMPLT | 2024-04-19 11:05 | 2024-04-19 11:05 | Full team meeting | (TC 04/19/2024 11:04:47 AM, Baghirzade, Nara) Full team meeting
1048645 | 206320 | PM | CMPLT | CMPLT | 2024-04-19 11:05 | 2024-04-19 11:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/15/2024 05:17:46 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048646 | 206321 | PM | CMPLT | CMPLT | 2024-04-19 11:06 | 2024-04-19 11:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:18:38 AM, Huang, Jessica) replaced the female iui connector completed PM
1048653 | 206322 | PM | CMPLT | CMPLT | 2024-04-19 11:13 | 2024-04-19 11:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:03:18 AM, Huang, Jessica) replaced the female iui connector completed PM
1048658 | 206323 | PM | CMPLT | CMPLT | 2024-04-19 11:18 | 2024-04-19 11:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048660 | 206324 | PM | CMPLT | CMPLT | 2024-04-19 11:19 | 2024-04-19 11:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69578 | 206325 | PM | CMPLT | CMPLT | 2024-04-19 11:24 | 2024-04-19 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:26:20 AM, Huang, Jessica) replaced the female iui connector completed PM
69548 | 206326 | PM | CMPLT | CMPLT | 2024-04-19 11:25 | 2024-04-19 11:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
69241 | 206327 | PM | CMPLT | CMPLT | 2024-04-19 11:26 | 2024-04-19 11:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:23:48 AM, Huang, Jessica) replaced the female iui connector completed PM
69353 | 206328 | PM | CMPLT | CMPLT | 2024-04-19 11:27 | 2024-04-19 11:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:33:04 AM, Huang, Jessica) replaced the female iui connector completed PM
68635 | 206329 | CM | CMPLT | CMPLT | 2024-04-19 11:29 | 2024-04-19 11:30 | Completed BME incident report for SAFER #72895 | (TC 04/19/2024 11:28:38 AM, Baghirzade, Nara) Completed BME incident report for SAFER #72895 and saved to T Drive. For work completed to mitigate the incident refer to WO 206289. Followed up with Samir, Wendy, Eden, and Emmerson.
1048668 | 206330 | PM | CMPLT | CMPLT | 2024-04-19 11:37 | 2024-04-19 11:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1070483 | 206331 | CM | CMPLT | CMPLT | 2024-04-19 11:37 | 2024-04-19 11:49 | Received original unit back; sending loaner back to vendor | (TC 04/19/2024 11:46:11 AM, Baghirzade, Nara) Sending loaner back to vendor. KJV# 428936
1048670 | 206332 | PM | CMPLT | CMPLT | 2024-04-19 11:42 | 2024-04-19 11:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1048674 | 206333 | PM | CMPLT | CMPLT | 2024-04-19 11:46 | 2024-04-19 11:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1066917 | 206334 | CM | CMPLT | CMPLT | 2024-04-19 13:19 | 2024-04-19 13:21 | error 00.04 | (TC 04/19/2024 01:19:27 PM, Cortez, John) Confirmed issue, error seems to be software issue Called Stryker to confirm. Suggestion was to power cycle bed. Issue seems to resolve. Informed nurse and Marlene
1048678 | 206335 | PM | CMPLT | CMPLT | 2024-04-19 13:28 | 2024-04-19 13:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 09:06:10 AM, Huang, Jessica) replaced the female iui connector completed PM
1048681 | 206336 | PM | CMPLT | CMPLT | 2024-04-19 13:31 | 2024-04-19 13:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/17/2024 09:36:52 AM, Huang, Jessica) replaced the female iui connector completed PM
1048683 | 206337 | PM | CMPLT | CMPLT | 2024-04-19 13:33 | 2024-04-19 13:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1070771 | 206338 | CM | CMPLT | CMPLT | 2024-04-19 13:40 | 2024-04-19 13:41 | No display power | (TC 04/19/2024 01:40:24 PM, Chukri, Samir) Reset the battery tested working fine. returned to HP Tech Andrew
1048693 | 206339 | PM | CMPLT | CMPLT | 2024-04-19 13:42 | 2024-04-19 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 05:24:35 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1048695 | 206340 | PM | CMPLT | CMPLT | 2024-04-19 13:44 | 2024-04-19 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1061212 | 206341 | CM | CMPLT | CMPLT | 2024-04-19 16:22 | 2024-04-24 14:30 | coriomatrix servicing | (TC 04/24/2024 08:46:57 AM, Huang, Jessica) see the wo 206359
1055841 | 206342 | PM | CMPLT | CMPLT | 2024-04-22 08:27 | 2024-04-22 08:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055842 | 206343 | PM | CMPLT | CMPLT | 2024-04-22 08:28 | 2024-04-22 08:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/16/2024 05:22:32 PM, Huang, Jessica) replaced the female IUI connectors PM completed
1055840 | 206344 | PM | CMPLT | CMPLT | 2024-04-22 08:29 | 2024-04-22 08:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1055839 | 206345 | PM | CMPLT | CMPLT | 2024-04-22 08:30 | 2024-04-22 08:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
| 206346 | OC | CMPLT | CMPLT | 2024-04-22 08:32 | 2024-04-29 08:56 | On-Call week of April 21-26 2024 | (TC 04/22/2024 08:32:03 AM, Scero, Tomlyn) Daily checks complete and email sent (TC 04/23/2024 08:27:42 AM, Scero, Tomlyn) On call daily checks complete (TC 04/25/2024 08:54:55 AM, Scero, Tomlyn) Daily checks complete (TC 04/24/2024 08:55:26 AM, Scero, Tomlyn) daily checks complete
| 206347 | RDS | CMPLT | CMPLT | 2024-04-22 08:33 | 2024-04-22 15:04 | Weekly checks in CCU | (TC 04/22/2024 03:03:01 PM, Scero, Tomlyn) weekly checks complete. 30J tests passed. checked printers as well.
1061216 | 206348 | CM | CMPLT | CMPLT | 2024-04-22 10:42 | 2024-04-22 10:44 | Monitor not moving | call re: monitor on the arm is not able to move. Went to check. Found Nurse did not know how to move monitor using the arm. Taught her how to move the monitor around using the monitor arm. Tested ok.
65698 | 206349 | CM | CMPLT | CMPLT | 2024-04-22 10:45 | 2024-04-22 10:46 | Not working | Call re: not working. Went to check and brought it down to biomed. Plugged the unit to power. Unti turned on. Checked fucntionality. No Issue found. Removed and cleaned air filters on both side of the unit. Tested ok. Returned to OR.
1061316 | 206350 | CM | CMPLT | CMPLT | 2024-04-22 13:43 | 2024-04-22 13:45 | Patient CD request | (TC 04/22/2024 01:43:26 PM, Fan, Tak) Pulled and copied 6 dcm & jpg/mp4 exams to a DVD.
67921 | 206351 | CM | CMPLT | CMPLT | 2024-04-22 14:38 | 2024-04-22 14:39 | Humidifier not working | (TC 04/22/2024 02:38:15 PM, Rhuland, Courtney) Ran incubator for an hour and humidity did not exceed 50% Incubator to be retired and replaced with new omnibed rather than repaired.
1061214 | 206352 | CM | CMPLT | CMPLT | 2024-04-22 16:03 | 2024-04-22 16:05 | Monitor Issue | Call re" Not working. Went to check. Found PACS computer and Nursing PC not being able to display on Olympus monitor. Rebooted VMC a couple of times to ensure VMC fucntioanlity. Checked and verified all connections. Found wall DVI port broken for PACS monitor and DVI/Display port adaptor broken on nursing PC. ADvised staff to contact help desk. Tested ok.
1061205 | 206353 | CM | CMPLT | CMPLT | 2024-04-22 16:08 | 2024-04-22 16:09 | VMC Issue | Call re" Not working. Went to check. Found Olympus monitor not displaying selected inputs on touchscreen. Rebooted VMC. Checked and verified all connections. Issue persisted. Rebooted Coriomatrix. Checked functionality. Tested ok.
1061208 | 206354 | CM | CMPLT | CMPLT | 2024-04-23 08:44 | 2024-04-23 08:45 | Missing Coriomatrix | (TC 04/23/2024 08:43:59 AM, Reid, Emmerson ) Was called that the video routing was not working. I reset the VMC, this did not work. Then reset the Coriomatrix, after a few minutes the video routing was working.
| 206355 | CM | CMPLT | CMPLT | 2024-04-23 08:49 | 2024-05-28 10:42 | Cuffs for NICU | (TC 04/23/2024 08:53:17 AM, Rhuland, Courtney) Ordered 20 boxes of sizes 1-4 NIBP cuffs for NICU Requested by Tara Lee from Materials Management for backup stock (TC 05/28/2024 10:42:22 AM, Baghirzade, Nara) Received 20 boxes of size 3; contacted Tara Lee and had materials management deliver the boxes to her. She confirmed that she had received it
66058 | 206356 | CM | CMPLT | CMPLT | 2024-04-23 10:02 | 2024-04-23 10:03 | Missing handle | (TC 04/23/2024 10:02:34 AM, Rhuland, Courtney) Handle replaced Functional check - OK
1061211 | 206357 | CM | CMPLT | CMPLT | 2024-04-23 10:07 | 2024-04-23 10:07 | error message on the olympus touch panel | (TC 04/19/2024 10:07:09 AM, Huang, Jessica) received the call from OR 10 3D video image was not showing on the monitors and "the video matix is busy and it cant perform the requested change". Restarted the VMC and checked other inputs. except 4K system none of them worked. informed the nurse and was told they would only use the 4k system on Friday and the room would be used 3D system on Monday. came back to the room when the first case was completed. restarted the VMC again and it did not fix the issue. restarted coriomatrix after the second case was done, restarted the VMC three times and finally got everything worked. informed Maureen and Rose, tested the 3D system with the 3D scope
1061215 | 206358 | CM | CMPLT | CMPLT | 2024-04-23 10:15 | 2024-04-23 10:15 | error message on the olympus touch panel | (TC 04/19/2024 10:14:40 AM, Huang, Jessica) received the call from Maureen there was the issue in OR 14. none of the inputs was able to display on the monitors. error message on the touch panel. " the video matrix is busy and it cant perfrom the requested change". restarted the VMC and it did not the issue. informed Maureen and they might need to switch the second case to a different room because they would need the monitors to dispaly the image went back to the room after the first case was completed, restarted Coriomatrix and restarted VMC twice, tested evertyhing, the room was functional.
1061212 | 206359 | CM | CMPLT | CMPLT | 2024-04-23 10:20 | 2024-04-23 10:20 | error message on the olympus touch panel | (TC 04/19/2024 10:30:19 AM, Huang, Jessica) received the email from Tory and there was an error message on the olympus touch panel. went to the room and none of the inputs was displaying on the monitors. restarted the VMC and did not fix the issue. contacted TVONE for tech support beacause the LED was red on the coriomatrix TVONE helped to re-install the firmware; the error message was gone and the touch panel was working; however, the monitors were not displaying anything. contacted Oscar and was informed someone would need to be on-site to take a look at the issue. informed the OR team, worked on the PO and sent it to Olympus (TC 04/22/2024 08:45:03 AM, Huang, Jessica) called Derrick to get support of changing the settings and configurations on the coriomatrix tested and confirmed everything worked in the room informed Maureen and Tory Tory came to confirm everything worked
68632 | 206360 | PM | CMPLT | CMPLT | 2024-04-23 10:23 | 2024-04-23 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69568 | 206361 | PM | CMPLT | CMPLT | 2024-04-23 10:30 | 2024-04-23 10:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | PM completed by BD
32809 | 206362 | CM | CMPLT | CMPLT | 2024-04-23 10:39 | 2024-05-01 08:06 | MSH DR LLI 4' HOLDER AND CASSETTES | (TC 04/23/2024 10:39:27 AM, Dolinar, John) MSH DR LLI 4' HOLDER AND CASSETTES REPAIR CASSETTES FOR #3 POSITION (TC 04/23/2024 10:41:23 AM, Whittington, Scott) MSH DR LLI 4' HOLDER AND CASSETTES | MSH DR LLI 4' HOLDER AND CASSETTES
1070798 | 206363 | CM | CMPLT | CMPLT | 2024-04-23 10:42 | 2024-04-23 12:31 | Incoming inspection | (TC 04/23/2024 12:30:33 PM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070799 | 206364 | CM | CMPLT | CMPLT | 2024-04-23 10:45 | 2024-04-23 12:31 | Incoming inspection | (TC 04/23/2024 12:31:10 PM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070800 | 206365 | CM | CMPLT | CMPLT | 2024-04-23 10:46 | 2024-04-23 12:32 | Incoming inspection | (TC 04/23/2024 12:31:33 PM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070801 | 206366 | CM | CMPLT | CMPLT | 2024-04-23 10:47 | 2024-04-23 12:32 | Incoming inspection | (TC 04/23/2024 12:31:55 PM, Rhuland, Courtney) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1064713 | 206367 | CM | CMPLT | CMPLT | 2024-04-23 11:17 | 2024-04-23 11:20 | Touch panel error E333 | (TC 04/23/2024 11:17:00 AM, Reid, Emmerson ) Had a message for Toucg panel error E333 Rebooted device and it worked.
1053863 | 206368 | CM | CMPLT | CMPLT | 2024-04-23 11:18 | 2024-04-23 11:25 | "not working" | "nor working". Cleaned unit, cable and lead wires. Checked functionality and current delivery. Tested ok. No issue found. Returned to OR.
1050271 | 206369 | PM | CMPLT | CMPLT | 2024-04-23 11:26 | 2024-04-23 11:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1054521 | 206370 | CM | CMPLT | CMPLT | 2024-04-23 11:27 | 2024-04-23 11:29 | V5 not working | (TC 04/23/2024 11:27:36 AM, Canjimavadivel, Kamallesh ) Confirmed issue, replaced V4-V6 chest lead set. Tested ok with simulator
62696 | 206371 | CM | CMPLT | NFF | 2024-04-23 13:04 | 2024-04-23 13:07 | Sent with yellow tag, no message | (TC 04/23/2024 01:04:23 PM, Canjimavadivel, Kamallesh ) Turned on and tested leads with simulator, no issues with signal.
62988 | 206372 | CM | CMPLT | NFF | 2024-04-23 13:49 | 2024-04-23 13:51 | "Light weak" | (TC 04/23/2024 01:49:13 PM, Rhuland, Courtney) Tightened loose bolt on stand High intensity - 31.5 uW/cm^2/nm Low intensity - 13.1 31.5 uW/cm^2/nm No fault found
68369 | 206373 | PM | CMPLT | CMPLT | 2024-04-23 14:02 | 2024-04-25 14:13 | PM-MAJOR - Annual PM For: Anesthesia Units | (TC 04/22/2024 02:02:26 PM, Scero, Tomlyn) 6 year PM parts replaced including batteries. (TC 04/23/2024 02:02:59 PM, Scero, Tomlyn) PM tests started as per IPM-L. (TC 04/24/2024 02:59:48 PM, Scero, Tomlyn) PM tests continued (TC 04/25/2024 02:11:34 PM, Scero, Tomlyn) All PM tests complete and passed. Test results sheet attached to asset.
1046439 | 206374 | CM | CMPLT | CMPLT | 2024-04-23 14:14 | 2024-04-23 14:15 | Error 395 | Call re: Cautery not working". Went to check. Was informed that they got Error code 395. Rebooted teh cautery unit. Issue resolved. Waited around to see if the issue comes up again. Tested ok.
1064745 | 206375 | CM | CMPLT | CMPLT | 2024-04-23 14:16 | 2024-04-23 14:18 | Repair | Call re: Not working". Went to check. Waited for access. Found monitor was displaying Error E226 - Scope coomunication error. Checked all connections. Unplugged the scope and image processor. Cleaned connectors and plugged it back. Rebooted. Issue resolved. Waited around to ensure functionality. Tested ok.
65831 | 206376 | CM | CMPLT | CMPLT | 2024-04-23 14:19 | 2024-04-29 10:51 | Power Cord Issue | Apr 23 - "not working". Went to check and brought it down to biomed. Foudn power cord damaged. Unit requires a Hospital grade Left angle IEC C13 power cord. Could not find in stock. Ordered power cord from AC Works. Apr 29 - Replaced and secured new power cord. Checked functionality. Tested ok. DElivered to OR.
32520 | 206377 | CM | CMPLT | CMPLT | 2024-04-23 14:29 | 2024-05-01 08:05 | MSH DR2 115 GRID STORAGE POSITION WILOL NOT ACCEPT GRID FULLY | (TC 04/24/2024 04:11:27 PM, Whittington, Scott) MSH DR2 115 GRID STORAGE POSITION WILOL NOT ACCEPT GRID FULLY (TC 05/01/2024 08:03:46 AM, Whittington, Scott) MSH DR2 115 GRID STORAGE POSITION WILL NOT ACCEPT GRID FULLY PBSTRUCTION CLEARED REPAIRED DETECTOR CLAMSHELL AND LEFT CLAMSHELL LOCK | MSH DR2 115 GRID STORAGE POSITION WILOL NOT ACCEPT GRID FULLY
| 206378 | CM | CMPLT | CMPLT | 2024-04-23 15:56 | 2024-04-23 15:58 | Suction Tree Repair | Call re: not working. Went to check and brought it down to biomed. Checked all connections. Found Suction regulator sitiing loose on the tree and suction hose connection loose as well. Tightned it. Checked functionality. Found internal leak in suction. Retured suction, secured the rest of the connections and delivered back to OR. Advised OR staff to order new suction regulator.
1069319 | 206379 | CM | CMPLT | CMPLT | 2024-04-24 09:10 | 2024-04-24 09:11 | Mounting on rolling stand | (TC 04/24/2024 09:10:14 AM, Rhuland, Courtney) Mounted device on rolling stand Added basket and handle from spare
1064787 | 206380 | CM | CMPLT | CMPLT | 2024-04-24 09:34 | 2024-06-21 14:26 | Video Cart Parts Order | Apr 24 - Call re: back panel falling off. Went to check and brought it dowm to biomed. Found back panel cover grommet missing. Contacted Olympus for part# and pricing. Received. Created cart and sent it to SC for approval. Awaiting delivery. MAy 06 - Followed up with Olympus twice. No response from customer service. Contacted sales rep for ETA. May 07 - Followed up. Wasn informed ETA is 3 weeks. June 06 - Followed up again. Was informed they are experiencing a delay. New ETA is Mid June., June 21 - Parts received. Installed it on the cart. Checked functionality. Tested ok. Delivered to OR.
68670 | 206381 | PM | CMPLT | CMPLT | 2024-04-24 09:35 | 2024-04-24 09:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69219 | 206382 | PM | CMPLT | CMPLT | 2024-04-24 09:39 | 2024-04-24 09:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 206383 | PA | CMPLT | CMPLT | 2024-04-24 09:40 | 2024-05-09 09:17 | Product validation trial | (TC 04/24/2024 09:55:24 AM, Canjimavadivel, Kamallesh ) Requested No Charge PO for trial items as per Becky
69591 | 206384 | PM | CMPLT | CMPLT | 2024-04-24 09:44 | 2024-04-24 09:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69624 | 206385 | PM | CMPLT | CMPLT | 2024-04-24 09:48 | 2024-04-24 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1054537 | 206386 | PM | CMPLT | CMPLT | 2024-04-24 09:59 | 2024-04-24 10:01 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/23/2024 09:59:48 AM, Cortez, John) Completed cleaning cycle
1061212 | 206387 | CM | CMPLT | CMPLT | 2024-04-24 10:06 | 2024-04-24 10:08 | Monitor Configuration Issue | Call Re; Monitor Dispaly mal function. Went to check. Waited for access. Found Monitor 1 display configuration was changed to display two simultaneous inputs. Changed configuration as per staff request and locked key control on all monitors so it diesnt happen again. Tested ok.
1061220 | 206388 | CM | CMPLT | CMPLT | 2024-04-24 10:10 | 2024-04-24 10:10 | bluetooth not working | (TC 04/19/2024 10:10:10 AM, Huang, Jessica) went to the room to fix the bluetooth restarted VMC and working fine
1054412 | 206389 | CM | CMPLT | CMPLT | 2024-04-24 10:21 | 2024-04-24 10:23 | Battery low | (TC 04/24/2024 10:22:10 AM, Canjimavadivel, Kamallesh ) Replaced 6x C batteries
1054813 | 206390 | PM | CMPLT | CMPLT | 2024-04-24 10:22 | 2024-04-24 10:28 | PM-MAJOR - Annual PM For: Ventilators | (TC 04/23/2024 10:24:20 AM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 29541hrs Total vent time: 16366hrs Membrane capacity: 74%
| 206391 | PA | CMPLT | CMPLT | 2024-04-24 10:28 | 2024-06-18 08:48 | H-1200 PM stock parts | (TC 04/24/2024 10:45:43 AM, Canjimavadivel, Kamallesh ) Placed order for stocks parts (TC 05/06/2024 01:49:18 PM, Canjimavadivel, Kamallesh ) Items delivered were incorrect, did not match PO. Went to receiving to inquire but was not able to find anyone. (TC 06/18/2024 08:48:14 AM, Canjimavadivel, Kamallesh ) Items were lost be receiving department, reordered shopping cart | Charge back as used
1067353 | 206392 | CM | CMPLT | CMPLT | 2024-04-24 10:57 | 2024-05-06 14:49 | Mounting arm cover order | Apr 24 - Found covers damaged/missing in all 4 rooms on mounting stand. Contacted Olympus for part# and pricing. Received. Created cart and sent it to SC for approval. Awaiting delivery. May 06 - Received. Will coordinate with clinical staff and replace the covers.
| 206393 | RDS | CMPLT | CMPLT | 2024-04-24 11:04 | 2024-04-24 13:08 | Weekly check for ED
| 206394 | RDS | CMPLT | CMPLT | 2024-04-24 11:04 | 2024-04-30 14:46 | Weekly check for ICU
1062506 | 206395 | CM | CMPLT | CMPLT | 2024-04-24 11:43 | 2024-04-24 11:43 | not working | reset the unit. tested. Functional. returned back to service.
1061991 | 206396 | CM | CMPLT | CMPLT | 2024-04-24 11:46 | 2024-04-24 11:46 | not working | reset the unit. tested. Functional. returned back to service.
| 206397 | RDS | CMPLT | CMPLT | 2024-04-24 11:50 | 2024-04-24 11:52 | weekly check for 14 North and 17M | completed
1057961 | 206398 | CM | CMPLT | CMPLT | 2024-04-24 12:47 | 2024-04-24 12:48 | Membrane change date | (TC 04/24/2024 12:47:33 PM, Rhuland, Courtney) Membrane date change updated from 28 to 31 days
1057675 | 206399 | CM | CMPLT | CMPLT | 2024-04-24 12:50 | 2024-04-24 12:50 | Membrane date change | (TC 04/24/2024 12:49:42 PM, Rhuland, Courtney) Membrane date change updated from 28 to 31 days
1054044 | 206400 | CM | CMPLT | CMPLT | 2024-04-24 12:51 | 2024-04-24 12:51 | Membrane date change | (TC 04/24/2024 12:50:56 PM, Rhuland, Courtney) Membrane date change updated from 28 to 31 days
| 206401 | CM | CMPLT | CMPLT | 2024-04-24 13:17 | 2024-04-24 13:18 | Light not working | (TC 04/24/2024 01:17:14 PM, Rhuland, Courtney) Replaced bulb in exam light at OPG Functional check - OK
1059778 | 206402 | FC | CMPLT | CMPLT | 2024-04-24 13:25 | 2024-04-24 13:26 | QC check | (TC 04/24/2024 01:25:08 PM, Rhuland, Courtney) QC checklist completed Returned to 11S
69288 | 206403 | PM | CMPLT | CMPLT | 2024-04-24 14:18 | 2024-04-24 14:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 206404 | CM | CMPLT | CMPLT | 2024-04-24 14:28 | 2024-04-24 14:28 | meeting | meeting
66064 | 206405 | PM | CMPLT | CMPLT | 2024-04-24 14:45 | 2024-04-24 14:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68781 | 206406 | PM | CMPLT | CMPLT | 2024-04-24 14:52 | 2024-04-24 14:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1062999 | 206407 | PM | CMPLT | CMPLT | 2024-04-24 15:05 | 2024-04-24 15:06 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/24/2024 03:05:32 PM, Cortez, John) Completed cleaning cycle
501064 | 206408 | CM | CMPLT | CMPLT | 2024-04-24 16:10 | 2024-09-25 09:41 | MSH CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8 | (TC 04/24/2024 04:10:02 PM, Whittington, Scott) CEOU RM 7 RONGE REPORTED PROBE ISSUE? (TC 04/24/2024 04:10:02 PM, Whittington, Scott) CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8 (TC 04/29/2024 10:47:39 AM, Whittington, Scott) CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8 (TC 04/29/2024 10:47:57 AM, Dolinar, John) CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8 (TC 06/10/2024 08:12:00 AM, Whittington, Scott) MSH CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8 & 9 OBTAIN EXTENSIONS CORDS AND RUN TEMP POWER TO ROOM 7 FROM ROOM 5 CONTACT DWAYNE ANDERSON TRIOVEST OPG BUILDING MANAGER REGARDING INTERFERENCE ISSUES (TC 06/10/2024 10:14:18 AM, Whittington, Scott) MSH CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8 & 9 OBTAIN EXTENSIONS CORDS AND RUN TEMP POWER TO ROOM 7 FROM ROOM 5 CONTACT DWAYNE ANDERSON TRIOVEST OPG BUILDING MANAGER REGARDING INTERFERENCE ISSUES (TC 06/13/2024 12:50:12 PM, Whittington, Scott) OBTAIN EXTENSIONS CORDS AND RUN TEMP POWER TO ROOM 7 FROM ROOM 5 CONTACT DWAYNE ANDERSON TRIOVEST OPG BUILDING MANAGER REGARDING INTERFERENCE ISSUES FOLLOW UP ANGELA TO HAVE TECHS RECORD OCCURENCES (TC 06/13/2024 12:50:12 PM, Whittington, Scott) OBTAIN EXTENSIONS CORDS AND RUN TEMP POWER TO ROOM 7 FROM ROOM 5 CONTACT DWAYNE ANDERSON TRIOVEST OPG BUILDING MANAGER REGARDING INTERFERENCE ISSUES FOLLOW UP ANGELA TO HAVE TECHS RECORD OCCURENCES (TC 06/10/2024 08:12:00 AM, Whittington, Scott) MSH CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8 & 9 OBTAIN EXTENSIONS CORDS AND RUN TEMP POWER TO ROOM 7 FROM ROOM 5 CONTACT DWAYNE ANDERSON TRIOVEST OPG BUILDING MANAGER REGARDING INTERFERENCE ISSUES (TC 06/13/2024 12:50:12 PM, Whittington, Scott) OBTAIN EXTENSIONS CORDS AND RUN TEMP POWER TO ROOM 7 FROM ROOM 5 CONTACT DWAYNE ANDERSON TRIOVEST OPG BUILDING MANAGER REGARDING INTERFERENCE ISSUES FOLLOW UP ANGELA TO HAVE TECHS RECORD OCCURENCES | MSH CEOU RM 7 RONGE REPORTED PROBE ISSUE? ENVIRONMENTAL ELECTRICAL NOISE? ALSO EXPERIENCED IN ROOM 8
1061214 | 206409 | CM | CMPLT | CMPLT | 2024-04-25 10:00 | 2024-04-25 10:03 | Bluetooth Issue | Call re: speaker not working. Went to check. Found device not able to recognize bluetooth device when connected. Rebooted VMC and computer. Issue resolved. Tested ok.
1064712 | 206410 | CM | CMPLT | CMPLT | 2024-04-25 10:04 | 2024-04-25 10:05 | Data Display Issue | Call re: need to hide data info. Went tot check. Waited for access. Was informed that the surgeon wants to get rid of the patient info on the screen. Went through menus and turned the data info off. Tested ok.
| 206411 | CM | CMPLT | CMPLT | 2024-04-25 10:53 | 2024-04-25 10:54 | VAC Ulta rental unit issue | OR sent a VAC Ulta rental unit to biomed for repair. Inspected unit and Found unit had a broken power supply. Consulted SC and was advised that rental units are handled by the dept. Delivered unit back to OR and communicated with staff to send it to 3M for repair.
1069775 | 206412 | CM | CMPLT | CMPLT | 2024-04-25 11:02 | 2024-05-15 10:56 | Cutting in and out | (TC 04/25/2024 11:04:32 AM, Rendon Lopez , Elizabeth) Signal cutting, previously repaired. semding out under warranty repair (TC 05/15/2024 10:56:05 AM, Rendon Lopez , Elizabeth) Unable to repair. Received device back, retired
1054419 | 206413 | CM | CMPLT | CMPLT | 2024-04-25 11:08 | 2024-05-15 10:58 | Case open | (TC 04/25/2024 11:08:41 AM, Rendon Lopez , Elizabeth) Confirmed. Sending out for repair under warranty (TC 05/15/2024 10:57:13 AM, Rendon Lopez , Elizabeth) Unable to repair. Received device back, retired
1048003 | 206414 | CM | CMPLT | CMPLT | 2024-04-25 11:11 | 2024-04-25 11:11 | Doesn't register activity on monitor | (TC 04/25/2024 11:10:46 AM, Rendon Lopez , Elizabeth) Tested with different weights, NFF. Returned to L&D
1050481 | 206415 | CM | CMPLT | CMPLT | 2024-04-25 11:16 | 2024-05-07 09:46 | "leakage too high" alarm than ventilation halted. no leak found in circuit or interface | Unit passed pre-use check. collected logs and sent them to Getinge. Running unit overnight. Got the response from Alex-Getinge "There are no technical alarms in the logs. If the device can pass the pre-use check at each attempt and if the simulated ventilation using a test lung does not generate any alarms, then you may proceed to return this device to clinical use." Then returned it back to service.
67958 | 206416 | CM | CMPLT | CMPLT | 2024-04-25 11:22 | 2024-04-25 11:23 | Broken flip door | (TC 04/25/2024 11:22:04 AM, Rhuland, Courtney) Replaced flip door with spare Replaced bent O2 tank holder Calibrated RH Functional check - OK
1070839 | 206417 | ACC | CMPLT | CMPLT | 2024-04-25 11:27 | 2024-04-25 14:15 | Incoming inspection | (TC 04/25/2024 02:14:44 PM, Baghirzade, Nara) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070841 | 206418 | ACC | CMPLT | CMPLT | 2024-04-25 11:28 | 2024-04-25 14:16 | Incoming inspection | (TC 04/25/2024 02:15:34 PM, Baghirzade, Nara) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070842 | 206419 | ACC | CMPLT | CMPLT | 2024-04-25 11:29 | 2024-04-25 14:17 | Incoming inspection | (TC 04/25/2024 02:16:11 PM, Baghirzade, Nara) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1070843 | 206420 | ACC | CMPLT | CMPLT | 2024-04-25 11:31 | 2024-04-25 14:14 | Incoming inspection | (TC 04/25/2024 02:13:37 PM, Baghirzade, Nara) Incoming inspection completed. Adjusted the light height and removed the adjustable knob. Removed outlet guard at the back of the omnibed. Incubator is functional.
1066474 | 206421 | CM | CMPLT | CMPLT | 2024-04-25 13:25 | 2024-06-11 14:21 | Dirty Filter | (TC 04/25/2024 01:24:56 PM, Reid, Emmerson ) Inspected all the Neptunes with Stryker. Found that this unit was also contaminated. Created safer and let the team know. Risk File ID 74281 (TC 04/29/2024 09:57:35 AM, Reid, Emmerson ) Created biomed safer report (TC 04/29/2024 02:34:42 PM, Reid, Emmerson ) Forwarded the repair quote to Reeba, also went to look for the last neptune but was unable to find it. (TC 06/05/2024 03:15:49 PM, Reid, Emmerson ) Neptune followup meeting (TC 06/11/2024 02:21:17 PM, Reid, Emmerson ) Neptune meeting with Witterick
1056125 | 206422 | PM | CMPLT | CMPLT | 2024-04-25 13:26 | 2024-04-25 13:27 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/25/2024 01:26:41 PM, Canjimavadivel, Kamallesh ) Cleaned hoses and nozzles. Cleaning cycle complete
69467 | 206423 | CM | CMPLT | CMPLT | 2024-04-25 13:39 | 2024-04-25 13:41 | Defective patient-side sensor; error 241.4140 | (TC 04/25/2024 01:39:36 PM, Baghirzade, Nara) Downloaded error log; noted error 241.4140, low voltage on patient side pressure sensor. Ran analog test and confirmed the error. Replaced patient side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
69370 | 206424 | CM | CMPLT | CMPLT | 2024-04-25 13:42 | 2024-04-25 13:42 | Defective patient-side sensor; error 241.4140 | (TC 04/25/2024 01:42:02 PM, Baghirzade, Nara) Downloaded error log; noted error 241.4140, low voltage on patient side pressure sensor. Ran analog test and confirmed the error. Replaced patient side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
69325 | 206425 | CM | CMPLT | CMPLT | 2024-04-25 13:43 | 2024-04-25 13:44 | Defective bottle-side pressure sensor; error 240.4150 | (TC 04/25/2024 01:43:16 PM, Baghirzade, Nara) Downloaded error log; noted error 240.4150, high voltage on bottle side pressure sensor. Ran analog test and confirmed the error. Replaced patient side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. (TC 04/25/2024 01:43:16 PM, Baghirzade, Nara) Downloaded error log; noted error 240.4150, high voltage on bottle side pressure sensor. Ran analog test and confirmed the error. Replaced patient side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
1061333 | 206426 | CM | CMPLT | CMPLT | 2024-04-25 13:46 | 2024-04-25 13:47 | Replaced mounting mechanism and QC | (TC 04/25/2024 01:45:54 PM, Baghirzade, Nara) Replaced mount mechanism and placed instructional sticker. QC checklist completed. Returned to 11 North.
1065976 | 206427 | CM | CMPLT | CMPLT | 2024-04-25 13:47 | 2024-04-25 13:48 | Functional check | (TC 04/25/2024 01:47:15 PM, Baghirzade, Nara) User reported channel error. Downloaded error log; no errors noted. Check-in tests passed and completed. IUI connectors and device cleaned.
69236 | 206428 | CM | CMPLT | CMPLT | 2024-04-25 13:48 | 2024-04-25 14:09 | Safety clamp error | (TC 04/25/2024 02:08:14 PM, Baghirzade, Nara) Failed patient occlusion during check-in tests. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. Checked safety clamp; no issue noted.
69013 | 206429 | CM | CMPLT | CMPLT | 2024-04-25 14:10 | 2024-04-25 14:11 | No issue stated on yellow biomed tag | (TC 04/25/2024 02:10:18 PM, Baghirzade, Nara) Downloaded error log; no errors noted. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 108 mohms Leakage current: 74.0 uA
68689 | 206430 | CM | CMPLT | CMPLT | 2024-04-25 14:11 | 2024-04-25 14:12 | No error reported on yellow biomed tag | (TC 04/25/2024 02:11:23 PM, Baghirzade, Nara) Downloaded error log; no errors noted. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 165 mohms Leakage current: 76.8 uA
1065567 | 206431 | CM | CMPLT | CMPLT | 2024-04-25 14:19 | 2024-04-25 14:20 | 3 year PM | (TC 04/25/2024 02:19:30 PM, Scero, Tomlyn) 3 year PM complete after it came down for a leak.
1063000 | 206432 | PM | CMPLT | CMPLT | 2024-04-25 14:26 | 2024-04-25 14:28 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 04/25/2024 02:26:22 PM, Cortez, John) completed cleaning cycle
1064312 | 206433 | CM | CMPLT | CMPLT | 2024-04-25 14:31 | 2024-04-25 14:56 | 6 year PM | (TC 04/25/2024 02:55:39 PM, Scero, Tomlyn) All PM parts replaced and PM test complete. Self test passed with no leaks.
| 206434 | CM | CMPLT | CMPLT | 2024-04-25 15:51 | 2024-05-06 14:59 | Elmed Insulation Testeer Issue from MDRD | Unit was dropped off to biomed by MDRD asking for help in troubleshooting. Upon diagnisis, found unit was making beep sound even when no probe in inserted into the testing hole. Contacted OEM for assistance. Found unit requires a new IC. Searched and ordered a new IC. Awaiting delivery. Parts received. Replced IC. Checked funcitonlity. Contacted MDRD and delivered the unit. Tested ok.
1070798 | 206435 | CM | CMPLT | CMPLT | 2024-04-26 08:24 | 2024-04-26 08:25 | Broken end door | (TC 04/26/2024 08:24:39 AM, Rhuland, Courtney) Re-attached end door Functional check - OK
1047044 | 206436 | PM | CMPLT | CMPLT | 2024-04-26 08:46 | 2024-03-15 08:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065058 | 206437 | PM | CMPLT | CMPLT | 2024-04-26 08:50 | 2024-04-26 08:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69323 | 206438 | PM | CMPLT | CMPLT | 2024-04-26 09:07 | 2024-04-26 09:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 206439 | CM | CMPLT | CMPLT | 2024-04-26 09:09 | 2024-06-20 08:32 | Order bed locator charger | (TC 04/26/2024 09:15:45 AM, Chukri, Samir) ordered 322 ea adapter no charge please see attched. (TC 06/20/2024 08:32:09 AM, Chukri, Samir) Received parts.
69583 | 206440 | PM | CMPLT | CMPLT | 2024-04-26 09:11 | 2024-04-26 09:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69164 | 206441 | PM | CMPLT | CMPLT | 2024-04-26 09:21 | 2024-04-26 09:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69250 | 206442 | PM | CMPLT | CMPLT | 2024-04-26 09:25 | 2024-04-26 09:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69515 | 206443 | PM | CMPLT | CMPLT | 2024-04-26 09:34 | 2024-04-26 09:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69329 | 206444 | PM | CMPLT | CMPLT | 2024-04-26 09:40 | 2024-04-26 09:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68668 | 206445 | PM | CMPLT | CMPLT | 2024-04-26 09:48 | 2024-04-26 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1067511 | 206446 | CM | CMPLT | CMPLT | 2024-04-26 09:51 | 2024-04-26 09:51 | Hose damaged | (TC 04/25/2024 09:50:58 AM, Rendon Lopez , Elizabeth) Replaced hose from spare. Tested-OK
69627 | 206447 | PM | CMPLT | CMPLT | 2024-04-26 09:52 | 2024-04-26 09:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68865 | 206448 | CM | CMPLT | CMPLT | 2024-04-26 09:54 | 2024-04-26 09:56 | Not working | (TC 04/25/2024 09:54:30 AM, Rendon Lopez , Elizabeth) Downloaded error logs, no issues found. Channel errow SW compatibility issue, no channels sent down. Check in tests, all passed
67395 | 206449 | CM | CMPLT | CMPLT | 2024-04-26 09:57 | 2024-04-26 09:59 | Not heating up to full temperature | (TC 04/24/2024 09:57:12 AM, Rendon Lopez , Elizabeth) Running for a long period of time, no issues found. Checked error logs, none. Heat lamp check- OK (TC 04/25/2024 09:57:12 AM, Rendon Lopez , Elizabeth) Tested, NFF, returned to L&D as spare
68647 | 206450 | PM | CMPLT | CMPLT | 2024-04-26 09:57 | 2024-04-26 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1066212 | 206451 | CM | CMPLT | CMPLT | 2024-04-26 10:12 | 2024-04-26 10:13 | Not working | (TC 04/24/2024 10:12:37 AM, Rendon Lopez , Elizabeth) Tested, for a long period, waming up no issues. NFF
1050725 | 206452 | CM | CMPLT | CMPLT | 2024-04-26 10:14 | 2024-04-26 10:15 | Not activating | (TC 04/23/2024 10:14:15 AM, Rendon Lopez , Elizabeth) Confirmed issue, defective activation cable.Replaced cable, tested-OK
69363 | 206453 | PM | CMPLT | CMPLT | 2024-04-26 10:19 | 2024-04-26 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1051371 | 206454 | PM | CMPLT | CMPLT | 2024-04-26 10:22 | 2024-04-26 10:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058228 | 206455 | PM | CMPLT | CMPLT | 2024-04-26 10:26 | 2024-04-26 10:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059882 | 206456 | PM | CMPLT | CMPLT | 2024-04-26 10:28 | 2024-04-26 10:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1061209 | 206457 | CM | CMPLT | CMPLT | 2024-04-26 10:52 | 2024-04-26 10:55 | Surgical Light Issue | Call re: not working. Went to check and waited for access. Patient on the table. Foudn one of the light had several segments turned off and two of the lights were unable to control via ipad. Wne t through the settings. Rectified the issue and adjusted settign as per surgeon's recommendation. FOund one of the light had wifi disabled. Enabled the wifi. Checked fucntionality. Tested ok.
1056441 | 206458 | CM | CMPLT | CMPLT | 2024-04-26 10:58 | 2024-04-26 10:59 | Defective SpO2 sensor | (TC 04/26/2024 10:57:59 AM, Baghirzade, Nara) SpO2 sensor does not read; replaced the SpO2 sensor. Root battery was fully depleted; recharged the device. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 206459 | CM | CMPLT | CMPLT | 2024-04-26 11:00 | 2024-04-26 11:01 | On-call daily checks | (TC 04/26/2024 11:00:12 AM, Baghirzade, Nara) On-call daily checks for Philips, Masimo, defibs, and Stryker beds systems.
66066 | 206460 | PM | CMPLT | CMPLT | 2024-04-26 12:15 | 2024-03-18 12:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66051 | 206461 | PM | CMPLT | CMPLT | 2024-04-26 12:18 | 2024-03-18 12:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66009 | 206462 | PM | CMPLT | CMPLT | 2024-04-26 12:22 | 2024-03-18 12:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66030 | 206463 | PM | CMPLT | CMPLT | 2024-04-26 12:24 | 2024-03-18 12:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66006 | 206464 | PM | CMPLT | CMPLT | 2024-04-26 12:27 | 2024-03-18 12:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1059763 | 206465 | CM | CMPLT | CMPLT | 2024-04-26 12:28 | 2024-09-25 09:29 | MSH 15 L&D TE-7 D WILL NOT BOOT | (TC 04/26/2024 12:30:34 PM, Whittington, Scott) From: Ngo, Kim <Kim.Ngo@sinaihealth.ca> Sent: April 25, 2024 2:40 PM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Bell Webster, Nicole - Sinai Health System <Nicole.BellWebster@sinaihealth.ca> Cc: Dolinar, John <John.Dolinar@uhn.ca> Subject: RE: [External] RE: MSH L&D TE-Y POWER ADAPTER Hello, The cord has finally arrived. It is in my office on 15-212. I am away tomorrow but my co-worker Gill will be in the office. Kim Kim Ngo, RN, BScN Clinical Co-ordinator, Labour and Delivery T 416-586-4800 x 5450 kim.ngo@sinaihealth.ca 600 University Avenue, Toronto, ON M5G 1X5 A patient care, teaching and research centre affiliated with University of Toronto. The information in this email may be confidential. This email is intended be viewed by the individual(s) or organization named above. If you are not the intended recipient or an authorized representative of the intended recipient, please be notified that any review, dissemination or copying of this email and its attachments (if any) is strictly prohibited. If you have received this email in error, please immediately notify the sender by return email and delete this email from your system, including from the deleted items folder. Thank you for your co-operation. A patient care, teaching and research centre affiliated with University of Toronto. From: Whittington, Scott [mailto:Scott.Whittington@uhn.ca] Sent: March 04, 2024 3:42 PM To: Bell Webster, Nicole <Nicole.BellWebster@sinaihealth.ca> Cc: Ngo, Kim <Kim.Ngo@sinaihealth.ca>; Dolinar, John - UHN - University Health Network <John.Dolinar@uhn.ca> Subject: Fwd: [External] RE: MSH L&D TE-Y POWER ADAPTER CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Kim Could you please order the power supply and cord indicated on the quote. This is for the MINDRAY TE7 that Kim reported to us on Friday Thanks Scott 416-402-2462 SCOTT WHITTINGTON SINAI HEALTH SYSTEM TL MEDICAL IMAGING ENGINEERING DIGITAL RADIOGRAPHY, ULTRASOUND, MAMMOGRAPHY UHN, SHS, WCH 416-586-4800 x2113 OFFICE PHONE/SERVICE LINE 416-402-2462 MOBILE scott.whittington@uhn.ca MOUNT SINAI HOSPITAL 600 UNIVERSITY AVE, TORONTO ON ROOM MSH 5-152 M5G 1X5 Get Outlook for iOS ________________________________________ From: Lisa Petteplace <lpetteplace@bomimed.com> Sent: Monday, March 4, 2024 2:33:02 PM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Michelle Lee <MLee@bomimed.com> Cc: Dolinar, John <John.Dolinar@uhn.ca> Subject: [External] RE: MSH L&D TE-Y POWER ADAPTER HI Scott Please see the attached quotation for the power cord on the TE7. You will need both pieces, the adaptor and then the cord from the adaptor to the wall. Thank you, Lisa Petteplace RDMS, CMRITO, CRGS ULTRASOUND SALES REPRESENTATIVE 1-100 IRENE STREET WINNIPEG, MB R3T 4E1 C. 226.749.2926 T. 800.667.6276 F: 877.435.6984 www.bomimed.ca *PLEASE THINK OF THE ENVIRONMENT BEFORE PRINTING THIS EMAIL. This e-mail and any attachment transmitted with it are confidential and intended solely for the person or entity to whom it is addressed. If you receive this message in error please contact the sender immediately. Any unauthorized dissemination of this e-mail or its attachments and any use of any information contained in it is strictly prohibited. Any views expressed in this e-mail are those of the individual sender and do not necessarily reflect the views of BOMImed. Whilst we use reasonable endeavors to check and virus scan all e-mails leaving the Company, BOMImed accepts no responsibility for any computer virus or for any errors or omissions in this e-mail. Please undertake your own virus checking. Please note that BOMImed reserves the right to monitor e-mail communications through its networks at all times. From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Friday, March 1, 2024 3:56 PM To: Lisa Petteplace <lpetteplace@bomimed.com>; Michelle Lee <MLee@bomimed.com> Cc: Dolinar, John <John.Dolinar@uhn.ca> Subject: MSH L&D TE-Y POWER ADAPTER CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Lisa Could you please provide a price for: Adapter 022-000187-00 I am assuming from the service manual this is the adapter to power the te-7 independent of the Trolley. Thanks, Scott 416-402-2462 (TC 04/30/2024 01:02:06 PM, Whittington, Scott) MSH 15 L&D TE-7 D WILL NOT BOOT REPLACEMENT TRANSFORMER AND WIRE HARNESS ARRIVED ON SITE | MSH 15 L&D TE-7 D WILL NOT BOOT STAND POWER SUPPLY FAULTY 20. Adapter 022-000187-00 2121 Trolley power adapter 19 V 150 W 022-000194-00
68892 | 206466 | PM | CMPLT | CMPLT | 2024-04-26 12:47 | 2024-03-18 12:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68961 | 206467 | PM | CMPLT | CMPLT | 2024-04-26 12:49 | 2024-03-18 12:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68706 | 206468 | PM | CMPLT | CMPLT | 2024-04-26 13:05 | 2024-03-18 13:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68973 | 206469 | PM | CMPLT | CMPLT | 2024-04-26 13:10 | 2024-03-18 13:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
500732 | 206470 | CM | CMPLT | CMPLT | 2024-04-26 13:19 | 2024-04-26 13:19 | MSH-20-470 (Sports Clinic) APLIO XG REMOVED FROM CLINICAL USE IN SPORTS CLINIC | (TC 04/26/2024 01:18:58 PM, Whittington, Scott) MSH-20-470 (Sports Clinic) APLIO XG REMOVED FROM CLINICAL USE IN SPORTS CLINIC | MSH-20-470 (Sports Clinic) APLIO XG REMOVED FROM CLINICAL USE IN SPORTS CLINIC
69457 | 206471 | PM | CMPLT | CMPLT | 2024-04-26 13:28 | 2024-03-18 13:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065954 | 206472 | CM | CMPLT | CMPLT | 2024-04-26 13:32 | 2024-04-26 13:33 | Door not closing properly | (TC 04/26/2024 01:32:06 PM, Baghirzade, Nara) Took the screw from the latch out, realigned the spring and put the latch back on; door closing properly and easily. During check-in, the patient occlusion was failing as it was reading high. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
1063369 | 206473 | CM | CMPLT | CMPLT | 2024-04-26 13:34 | 2024-05-02 10:16 | User reported that the measurement was not accurate; reading too high | (TC 04/26/2024 01:34:05 PM, Baghirzade, Nara) Sarabeth reported that there is a large discrepancy in the output measurement (nearly 500mL) between this bladder scanner and another primeplus and old BVI that was borrowed from 11S. Brought the device down, completed the caliscan without issue. Contacted CHS for troubleshooting and next steps. (TC 04/30/2024 02:25:57 PM, Baghirzade, Nara) Followed up with CHS; the tech who completed the recent calibration at CHS sent the verification/measurements that he had from the phantom tests. The measurements were very close to the phantom measurements and he had confirmed that there was no issue with the device or the probe. Contacted Sarabeth and asked her to repeat measurements on a patient with this bladder scanner and with the 11S bladder scanner for comparison. Updated CHS of our plans and will proceed after Sarabeth provides an update on the measurements. (TC 05/02/2024 10:13:36 AM, Baghirzade, Nara) Met with Sarabeth to discuss recent bladder scanner measurement results. Results were measured against the 11 South bladder scanner and there were very minimal to no discrepancies. Sarabeth has deemed that the clinical staff require in-service training as there has been a high staff turnover. Updated Jagdish from CHS that bladder scanner is functional and doesn't require further servicing. Reached out to Naz to set up an in-service for 11th fl. clinical staff.
69331 | 206474 | PM | CMPLT | CMPLT | 2024-04-26 13:38 | 2024-03-18 13:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69412 | 206475 | PM | CMPLT | CMPLT | 2024-04-26 13:45 | 2024-03-18 13:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69392 | 206476 | PM | CMPLT | CMPLT | 2024-04-26 13:49 | 2024-03-18 13:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69593 | 206477 | PM | CMPLT | CMPLT | 2024-04-26 13:59 | 2024-03-18 13:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69060 | 206478 | PM | CMPLT | CMPLT | 2024-04-26 14:01 | 2024-03-18 14:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69536 | 206479 | PM | CMPLT | CMPLT | 2024-04-26 14:04 | 2024-03-18 14:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1061333 | 206480 | FC | CMPLT | CMPLT | 2024-04-26 14:07 | 2024-04-26 14:07 | QC checklist | (TC 04/26/2024 02:06:52 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
| 206481 | RDS | CMPLT | CMPLT | 2024-04-26 14:11 | 2024-04-26 14:13 | Weekly Rounds: 11 North | (TC 04/26/2024 02:11:41 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
69097 | 206482 | PM | CMPLT | CMPLT | 2024-04-26 14:12 | 2024-03-18 14:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 206483 | RDS | CMPLT | CMPLT | 2024-04-26 14:14 | 2024-04-26 14:15 | Weekly Rounds: 11 South | (TC 04/26/2024 02:13:47 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
68855 | 206484 | PM | CMPLT | CMPLT | 2024-04-26 14:36 | 2024-03-20 14:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68675 | 206485 | PM | CMPLT | CMPLT | 2024-04-26 14:42 | 2024-03-20 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69020 | 206486 | PM | CMPLT | CMPLT | 2024-04-26 14:47 | 2024-03-20 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69405 | 206487 | PM | CMPLT | CMPLT | 2024-04-26 15:58 | 2024-04-02 15:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69231 | 206488 | PM | CMPLT | CMPLT | 2024-04-26 16:03 | 2024-04-02 16:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69483 | 206489 | PM | CMPLT | CMPLT | 2024-04-26 16:17 | 2024-04-02 16:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69483 | 206490 | PM | CANCL | DUP | 2024-04-26 16:17 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69142 | 206491 | PM | CMPLT | CMPLT | 2024-04-26 16:21 | 2024-04-02 16:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69524 | 206492 | PM | CMPLT | CMPLT | 2024-04-26 16:24 | 2024-04-02 16:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69605 | 206493 | PM | CMPLT | CMPLT | 2024-04-26 16:28 | 2024-04-02 16:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69340 | 206494 | PM | CMPLT | CMPLT | 2024-04-26 16:47 | 2024-03-11 16:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69176 | 206495 | PM | CMPLT | CMPLT | 2024-04-26 16:52 | 2024-03-11 16:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69181 | 206496 | PM | CMPLT | CMPLT | 2024-04-26 16:57 | 2024-03-11 16:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69596 | 206497 | PM | CMPLT | CMPLT | 2024-04-26 18:48 | 2024-03-11 18:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69066 | 206498 | PM | CMPLT | CMPLT | 2024-04-26 18:52 | 2024-03-11 18:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69271 | 206499 | PM | CMPLT | CMPLT | 2024-04-26 18:58 | 2024-03-11 18:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69582 | 206500 | PM | CMPLT | CMPLT | 2024-04-26 20:02 | 2024-03-19 20:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69564 | 206501 | PM | CMPLT | CMPLT | 2024-04-26 20:08 | 2024-03-19 20:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69572 | 206502 | PM | CMPLT | CMPLT | 2024-04-26 20:10 | 2024-03-19 20:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69186 | 206503 | PM | CMPLT | CMPLT | 2024-04-26 20:12 | 2024-03-19 20:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69169 | 206504 | PM | CMPLT | CMPLT | 2024-04-26 20:20 | 2024-03-19 20:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69249 | 206505 | PM | CMPLT | CMPLT | 2024-04-26 20:24 | 2024-03-19 20:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69395 | 206506 | PM | CMPLT | CMPLT | 2024-04-26 20:34 | 2024-03-19 20:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68650 | 206507 | PM | CMPLT | CMPLT | 2024-04-26 20:38 | 2024-03-19 20:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69428 | 206508 | PM | CMPLT | CMPLT | 2024-04-26 20:44 | 2024-03-19 20:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69102 | 206509 | PM | CMPLT | CMPLT | 2024-04-26 20:52 | 2024-03-19 20:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69109 | 206510 | PM | CMPLT | CMPLT | 2024-04-26 21:01 | 2024-03-19 21:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69655 | 206511 | PM | CMPLT | CMPLT | 2024-04-26 21:14 | 2024-03-19 21:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69678 | 206512 | PM | CMPLT | CMPLT | 2024-04-26 21:20 | 2024-03-19 21:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69140 | 206513 | PM | CMPLT | CMPLT | 2024-04-26 21:37 | 2024-03-19 21:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69588 | 206514 | PM | CMPLT | CMPLT | 2024-04-26 21:40 | 2024-03-19 21:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69368 | 206515 | PM | CMPLT | CMPLT | 2024-04-27 13:20 | 2024-03-19 13:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69279 | 206516 | PM | CMPLT | CMPLT | 2024-04-27 13:23 | 2024-03-19 13:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69285 | 206517 | PM | CMPLT | CMPLT | 2024-04-27 13:28 | 2024-03-19 13:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69439 | 206518 | PM | CMPLT | CMPLT | 2024-04-27 13:34 | 2024-03-19 13:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69643 | 206519 | PM | CMPLT | CMPLT | 2024-04-27 13:48 | 2024-03-19 13:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050039 | 206520 | PM | CMPLT | CMPLT | 2024-04-27 13:53 | 2024-03-19 13:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058227 | 206521 | PM | CMPLT | CMPLT | 2024-04-27 13:56 | 2024-03-19 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058237 | 206522 | PM | CMPLT | CMPLT | 2024-04-27 13:59 | 2024-03-19 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058248 | 206523 | PM | CMPLT | CMPLT | 2024-04-27 14:01 | 2024-03-19 14:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68294 | 206524 | PM | CMPLT | CMPLT | 2024-04-27 18:06 | 2024-03-21 18:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69456 | 206525 | PM | CMPLT | CMPLT | 2024-04-27 18:09 | 2024-03-21 18:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69126 | 206526 | PM | CMPLT | CMPLT | 2024-04-27 18:13 | 2024-03-21 18:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69400 | 206527 | PM | CMPLT | CMPLT | 2024-04-27 18:27 | 2024-03-21 18:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69464 | 206528 | PM | CMPLT | CMPLT | 2024-04-27 18:32 | 2024-03-21 18:32 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69135 | 206529 | PM | CMPLT | CMPLT | 2024-04-27 18:36 | 2024-03-21 18:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69562 | 206530 | PM | CMPLT | CMPLT | 2024-04-27 20:29 | 2024-03-21 20:29 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69178 | 206531 | PM | CMPLT | CMPLT | 2024-04-27 20:36 | 2024-03-21 20:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69258 | 206532 | PM | CMPLT | CMPLT | 2024-04-27 20:47 | 2024-03-21 20:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69632 | 206533 | PM | CMPLT | CMPLT | 2024-04-27 20:57 | 2024-03-21 20:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69064 | 206534 | PM | CMPLT | CMPLT | 2024-04-27 20:59 | 2024-03-21 20:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69061 | 206535 | PM | CMPLT | CMPLT | 2024-04-27 21:06 | 2024-03-21 21:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69034 | 206536 | PM | CMPLT | CMPLT | 2024-04-28 17:59 | 2024-03-21 17:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69665 | 206537 | PM | CMPLT | CMPLT | 2024-04-28 18:08 | 2024-03-21 18:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050260 | 206538 | PM | CMPLT | CMPLT | 2024-04-28 18:16 | 2024-03-21 18:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050269 | 206539 | PM | CMPLT | CMPLT | 2024-04-28 18:18 | 2024-03-21 18:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1058230 | 206540 | PM | CMPLT | CMPLT | 2024-04-28 18:22 | 2024-03-21 18:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050071 | 206541 | CM | CMPLT | CMPLT | 2024-04-29 08:08 | 2024-04-29 08:31 | Thermometer not working | (TC 04/29/2024 08:27:47 AM, Canjimavadivel, Kamallesh ) Repalced thermometer, tested ok Battery replacement overdue, replaced battery
541061 | 206542 | CM | CMPLT | CMPLT | 2024-04-29 08:48 | 2024-04-29 08:48 | not working | (TC 04/23/2024 08:48:49 AM, Huang, Jessica) replaced the SPO2 sensor and tested it working fine
120021 | 206543 | CM | CMPLT | CMPLT | 2024-04-29 08:50 | 2024-04-29 08:50 | not working | (TC 04/23/2024 08:49:54 AM, Huang, Jessica) replaced the SPO2 sensor and tested it working fine
541266 | 206544 | CM | CMPLT | CMPLT | 2024-04-29 08:51 | 2024-04-29 08:51 | not working | (TC 04/23/2024 08:54:46 AM, Huang, Jessica) replaced the SPO2 sensor
1046463 | 206545 | CM | CMPLT | CMPLT | 2024-04-29 08:56 | 2024-04-29 08:56 | not working | (TC 04/23/2024 08:55:52 AM, Huang, Jessica) replaced the SPO2 sensor
541307 | 206546 | CM | CMPLT | CMPLT | 2024-04-29 08:57 | 2024-04-29 08:57 | not working | (TC 04/23/2024 08:57:04 AM, Huang, Jessica) replaced the SPO2 sensor
1069393 | 206547 | CM | CMPLT | CMPLT | 2024-04-29 08:58 | 2024-05-02 14:34 | "Maintenance past due" message | (TC 04/29/2024 09:43:27 AM, Canjimavadivel, Kamallesh ) Confirmed issue upon startup, brought machine down to Biomed. Contacted vendor (TC 05/02/2024 02:32:47 PM, Canjimavadivel, Kamallesh ) Spoke with vendor and was provided instructions how to change the PM date. Changed schedule and brought machine back to ICU. Check other devices that were available and changed schedule as well
541061 | 206548 | CM | CMPLT | CMPLT | 2024-04-29 08:58 | 2024-04-29 08:58 | removed from service
1048046 | 206549 | CM | CMPLT | CMPLT | 2024-04-29 09:00 | 2024-04-29 09:00 | power connector broken | (TC 04/23/2024 08:59:43 AM, Huang, Jessica) shipped the device for repair
1063116 | 206550 | CM | CMPLT | CMPLT | 2024-04-29 09:01 | 2024-04-29 09:01 | not working | (TC 04/22/2024 09:01:43 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1069389 | 206551 | CM | CMPLT | CMPLT | 2024-04-29 09:02 | 2024-05-07 13:45 | "System failure" message | (TC 04/29/2024 09:44:34 AM, Canjimavadivel, Kamallesh ) Contacted vendor. Brought machine down to Biomed. Turned on and performed self tests, no alarms present. (TC 05/07/2024 01:37:20 PM, Canjimavadivel, Kamallesh ) Vendor came on-site and tested machine, all values within spec. Service report attached. Machiens taken back to ICU (TC 05/07/2024 01:43:18 PM, Canjimavadivel, Kamallesh ) Vendor came on-site and tested machine, all values within spec. Service report attached. Machiens taken back to ICU
1069394 | 206552 | CM | CMPLT | CMPLT | 2024-04-29 09:02 | 2024-05-07 13:42 | "System failure" message | (TC 04/29/2024 09:45:38 AM, Canjimavadivel, Kamallesh ) Contacted vendor. Brought machine down to Biomed. Turned on and performed self tests, no alarms present.
| 206553 | RDS | CMPLT | CMPLT | 2024-04-29 09:03 | 2024-04-29 14:03 | Weekly checks in CCU | (TC 04/29/2024 02:02:44 PM, Scero, Tomlyn) weekly checks complete. 30J tests passed.
1063114 | 206554 | CM | CMPLT | CMPLT | 2024-04-29 09:04 | 2024-04-29 09:04 | not working | (TC 04/23/2024 09:03:56 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1063115 | 206555 | CM | CMPLT | CMPLT | 2024-04-29 09:04 | 2024-04-29 09:04 | not working | (TC 04/23/2024 09:04:41 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1063123 | 206556 | CM | CMPLT | CMPLT | 2024-04-29 09:05 | 2024-04-29 09:05 | not working | (TC 04/23/2024 09:05:29 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1054878 | 206557 | CM | CMPLT | CMPLT | 2024-04-29 09:06 | 2024-04-29 09:06 | not working | (TC 04/24/2024 09:07:05 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1063120 | 206558 | CM | CMPLT | CMPLT | 2024-04-29 09:08 | 2024-04-29 09:08 | not working | (TC 04/24/2024 09:07:46 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1054884 | 206559 | CM | CMPLT | CMPLT | 2024-04-29 09:08 | 2024-04-29 09:08 | not working | (TC 04/24/2024 09:08:27 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1054887 | 206560 | CM | CMPLT | CMPLT | 2024-04-29 09:09 | 2024-04-29 09:09 | not working | (TC 04/24/2024 09:09:08 AM, Huang, Jessica) replaced the batteries and cleaned the sensor
1062964 | 206561 | CM | CMPLT | CMPLT | 2024-04-29 09:13 | 2024-04-29 09:13 | not working | (TC 04/24/2024 09:13:08 AM, Huang, Jessica) received the call in OR 8 and was informed the light source was not working took the light source out from the room and noticed the power cord was loose from the rear case tested it, working fine and sent it back to the OR desk
63507 | 206562 | CM | CMPLT | CMPLT | 2024-04-29 09:22 | 2024-04-29 09:22 | foot pedal not working | (TC 04/25/2024 09:22:32 AM, Huang, Jessica) informed Adeel and Samir the foot pedal was not working contacted the thrid party to require the quote for the foot pedal; received it and went to Reeba's office to discuss it
1057051 | 206563 | CM | ACTIV | ISSUE | 2024-04-29 09:33 |  | not working
| 206564 | CM | CMPLT | CMPLT | 2024-04-29 10:39 | 2024-04-29 10:39 | missing medical devices from MBU | (TC 04/24/2024 10:41:15 AM, Huang, Jessica) Shermaine sent email in regarded to discuss the missing devices in MBU went to the unit to discuss the issue
1061203 | 206565 | CM | CMPLT | CMPLT | 2024-04-29 10:43 | 2024-04-29 10:43 | the x-ray was not displaying on the monitors | (TC 04/24/2024 10:46:11 AM, Huang, Jessica) Amanda called from OR 2 and they were not able to display the image from X-Ray device to the monitors replaced the DVI cable and it did not fix the issue the connector on the x-ray machine was defective and informed Amanda to contact JDMI
1047042 | 206566 | CM | CMPLT | CMPLT | 2024-04-29 10:49 | 2024-04-29 10:52 | No error stated on biomed tag | (TC 04/29/2024 10:52:08 AM, Baghirzade, Nara) Downloaded error logs; no errors to note. During check-in test, the pressure was higher than nominal range. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
| 206567 | CM | CMPLT | CMPLT | 2024-04-29 10:51 | 2024-04-29 10:52 | MSJ OPG SPP C2 REMOTE WALL MONITOR HOOK UP; TELECONFERENCE HOOKUP ASSISTANCE | (TC 04/29/2024 10:51:47 AM, Whittington, Scott) MSJ OPG SPP C2 REMOTE WALL MONITOR HOOK UP TELECONFERENCE HOOKUP ASSISTANCE | MSJ OPG SPP C2 REMOTE WALL MONITOR HOOK UP TELECONFERENCE HOOKUP ASSISTANCE
68633 | 206568 | CM | CMPLT | CMPLT | 2024-04-29 10:53 | 2024-04-29 10:55 | Defective bottle-side pressure sensor | (TC 04/29/2024 10:53:08 AM, Baghirzade, Nara) Downloaded error log; noted error 240.4150, high voltage on bottle-side pressure sensor. Ran analog test and confirmed the error. Replaced bottle-side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared.
67913 | 206569 | CM | CMPLT | CMPLT | 2024-04-29 10:55 | 2024-04-29 10:56 | Broken porthole latch | (TC 04/29/2024 10:55:10 AM, Baghirzade, Nara) Replaced porthole latch on west door. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74%RH Incubator is fully functional.
| 206570 | RDS | CMPLT | CMPLT | 2024-04-29 10:57 | 2024-04-29 11:06 | Weekly Checks For PACU | (TC 04/29/2024 11:05:06 AM, Reid, Emmerson ) Completed procedure. Went around and turned on MX800 monitors that were off.
67957 | 206571 | CM | CMPLT | CMPLT | 2024-04-29 10:58 | 2024-04-29 10:58 | Broken door hinge and worn out porthole seals | (TC 04/29/2024 10:57:42 AM, Baghirzade, Nara) Replaced west door hinge. Replaced both seals on east door. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74.5%RH Incubator is fully functional.
| 206572 | RDS | CMPLT | CMPLT | 2024-04-29 11:10 | 2024-04-29 11:17 | Weekly check for ED
1066460 | 206573 | FC | CMPLT | CMPLT | 2024-04-29 11:23 | 2024-04-29 11:23 | QC checklist | (TC 04/29/2024 11:23:07 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North as per Sarabeth's request
62500 | 206574 | CM | CMPLT | CMPLT | 2024-04-29 12:56 | 2024-04-29 12:58 | IAD leg wheels lossing | (TC 04/29/2024 12:55:57 PM, Chukri, Samir) Repaired the wheel and check the rest of the leg wheels working fine. (TC 04/29/2024 12:56:53 PM, Ferreira, Qwynn) Helping samir for the repair
1047851 | 206575 | CM | CMPLT | CMPLT | 2024-04-29 13:44 | 2024-04-29 13:46 | Motor issue reported | (TC 04/29/2024 01:44:10 PM, Baghirzade, Nara) No issues or errors when booting up. Ran a functional check on different rates; no errors with the motor. Ran occlusion test; pump occluded and alarmed with no issue. Ran an infusion for the second time at varying rates, no errors observed.
1055974 | 206576 | CM | CMPLT | CMPLT | 2024-04-29 13:47 | 2024-04-29 13:47 | Battery replacement | (TC 04/29/2024 01:47:01 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055915 | 206577 | CM | CMPLT | CMPLT | 2024-04-29 13:47 | 2024-04-29 13:48 | Damaged power cord | (TC 04/29/2024 01:47:37 PM, Baghirzade, Nara) Power cord prongs were damaged; replaced power cord. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 206578 | CM | CMPLT | CMPLT | 2024-04-29 13:48 | 2024-04-29 13:49 | PM huddle | (TC 04/29/2024 01:48:32 PM, Baghirzade, Nara) PM huddle biomed
1064198 | 206579 | CM | CMPLT | CMPLT | 2024-04-29 13:57 | 2024-05-09 16:56 | Error 1591 | (TC 04/29/2024 01:57:01 PM, Cortez, John) Checked service manual, error 1591 refers to a "fault with a heaterbase PCB" Checked alarm history and retrieved logs, updated software to 6.1 Unable to view logs on PC. Kam reached out to FP rep (TC 05/07/2024 04:54:54 PM, Cortez, John) Ran humidifier, no issues. Returned to ICU
| 206580 | CM | CMPLT | CMPLT | 2024-04-29 14:14 | 2024-04-29 14:14 | meeting | meeting
| 206581 | CM | CMPLT | CMPLT | 2024-04-29 14:17 | 2024-04-29 14:17 | worked on TMS | worked on TMS
| 206582 | RDS | CMPLT | CMPLT | 2024-04-29 14:19 | 2024-04-29 14:19 | weekly check for 14S and 17M | weekly check for 14S and 17M
1059778 | 206583 | FC | CMPLT | CMPLT | 2024-04-29 14:19 | 2024-04-29 14:21 | QC checklist | (TC 04/29/2024 02:20:17 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
| 206584 | CM | CMPLT | CMPLT | 2024-04-29 14:51 | 2024-05-03 11:19 | ICU incident, Alaris LVP, SAFER #74195 | (TC 04/29/2024 03:19:43 PM, Cortez, John) Reviewed SAFER report. LVP #69440 is listed in report, but event logs for this pump dont match description of incident on listed date (Apr22). Last event for this pump occurred on Apr 5. However, 4 other pumps were sent down from ICU April 24, and one the pump's event logs might match events. (biomed report is attached) Completed check-in successfully on all pumps. However, with given information, unable to confirm which pump set was stuck on the door of one of the other pumps. (TC 05/02/2024 11:16:02 AM, Cortez, John) Forwarded report to Marlene
1064311 | 206585 | CM | CMPLT | CMPLT | 2024-04-29 14:54 | 2024-04-29 14:54 | 6 year PM | (TC 04/29/2024 02:53:48 PM, Scero, Tomlyn) 6 year PM parts replaced and tests complete and passed.
64260 | 206586 | CM | CMPLT | CMPLT | 2024-04-29 15:36 | 2024-04-30 09:03 | Touch screen not working | (TC 04/29/2024 03:36:34 PM, Rhuland, Courtney) Went up and looked at monitor and touch screen was unresponsive Brought it to biomed and tried calibrating screen and was unable to do so Replaced monitor with biomed spare with NICU configuration for overnight (TC 04/30/2024 09:01:54 AM, Rhuland, Courtney) Replaced touch screen with spare Calibrated touch Swapped monitors back and verified connectivity with central monitor
32520 | 206587 | CM | CMPLT | CMPLT | 2024-04-29 15:42 | 2024-09-25 09:28 | MSH DR2 DETECTOR | (TC 04/29/2024 03:42:33 PM, Whittington, Scott) msh dr2 detector | MSH DR2 DETECTOR
1064352 | 206588 | CM | CMPLT | CMPLT | 2024-04-29 15:46 | 2024-04-30 09:22 | MSH ICU SPARQ POD A SYSTEM CRASHED | (TC 04/29/2024 03:46:44 PM, Whittington, Scott) MSH ICU SPARQ POD A SYSTEM CRASHED | MSH ICU SPARQ POD A SYSTEM CRASHED
1049902 | 206589 | CM | CMPLT | CMPLT | 2024-04-29 15:59 | 2024-04-29 16:00 | MSH R62 DETECTOR BATTERY CONTACTS AND DETECTOR BATTERY CHARGER CONTACTS | (TC 04/29/2024 04:00:25 PM, Dolinar, John) MSH R62 DETECTOR BATTERY CONTACTS AND DETECTOR BATTERY CHARGER CONTACTS | MSH R62 DETECTOR BATTERY CONTACTS AND DETECTOR BATTERY CHARGER CONTACTS
| 206590 | PA | CMPLT | CMPLT | 2024-04-30 08:50 | 2024-06-28 12:20 | Isopropyl Alcohol 99%
| 206591 | CM | CMPLT | CMPLT | 2024-04-30 09:26 | 2024-05-07 08:56 | V500 battery order for PMs | (TC 04/30/2024 09:29:00 AM, Scero, Tomlyn) Ordered batteries on Ereq (TC 05/07/2024 08:55:42 AM, Scero, Tomlyn) batteries received
| 206592 | CM | CMPLT | CMPLT | 2024-04-30 09:33 | 2024-05-13 15:11 | Hamilton Vent battery order + pm kit | (TC 04/30/2024 10:01:49 AM, Cortez, John) Ordered 4x (TC 04/30/2024 10:34:54 AM, Cortez, John) Ordered 2x MR1 PM kits (TC 05/03/2024 04:29:03 PM, Cortez, John) PM kits arrived (TC 05/13/2024 03:08:59 PM, Cortez, John) batteries arrived
1070774 | 206593 | CM | CMPLT | CMPLT | 2024-04-30 09:51 | 2024-04-30 09:51 | new incoming | (TC 04/23/2024 09:51:09 AM, Huang, Jessica) unboxed the parts and assembled them new incoming inspection
63801 | 206594 | CM | CMPLT | CMPLT | 2024-04-30 10:37 | 2024-05-29 10:02 | Inspection/PM after Storage | Unit was being stored in 6th floor storage. NB reached out for this unit to be put into service in 12N. Went and brought the unit down to biomed. Cleaned unit.Not familier with the equipment. Took a short crash course from T drive document. Found many accesstories missing including lead wires, PIM interface cable, PIM mount, barcode scanner. Found battery due for replacement. Reached out to vendor to inquire about part# nad pricing. May 01 - Foudn some accessories in biomed stock. Went to storage to get a spare PIM mount. Installed it. May 02 - Followed up on with vendor about parts info. received. SC Created. Awaiting delivery. May 29 - Parts received. Installed. Configured ECG machine. Performed diagnostic test. Checked functionlity. Tested ok. Delivered unit to 12N.
| 206595 | PA | CMPLT | CMPLT | 2024-04-30 10:40 | 2024-05-10 14:06 | Formlabs Order, wash and resin
1064356 | 206596 | CM | CMPLT | CMPLT | 2024-04-30 10:59 | 2024-05-03 07:54 | V1-V6 leads not working | (TC 04/30/2024 11:18:39 AM, Canjimavadivel, Kamallesh ) V4-V6 chest lead had intermittent signal loss. Replaced and tested ok.
| 206597 | CM | CMPLT | CMPLT | 2024-04-30 13:05 | 2024-04-30 14:08 | On-call daily checks | (TC 04/30/2024 02:07:48 PM, Baghirzade, Nara) Completed on-call daily checks
68688 | 206598 | PM | CMPLT | CMPLT | 2024-04-30 13:30 | 2024-04-30 13:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 04/30/2024 01:29:46 PM, Rhuland, Courtney) Sent out for PM - completed by BD Loaded network and data set
65962 | 206599 | PM | CMPLT | CMPLT | 2024-04-30 13:32 | 2024-04-30 13:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1055947 | 206600 | CM | CMPLT | CMPLT | 2024-04-30 13:41 | 2024-05-14 12:56 | when turned on , it shuts itself off and alarming | "when turned on , it shuts itself off and alarming" getting a PO and shipping to Masimo (TC 05/14/2024 12:55:36 PM, Baghirzade, Nara) Received replacement Root. PO 7500205987
502044 | 206601 | RQ | CMPLT | CMPLT | 2024-04-30 13:42 | 2024-05-26 10:40 | maintenance due May26th warning ERGO | (TC 05/26/2024 10:40:26 AM, Green, Paul) This was done, but the date was not changed on system. Changed PM date on system. | maintenance due May26th warning ERGO
1055939 | 206602 | CM | CMPLT | CMPLT | 2024-04-30 13:47 | 2024-04-30 13:48 | BP not working | (TC 04/30/2024 01:46:56 PM, Ferreira, Qwynn) Tested the BP and heard a "hiss" noise. BP tube connecting to the masimo had a tiny hole. Replaced tubing and tested - Working as normal
| 206603 | CM | CMPLT | CMPLT | 2024-04-30 13:48 | 2024-05-01 13:08 | On-call daily check | (TC 05/01/2024 01:07:21 PM, Baghirzade, Nara) Completed on-call daily checks
1058242 | 206604 | PM | CMPLT | CMPLT | 2024-04-30 14:00 | 2024-04-30 14:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
| 206605 | CM | ACTIV | ISSUE | 2024-04-30 14:29 |  | Upgrading Expair from version 1.34 to 1.35 | (TC 04/30/2024 02:29:33 PM, Baghirzade, Nara) Went up to CPET; copied over the most recent upgrade file for version 1.35 on to a USB. Using the expair on my laptop for a test run. Took the network file off the expair folder in c drive and copied to desktop. Renamed the current expair folder. Ran the upgrade file for expair. Successfully created a new expair folder on C drive however when trying to open expair, there are configuration and database errors. Tried to run several times but running into errors. Reverted expair back to version 1.34 and renamed the expair folders. Will follow up with Chris from Novus regarding this. (TC 05/02/2024 11:12:39 AM, Baghirzade, Nara) Copied the Expair folder from C:/ drive from the CPET device. Dropped it on my laptop C:/ drive; completed the upgrade from 1.34 to 1.35. Expair works and the predicteds tab is no longer greyed out. Went up to PFT lab and copied the Expair folder from Hypair. Completed the same steps on my laptop. Was able to upgrade the expair software and use the new predicteds. Will update Novus and proceed with preparing a plan to upgrade all stations and review stations. (TC 10/22/2024 02:10:21 PM, Baghirzade, Nara) Tried upgrading again however getting the same DB errors and unable to boot up expair. Reached out to Chris from Novus for remote support. (TC 10/29/2024 02:01:36 PM, Baghirzade, Nara) Coordinating the upgrade with Gaby, Steve, and physicians. (TC 11/27/2024 08:49:58 AM, Baghirzade, Nara) Sent out email to all physicians regarding upgrade. Coordinated with helpdesk/IT for admin rights to all computers. Picked up the remote laptop from PFT. (TC 11/28/2024 07:43:29 AM, Baghirzade, Nara) Created 2 USBs with installation files for expair upgrade and WinSpiro software, and predicteds file. Checked the predicted file on expair to ensure that it is the updated 2021 GLI volumes.
63438 | 206606 | CM | CMPLT | CMPLT | 2024-04-30 14:37 | 2024-04-30 14:38 | Reported that it errors during feed | (TC 04/30/2024 02:36:48 PM, Baghirzade, Nara) Found an issue with the battery and the power supply (defective, intermittently charging); replaced both battery and power supply. Ran a feed for about 3 hours at varying rates; no errors observed during the feed. Pump is functional.
1051686 | 206607 | CM | CMPLT | CMPLT | 2024-04-30 14:41 | 2024-07-30 15:23 | Printer and touchscreen not working | (TC 04/30/2024 02:41:08 PM, Baghirzade, Nara) Checked printer tray, paper is correctly in and no paper jams. During troubleshooting of the printer tray, the touchscreen stopped working. Reached out to Scott from bomimed for further troubleshooting. (TC 05/06/2024 02:03:49 PM, Baghirzade, Nara) Sent follow up email to Mike from Bomimed (TC 05/28/2024 01:36:25 PM, Baghirzade, Nara) Sent a follow up email to Scott. (TC 07/30/2024 03:22:33 PM, Baghirzade, Nara) Followed up with Steve and he does not want to proceed with the repair. Unit will be used for parts.
1051351 | 206608 | CM | CMPLT | CMPLT | 2024-04-30 16:04 | 2024-05-03 15:46 | Retire From Service | Call re: not working. Wnet to check and brought it down to biomed. Cleaned unit. Unit on consignment. Checked functionality. FOund filter due for replacement. Contacted vendor for info. Vendor proposed that they will replace it with a new model under no charge. Awaiting replament. MAy 03 - Followed up re: returning this old unit. Was advised to dispose the unit on site. Removed all stickers and accessories. Retired and disposed the unit.
| 206609 | CM | CMPLT | CMPLT | 2024-04-30 16:27 | 2024-04-30 16:27 | Site coverage.
51580 | 206610 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52232 | 206611 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-08-14 11:27 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/14/2024 11:26:39 AM, Rhuland, Courtney) Replaced: Rechargeable 9V battery Side wall latch Flip door Inside wall Calibrated: VREF & MAINS voltage Scale Functional check - OK Set 35C, 75% humidity Read 35C, 76.3% humidity Scale - OK Electrical safety - OK Ground resistance - 128 mOhms Leakage current - 65 uA
52244 | 206612 | PM | CANCL | CANCL | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
53007 | 206613 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
53009 | 206614 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
53013 | 206615 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
62226 | 206616 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-10 12:56 | PM-MAJOR - Annual PM For: Simulators, Multiparameter | (TC 06/10/2024 12:56:17 PM, Chukri, Samir) Performed calibration by Getinge Tech , Please see attached.
62369 | 206617 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-10 13:09 | PM-MAJOR - Annual PM For: Simulators, Transducer Output, Invasive Blood Pressure | (TC 06/10/2024 01:08:37 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached.
62974 | 206618 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-19 09:47 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/19/2024 09:46:45 AM, Rhuland, Courtney) PM complete Light intensity check - OK Electrical safety - OK
62978 | 206619 | PM | CANCL | CANCL | 2024-05-01 03:27 | 2024-07-31 07:25 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62979 | 206620 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-19 09:49 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/19/2024 09:48:55 AM, Rhuland, Courtney) PM complete Light intensity check - OK Electrical safety - OK
62980 | 206621 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-19 09:47 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/19/2024 09:47:30 AM, Rhuland, Courtney) PM complete Light intensity check - OK Electrical safety - OK
62983 | 206622 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-09-13 10:09 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 09/13/2024 10:09:22 AM, Alvarado, Karla) Timer was defective; replaced timer 11619.4 hours Intensity light: - High: 33.7 uW/cm2/nm - Low:14.4 uW/cm2/nm Target light is functional GND Resistance: 0.142 ohms Leakage current: 49.7 uA PM completed.
62986 | 206623 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-19 09:38 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/19/2024 09:38:08 AM, Rhuland, Courtney) PM complete Light intensity check - OK Electrical safety - OK
62987 | 206624 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-10-02 11:50 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/02/2024 11:46:32 AM, Alvarado, Karla) Replaced timer. Rotated/resited fan (left side). 0.0 hours Intensity light: - High: 34 uW/cm2/nm - Low: 13.8 uW/cm2/nm Target light is functional GND Resistance: 0.170 ohms Leakage current: 47.8 uA PM completed. (TC 10/02/2024 11:49:32 AM, Rhuland, Courtney) Supervised
62988 | 206625 | PM | CANCL | CANCL | 2024-05-01 03:27 | 2024-07-17 10:47 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62990 | 206626 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-19 09:48 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/19/2024 09:48:10 AM, Rhuland, Courtney) PM complete Light intensity check - OK Electrical safety - OK
63010 | 206627 | PM | CANCL | CANCL | 2024-05-01 03:27 | 2024-06-27 09:59 | PM-MAJOR - Annual PM For: Radiometer | its is for the old Phototherapy lights.
63036 | 206628 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-05 08:53 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/05/2024 08:52:30 AM, Rhuland, Courtney) PM complete Irradiance - 55.4 uW/cm^2/nm Hours - 7328 Leakage current - 13.2 uA Ground resistance - 47 mOhms
63175 | 206629 | PM | CNP | ENL | 2024-05-01 03:27 | 2024-06-27 09:48 | PM-MAJOR - Annual PM For: Simulators, Cardiac, Electrocardiography | (TC 06/27/2024 09:47:56 AM, Chukri, Samir) Peter from Getinge not located the tester during the company performed the certificate
63276 | 206630 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-25 08:05 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/25/2024 08:04:55 AM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
63447 | 206631 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR-Annual PM for: Oscilloscope
63712 | 206632 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-10 13:13 | PM-MAJOR- Annual PM for: Tester Electrical, Safety | (TC 06/10/2024 01:13:16 PM, Chukri, Samir) Performed Calibration by Getinge tech. calibration certificate attached.
63726 | 206633 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-10 13:19 | PM-MAJOR - Annual PM For: Testers | (TC 06/10/2024 01:18:46 PM, Chukri, Samir) Performed calibration by Getinge tech, Calibration certificate attached.
64920 | 206634 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-11 09:17 | PM-MAJOR - Annual PM For: multimeter | (TC 06/11/2024 09:17:11 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration Certificate attached.
64997 | 206635 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-18 11:56 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/18/2024 11:56:07 AM, DeGuzman, Leila) PM completed. All tests passed.
65117 | 206636 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-11 09:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/11/2024 09:05:58 AM, DeGuzman, Leila) PM completed. See procedures for details.
65118 | 206637 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65119 | 206638 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65122 | 206639 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65128 | 206640 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65130 | 206641 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65149 | 206642 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65150 | 206643 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65152 | 206644 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65156 | 206645 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65157 | 206646 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65162 | 206647 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65166 | 206648 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-28 08:58 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/28/2024 08:58:19 AM, Reid, Emmerson ) PM Complete, see procedure for details.
65168 | 206649 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-05-13 10:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/13/2024 10:46:06 AM, Reid, Emmerson ) PM complete, see procedure for details.
65171 | 206650 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65172 | 206651 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65864 | 206652 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-25 08:00 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/25/2024 07:59:57 AM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
65868 | 206653 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-25 07:58 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/25/2024 07:57:31 AM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
65952 | 206654 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-05-02 14:54 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 05/02/2024 02:54:03 PM, Rhuland, Courtney) Replaced front case Replaced mylar gasket Calibrated syringe Check-in passed Refer to ASM for details
65973 | 206655 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-25 11:48 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
66004 | 206656 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66345 | 206657 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-11 09:24 | PM-MAJOR - Annual PM For: Testers | (TC 06/11/2024 09:24:16 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration Certificate attached.
66453 | 206658 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-28 09:35 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced V-battery and air filter. Run battery check. 220 minutes, 19 Ah. Tested. Functional. SW2.51. Leakage: 84uA and 89ohm.
66703 | 206659 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-28 08:05 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/28/2024 08:05:14 AM, Rhuland, Courtney) Repaired side door hinge Repaired 2 porthole doors PM completed Replaced 9V rechargeable battery Calibrated: Scale Main & LV voltage RH Functional check - OK Set 35C & 75% humidity Read 35C & 75.8% humidity Electrical safety - OK Current leakage - 89.2 uA Ground resistance - 92 mOhms
66760 | 206660 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66764 | 206661 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66765 | 206662 | PM | ACTIV | READY | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66815 | 206663 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-09-26 08:34 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/26/2024 08:34:18 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
66827 | 206664 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-18 11:11 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/18/2024 11:11:14 AM, DeGuzman, Leila) PM completed. See procedures for details.
66834 | 206665 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-08-28 10:46 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/28/2024 10:46:25 AM, DeGuzman, Leila) PM passed and completed. NIBP calibrated. See procedure for details.
66835 | 206666 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-30 14:48 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/09/2024 09:09:52 AM, Alvarado, Karla) PM Completed. All tests passed See Procedure for details. (TC 09/09/2024 09:12:30 AM, Rhuland, Courtney) Supervised Karla for PM
66837 | 206667 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-09-26 08:14 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/26/2024 08:13:55 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
66839 | 206668 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66840 | 206669 | PM | CANCL | CANCL | 2024-05-01 03:27 |  | BP Check
66848 | 206670 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66850 | 206671 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-08-14 09:36 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/14/2024 09:36:19 AM, DeGuzman, Leila) PM passed and completed. Calibrated NIBP See procedure for details.
66861 | 206672 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67204 | 206673 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-09-09 10:02 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/09/2024 09:59:46 AM, Alvarado, Karla) PM completed All testes passed See procedure for details (TC 09/09/2024 10:02:26 AM, Rhuland, Courtney) Supervised Karla for PM
67760 | 206674 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67762 | 206675 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67767 | 206676 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67768 | 206677 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67769 | 206678 | PM | ACTIV | ISSUE | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67885 | 206679 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-06-14 13:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/14/2024 01:54:48 PM, DeGuzman, Leila) PM completed. All tests passed
67915 | 206680 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-07-31 14:50 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/31/2024 02:49:32 PM, Rhuland, Courtney) PM completed Replaced porthole door Replaced middle porthole hinge (x2) Replaced rechargeable 9V battery Calibrated: LV and MAINS voltage Scale Functional check - OK Set - 35C & 75% humidity Read 35C & 74.8% humidity Electrical safety - OK Leakage current - 86.7uA Ground resistance - 123mOhms
67947 | 206681 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-09-04 11:58 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/04/2024 11:58:13 AM, Baghirzade, Nara) Retrieved the unit from NICU and brought it to biomed. Worked with the coop student, Carla, on the PM. Demonstrated all PM steps and explained the PM process. Replaced power cord as it was damaged. Replaced hinge on door. PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 83.0 uA Power cord resistance = 110 mohms
67954 | 206682 | PM | CANCL | CANCL | 2024-05-01 03:27 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
67956 | 206683 | PM | CANCL | CANCL | 2024-05-01 03:27 | 2024-06-04 10:10 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
67963 | 206684 | PM | CMPLT | CMPLT | 2024-05-01 03:27 | 2024-08-07 12:29 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/07/2024 12:28:52 PM, Rhuland, Courtney) Replaced: Red humidifier gasket Reservoir switch Side door latch Rechargeable 9V battery Calibrated: Mains & VREF voltage Scale RH Functional check - OK Set 35C, 75% humidity Read 35C, 75.3% humidity Electrical safety - OK Leakage current - 93.2 uA Ground resistance - 108 mOhms
68301 | 206685 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 12:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68376 | 206686 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68793 | 206687 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68830 | 206688 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-03 13:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:55:03 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered Transferred network configuration & data set
68872 | 206689 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68874 | 206690 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68898 | 206691 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-22 11:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/22/2024 11:25:40 AM, Baghirzade, Nara) Battery was replaced this year. PM passed and completed. Electrical Safety - Ground resistance: 108 mohms Leakage current: 65.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
68924 | 206692 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 14:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68979 | 206693 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-23 12:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/22/2024 11:29:10 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 205 mohms Leakage current: 78.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 05/23/2024 12:12:19 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3762 mamphr PM completed.
68981 | 206694 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 14:37 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68985 | 206695 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-22 11:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/21/2024 02:11:06 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 167 mohms Leakage current: 63.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 05/22/2024 11:20:49 AM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3761 mamphr PM completed.
68989 | 206696 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 14:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69006 | 206697 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69016 | 206698 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-23 12:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/23/2024 12:15:05 PM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 110 mohms Leakage current: 72.6 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 05/23/2024 12:15:40 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3930 mamphr PM completed.
69083 | 206699 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69200 | 206700 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-09 14:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | By BD
69212 | 206701 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-09 14:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | By BD
69316 | 206702 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-09 12:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:59:09 PM, Patel, Rutvik) by BD
69401 | 206703 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 15:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69418 | 206704 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-16 11:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/16/2024 11:02:21 AM, DeGuzman, Leila) PM completed - all tests passed, cleaned.
69514 | 206705 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-11-21 11:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/21/2024 11:02:20 AM, Cortez, John) Updated software PM completed successfully, results logged in ASM
69519 | 206706 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-08 14:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/08/2024 02:46:45 PM, Rhuland, Courtney) Replaced door spring PM completed All tests passed Refer to ASM for details
69557 | 206707 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-15 15:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/15/2024 03:09:55 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69619 | 206708 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-03 13:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:48:40 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered
32520 | 206709 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-17 14:11 | MSH DR2 PM - Semi-Annual | (TC 06/17/2024 02:11:03 PM, Whittington, Scott) MSH DR2 PM | MSH DR2 PM
32520 | 206710 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-02 21:30 | MSH DR2 HARP - Semi Annual | (TC 06/02/2024 12:29:00 PM, Whittington, Scott) MSH DR2 HARP - Semi Annual (TC 06/17/2024 02:12:33 PM, Whittington, Scott) MSH DR2 HARP FOLLOWUP (TC 06/02/2024 12:29:00 PM, Whittington, Scott) MSH DR2 HARP - Semi Annual | MSH DR2 HARP - Semi Annual
32540 | 206711 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-10 15:38 | PM - Monthly | PM completed
32560 | 206712 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-04 15:35 | PM - Monthly | PM completed Mechanical PM done sc-1437708 TOS phantom a(-991.0,10.3) b(333.3,14.0) c(133.4,13.3) d(96.4,12.9) e(-107.4,11.0) f(-0.2,12.9) -passed all phantom tests
32580 | 206713 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-05 15:23 | PM - Monthly | PM done Side cover and meshes removed to clean inside Error log collected from GTS for Gantry Rotation Error - Servo Amp ordered - Rotational Slice count - 249310 TOS phantom a(-992.5, 10.1) b(333.1, 14.3) c(134.0, 14.4) d(96.9, 13.2) e(-106.7, 13.9) f(1.2, 13.6) -System passed all tests
501640 | 206714 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-26 14:12 | PM - Annual | (TC 07/26/2024 02:12:17 PM, Green, Paul) Performed a PM on system. Fully functional.
1046457 | 206715 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 09:33 | PM-MAJOR-Annual-PM for : Pressure meter | (TC 06/11/2024 09:33:04 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached.
1046458 | 206716 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 10:44 | Annual PM | (TC 06/11/2024 10:44:18 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1046612 | 206717 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1046781 | 206718 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-09-20 17:51 | PM - Semi-Annual | PM was done.
1046781 | 206719 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-09-27 18:52 | HARP - Semi Annual | HARP
1046789 | 206720 | PM | CANCL | CANCL | 2024-05-01 03:28 |  | Annual PM | Laser end of support, no CM or PM available. In process of acquiring a new laser as its replacement.
1047860 | 206721 | PM | CANCL | CANCL | 2024-05-01 03:28 |  | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe
1048377 | 206722 | PM | CANCL | CANCL | 2024-05-01 03:28 |  | PM
1048390 | 206723 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-22 10:08 | PM-MAJOR-Annual PM For: Multimeter | (TC 06/04/2024 10:08:30 AM, Chukri, Samir) Performed PM and certificate by Getinge Tech -Please see attached
1048603 | 206724 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM
1049204 | 206725 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-19 14:07 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/19/2024 02:07:28 PM, Canjimavadivel, Kamallesh ) Rebooted to finish patch install Vacuumed PC
1049215 | 206726 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-19 14:08 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/19/2024 02:08:26 PM, Canjimavadivel, Kamallesh ) Rebooted to finish patch install Vacuumed PC
1049217 | 206727 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-02 14:48 | PM-MAJOR - Annual PM For: Central Monitor | (TC 08/02/2024 02:47:32 PM, Scero, Tomlyn) SW patch done and PM complete
1049218 | 206728 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-02 14:47 | PM-MAJOR - Annual PM For: Central Monitor | (TC 08/02/2024 02:46:40 PM, Scero, Tomlyn) SW patch done and PM complete
1049221 | 206729 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-01 14:37 | PM-MAJOR - Annual PM For: Central Monitor | (TC 08/01/2024 02:36:53 PM, Scero, Tomlyn) PM complete during SW patch and reboot
1049222 | 206730 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-02 14:46 | PM-MAJOR - Annual PM For: Central Monitor | (TC 08/02/2024 02:45:44 PM, Scero, Tomlyn) SW patch complete. Restarted CPU. PM complete
1049410 | 206731 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Radiometer
1049901 | 206732 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-21 14:19 | HARP - Semi Annual | (TC 06/21/2024 02:19:08 PM, Chau, Ken) HARP Completed
1049901 | 206733 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-21 14:18 | PM | (TC 06/21/2024 02:18:15 PM, Chau, Ken) Replaced powercord reel Replaced side handles Replaced column sensor and broken cables PM completed
1049902 | 206734 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-25 08:46 | HARP - Semi Annual | (TC 06/25/2024 08:46:17 AM, Fan, Tak) HARP completed. Uploaded to sharepoint.
1049902 | 206735 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-03 11:33 | PM | (TC 07/03/2024 11:33:00 AM, Fan, Tak) PM completed.
1049903 | 206736 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-27 11:04 | HARP - Semi Annual | (TC 06/27/2024 11:04:24 AM, Fan, Tak) HARP completed and uploaded.
1049903 | 206737 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-27 11:05 | PM | (TC 06/27/2024 11:05:18 AM, Fan, Tak) PM Completed. Will follow up for wearpad and drive motor replacement.
1049904 | 206738 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-17 08:29 | HARP - Semi Annual | (TC 06/17/2024 08:29:00 AM, Fan, Tak) HARP completed.
1049904 | 206739 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-17 08:28 | PM | (TC 06/17/2024 08:28:00 AM, Fan, Tak) PM as per checklist. Replaced 3 det corners, wearpads.
1049905 | 206740 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-25 14:23 | HARP - Semi Annual | (TC 06/25/2024 02:22:52 PM, Fan, Tak) Completed. Uploaded.
1049905 | 206741 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-25 14:23 | PM | (TC 06/25/2024 02:23:13 PM, Fan, Tak) PM completed. Detector calibrated.
1050259 | 206742 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-13 08:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1050576 | 206743 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 12:37 | PM-MAJOR- Annual PM for : Scale, Tester | (TC 06/11/2024 12:36:49 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1051356 | 206744 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
1051357 | 206745 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-10-02 08:46 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/02/2024 08:45:58 AM, Alvarado, Karla) Timer was defective; replaced timer 8452.2 hours Intensity light: - High: 33.2 uW/cm2/nm - Low:14.5 uW/cm2/nm Target light is functional GND Resistance: 0.057 ohms Leakage current: 57.7 uA PM completed.
1051359 | 206746 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-03 08:38 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 06/03/2024 08:37:35 AM, Rhuland, Courtney) PM completed. All tests passed. Fan filters vacuumed Replaced cord retainer Hours: 6637.1 Ground Resistance: 101 mOhms Leakage Current: 50.7 uA High Iridescence: 31.4 uW/cm^2/nm Low Iridescence: 12.6 uW/cm^2/nm
1051613 | 206747 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 10:52 | PM-MAJOR - Annual PM For: Testers | (TC 06/11/2024 10:52:30 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1051830 | 206748 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PLANNED MAINTENANCE-VACUUM PC
1051831 | 206749 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Central Monitor
1051840 | 206750 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR - PM For Central Monitor PC
1053213 | 206751 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 10:58 | PM-MAJOR - Annual PM For: Testers | (TC 06/11/2024 10:58:25 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1053384 | 206752 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 12:32 | PM-MAJOR-Annual PM for: Scale, tester | (TC 06/11/2024 12:32:02 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1053439 | 206753 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 11:06 | PM-MAJOR - Annual PM For: Simulators, Multiparameter | (TC 06/11/2024 11:05:47 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1053440 | 206754 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 11:08 | PM-MAJOR - Annual PM For: Testers, Pulse Oximetry Monitor | (TC 06/11/2024 11:07:50 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1053441 | 206755 | PM | CNP | ENL | 2024-05-01 03:28 | 2024-06-27 09:52 | PM-MAJOR - Annual PM For: Testers, Pulse Oximetry Monitor | (TC 06/27/2024 09:51:39 AM, Chukri, Samir) Peter not located the unit during the certificated time.
1053447 | 206756 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR-Stretchers, Mobile, Hospital, Patient Transfer
1053475 | 206757 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 11:10 | Annual PM | (TC 06/11/2024 11:09:58 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1054591 | 206758 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR-Annual Philips Monitoring, Central Station
1054599 | 206759 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-09 10:40 | PM MAJOR-Annual Philips Monitoring, Central Station | (TC 05/09/2024 10:39:44 AM, Baghirzade, Nara) PM completed.
1055771 | 206760 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
1055774 | 206761 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-12 10:23 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/12/2024 10:22:31 AM, Canjimavadivel, Kamallesh ) All tests pass PM complete
1055793 | 206762 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-05 16:41 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/05/2024 04:41:18 PM, Cortez, John) PM passed, see procedure for results
1055794 | 206763 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-11 14:21 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/10/2024 11:18:41 AM, Cortez, John) Waiting to do IBP test, tester sent out for maintenance. Rest of PM passed, see procedure for results (TC 07/11/2024 02:21:16 PM, Cortez, John) Checked IBP functionality
1055795 | 206764 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-11 14:23 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/25/2024 04:22:49 PM, Cortez, John) IBP test on hold, tester away for maintenance Rest of PM passed, see procedure for results (TC 07/11/2024 02:22:45 PM, Cortez, John) Checked IBP functionality
1055796 | 206765 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-11 14:25 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/20/2024 04:09:34 PM, Cortez, John) Unable to test IBP, Rigel tester out for maintenance. Will come back and do test. Rest of PM passed, see procedure for results (TC 07/11/2024 02:24:51 PM, Cortez, John) Checked IBP functionality
1055797 | 206766 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-06 16:17 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/06/2024 04:16:39 PM, Cortez, John) PM passed, see procedure for results
1055798 | 206767 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 11:02 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 08/16/2024 11:02:14 AM, Canjimavadivel, Kamallesh ) All tests pass PM complete
1055802 | 206768 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-09 08:14 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/09/2024 08:14:27 AM, Canjimavadivel, Kamallesh ) Finished up PM and returned device to unit
1055803 | 206769 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-05 11:14 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/05/2024 11:14:18 AM, DeGuzman, Leila) PM completed.
1055804 | 206770 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-05 14:03 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/05/2024 02:02:47 PM, DeGuzman, Leila) PM completed.
1055805 | 206771 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-05 09:48 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/05/2024 09:47:49 AM, Canjimavadivel, Kamallesh ) PM complete
1055806 | 206772 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-05 14:14 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/05/2024 02:14:09 PM, Cortez, John) PM passed, see procedure for results
1055807 | 206773 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-04 14:22 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/04/2024 02:21:56 PM, Canjimavadivel, Kamallesh ) Replaced ECG leads All tests pass
1055808 | 206774 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-09 10:30 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/09/2024 10:30:18 AM, DeGuzman, Leila) PM complete, all tests pass
1055809 | 206775 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-09 11:26 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 07/09/2024 11:25:33 AM, DeGuzman, Leila) PM completed.
1055810 | 206776 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-04 16:52 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/04/2024 04:51:37 PM, Cortez, John) PM passed, see procedure for results
1055811 | 206777 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-03 16:59 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/03/2024 04:59:16 PM, Cortez, John) PM passed, see procedure for results
1055812 | 206778 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-06 16:21 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/06/2024 04:20:54 PM, Cortez, John) PM passed, see procedure for results
1056037 | 206779 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 11:12 | PM-MAJOR-Annual-PM for : Tester | (TC 06/11/2024 11:11:59 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1056190 | 206780 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-24 14:43 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/24/2024 02:42:50 PM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch install Vacuumed PC
1056895 | 206781 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR_ Annual PM for: Microscope -Surgical
1057302 | 206782 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-08 08:24 | PM - Annual | (TC 07/08/2024 08:24:37 AM, Fan, Tak) PM Completed. Gnd R = 120 mohm Max A = 1.7 / 116.5 / 191.8 uA C1-6 (363058YP9) = 0.8 / 2.8 uA C2-9 (359931YP3) = 0.7 / 2.3 uA RM6C (238232KR3) = 0.8 / 4.4 uA
1057699 | 206783 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-05 08:28 | Battery Cycling | (TC 06/05/2024 08:27:49 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057700 | 206784 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-10 09:59 | Battery Cycling | (TC 06/10/2024 09:58:57 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057710 | 206785 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-13 08:33 | Battery Cycling | (TC 06/13/2024 08:33:03 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057711 | 206786 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-10 10:02 | Battery Cycling | (TC 06/10/2024 10:01:56 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057713 | 206787 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-13 08:29 | Battery Cycling | (TC 06/13/2024 08:29:10 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057714 | 206788 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-05 08:30 | Battery Cycling | (TC 06/05/2024 08:29:32 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057715 | 206789 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-07 08:10 | Battery Cycling | (TC 06/07/2024 08:09:34 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057716 | 206790 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-03 09:01 | Battery Cycling | (TC 06/03/2024 09:01:26 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1057717 | 206791 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-06 08:46 | Battery Cycling | (TC 06/06/2024 08:46:15 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capactiy 3/5 LEDs
1057718 | 206792 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-07 08:11 | Battery Cycling | (TC 06/07/2024 08:11:19 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057719 | 206793 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-10 10:03 | Battery Cycling | (TC 06/10/2024 10:02:48 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057720 | 206794 | PM | CNP | ENL | 2024-05-01 03:28 | 2024-07-22 07:49 | Battery Cycling
1057721 | 206795 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-06 08:40 | Battery Cycling | (TC 06/06/2024 08:39:31 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capactiy 4/5 LEDs
1057722 | 206796 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-03 09:05 | Battery Cycling | (TC 06/03/2024 09:05:01 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057723 | 206797 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-10 10:04 | Battery Cycling | (TC 06/10/2024 10:03:49 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057724 | 206798 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 10:28 | Battery Cycling | (TC 06/11/2024 10:28:07 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057726 | 206799 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-04 08:56 | Battery Cycling | (TC 06/04/2024 08:55:43 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057727 | 206800 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-04 08:25 | Battery Cycling | (TC 06/04/2024 08:25:13 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1057730 | 206801 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-26 14:41 | Battery Cycling | (TC 06/26/2024 02:40:37 PM, Reid, Emmerson ) Battery cycling complete. 3/5 lights shown for capacity. Pass
1057732 | 206802 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-26 14:38 | Battery Cycling | (TC 06/26/2024 02:37:39 PM, Reid, Emmerson ) Battery cycling complete. 3/5 lights shown for capacity. Pass
1057733 | 206803 | PM | ACTIV | READY | 2024-05-01 03:28 |  | Battery Cycling
1057734 | 206804 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-26 14:39 | Battery Cycling | (TC 06/26/2024 02:39:13 PM, Reid, Emmerson ) Battery cycling complete. 3/5 lights shown for capacity. Pass
1057735 | 206805 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-28 09:30 | Battery Cycling | (TC 06/28/2024 09:30:23 AM, Reid, Emmerson ) Battery conditioning complete. Battery capacity shows 3/5 LEDs. Pass
1057736 | 206806 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-28 09:32 | Battery Cycling | (TC 06/28/2024 09:31:45 AM, Reid, Emmerson ) Battery conditioning complete. Battery capacity shows 4/5 LEDs. Pass
1057737 | 206807 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-27 08:35 | Battery Cycling | (TC 06/27/2024 08:35:29 AM, Reid, Emmerson ) Conditioning complete. Full capacity shows 3/5 LEDs.
1057738 | 206808 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-28 09:33 | Battery Cycling | (TC 06/28/2024 09:33:19 AM, Reid, Emmerson ) Battery conditioning complete. Battery capacity shows 3/5 LEDs. Pass
1057739 | 206809 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-09-19 09:04 | Battery Cycling | (TC 09/18/2024 02:10:47 PM, Alvarado, Karla) Battery reconditioning. (TC 09/19/2024 09:04:21 AM, Alvarado, Karla) Battery conditioning complete
1057740 | 206810 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-27 08:33 | Battery Cycling | (TC 06/27/2024 08:32:58 AM, Reid, Emmerson ) Conditioning complete. Full capacity shows 3/5 LEDs.
1057758 | 206811 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-24 14:55 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/24/2024 02:55:07 PM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch install Vacuumed PC
1057760 | 206812 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-24 14:56 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/24/2024 02:55:51 PM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch install Vacuumed PC
1057761 | 206813 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-24 14:57 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/24/2024 02:56:34 PM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch install Vacuumed PC
1057764 | 206814 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-23 09:41 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/23/2024 09:40:46 AM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch Vacuumed PC PM complete
1057837 | 206815 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1058169 | 206816 | PM | CMPLT | PMFAIL | 2024-05-01 03:28 | 2024-07-05 10:05 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/05/2024 10:04:42 AM, Reid, Emmerson ) Not working, need to send out to repiar. Had an issue while upgrading and the firmware is failing.
1058181 | 206817 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-20 09:22 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 08/20/2024 09:22:30 AM, Reid, Emmerson ) PM Complete, see procedure for details.
1058183 | 206818 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058184 | 206819 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058186 | 206820 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058188 | 206821 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058190 | 206822 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058191 | 206823 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-11 11:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/11/2024 11:54:50 AM, DeGuzman, Leila) PM completed. See procedure for details.
1058193 | 206824 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058569 | 206825 | PM | ACTIV | READY | 2024-05-01 03:28 |  | Annual PM
1059330 | 206826 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-09 10:41 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:40:31 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:40:43 AM, Rhuland, Courtney) PM completed.
1059530 | 206827 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059531 | 206828 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059532 | 206829 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059533 | 206830 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059534 | 206831 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/24/2024 11:11:30 AM, Canjimavadivel, Kamallesh ) Removed usb from main board and rebuit PC. Restored configuration
1059549 | 206832 | PM | ACTIV | ISSUE | 2024-05-01 03:28 |  | PM-MAJOR - Annual PM For: Central Monitor
1059550 | 206833 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-09 10:54 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:54:06 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:54:13 AM, Rhuland, Courtney) PM completed.
1059551 | 206834 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-23 14:23 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/23/2024 02:22:37 PM, Rendon Lopez , Elizabeth) Vacuumed central, PM completed
1059552 | 206835 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-09 10:55 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:54:42 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:54:48 AM, Rhuland, Courtney) PM completed.
1059555 | 206836 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-09 10:58 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:57:51 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:57:59 AM, Rhuland, Courtney) PM completed.
1059874 | 206837 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-08-16 10:47 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD.
1059879 | 206838 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-10-28 11:25 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 10/28/2024 11:24:36 AM, Cortez, John) Replaced rear case, Completed calibration and PM
1060155 | 206839 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-03 08:57 | Battery Cycling | (TC 06/03/2024 08:56:56 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060156 | 206840 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 10:30 | Battery Cycling | (TC 06/11/2024 10:30:14 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060157 | 206841 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-03 08:58 | Battery Cycling | (TC 06/03/2024 08:58:13 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060195 | 206842 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-05 08:31 | Battery Cycling | (TC 06/05/2024 08:31:17 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1060198 | 206843 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-07 08:05 | Battery Cycling | (TC 06/07/2024 08:05:09 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060333 | 206844 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 11:15 | PM-MAJOR- Annual PM For: Patient Simulator | (TC 06/11/2024 11:14:48 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1060336 | 206845 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-14 08:34 | Battery Cycling | (TC 06/14/2024 08:34:05 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1060337 | 206846 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-13 08:31 | Battery Cycling | (TC 06/13/2024 08:30:45 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1060338 | 206847 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-06 08:44 | Battery Cycling | (TC 06/06/2024 08:43:30 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capactiy 3/5 LEDs
1060579 | 206848 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-22 10:35 | PM-MAJOR - Annual PM For: multimeter | (TC 06/04/2024 10:35:21 AM, Chukri, Samir) Performed PM by Getinge Tech. attached Certificate .
1060596 | 206849 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-05-22 10:39 | PM-MAJOR- Annual PM For: Multimeter | (TC 06/04/2024 10:39:31 AM, Chukri, Samir) Performed PM by Getinge Tech. Attached certificate
1060609 | 206850 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-05 10:59 | PM-MAJOR-Annual PM for ESU Tester | (TC 07/05/2024 10:59:23 AM, Chukri, Samir) Received from Clare Instrument, Certificate of Calibration Attached.
1060834 | 206851 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 12:34 | PM-MAJOR- Annual PM for: Scale | (TC 06/11/2024 12:34:25 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1060901 | 206852 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-05 08:25 | Battery Cycling | (TC 06/05/2024 08:25:19 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1060902 | 206853 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 10:25 | Battery Cycling | (TC 06/11/2024 10:25:14 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 3/5 LEDs
1061254 | 206854 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-06 08:45 | Battery Cycling | (TC 06/06/2024 08:44:52 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capactiy 3/5 LEDs
1063201 | 206855 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 11:21 | PM-MAJOR- Annual PM for: Pressure, test digital | (TC 06/11/2024 11:20:31 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1063358 | 206856 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-23 09:42 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/23/2024 09:42:10 AM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch Vacuumed PC PM complete
1064496 | 206857 | PM | ACTIV | READY | 2024-05-01 03:28 |  | Battery Cycling
1064497 | 206858 | PM | ACTIV | READY | 2024-05-01 03:28 |  | Battery Cycling | (TC 09/19/2024 11:27:53 AM, Alvarado, Karla) Battery reconditioning. (TC 09/19/2024 11:27:53 AM, Rendon Lopez , Elizabeth) Supervised
1064572 | 206859 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-11 12:21 | PM-MAJOR- Annual PM for: Multimeter | (TC 06/11/2024 12:20:50 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1066939 | 206860 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-19 14:09 | PM-MAJOR - Annual PM For: Central Monitor | (TC 07/19/2024 02:08:59 PM, Canjimavadivel, Kamallesh ) Rebooted to finish patch install Vacuumed PC
1067968 | 206861 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-07-23 09:42 | PM MAJOR-Monitors, Physiologic, Central Station | (TC 07/23/2024 09:41:45 AM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch Vacuumed PC PM complete
1070257 | 206862 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-13 12:48 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 06/13/2024 12:48:04 PM, Canjimavadivel, Kamallesh ) Al ltests pass PM complete
1070268 | 206863 | PM | CMPLT | CMPLT | 2024-05-01 03:28 | 2024-06-04 08:29 | Battery Cycling | (TC 06/04/2024 08:29:15 AM, Canjimavadivel, Kamallesh ) Reconditioning complete Capacity 4/5 LEDs
1047866 | 206864 | PA | CMPLT | CMPLT | 2024-05-01 07:51 | 2024-05-06 14:17 | Ordering parts for repair | (TC 05/01/2024 07:57:01 AM, Baghirzade, Nara) SC 1003221460 (TC 05/06/2024 02:06:58 PM, Baghirzade, Nara) PO 7500205981
32510 | 206865 | CM | CMPLT | CMPLT | 2024-05-01 07:51 | 2024-09-25 09:30 | MSH DR1 OGP ERRORS, WILL NOT MOVE, WILL NOT LINK | (TC 04/30/2024 07:52:00 PM, Whittington, Scott) MSH DR1 OGP ERRORS, WILL NOT MOVE, WILL NOT LINK FOUND MAJOR ISSUE SAFETY RELAY INTERMITTENT REPLACED FROM LOCAL STOCK (TC 04/30/2024 07:52:00 PM, Whittington, Scott) MSH DR1 OGP ERRORS, WILL NOT MOVE, WILL NOT LINK FOUND MAJOR ISSUE SAFETY RELAY INTERMITTENT REPLACED FROM LOCAL STOCK | MSH DR1 OGP ERRORS, WILL NOT MOVE, WILL NOT LINK
65572 | 206866 | CM | CMPLT | CMPLT | 2024-05-01 09:54 | 2024-05-01 09:56 | Functional test | (TC 05/01/2024 09:54:29 AM, Baghirzade, Nara) As per Reeba's request, we did a functional test of 10 W and 20 W; no errors observed. No issue at boot up. Laser is functional. Updated Reeba.
| 206867 | CM | CMPLT | CMPLT | 2024-05-01 10:02 | 2024-05-01 10:03 | Managing assets, old WOs, and updating contacts on respective NICU equipment and lasers | (TC 05/01/2024 10:02:29 AM, Baghirzade, Nara) Managing assets, old WOs, and updating contacts on respective NICU equipment and lasers. Removed old contacts from WOs, contacts, and PM scheduling. Updated all contacts and closed old WOs for NICU and lasers.
1053261 | 206868 | CM | CMPLT | CMPLT | 2024-05-01 10:29 | 2024-05-01 10:30 | static noise | (TC 05/01/2024 10:29:25 AM, Rendon Lopez , Elizabeth) Confirmed issue. Repaired beofre, no needed, retired from service
66527 | 206869 | CM | CMPLT | CMPLT | 2024-05-01 10:31 | 2024-05-15 11:00 | Not reading SpO2 | (TC 04/30/2024 10:51:09 AM, Rendon Lopez , Elizabeth) Confirmed issue in Rm 10, SpO2 not working with different probes. Removed monitor, replaced with their spare. (TC 05/01/2024 10:51:09 AM, Rendon Lopez , Elizabeth) Replaced SpO2 board, cleaned contacts on API board, tested-OK. Replaced enclosure gasket, and cleaned recorder (TC 05/02/2024 10:59:27 AM, Rendon Lopez , Elizabeth) Returned to L&D room 10. Tested-Ok
| 206870 | PA | CMPLT | CMPLT | 2024-05-01 10:45 | 2024-05-07 14:19 | Double sided tape for biomed | (TC 05/01/2024 11:19:13 AM, Canjimavadivel, Kamallesh ) Placed order for double-sided tape
1063808 | 206871 | CM | CMPLT | CMPLT | 2024-05-01 11:15 | 2024-05-02 10:14 | MSH 15 SPP PROCEDURE Ultrasound GE VOLUSON 10 C2-9 PROBE ERROR REPORTED | (TC 05/01/2024 11:15:27 AM, Whittington, Scott) MSH 15 SPP PROCEDURE Ultrasound GE VOLUSON 10 C2-9 PROBE ERROR REPORTED | MSH 15 SPP PROCEDURE Ultrasound GE VOLUSON 10 C2-9 PROBE ERROR REPORTED
1054278 | 206872 | CM | CMPLT | CMPLT | 2024-05-01 11:18 | 2024-05-02 10:14 | MSH OPG SPP VOLUSON PROBE C2-9-D BORROWED FOR MSH 15 SPP PROCEDURE | (TC 05/01/2024 11:18:34 AM, Whittington, Scott) MSH OPG SPP VOLUSON PROBE C2-9-D BORROWED FOR MSH 15 SPP PROCEDURE | MSH OPG SPP VOLUSON PROBE C2-9-D BORROWED FOR MSH 15 SPP PROCEDURE
1050480 | 206873 | CM | CMPLT | CMPLT | 2024-05-01 12:22 | 2024-05-06 14:10 | ordering PM kits for Servo-Ns | received 10 x Servo-N PM kit. PO#7500206013
| 206874 | CM | CMPLT | CMPLT | 2024-05-01 12:39 | 2024-05-01 12:39 | worked on TMS | worked on TMS
| 206875 | CM | CMPLT | CMPLT | 2024-05-01 12:58 | 2024-05-01 13:07 | Communications meeting | (TC 05/01/2024 12:57:42 PM, Baghirzade, Nara) Communications committee meeting
1061333 | 206876 | FC | CMPLT | CMPLT | 2024-05-01 13:08 | 2024-05-01 13:08 | QC checklist | (TC 05/01/2024 01:07:56 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1053391 | 206877 | CM | CMPLT | CMPLT | 2024-05-01 13:09 | 2024-05-01 13:10 | Check before/after patient use and filter change | (TC 05/01/2024 01:08:57 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 126 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
66431 | 206878 | CM | CMPLT | NFF | 2024-05-01 13:19 | 2024-05-01 13:20 | Lead 1 not working nurse said | (TC 05/01/2024 01:19:11 PM, Scero, Tomlyn) Installed batteries and ran Tele pack on a simulator for a few hours. Monitored saves on biomed central. No issues were found with any of the leads.
| 206879 | CM | CMPLT | CMPLT | 2024-05-01 13:55 | 2024-05-01 13:59 | Rounds - NICU
| 206880 | CM | CMPLT | CMPLT | 2024-05-01 14:01 | 2024-05-01 14:03 | Rounds - RESUS
1050580 | 206881 | CM | CMPLT | CMPLT | 2024-05-01 14:07 | 2024-05-01 14:09 | Caster wheel broken | call re: broken. Went to check and brought it down to biomed. Cleaned unit. Found one of the caster wheel broke. Replaced it. Checked funtionality. Tested ok. Returned to OR.
| 206882 | CM | CMPLT | CMPLT | 2024-05-01 14:10 | 2024-05-01 14:13 | Suction Tree Issue | Call re: not working. Went to check and brought ir down to biomed. Cleaned unit, Found suction not working. Checked all connections and secured it. Worked but very slow. Removed all cannisters and tubing. Checked suction regulator only and found that it suction is working fine but cannister liners and tubing has issues. Reassembled everything and dropped it back to OR. Advised OR staff to repalce all cannisters and tubing. Tested ok.
1070743 | 206883 | CM | CMPLT | CMPLT | 2024-05-01 15:26 | 2024-05-01 15:27 | Update serial number | (TC 05/01/2024 03:25:49 PM, Canjimavadivel, Kamallesh ) Device was swapped out by company, updated serial number in TMS and added asset tag
| 206884 | CM | CMPLT | CMPLT | 2024-05-02 09:08 | 2024-05-02 13:20 | Boom inspection | (TC 05/02/2024 09:08:06 AM, Rhuland, Courtney) Checked booms in NICU for issues Noted serial numbers of the ones that need repair (TC 05/02/2024 10:04:17 AM, Baghirzade, Nara) Inspected the booms in zone C and partial zone B for defects and broken parts; noted 6 rooms in total.
1060000 | 206885 | CM | CMPLT | CMPLT | 2024-05-02 10:03 | 2024-05-02 10:04 | Screen was frozen | (TC 05/02/2024 10:03:24 AM, Baghirzade, Nara) Rebooted the PSN. No errors or issues during boot up. All sectors loaded properly.
| 206886 | CM | CMPLT | CMPLT | 2024-05-02 10:06 | 2024-05-02 10:06 | On-call daily checks | (TC 05/02/2024 10:05:55 AM, Baghirzade, Nara) Completed the daily on-call checks.
501070 | 206887 | CM | CMPLT | CMPLT | 2024-05-02 13:29 | 2024-05-02 13:30 | Replaced casters.
501063 | 206888 | CM | CMPLT | CMPLT | 2024-05-02 13:30 | 2024-05-02 13:30 | Replaced casters.
1055585 | 206889 | CM | CMPLT | CMPLT | 2024-05-02 13:41 | 2024-05-02 13:45 | Created a spark when plugged into the wall | (TC 05/02/2024 01:41:38 PM, Ferreira, Qwynn) Masimo had a tag saying there was a spark when the device was plugged in. Viewed the room and saw the plate was loose and the was shorting the plug when inserted. Relpaced the power cord, and tested the device - Working. Electrical Safety passed Functional Test passed Reviewed with the unit
1070989 | 206890 | CM | CMPLT | CMPLT | 2024-05-02 13:42 | 2024-05-02 13:43 | Created asset on TMS | (TC 05/02/2024 01:42:28 PM, Baghirzade, Nara) Created asset on TMS; last cal was April 2020. Informed Samir that it requires calibration and added it to the tester equipment list.
500196 | 206891 | CM | CMPLT | CMPLT | 2024-05-02 13:47 | 2024-05-06 14:57 | Entegra won't boot | (TC 05/02/2024 01:46:47 PM, Green, Paul) Spare 330 onsite. Swapped over hard drive and cards. Had problems powering on but were able to resolve. Problems now with video card settings.
| 206892 | PA | CMPLT | CMPLT | 2024-05-02 13:56 | 2024-05-10 14:04 | Transport Incubator Battery | (TC 07/17/2024 11:29:39 AM, Chukri, Samir) received batteries.
69103 | 206893 | PM | CMPLT | CMPLT | 2024-05-02 14:26 | 2024-05-02 14:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/02/2024 02:27:22 PM, Baghirzade, Nara) Replaced right IUI. Replaced both pressure sensors. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 206894 | CM | CMPLT | CMPLT | 2024-05-02 14:27 | 2024-05-31 16:19 | misc biomed - May 2024 | (TC 05/01/2024 02:32:56 PM, Cortez, John) Ivy(CNS, radiology) reached out asking for extra cuffs. When I went to deliver, it turns out they need a device as well as, their previous handheld BP device failed Becky and I are looking into providing a sphyg on rollstand, or a spare electronic monitor (TC 05/02/2024 02:35:13 PM, Cortez, John) 10N nurse station meeting with IT, human factors and Natalia Masimo to be moved to make room for printers, timeline TBA (TC 05/03/2024 03:32:34 PM, Cortez, John) 10N nurse station meeting (TC 05/03/2024 03:57:49 PM, Cortez, John) Amarjeet from 10N is asking if we have spare cuffs. Provided ordering info for Masimo cuffs. Also mentioned we have two in case of emergency (TC 04/29/2024 04:04:29 PM, Cortez, John) huddle (TC 05/09/2024 04:44:59 PM, Cortez, John) Inventory of vital signs monitors on 10N/10S (TC 05/08/2024 04:50:40 PM, Cortez, John) Helped Becky redistribute vital signs monitors from 10N/10S (TC 05/07/2024 04:52:27 PM, Cortez, John) 2 APs down in Emerg triage Checked SW closet. Reseated patch cables in ports for both APs Initially, APs didnt show up, but showed up next day (TC 05/10/2024 01:03:11 PM, Cortez, John) 10N nurse station meeting, Masimo viewstation appears to be staying in same spot (TC 05/10/2024 01:03:11 PM, Cortez, John) 10N nurse station meeting, Masimo viewstation appears to be staying in same spot (TC 05/10/2024 01:04:26 PM, Cortez, John) Spoke with Barb about vital signs repurposing and impact to work flow We talked about equipping each monitor with 2 adult cuffs, so they can be swapped between patients provided ordering info to Barb (TC 05/10/2024 04:15:22 PM, Cortez, John) Grabbed SCD from ICU for Samir (TC 05/14/2024 04:49:28 PM, Cortez, John) Root inventory on 10N/10S (TC 05/13/2024 04:49:58 PM, Cortez, John) Removed Roots in safetynet that will be moved to 12N (TC 05/17/2024 02:43:08 PM, Cortez, John) Focal point alerted that 12th floor switch and APs are unreachable. Visited closet and switch was unplugged by accident by contractor, who was working in closet (TC 05/23/2024 02:24:02 PM, Cortez, John) Masimo 16N1 and 16N2 offline. Went with Jessica to resolve. 16N1 was powered off and 16N2 was found on 12N powered off. Powered both on and issue was resolved (TC 05/31/2024 04:19:07 PM, Cortez, John) Looked into wireless LAN options present in ECG carts, for Kamran. Looks like we have a mix of SDC-CF10G and SDC-CF20G cards. The latter is option D21 (TC 05/30/2024 04:30:36 PM, Cortez, John) Amanda from ICU asked for ordering info for several Zoll defib accessories. Provided info
1055907 | 206895 | CM | CMPLT | CMPLT | 2024-05-02 14:44 | 2024-05-02 14:48 | "BP 0 working" | (TC 05/02/2024 02:44:22 PM, Cortez, John) Replaced hose, added Radical battery Cleaned filter Functional check ok
501172 | 206896 | CM | CMPLT | CMPLT | 2024-05-02 15:14 | 2024-09-25 09:30 | MSH 15 SPP 15-217 15th Floor BLANCHE - Murray Side Blanche Small Procedure Room 501172 PROBE C5-1 MIE 501596 | (TC 05/02/2024 03:14:01 PM, Whittington, Scott) MSH 15 SPP 15-217 15th Floor BLANCHE - Murray Side Blanche Small Procedure Room 501172 PROBE C5-1 MIE 501596 FAULTY PROBE MIE 501596 TEMP REPLACEMENT FROM LOCAL STOCK MIE 1047613 TGH | MSH 15 SPP 15-217 15th Floor BLANCHE - Murray Side Blanche Small Procedure Room 501172 PROBE C5-1 MIE 501596 FAULTY PROBE MIE 501596 TEMP REPLACEMENT FROM LOCAL STOCK MIE 1047613 TGH
501596 | 206897 | CM | CMPLT | CMPLT | 2024-05-02 15:16 | 2024-09-25 09:31 | MSH 15 SPP 15-217 15th Floor BLANCHE - Murray Side Blanche Small Procedure Room 501172 PROBE C5-1 MIE 501596 | (TC 05/02/2024 03:16:36 PM, Whittington, Scott) MSH 15 SPP 15-217 15th Floor BLANCHE - Murray Side Blanche Small Procedure Room 501172 PROBE C5-1 MIE 501596 FAULTY PROBE MIE 501596 TEMP REPLACEMENT FROM LOCAL STOCK MIE 1047613 TGH mie 501896 tested appeared ok, will check it again prior to return | MSH 15 SPP 15-217 15th Floor BLANCHE - Murray Side Blanche Small Procedure Room 501172 PROBE C5-1 MIE 501596 FAULTY PROBE MIE 501596 TEMP REPLACEMENT FROM LOCAL STOCK MIE 1047613 TGH
1053348 | 206898 | CM | CMPLT | CMPLT | 2024-05-02 15:41 | 2024-09-25 09:32 | MSH BI 1272 H3 CONSOLE UP/DOWN CTRL PN# MEL-00390 INSTALL WITH RYAN | (TC 05/02/2024 03:41:35 PM, Whittington, Scott) MSH BI 1272 H3 CONSOLE UP/DOWN CTRL PN# MEL-00390 INSTALL WITH RYAN (TC 05/02/2024 03:41:35 PM, Whittington, Scott) MSH BI 1272 H3 CONSOLE UP/DOWN CTRL PN# MEL-00390 INSTALL WITH RYAN | MSH BI 1272 H3 CONSOLE UP/DOWN CTRL PN# MEL-00390 INSTALL WITH RYAN
1066454 | 206899 | CM | CMPLT | CMPLT | 2024-05-03 07:26 | 2024-05-28 07:19 | instill pump not working | shipping VAC for repair estimate. We will update the cost after the repair quote. (TC 05/21/2024 07:21:50 AM, Baghirzade, Nara) Repair quote approved by Ella Sarvi. E-mailed Fan the signed copy of the quote approval and requested the PO be adjusted to the amount on the quote. Once received, will email the PO to the vendor to initiate the device repair. (TC 05/21/2024 10:22:04 AM, Baghirzade, Nara) Received the revised PO from Fan. Sent it to vendor. (TC 05/28/2024 07:18:57 AM, Baghirzade, Nara) Received VAC from repair. Attached service report/QC. Returned to 14 South
1060900 | 206900 | CM | CMPLT | CMPLT | 2024-05-03 08:35 | 2024-05-07 08:27 | Broken panoptic ophthalmoscope handle and head not working | (TC 05/03/2024 09:51:44 AM, Canjimavadivel, Kamallesh ) Picked up device from ED MD office. Handles was broken. Found part number online and contacted distributer for pricing. Placed order for handle and bulb. (TC 05/06/2024 01:48:22 PM, Canjimavadivel, Kamallesh ) Received handle and bulb, unboxed and put handle on charge. (TC 05/07/2024 08:26:28 AM, Canjimavadivel, Kamallesh ) Turned on and tested ok. Returned to ED physicians' office
| 206901 | PA | CMPLT | CMPLT | 2024-05-03 10:32 | 2024-06-04 11:07 | Rechargeable battery pack for stock | (TC 05/03/2024 10:36:58 AM, Canjimavadivel, Kamallesh ) Placed order for batteries (TC 06/04/2024 11:07:25 AM, Canjimavadivel, Kamallesh ) Received batteries
63873 | 206902 | CM | CMPLT | CMPLT | 2024-05-03 10:55 | 2024-05-03 11:03 | Broken tilt plate assembly | (TC 05/03/2024 10:55:58 AM, Rhuland, Courtney) Repaired tilt platform Replaced brake pad and ball insert Calibrated RH Functional check - OK
68295 | 206903 | PM | CMPLT | CMPLT | 2024-05-03 12:57 | 2024-05-03 13:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/03/2024 12:57:24 PM, Baghirzade, Nara) Replaced right IUI. Replaced bottle-side pressure sensor first. During pressure calibration, patient-side pressure sensor was failing. Replaced patient-side pressure sensor. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 206904 | CM | CMPLT | CMPLT | 2024-05-03 13:07 | 2024-05-03 13:08 | 1 on 1 meeting with Adeel | (TC 05/03/2024 01:07:17 PM, Baghirzade, Nara) 1-1 with Adeel
| 206905 | RDS | CMPLT | CMPLT | 2024-05-03 13:08 | 2024-05-03 13:10 | Weekly Rounds: 11 North | (TC 05/03/2024 01:08:02 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 206906 | RDS | CMPLT | CMPLT | 2024-05-03 13:10 | 2024-05-03 13:11 | Weekly Rounds: 11 South | (TC 05/03/2024 01:10:00 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 206907 | CM | CMPLT | CMPLT | 2024-05-03 13:11 | 2024-05-03 13:14 | Weekly Rounds: NICU & RESUS | (TC 05/03/2024 01:11:30 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
| 206908 | CM | CMPLT | CMPLT | 2024-05-03 13:15 | 2024-05-03 14:26 | Closing old WOs for all PCUs, syringes, and LVPs on "scheduled service" tab | (TC 05/03/2024 02:25:26 PM, Baghirzade, Nara) Closing old WOs for all PCUs, syringes, and LVPs on "scheduled service" tab. Put "equipment not located" for PM WO's from year 2019-2023
1053347 | 206909 | CM | CMPLT | CMPLT | 2024-05-03 13:45 | 2024-09-25 09:32 | MSH BI 1270 H1 COMPRESSION PADDLE MOTOR FAULTY ROOM DOWN | (TC 05/03/2024 01:44:58 PM, Whittington, Scott) MSH BI 1270 H1 COMPRESSION PADDLE MOTOR FAULTY ROOM DOWN (TC 05/07/2024 02:22:33 PM, Whittington, Scott) MSH BI 1270 H1 COMPRESSION PADDLE MOTOR FAULTY ROOM DOWN DETECTOR FIBRE CABLE CONNECTION REPAIRED BY RYAN & NOOR | MSH BI 1270 H1 COMPRESSION PADDLE MOTOR FAULTY ROOM DOWN
| 206910 | CM | CMPLT | CMPLT | 2024-05-03 14:07 | 2024-05-03 14:08 | Otoscope/ophthalmoscope not turning on | (TC 05/03/2024 02:07:01 PM, Rhuland, Courtney) Replaced bulbs in 3 scopes in family med Checked all other scopes for functionality
67961 | 206911 | CM | CMPLT | CMPLT | 2024-05-03 14:50 | 2024-05-03 14:52 | Air probe failure | (TC 05/03/2024 02:50:17 PM, Rhuland, Courtney) Ran incubator and 30 minutes in, air probe failure alarm sounded Replaced scale as it was not reading Pushed pin back into place on back of scale connection Was unable to calibrate RH, so RH sensor was replaced Functional check - OK
69514 | 206912 | CM | CMPLT | CMPLT | 2024-05-03 16:25 | 2024-05-03 16:27 | 245.302 error | (TC 05/03/2024 04:26:05 PM, Cortez, John) Completed check-in successfully Need to upgrade and complete PM
502231 | 206913 | CM | CMPLT | CMPLT | 2024-05-03 16:57 | 2024-05-03 16:56 | Replacement of the cable hose in the C-Arm. | Hose was replaced.
1063808 | 206914 | CM | CMPLT | CMPLT | 2024-05-03 21:07 | 2024-09-25 09:43 | MSH 15 SPP Ultrasound GE VOLUSON 10 H PROCEDURE ROOM CLOSED FOR MAINTENANCE, SYSTEM MOVED TO SMALL SCAN ROOM, WILL NOT CONNECT TO NETWORK | (TC 05/03/2024 09:11:41 PM, Whittington, Scott) The issue is that the GE is setup to connect to HSC, (although that is not working ) and the network jack in the procedure room is specifically setup for that. So it does not plug & play in another network jack. If the wireless works, then it shouldn't be an issue to connect to Astraia and JDMIPACS CORAL. The wireless should be active whenever the network cable is NOT plugged in. I will look into the wireless while here. I Will stop in first thing Monday morning to return GE to procedure room. (TC 05/03/2024 09:11:41 PM, Whittington, Scott) MSH 15 SPP Ultrasound GE VOLUSON 10 H PROCEDURE ROOM CLOSED FOR MAINTENANCE, SYSTEM MOVED TO SMALL SCAN ROOM, WILL NOT CONNECT TO NETWORK The issue is that the GE is setup to connect to HSC, (although that is not working ) and the network jack in the procedure room is specifically setup for that. So it does not plug & play in another network jack. If the wireless works, then it shouldn't be an issue to connect to Astraia and JDMIPACS CORAL. The wireless should be active whenever the network cable is NOT plugged in. I will look into the wireless while here. I Will stop in first thing Monday morning to return GE to procedure room. (TC 05/03/2024 09:11:41 PM, Whittington, Scott) MSH 15 SPP Ultrasound GE VOLUSON 10 H PROCEDURE ROOM CLOSED FOR MAINTENANCE, SYSTEM MOVED TO SMALL SCAN ROOM, WILL NOT CONNECT TO NETWORK The issue is that the GE is setup to connect to HSC, (although that is not working ) and the network jack in the procedure room is specifically setup for that. So it does not plug & play in another network jack. If the wireless works, then it shouldn't be an issue to connect to Astraia and JDMIPACS CORAL. The wireless should be active whenever the network cable is NOT plugged in. I will look into the wireless while here. I Will stop in first thing Monday morning to return GE to procedure room. | MSH 15 SPP Ultrasound GE VOLUSON 10 H PROCEDURE ROOM CLOSED FOR MAINTENANCE, SYSTEM MOVED TO SMALL SCAN ROOM, WILL NOT CONNECT TO NETWORK
1059536 | 206915 | CM | ACTIV | ISSUE | 2024-05-03 21:53 |  | MSH OB ANAES 15th Floor C5-2 PROBE MEMBRANE ARTIFACT? |  | MSH OB ANAES 15th Floor C5-2 PROBE MEMBRANE ARTIFACT?
| 206916 | RDS | CMPLT | CMPLT | 2024-05-06 09:02 | 2024-05-06 09:09 | Weekly Checks For PACU | (TC 05/06/2024 09:03:56 AM, Reid, Emmerson ) Weekly round complete. Brought down two CADD pumps due for PM
| 206917 | RDS | CMPLT | CMPLT | 2024-05-06 09:18 | 2024-05-06 14:29 | Weekly checks in CCU | (TC 05/06/2024 02:27:10 PM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
1054965 | 206918 | CM | CMPLT | CMPLT | 2024-05-06 09:28 | 2024-05-06 09:28 | battery defective | (TC 04/29/2024 09:29:10 AM, Huang, Jessica) the battery indicator showed fully charged however, the monitor would turn off. contacted Masimo and sent the monitor for repair received it back in Apr 2024 SN 2000058401 replaced by SN 2000196313
1057541 | 206919 | PM | CMPLT | CMPLT | 2024-05-06 09:35 | 2024-05-06 13:51 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/06/2024 01:50:28 PM, Scero, Tomlyn) Battery replaced. PM checks complete and passed.
1070852 | 206920 | CM | CMPLT | CMPLT | 2024-05-06 09:41 | 2024-05-06 09:41 | new incoming inspection | (TC 04/30/2024 09:41:36 AM, Huang, Jessica) unboxed everything and assembled the monitor completed new incoming inspection
1070854 | 206921 | CM | CMPLT | CMPLT | 2024-05-06 09:44 | 2024-05-06 09:44 | new incoming inspection | (TC 04/30/2024 09:44:47 AM, Huang, Jessica) unboxed everything and assembled the monitor completed new incoming inspection
1070860 | 206922 | CM | CMPLT | CMPLT | 2024-05-06 09:47 | 2024-05-06 09:47 | new incoming inspection
1070586 | 206923 | CM | CMPLT | CMPLT | 2024-05-06 09:48 | 2024-05-06 09:48 | not working | (TC 04/30/2024 09:48:08 AM, Huang, Jessica) the power cord was loose so the battery could not get charged charged the battery, working fine
1070849 | 206924 | CM | CMPLT | CMPLT | 2024-05-06 09:51 | 2024-05-06 09:51 | new incoming inspection | (TC 05/02/2024 09:51:31 AM, Huang, Jessica) unboxed everything and assembled the monitor completed new incoming inspection dropped off the additonal accessories in Tanya's office
1064355 | 206925 | CM | CMPLT | CMPLT | 2024-05-06 09:54 | 2024-05-06 09:58 | Not picking up leads and not connecting to network | (TC 05/06/2024 09:55:11 AM, Canjimavadivel, Kamallesh ) Tested with simulator, all leads working. WIFI connected while testing in biomed, was able to transfer order and pull up worklist
1070850 | 206926 | CM | CMPLT | CMPLT | 2024-05-06 10:04 | 2024-05-06 10:04 | new incoming inspection | (TC 05/02/2024 10:04:15 AM, Huang, Jessica) unboxed everything and assembled the monitor completed new incoming inspection dropped off the additonal accessories in Tanya's office
1070235 | 206927 | CM | CMPLT | CMPLT | 2024-05-06 10:06 | 2024-05-06 10:06 | low battery | (TC 05/03/2024 10:06:20 AM, Huang, Jessica) replaced batteries and tested it working fine
1048550 | 206928 | CM | CMPLT | CMPLT | 2024-05-06 10:10 | 2024-05-06 10:10 | error code E0016 | (TC 05/03/2024 10:10:28 AM, Huang, Jessica) picked up the device from OR desk and contacted the vendor regarding the error code. Was informed the device needed to be sent for repair and informed OR. OR requested the loaner sent the device to the vendor and sent the loaner to OR received the device back and swapped it shipped the loaner back to the company and sent the device back to OR
1070879 | 206929 | CM | CMPLT | CMPLT | 2024-05-06 10:30 | 2024-05-06 10:30 | new incoming inspection
1070880 | 206930 | CM | CMPLT | CMPLT | 2024-05-06 10:31 | 2024-05-06 10:31 | new incoming inspection
| 206931 | RDS | CMPLT | CMPLT | 2024-05-06 11:43 | 2024-05-06 11:43 | weekly check for 14S and 17M | weekly check for 14S and 17M
500196 | 206932 | CM | CMPLT | CMPLT | 2024-05-06 13:32 | 2024-05-06 13:33 | Entegra Computer Wont Boot | (TC 05/06/2024 01:31:36 PM, Paz, Ramon (Tino)) Swapped PS and installed all necessary components. Replaced CMOS battery as well.
1059778 | 206933 | FC | CMPLT | CMPLT | 2024-05-06 13:53 | 2024-05-06 13:54 | QC checklist | (TC 05/06/2024 01:53:04 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
66431 | 206934 | CM | CMPLT | CMPLT | 2024-05-06 13:57 | 2024-05-06 14:35 | LL not working, showing only 1 wave on patient | (TC 05/06/2024 02:29:25 PM, Scero, Tomlyn) Central station showed LL lead off alarm. I brought the L&D nurse a new set of leads and helped her change them and showed her the waveforms on the central station.
65416 | 206935 | CM | CMPLT | CMPLT | 2024-05-06 13:58 | 2024-05-06 13:59 | User reported that the baby probe does not measure when on fan mode | (TC 05/06/2024 01:58:04 PM, Baghirzade, Nara) Ran a full functional test; checked on baby probe 1 and 2, both measure when on fan mode. Left incubator to run for 40 mins and checked the baby probe throughout the run time. Most likely the user had a defective probe that was not reading. Found no issue during functional test.
65978 | 206936 | CM | CMPLT | CMPLT | 2024-05-06 14:02 | 2024-05-06 14:03 | Calibration required | (TC 05/06/2024 02:02:28 PM, Baghirzade, Nara) Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
68367 | 206937 | CM | CMPLT | CMPLT | 2024-05-06 14:37 | 2024-05-06 14:39 | Data cable damaged | (TC 05/06/2024 02:37:15 PM, Scero, Tomlyn) Replaced the EC10 cable to the AGM. Cable was damged. New cable is a 90 degree cable.
1061202 | 206938 | CM | CMPLT | CMPLT | 2024-05-07 09:13 | 2024-08-01 14:57 | OR Light casing damaged | (TC 05/07/2024 09:19:26 AM, Reid, Emmerson ) One of the plastic cover pieces was coming off. I replaced with one from OR12 and will need to order a replacement. (TC 07/02/2024 08:27:36 AM, Reid, Emmerson ) Placed the order on erec (TC 08/01/2024 02:57:10 PM, Reid, Emmerson ) Recieved and Installed on the light. Working now.
1049904 | 206939 | CM | CMPLT | CMPLT | 2024-05-07 10:46 | 2024-05-07 10:47 | MSH R61 Revo - Broken Top Charger Prongs | (TC 05/07/2024 10:46:12 AM, Chau, Ken) Replaced top charging pins and tested okay
1056125 | 206940 | PM | CMPLT | CMPLT | 2024-05-07 11:39 | 2024-05-07 11:40 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/07/2024 11:39:22 AM, Canjimavadivel, Kamallesh ) Cleaned nozzles and hoses. Cleaning cycle complete.
| 206941 | CM | CMPLT | CMPLT | 2024-05-07 13:06 | 2024-05-16 13:39 | Parts for incubators | (TC 05/07/2024 02:53:01 PM, Rhuland, Courtney) Inventoried incubator parts and placed order for parts that were running low (TC 05/16/2024 01:38:01 PM, Baghirzade, Nara) PO 7500206235
1059778 | 206942 | FC | CMPLT | CMPLT | 2024-05-07 13:19 | 2024-05-07 13:20 | Sri reported that it wasn't suctioning | (TC 05/07/2024 01:19:07 PM, Baghirzade, Nara) Ran 1st QC; no issues during QC. Pump is working and could feel the suction from the tubing. Booted down the device and ran it on VAC therapy mode for some time; no issues with suction. Ran 2nd QC; QC checklist completed. No issues found with the VAC. Returned to 11 South and followed up with Sri.
1046035 | 206943 | CM | CMPLT | CMPLT | 2024-05-07 13:20 | 2024-05-07 13:21 | MSH ER Evo - Replaced table bucky microswitch
1054694 | 206944 | CM | CMPLT | CMPLT | 2024-05-07 13:25 | 2024-05-07 13:27 | User reported error | (TC 05/07/2024 01:25:34 PM, Baghirzade, Nara) Worked with Laila, coop student, on this. She ran a functional test; no issues or errors during the feed. Checked the pump and its fully operational.
67937 | 206945 | CM | CMPLT | CMPLT | 2024-05-07 13:29 | 2024-05-07 13:30 | Setup new printer | (TC 05/07/2024 01:29:14 PM, Rhuland, Courtney) New printer was ordered to replace broken one Installed all necessary drivers and tested printer All tests passed unit is functional
| 206946 | CM | CMPLT | CMPLT | 2024-05-07 13:33 | 2024-05-09 14:34 | Closing old WOs for PCUs, syringes, and LVPs | (TC 05/07/2024 01:37:31 PM, Baghirzade, Nara) Closing old WOs for PCUs, syringes, and LVPs from years 2019-2023. (TC 05/08/2024 01:10:34 PM, Baghirzade, Nara) Closing old WOs for PCUs, syringes, and LVPs from years 2019-2023. (TC 05/09/2024 02:33:42 PM, Baghirzade, Nara) Completed closing old WOs for PCUs, syringes, and LVPs from years 2019-2023. Only 2024 WOs are active. Confirmed with Samir.
| 206947 | CM | CMPLT | CMPLT | 2024-05-07 13:38 | 2024-05-09 12:09 | X2 batteries | (TC 05/07/2024 02:53:36 PM, Rhuland, Courtney) Reached out to multiple vendors for quotes, placed order with Dixon Med for 20 X2 batteries
1057553 | 206948 | PM | CMPLT | CMPLT | 2024-05-07 14:17 | 2024-05-07 14:19 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/07/2024 02:16:54 PM, Scero, Tomlyn) PM checks complete and passed. battery replaced. Broken bezel replaced.
| 206949 | CM | CMPLT | CMPLT | 2024-05-07 14:35 | 2024-05-10 11:20 | U Line order | Searched and ordered biomed items from uline. May 10 - Order received. Removed and orgnazed everything in the work area. Cut and disposed boxes, packing material.
1071009 | 206950 | CM | CMPLT | CMPLT | 2024-05-08 07:11 | 2024-05-08 12:28 | Incoming inspection | (TC 05/08/2024 12:27:29 PM, Rhuland, Courtney) Unit tagged and serial numbers noted Functional check - OK Electrical safety - OK
1071010 | 206951 | CM | CMPLT | CMPLT | 2024-05-08 07:14 | 2024-05-08 12:30 | Incoming inspection | (TC 05/08/2024 12:29:31 PM, Rhuland, Courtney) Unit tagged and serial numbers noted Functional check - OK Electrical safety - OK
| 206952 | RDS | CMPLT | CMPLT | 2024-05-08 09:18 | 2024-05-08 10:55 | Weekly check for ICU
| 206953 | RDS | CMPLT | CMPLT | 2024-05-08 09:18 | 2024-05-21 11:44 | Weekly check for ED
67969 | 206954 | CM | CMPLT | CMPLT | 2024-05-08 09:58 | 2024-06-04 10:11 | Not level | (TC 05/08/2024 10:05:20 AM, Rhuland, Courtney) Worked with Leila to fix the loose base plate During functional check, humidity sensor was found to be defective - running as a manual humidifier Waiting on order of new humidity sensors (TC 06/04/2024 10:10:32 AM, Baghirzade, Nara) Retiring this device as we received new fleet of omnibeds and incubators.
1066225 | 206955 | CM | CMPLT | CMPLT | 2024-05-08 10:27 | 2024-05-08 10:27 | error on the screen | tested unit for 2 days. functional.
66468 | 206956 | CM | CMPLT | CMPLT | 2024-05-08 10:28 | 2024-05-08 10:28 | Constantly alarming "pressure measurement inaccurate" | Passed device check. Running the unit without any issue for 2 days. Sent logs to Draeger. Got the response from Draeger: I would recommend the following: 1. Run the Device Check and Leakage Test, does the unit pass? If it does not, check for source of leakage or reason one of the steps fails. Passed the test. returned back to service.
63215 | 206957 | CM | CMPLT | CMPLT | 2024-05-08 11:50 | 2024-05-08 13:30 | Broken, missing loading door | (TC 05/09/2024 11:03:36 AM, DeGuzman, Leila) Missing loading door - replaced Ran the feed, no issue found. (TC 05/08/2024 01:35:00 PM, DeGuzman, Leila) Missing loading door - replaced Ran the feed, no issue found.
1047865 | 206958 | PM | CMPLT | PMFAIL | 2024-05-08 12:44 | 2024-05-08 12:46 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/08/2024 12:44:52 PM, Baghirzade, Nara) Failed PM during occlusion test. When testing the force sensor, had to calibrate the force sensor as it was not zeroing properly. Calibration passed however during the occlusion test, was consistently getting "coarse travel/check clutch" error. Will inspect clutch and replace if needed.
1047865 | 206959 | CM | CMPLT | CMPLT | 2024-05-08 12:46 | 2024-06-04 12:14 | Created by Failed PM WO # 206958 : PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/16/2024 01:51:45 PM, Baghirzade, Nara) Disassembled the device, replaced the clutch. Greased the guide screw. Assembled the device back however when running the pump, plunger was making a loud noise. Disassembled the device again down to the guide screw to check if there was anything in the way or misplaced; everything looked ok, did not find any issue. Assembled the device and will run functional tests. (TC 06/04/2024 12:13:11 PM, Baghirzade, Nara) Reseated the clutch. Ensured that the guide screw was greased. Assembled the device back and ran a functional with occlusion; no issues or errors observed. (TC 06/04/2024 12:14:13 PM, Baghirzade, Nara) Completed calibration for: -syringe size and position -force sensor -plunger position -syringe size sensor Completed alarm/speaker tests. PM completed. See procedure document for PM results. PM schedule updated. | (TC 05/08/2024 12:44:52 PM, Baghirzade, Nara) Failed PM during occlusion test. When testing the force sensor, had to calibrate the force sensor as it was not zeroing properly. Calibration passed however during the occlusion test, was consistently getting "coarse travel/check clutch" error. Will inspect clutch and replace if needed.
1047849 | 206960 | PM | CMPLT | CMPLT | 2024-05-08 12:47 | 2024-05-08 12:54 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/08/2024 12:47:20 PM, Baghirzade, Nara) PM completed. See procedure for test results. (TC 05/08/2024 12:47:20 PM, Baghirzade, Nara) PM completed. Replaced battery. Replaced power cord retainer. See procedure for test results.
1053391 | 206961 | CM | CMPLT | CMPLT | 2024-05-08 12:57 | 2024-05-08 12:58 | Before and after patient use | (TC 05/08/2024 12:56:37 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 51 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Device is functional. Returned to 11S.
1054687 | 206962 | CM | CMPLT | CMPLT | 2024-05-08 12:58 | 2024-05-08 13:10 | Power cord missing adapter and loading area was dislodged from its place | (TC 05/08/2024 12:58:28 PM, Baghirzade, Nara) Worked with Laila; she disassembled the device and placed the loading area back in to place. Replaced the adapter. Ran a feed and the pump is fully functional; no errors/issues observed during feed.
1062401 | 206963 | CM | CMPLT | CMPLT | 2024-05-08 13:22 | 2024-05-08 13:33 | User reported that the head portion does not move | (TC 05/08/2024 01:32:22 PM, Baghirzade, Nara) Staff reported bed not working; found all motion controls locked. Checked all motion controls including both side panels and patient panels. Followed up with Olcay. Bed is functional.
1061216 | 206964 | CM | CMPLT | CMPLT | 2024-05-08 13:48 | 2024-05-08 13:50 | Light iPad froze | (TC 05/08/2024 01:48:55 PM, Reid, Emmerson ) iPad was frozen. I rebooted the unit and it was working. I also had to pick up a megadyn pad with a rip in it.
1067048 | 206965 | CM | CMPLT | CMPLT | 2024-05-08 14:07 | 2024-05-10 16:24 | Exposed cable wire + intermittent bed exit issue | (TC 05/08/2024 02:09:36 PM, Canjimavadivel, Kamallesh ) Bed was reported having intermittent bed exit alarm overnight. Brought another bed to swap out. Tested bed exit functions, no issues observed. Found wire exposed on left foot side. Moved bed to 18th floor and put in service request with Baxter. (TC 05/10/2024 04:20:43 PM, Cortez, John) Bed repaired by Vince, and returned to service
| 206966 | CM | CMPLT | CMPLT | 2024-05-08 14:15 | 2024-05-08 14:15 | Closing old X2 WOs and adjusting the schedules and contacts | (TC 05/08/2024 02:14:51 PM, Baghirzade, Nara) Closing old X2 WOs and adjusting the schedules and contacts for NICU and SSDU
1066843 | 206967 | CM | CMPLT | CMPLT | 2024-05-08 14:24 | 2024-05-13 15:56 | Lamp replacement | May 08 - Call re: Light spource not working. Went to check. found lamp replacement required. Spoke with Anita. Was informed that endo may have spare one. Looked through dept inventory, Could not find any. Took out the bronc room lightsource lamp and replace it in room 1 temporarily. Contacted Olympus for purchase info. Created SC. Approved. Awaiting delivery.. May 13 - Lamp received. Coordianted with clinical staff for replacement. Retrieved original lamp from bronc room. Removed faulty bulb and replaced a new one. Installed it in room 1. Checked functionality. Tested ok. Installed bronc room lamp back from room 1 back in beonc room.
| 206968 | CM | CMPLT | CMPLT | 2024-05-08 14:33 | 2024-06-25 13:10 | Medfusion batteries
1057539 | 206969 | PM | CMPLT | CMPLT | 2024-05-08 14:36 | 2024-05-08 14:57 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/08/2024 02:56:19 PM, Scero, Tomlyn) PM tests complete and passed. Battery replaced.
| 206970 | CM | CMPLT | CMPLT | 2024-05-08 14:48 | 2024-06-06 10:57 | 12 North / 9 South Move and Installation | May 08 - Worked with Naz and moved all newly purchased equipment in boxes to 12 North for assembply and installation. May 09 - Toured 9 South phychiatry with BN to check the vital sign machine distribution. Counted number of VS machines and marked the once that need to be taken out of service. Retired units will be repalced with Masimo. May 15 - Moved 120 new IV pums from 6th floor storage to 12N for upcoming BD inspection. May 16 - Coordination with BD for IV pumps inspection. Arranged and delivered stuff with for clinical orientation. May 17 - BD Inspection completed. Arranged and distributed all IV pumps across dept. Delivered slings from storage as per NB's advise for clincial orientation. Inspected and checked Masimo units. May 21 - Updated all IV pump info on TMS. Moved furtniture and other stuff that we gave to BD to appropriate location. Unboxed and Arranged Bed pumps and accessories. May 29 - Masimo rep came in for network testing. Completed. Checked functionality of all masimo monitors as per NB's request and distributed them into rooms. May 31 - Three beds encountered foot board issue. Worked with stryker to resolve it. Two issues resolved. Awaiting stryker to resolve the 3rd. Jun 03 - Worked ans supported in 12N go live. Jun 06 - Connected with staff anf Ensured all is working fine. WO closed.
1061211 | 206971 | CM | CMPLT | CMPLT | 2024-05-08 14:53 | 2024-05-08 14:54 | Monitor Tilting Issue | CAll re: monitro tilting automatically. Went to check. Waited for access. Found when attempting to tilt the monitor at certain angle, it comes back. Issue with tension. Removed cover and tightned screws to adjust the tension. Issue resolved. Tested ok.
501583 | 206972 | CM | CMPLT | CMPLT | 2024-05-08 15:03 | 2024-05-08 15:03 | Monitor spring and power cord replacement
501106 | 206973 | CM | CMPLT | CMPLT | 2024-05-08 15:04 | 2024-05-08 15:05 | Gas spring replacement | (TC 05/08/2024 03:05:04 PM, Fan, Tak) Replaced gas spring.
1053348 | 206974 | CM | CMPLT | CMPLT | 2024-05-08 15:44 | 2024-09-25 09:43 | MSH BI 1272 H3 BIOPSY DECK WITH FAULTY RT SIDE LOCK MECHANISM RETURNED TO HOLOGIC | (TC 05/08/2024 04:08:56 PM, Whittington, Scott) MSH BI 1272 H3 BIOPSY DECK WITH FAULTY RT SIDE LOCK MECHANISM RETURNED TO HOLOGIC (TC 05/08/2024 04:09:05 PM, Dolinar, John) MSH BI 1272 H3 BIOPSY DECK WITH FAULTY RT SIDE LOCK MECHANISM RETURNED TO HOLOGIC | MSH BI 1272 H3 BIOPSY DECK WITH FAULTY RT SIDE LOCK MECHANISM RETURNED TO HOLOGIC
| 206975 | CON | CMPLT | CMPLT | 2024-05-08 16:25 | 2024-09-25 09:44 | MSH BI MAMOGRAPHY HARP REVISIONS | (TC 05/08/2024 04:25:08 PM, Dolinar, John) MSH BI MAMOGRAPHY HARP REVISIONS | MSH BI MAMOGRAPHY HARP REVISIONS
500668 | 206976 | CM | CMPLT | CMPLT | 2024-05-09 08:44 | 2024-07-31 09:22 | MSH CEOU-RM 6 GE VOLUSON 4 ANGELA HAS REQ FROM RADS TO USE THIS SYSTEM, WILL INVESTIGATE | (TC 05/09/2024 08:43:39 AM, Whittington, Scott) MSH CEOU-RM 6 GE VOLUSON 4 ANGELA HAS REQ FROM RADS TO USE THIS SYSTEM, WILL INVESTIGATE | MSH CEOU-RM 6 GE VOLUSON 4 ANGELA HAS REQ FROM RADS TO USE THIS SYSTEM, WILL INVESTIGATE
1071017 | 206977 | ACC | CMPLT | CMPLT | 2024-05-09 08:47 | 2024-05-09 08:49 | Incoming Inspection | (TC 05/09/2024 08:46:59 AM, Reid, Emmerson ) Connected tubing and installed the battery. Asset created. Incoming complete.
| 206978 | CM | CMPLT | CMPLT | 2024-05-09 08:51 | 2024-05-09 09:00 | Beds are not working | (TC 05/09/2024 08:51:15 AM, Rhuland, Courtney) Received email that multiple beds were not working Went over to OPG and found that both beds had loose power cords attached to bed Reseated power cords and inspected other beds for same issue Talked to reporter and she thinks it is from the cleaners Functional check of beds - OK
| 206979 | CM | CMPLT | CMPLT | 2024-05-09 09:02 | 2024-05-10 12:17 | Ordering supplies for biomed stock
1071018 | 206980 | ACC | CMPLT | CMPLT | 2024-05-09 09:12 | 2024-05-09 09:14 | Incoming Inspection | (TC 05/09/2024 09:13:37 AM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1071019 | 206981 | ACC | CMPLT | CMPLT | 2024-05-09 09:59 | 2024-05-09 10:00 | Incoming Inspection | (TC 05/09/2024 09:59:14 AM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1061212 | 206982 | CM | CMPLT | CMPLT | 2024-05-09 10:01 | 2024-06-24 13:53 | Wall monitor not showing video | (TC 05/09/2024 10:01:54 AM, Reid, Emmerson ) Went to OR11 and the wall monitor was not working. I tried manipulating the video source but it was not working. Loking at the VMC service mode, nothing seemed wrong. I wanted to restart the VMC but they wanted me to wait until the next morning. (TC 05/10/2024 09:01:16 AM, Reid, Emmerson ) Reset VMC multiple times and needed to reset the Corio. Afterwards the scope signals were working but the nurse PC and PACs were not sending their signals. NOTE: only after the first VMC reset an CORIO error appeared. (TC 05/10/2024 09:01:16 AM, Reid, Emmerson ) Reset VMC multiple times and needed to reset the Corio. Afterwards the scope signals were working but the nurse PC and PACs were not sending their signals. NOTE: only after the first VMC reset an CORIO error appeared. Also attended the laser in-service and grabbed 4 NICU pumps. (TC 06/11/2024 02:19:55 PM, Reid, Emmerson ) Still dealing with the issue. Will need to contact Olympus for this. (TC 06/12/2024 02:16:54 PM, Patel, Rutvik) Olympus rep came in to check. WOrked with him to test the monitor. Arranged fro a video tower to connect to the monitor. It worked but room inputs are still not showing up. Contacted integration specialist. Was informed they had a patient. Will work on this issue on Friday. (TC 06/24/2024 01:50:40 PM, Patel, Rutvik) checked OR multiple times within past week. Found Wall monitor issue was resolved itself. Tested ok.
1071020 | 206983 | ACC | CMPLT | CMPLT | 2024-05-09 10:10 | 2024-05-09 10:12 | Incoming Inspection | (TC 05/09/2024 10:10:27 AM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1071021 | 206984 | ACC | CMPLT | CMPLT | 2024-05-09 10:23 | 2024-05-09 10:24 | Incoming Inspection | (TC 05/09/2024 10:23:28 AM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1071022 | 206985 | ACC | CMPLT | CMPLT | 2024-05-09 10:31 | 2024-05-09 10:31 | Incoming Inspection | (TC 05/09/2024 10:30:58 AM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1059541 | 206986 | PM | CMPLT | CMPLT | 2024-05-09 10:42 | 2024-05-09 10:43 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:42:03 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:42:18 AM, Rhuland, Courtney) PM completed.
1059541 | 206987 | PM | CANCL | DUP | 2024-05-09 10:42 |  | PM-MAJOR - Annual PM For: Central Monitor
1059542 | 206988 | PM | CMPLT | CMPLT | 2024-05-09 10:44 | 2024-05-09 10:45 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:44:39 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:44:47 AM, Rhuland, Courtney) PM completed.
1059543 | 206989 | PM | CMPLT | CMPLT | 2024-05-09 10:46 | 2024-05-09 10:47 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:46:34 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:46:42 AM, Rhuland, Courtney) PM completed.
1059545 | 206990 | PM | CMPLT | CMPLT | 2024-05-09 10:48 | 2024-05-09 10:49 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:48:23 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:48:32 AM, Rhuland, Courtney) PM completed.
1059546 | 206991 | PM | CMPLT | CMPLT | 2024-05-09 10:50 | 2024-05-09 10:50 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:49:53 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:50:00 AM, Rhuland, Courtney) PM completed.
1059547 | 206992 | PM | CMPLT | CMPLT | 2024-05-09 10:51 | 2024-05-09 10:51 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:51:10 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:51:17 AM, Rhuland, Courtney) PM completed.
1059548 | 206993 | PM | CMPLT | CMPLT | 2024-05-09 10:52 | 2024-05-09 10:53 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:52:29 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:52:41 AM, Rhuland, Courtney) PM completed.
1071023 | 206994 | ACC | CMPLT | CMPLT | 2024-05-09 10:53 | 2024-05-09 10:53 | Incoming Inspection | (TC 05/09/2024 10:52:36 AM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1059553 | 206995 | PM | CMPLT | CMPLT | 2024-05-09 10:56 | 2024-05-09 10:56 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:55:47 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:55:54 AM, Rhuland, Courtney) PM completed.
1059554 | 206996 | PM | CMPLT | CMPLT | 2024-05-09 10:57 | 2024-05-09 10:57 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:57:02 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:57:09 AM, Rhuland, Courtney) PM completed.
1071024 | 206997 | ACC | CMPLT | CMPLT | 2024-05-09 10:58 | 2024-05-09 10:59 | Incoming Inspection | (TC 05/09/2024 10:58:34 AM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1059556 | 206998 | PM | CMPLT | CMPLT | 2024-05-09 10:59 | 2024-05-09 10:59 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:58:45 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 10:58:52 AM, Rhuland, Courtney) PM completed.
1059557 | 206999 | PM | CMPLT | CMPLT | 2024-05-09 11:00 | 2024-05-09 11:00 | PM-MAJOR - Annual PM For: Central Monitor | (TC 05/09/2024 10:59:59 AM, Baghirzade, Nara) PM completed. (TC 05/09/2024 11:00:09 AM, Rhuland, Courtney) PM completed.
1063746 | 207000 | CM | CMPLT | CMPLT | 2024-05-09 11:01 | 2024-05-09 11:02 | Cleaning | (TC 05/09/2024 11:01:14 AM, Baghirzade, Nara) Drained, disinfected, rinsed, and refilled with sterile water. (TC 05/09/2024 11:01:20 AM, DeGuzman, Leila) Drained, disinfected, rinsed, and refilled with sterile water.
1065805 | 207001 | CM | CMPLT | CMPLT | 2024-05-09 11:04 | 2024-05-09 11:05 | Functional check | (TC 05/09/2024 11:04:12 AM, Baghirzade, Nara) Showing Laila how to do the check-in/functional test for LVPs. Check-in tests passed and completed. IUI connectors and device cleaned.
| 207002 | CM | CMPLT | CMPLT | 2024-05-09 11:16 | 2024-05-09 11:27 | Rounds - NICU | (TC 05/09/2024 11:26:40 AM, Rhuland, Courtney) NICU rounds completed
| 207003 | CM | CMPLT | CMPLT | 2024-05-09 11:31 | 2024-05-09 11:45 | Rounds - RESUS | (TC 05/09/2024 11:31:46 AM, Rhuland, Courtney) RESUS rounds complete
502088 | 207004 | CM | CMPLT | CMPLT | 2024-05-09 11:45 | 2024-05-24 13:34 | Det2 hydration- Image Quality Evaluation | (TC 05/09/2024 11:44:41 AM, Green, Paul) Performed imaged quality evaluation for Siemens. This is in regards to MSH purchasing a new det2 for Symbia1 system. This is a 5/8 xtal. (TC 05/10/2024 01:29:03 PM, Green, Paul) From: Green, Paul Sent: May 10, 2024 9:27 AM To: FLORENTINO Jr, DANILO <danilo.florentino@siemens-healthineers.com> Cc: Paz, Ramon <Ramon.Paz@uhn.ca>; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca> Subject: MSH Symbia1 Site 9495 IMAGE QUALITY EVALUATION REQUEST with Det2 hydration Hi June, Here are the files for the MSH Symbia1 Site 9495 IMAGE QUALITY EVALUATION REQUEST with Det2 hydration. To download the files, password = journal Let us know if you need anything further. Regards, Paul From: NoReply.UHNFileShare@uhn.ca <NoReply.UHNFileShare@uhn.ca> Sent: May 10, 2024 9:23 AM To: Green, Paul <Paul.Green@uhn.ca> Subject: Your UHN File Share upload UHN File Share Notification Your file upload was successful. To have others download your file(s), please send them this web link: https://fileshare.uhn.ca/download/dc04eca2-65ff-4aff-ad7a-99de48dac0a6 Please note that this file upload will expire in 30 days. Please note that password(s) used for this upload cannot be recovered. If lost, you will need to repeat your upload again. (TC 05/23/2024 04:16:29 PM, Green, Paul) Removed and installed new det2. Powered up and performed tuning and 200million flood. (TC 05/24/2024 01:34:27 PM, Green, Paul) Confirmed detector monthly looks ok and all ok with new detector2 that was installed. Packed up and crated all old ps, boards and detector. Notified Siemens to pick up old detector.
65877 | 207005 | CM | CMPLT | CMPLT | 2024-05-09 11:53 | 2024-05-09 13:43 | Channel error | (TC 05/09/2024 01:41:03 PM, Rhuland, Courtney) Channel error indicating keyboard malfunciton Restart button would not register Replaced front case Replaced mylar gasket During check-in, barrel clamp - open position test failed Re-opened device and reseated park collar Test failed again Calibrated barrel clamp All tests passed post-calibration Refer to ASM for details
1064457 | 207006 | CM | CMPLT | CMPLT | 2024-05-09 12:21 | 2024-09-25 09:33 | MSH ER VENUE 40 STAND UPGRADE | (TC 05/09/2024 12:21:12 PM, Whittington, Scott) MSH ER VENUE 40 STAND UPGRADE | MSH ER VENUE 40 STAND UPGRADE
1064458 | 207007 | CM | CMPLT | CMPLT | 2024-05-09 12:23 | 2024-09-25 09:44 | MSH ER VENUE 40 STAND UPGRADE | (TC 05/09/2024 12:22:57 PM, Whittington, Scott) MSH ER VENUE 40 STAND UPGRADE | MSH ER VENUE 40 STAND UPGRADE
1064458 | 207008 | CM | CMPLT | CMPLT | 2024-05-09 12:24 | 2024-09-25 09:45 | MSH ER VENUE 40 PROBE FAULT | (TC 05/10/2024 09:11:32 AM, Whittington, Scott) MSH ER VENUE 40 PROBE FAULT C1-5-RS SN# 1094636WX6 PN REF# 5384874 MFG FEB 2022 POSSIBLE MEMBRANE ISSUE | MSH ER VENUE 40 PROBE FAULT C1-5-RS SN# 1094636WX6 PN REF# 5384874 MFG FEB 2022 POSSIBLE MEMBRANE ISSUE
| 207009 | CON | CMPLT | CMPLT | 2024-05-09 12:26 | 2024-09-25 09:46 | MSH BI GAID INSTALL | (TC 05/09/2024 12:26:38 PM, Whittington, Scott) MSH BI GAID INSTALL PROJECT INITIATED (TC 05/16/2024 02:34:10 PM, Whittington, Scott) MSH BI GAID INSTALL (TC 05/17/2024 08:27:00 PM, Whittington, Scott) MSH GAID INSTALL (TC 05/22/2024 12:28:31 PM, Fan, Tak) MSH BI GAID INSTALL CONSULT WITH MOE (TC 05/22/2024 12:28:56 PM, Whittington, Scott) MSH BI GAID INSTALL CONSULT WITH MOE (TC 05/17/2024 08:27:00 PM, Whittington, Scott) MSH GAID INSTALL | MSH BI GAID INSTALL
1066462 | 207010 | FC | CMPLT | CMPLT | 2024-05-09 12:46 | 2024-05-09 12:47 | QC checklist | (TC 05/09/2024 12:46:31 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1056450 | 207011 | CM | CMPLT | CMPLT | 2024-05-09 12:47 | 2024-05-09 12:49 | BP and temperature module not working | (TC 05/09/2024 12:47:46 PM, Baghirzade, Nara) No cuff was sent down with Masimo. Tested BP with multiple different cuffs and inspected the tubing; no damage to the tubing or connectors. Tested the temperature module on oral and adult aux; no issues with measuring the temperature. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1071026 | 207012 | ACC | CMPLT | CMPLT | 2024-05-09 13:18 | 2024-05-09 13:19 | Incoming Inspection | (TC 05/09/2024 01:18:39 PM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
1071027 | 207013 | ACC | CMPLT | CMPLT | 2024-05-09 13:25 | 2024-05-09 13:25 | Incming Inspection | (TC 05/09/2024 01:25:05 PM, Reid, Emmerson ) Created asset, installed tubing and battery. Incoming complete.
| 207014 | MEP | ACTIV | ISSUE | 2024-05-09 13:37 |  | CMMS Comittee | (TC 05/09/2024 01:36:59 PM, Reid, Emmerson ) Comittee meeting (TC 05/15/2024 02:14:52 PM, Reid, Emmerson ) Weekly meeting (TC 05/22/2024 02:20:07 PM, Reid, Emmerson ) CMMS Meeting (TC 06/03/2024 03:15:23 PM, Reid, Emmerson ) Meeting (TC 06/10/2024 03:03:02 PM, Reid, Emmerson ) CMMS meeting (TC 06/24/2024 09:11:21 AM, Reid, Emmerson ) CMMS meeting (TC 07/08/2024 03:17:34 PM, Reid, Emmerson ) Weekly meeting
1057548 | 207015 | PM | CMPLT | CMPLT | 2024-05-09 13:37 | 2024-05-09 13:39 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/09/2024 01:37:26 PM, Scero, Tomlyn) PM tests complete and passed. battery replaced.
1057555 | 207016 | PM | CMPLT | CMPLT | 2024-05-09 13:43 | 2024-05-09 13:44 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/09/2024 01:43:05 PM, Scero, Tomlyn) Pm test complete and passed. Battery replaced
| 207017 | CM | CMPLT | CMPLT | 2024-05-09 14:12 | 2024-06-20 08:22 | Biomedical test equipment annual Calibration and Certificate | (TC 05/09/2024 02:18:58 PM, Chukri, Samir) Replaced a PO for the biomed test equipment Cal and Cer (TC 06/20/2024 08:21:59 AM, Chukri, Samir) update the total price in the PO #7500206390 for 40 biomed test equipment . work done by Petter from Getinge .
| 207018 | CM | CMPLT | CMPLT | 2024-05-09 14:32 | 2024-05-09 14:32 | 11N & 11S VSM repurposing | (TC 05/09/2024 02:31:41 PM, Baghirzade, Nara) 11N & 11S VSM repurposing with Becky and Laila. (TC 05/09/2024 02:31:53 PM, DeGuzman, Leila) 11N & 11S VSM repurposing
| 207019 | CM | CMPLT | CMPLT | 2024-05-09 14:36 | 2024-05-09 14:37 | Closing old X2 and older NICU equipment active WOs | (TC 05/09/2024 02:36:32 PM, Baghirzade, Nara) Closed old X2 and other NICU equipment WOs from years 2019-2023; adjusted PM schedules. Confirmed with Samir.
| 207020 | CM | CMPLT | CMPLT | 2024-05-09 14:49 | 2024-06-06 15:51 | Yellow Fin repair | May 9 - Call re: not working. Went to check. Went to OR 12 and tested the unit. FOund cylinder making noise and stiff handle. Unit will need to be send to vendor for repair. Contacted vendor and requested RMA. Received. Searched in the dept and with receiving but could not find appropriate box to ship the unit. Ordered boxes from Uline and received it. May 10 - Created SC. Sent for approval. Received. Issued PO via MMC. Assembled box for shipping. May 17 - Acquired packaging material from different sources. Scheduled KJV Courier pick up. Packed and shipped the unit. June 06 - Unit received. Attached accessories. Checked functionality. Tested ok. Delivered back to OR.
1071025 | 207021 | CM | CMPLT | CMPLT | 2024-05-09 16:47 | 2024-05-09 16:50 | Converted for use in Mammography | (TC 05/09/2024 04:47:10 PM, Cortez, John) Ivy (supervisor, Radiology) reached out saying Mammography's portable BP device is broken. Becky took unit from 16N during vital signs redistribution. Replaced hoses, checked functionality, delivered to Mammography dept
1049905 | 207022 | CM | CMPLT | CMPLT | 2024-05-10 08:46 | 2024-09-25 09:47 | MSH R65 SYSTEM BATTERIES DEPLETED , CHARGING NOW | (TC 05/10/2024 08:45:54 AM, Whittington, Scott) MSH R65 SYSTEM BATTERIES DEPLETED , CHARGING NOW | MSH R65 SYSTEM BATTERIES DEPLETED , CHARGING NOW
68299 | 207023 | CM | CMPLT | CMPLT | 2024-05-10 09:28 | 2024-05-13 10:36 | Infusion rate accuracy | (TC 05/10/2024 11:29:00 AM, DeGuzman, Leila) Performed check in, analog sensor - passed Ran pump in basic infusion - VTBI=20mL, Rate=200mL, Scale reading= 19.420g
1066114 | 207024 | CM | CMPLT | CMPLT | 2024-05-10 10:12 | 2024-05-10 10:14 | Pokle mount stuck | (TC 05/10/2024 10:12:09 AM, Canjimavadivel, Kamallesh ) Opened up housing and repaired pole mount assembly.
| 207025 | CM | CMPLT | CMPLT | 2024-05-10 10:14 | 2024-05-10 10:15 | Power shutdown follow-up CCU | (TC 05/10/2024 10:14:15 AM, Canjimavadivel, Kamallesh ) Went to CCU to verify all systems ok while they switch over from emergency power to normal power in prep for overnight shutdown. All systems up and running.
1061204 | 207026 | CM | CMPLT | CMPLT | 2024-05-10 10:48 | 2024-05-10 11:01 | OR light not disconnected from WiFi | (TC 05/10/2024 10:47:48 AM, Reid, Emmerson ) Went to OR and the lights were not connecting to WiFi/ control panel. They told me they tried reseting the wifi on the lights. When I checked, light one's wifi was off. I turned it back on but was still having connection issues. I rebooted the router twice and it was working better after that.
1061210 | 207027 | CM | CMPLT | CMPLT | 2024-05-10 10:57 | 2024-05-09 15:15 | VMC Issue | Two times call from OR re: vitals not showing up on boom and wall monitors. Went to check. Found blue screen showing up when attepting to select vitals on the nursing desk monitor. Removed and reattached Philips monitor cable. Rebooted the VMC. Issue Resolved. Tested ok.
1071038 | 207028 | ACC | CMPLT | CMPLT | 2024-05-10 11:05 | 2024-05-16 13:23 | Incoming inspection | (TC 05/10/2024 11:23:30 AM, Canjimavadivel, Kamallesh ) Unboxed and added to database (TC 05/16/2024 01:22:24 PM, Canjimavadivel, Kamallesh ) Tested with cuff, functional test pass
1071039 | 207029 | ACC | CMPLT | CMPLT | 2024-05-10 11:16 | 2024-05-16 13:23 | Incoming inspection | (TC 05/10/2024 11:24:02 AM, Canjimavadivel, Kamallesh ) Unboxed and added to database (TC 05/16/2024 01:23:08 PM, Canjimavadivel, Kamallesh ) Tested with cuff, functional test pass
1061210 | 207030 | CM | CMPLT | CMPLT | 2024-05-10 11:17 | 2024-05-10 11:17 | VMC Issue | Call from OR re: vitals not showing up on boom and wall monitors. Went to check. Found blue screen showing up when attepting to select vitals on the nursing desk monitor. Removed and reattached Philips monitor cable. Rebooted the VMC. Issue Resolved. Tested ok.
| 207031 | CM | CMPLT | CMPLT | 2024-05-10 12:28 | 2024-05-10 12:28 | Thulium fiber laser in-service | (TC 05/10/2024 12:27:53 PM, Baghirzade, Nara) Thulium fiber laser in-service with Olympus
| 207032 | CM | CMPLT | CMPLT | 2024-05-10 12:29 | 2024-05-10 12:30 | Preparing thulium laser for use | (TC 05/10/2024 12:28:41 PM, Baghirzade, Nara) Preparing thulium fiber laser for use; printing new signage to be sent to Reeba for lamination. Olympus tech was on site and completed the output test as well as demonstrated the PM and some troubleshooting tips. Make-shifted a cover for the laser unless Reeba saved the old Holmium covers.
1059828 | 207033 | CM | CNP | EIU | 2024-05-10 12:42 |  | Cleaning not completed | (TC 05/10/2024 12:42:03 PM, Baghirzade, Nara) Went up to look for the Altrix 3rd day in a row; device is in use. No cleaning completed this week.
65896 | 207034 | CM | CMPLT | CMPLT | 2024-05-10 12:58 | 2024-05-10 12:58 | Calibration required | (TC 05/10/2024 12:57:49 PM, Rhuland, Courtney) Calibration and check in completed All tests passed Refer to ASM for details
| 207035 | RDS | CMPLT | CMPLT | 2024-05-10 13:32 | 2024-05-10 13:34 | Weekly Rounds: 11 North | (TC 05/10/2024 01:32:38 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 207036 | RDS | CMPLT | CMPLT | 2024-05-10 13:34 | 2024-05-10 13:37 | Weekly Rounds: 11 South | (TC 05/10/2024 01:34:32 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 207037 | RDS | CMPLT | CMPLT | 2024-05-10 13:37 | 2024-05-10 13:39 | Weekly Rounds: NICU & RESUS | (TC 05/10/2024 01:37:16 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
65993 | 207038 | CM | CMPLT | CMPLT | 2024-05-10 13:39 | 2024-05-10 13:39 | Returned from OR | (TC 05/10/2024 01:38:51 PM, Rhuland, Courtney) Returned from OR Check-in completed Returned to NICU
65898 | 207039 | CM | CMPLT | CMPLT | 2024-05-10 13:40 | 2024-05-10 13:41 | Returned from OR | (TC 05/10/2024 01:40:25 PM, Rhuland, Courtney) Returned from OR Missing asset tag - reprinted Check-in completed Returned to NICU
| 207040 | CM | CMPLT | CMPLT | 2024-05-10 13:46 | 2024-05-10 13:48 | Dental Equipment Issue | Call re: dental unit issue. Went to check. Found unit handle had a crack. Consulted ER and was informed that OR staff can reach out to vendor. Contacted OR staff and informed them.
| 207041 | CM | CMPLT | CMPLT | 2024-05-10 14:25 | 2024-05-29 09:58 | Masimo Biomed Stock Items | MAy 10 - Inventoried and found masimo stock items had been used up. Contacted Masimo for BP accessories purchase info. Created SC and sned it for approval. May 29 - Received order. Sorted and organized in shelved.
| 207042 | CM | CMPLT | CMPLT | 2024-05-10 14:44 | 2024-05-10 14:45 | YelloFin Repair | Call re: Yello Fin not working. Went to check and brought it down to biomed. Checked unit. Found handle not able to hold the leg position. Removed it and found screw threads were damaged. Repalced handle from an old unit that was retired. Checked functionality. Tested ok. Returned to OR.
| 207043 | CM | CMPLT | CMPLT | 2024-05-10 14:47 | 2024-05-10 14:49 | IVF Wheelchair Broken | Call from wendy re: Wheelchair not working. Consulted SC and ER. Was informed that we do not look after wheelchairs. Contacted Wendy and advised to connect with Building services to order a replacement.
1062471 | 207044 | CM | CMPLT | CMPLT | 2024-05-10 15:20 | 2024-05-16 13:01 | Head not lifting | (TC 05/10/2024 03:20:12 PM, Ferreira, Qwynn) Troubleshoot the bed to get teh head board to lift - did not lift Reset the bed - did not work Checked the pins - all were ok Motor issue potentially - calling stryker. (TC 05/16/2024 12:59:39 PM, Ferreira, Qwynn) Met with mark to discuss the bed situation. Seems the unit sent the bed to 11S then to 15M even though the bed was broken. Report has been attached to the asset. Bed is repaired
| 207045 | CM | CMPLT | CMPLT | 2024-05-10 16:13 | 2024-05-10 16:15 | PFT Hypair requires helium cal, yellow light during DLCO | (TC 05/10/2024 04:12:48 PM, Cortez, John) Ran cal in expairtec fast tlco tab. With tech's help, completed calibration in expair
| 207046 | PA | CMPLT | CMPLT | 2024-05-13 06:29 | 2024-05-28 07:38 | Ordering mattress hose assembly - stryker beds | (TC 05/13/2024 06:32:42 AM, Baghirzade, Nara) SC 1003229198 (TC 05/28/2024 07:38:21 AM, Baghirzade, Nara) PO 7500206487
1047006 | 207047 | CM | CMPLT | CMPLT | 2024-05-13 07:49 | 2024-05-16 11:29 | Unable to calibrate the pressure sensors | (TC 05/13/2024 07:48:55 AM, Baghirzade, Nara) Initially the patient-side pressure sensor was defective; this pressure sensor was replaced. During pressure calibration, the cal kept failing on both sensors. Replaced both pressure sensors however unable to calibrate and keeps failing. E-mailed BD to inquire. Will proceed to flash the LVP as a last resort. (TC 05/13/2024 12:05:43 PM, Baghirzade, Nara) Flashed the LVP module however both pressure sensors fail the pressure calibration. Waiting on a response from BD. (TC 05/16/2024 11:27:28 AM, Baghirzade, Nara) BD recommended to do the default pressure calibration in order to revert the pressure sensors back to factory settings and a rate calibration; Justin from BD mentioned that if the check-in or PM passes after this calibration, then the pump is functional. Default pressure calibration completed. Rate calibration completed. Check-in tests completed. IUI connectors and device cleaned.
68801 | 207048 | CM | CMPLT | CMPLT | 2024-05-13 07:52 | 2024-05-23 12:21 | Unable to download dataset | (TC 05/13/2024 07:51:45 AM, Baghirzade, Nara) Powered on/off the PCU several times to allow it to download the latest dataset however dataset won't download. Confirmed that network settings are correct. E-mailed BD to inquire about next steps without manually pushing the dataset. (TC 05/23/2024 12:20:36 PM, Baghirzade, Nara) BD tech support suggested to reload the network config and dataset files. Transferred network config and dataset. After power cycling the PCU, the correct dataset is showing. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 230 mohms Leakage current: 78.5 uA
| 207049 | RDS | CMPLT | CMPLT | 2024-05-13 08:12 | 2024-05-13 08:14 | Weekly Checks For PACU | (TC 05/13/2024 08:13:19 AM, Reid, Emmerson ) Rounds complete
1061219 | 207050 | CM | CMPLT | NFF | 2024-05-13 08:14 | 2024-05-13 08:16 | Called about a monitor not connecting | (TC 05/13/2024 08:14:52 AM, Reid, Emmerson ) Went up and they had already dealt with the issue.
1070853 | 207051 | CM | CMPLT | CMPLT | 2024-05-13 08:52 | 2024-05-13 08:52 | new incoming inspection | (TC 05/07/2024 08:51:45 AM, Huang, Jessica) unboxed everything and assembled the monitor completed new incoming inspection
1070855 | 207052 | CM | CMPLT | CMPLT | 2024-05-13 09:03 | 2024-05-13 09:03 | new incoming inspection | (TC 05/07/2024 09:03:27 AM, Huang, Jessica) unboxed everything and assembled the monitor completed new incoming inspection
1057607 | 207053 | CM | CMPLT | CMPLT | 2024-05-13 09:08 | 2024-05-13 09:08 | not working | (TC 05/06/2024 09:07:48 AM, Huang, Jessica) replaced the new temperature probe and charged the battery
541029 | 207054 | CM | CMPLT | CMPLT | 2024-05-13 09:10 | 2024-05-13 09:10 | out of service
120016 | 207055 | CM | CMPLT | CMPLT | 2024-05-13 09:12 | 2024-05-13 09:12 | out of service
120020 | 207056 | CM | CMPLT | CMPLT | 2024-05-13 09:13 | 2024-05-13 09:13 | out of service
1063122 | 207057 | CM | CMPLT | CMPLT | 2024-05-13 09:14 | 2024-05-13 09:14 | out of service
1071011 | 207058 | CM | CMPLT | CMPLT | 2024-05-13 09:16 | 2024-05-13 09:16 | new incoming | (TC 05/08/2024 09:16:09 AM, Huang, Jessica) received the skid and sent the device to Reeba's office (TC 05/10/2024 09:18:04 AM, Huang, Jessica) attended the inserive in OR and assisted Michael from Olympus to complete the new incoming inspection sent the device to the storage room and informed OR team
| 207059 | CM | CMPLT | CMPLT | 2024-05-13 09:23 | 2024-05-13 09:23 | requesting the table attachments quote for OR
1059828 | 207060 | PM | CMPLT | CMPLT | 2024-05-13 09:56 | 2024-05-13 09:57 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/13/2024 09:55:55 AM, DeGuzman, Leila) cleaned, drain, disinfected, refilled with sterile water. (TC 05/13/2024 10:06:05 AM, Baghirzade, Nara) Supervised Laila for the cleaning
| 207061 | RDS | CMPLT | CMPLT | 2024-05-13 10:02 | 2024-05-13 10:20 | Weekly check for ED
| 207062 | CM | CMPLT | CMPLT | 2024-05-13 10:04 | 2024-05-13 10:04 | OR Suction Tree Issue | Worked and rectified issue on Suction tree in OR
1063784 | 207063 | CM | CMPLT | CMPLT | 2024-05-13 10:05 | 2024-05-13 10:06 | Not working | Call re: Not working. Went to check and brought it down to biomed. Checked functionality. Could not find any issue. Checked focus and intensity adjustment. Tested ok. Returned to OR.
1046035 | 207064 | CM | CMPLT | CMPLT | 2024-05-13 10:07 | 2024-05-13 10:08 | Light field changing between 43x43 and 43x35 | (TC 05/13/2024 08:06:00 AM, Fan, Tak) System in 43x35. Checked switches in table bucky, they are working but portrait microswitch was not entirely being pushed down by detector. Adjusted angle of the switch. Tested, OK.
| 207065 | RDS | CMPLT | CMPLT | 2024-05-13 10:08 | 2024-06-07 08:34 | Weekly check for ICU
65155 | 207066 | CM | CMPLT | CMPLT | 2024-05-13 10:16 | 2024-05-13 10:17 | Upgrade to V4.3 | (TC 05/13/2024 10:16:33 AM, Reid, Emmerson ) Upgraded to V 4.3 and pushed APS IV PCA MAR 2022
65168 | 207067 | CM | CMPLT | CMPLT | 2024-05-13 10:20 | 2024-05-13 10:32 | Upgraded to V4.3 | (TC 05/13/2024 10:20:11 AM, Reid, Emmerson ) Upgraded to V4.3 and pushed APS IV PCA MAR 2022
1058192 | 207068 | CM | CMPLT | CMPLT | 2024-05-13 10:57 | 2024-05-13 11:02 | Upgraded to V4.3 | (TC 05/13/2024 10:58:19 AM, Reid, Emmerson ) Upgraded to V4.3 and pushed APS IV PCA MAR 2022
67956 | 207069 | CM | CMPLT | CMPLT | 2024-05-13 11:43 | 2024-05-13 12:00 | User reported that the temperature was not reading correctly/reading too high | (TC 05/13/2024 11:58:47 AM, Baghirzade, Nara) No issues found with the air probe temp. It was matching the temperature that the incubator was measuring. Tested on 35C and 37C. During functional test, the RH sensor was calibrated twice. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.6%RH Incubator is fully functional.
67920 | 207070 | CM | CMPLT | CMPLT | 2024-05-13 12:01 | 2024-05-13 12:02 | Wrong door placed and porthole seal damaged | (TC 05/13/2024 12:01:01 PM, Baghirzade, Nara) Swapped the door with a spare in biomed; initially 2 right doors were sent down with the incubator. Porthole seal was replaced. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.4%RH Incubator is fully functional.
1061333 | 207071 | FC | CMPLT | CMPLT | 2024-05-13 12:32 | 2024-05-13 12:34 | QC checklist | (TC 05/13/2024 12:34:05 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
1050632 | 207072 | CM | CMPLT | CMPLT | 2024-05-13 13:20 | 2024-05-13 13:34 | MSH 15 L&D TE7B NOT CHARGING | (TC 05/13/2024 01:20:03 PM, Whittington, Scott) MSH 15 L&D TE7B NOT CHARGING JOHN REPLACED CORD CAP (TC 05/13/2024 01:20:29 PM, Dolinar, John) MSH 15 L&D TE7B NOT CHARGING JOHN REPLACED CORD CAP | MSH 15 L&D TE7B NOT CHARGING
1064352 | 207073 | CM | CMPLT | CMPLT | 2024-05-13 13:37 | 2024-06-05 12:50 | MSH ICU POD A SPARQ. REPORTED THAT SYSTEM WOULD SHUTDOWN DURING USE | (TC 05/13/2024 01:38:27 PM, Whittington, Scott) MSH ICU POD A SPARQ. REPORTED THAT SYSTEM WOULD SHUTDOWN DURING USE RAN SELF TEST PASS EXPORTED LOGS (TC 05/13/2024 01:40:57 PM, Dolinar, John) MSH ICU POD A SPARQ. REPORTED THAT SYSTEM WOULD SHUTDOWN DURING USE RAN SELF TEST PASS EXPORTED LOGS (TC 05/22/2024 10:33:00 AM, Fan, Tak) Jillian requesting logs again. Ran system test again (passed). Corrected date and time. (TC 05/23/2024 02:34:09 PM, Fan, Tak) Jill ran self test. Passed. Was able to replicate problem. Reseating/removing battery seemed to not affect problem. System tries to start up by itself sometimes but fails. Re-installing battery and plugging in seemed to turn it back on. Problem became more frequent. Not sure why. PS was ordered but lost. Arrived at TGH but Fedex person took it back on the truck. (TC 06/03/2024 12:59:17 PM, Whittington, Scott) JILLIAN REPLACED EBOX (TC 06/03/2024 12:59:39 PM, Fan, Tak) JILLIAN REPLACED EBOX (TC 06/05/2024 12:59:17 PM, Whittington, Scott) JUNE 3 JILLIAN REPLACED EBOX SYSTEM RETURNED TO ICU AFTER 4 HOURS TESTING OK | MSH ICU POD A SPARQ. REPORTED THAT SYSTEM WOULD SHUTDOWN DURING USE
| 207074 | CM | CMPLT | CMPLT | 2024-05-13 13:52 | 2024-05-13 13:52 | PM huddle | (TC 05/13/2024 01:51:51 PM, Baghirzade, Nara) PM huddle
1062258 | 207075 | CM | CMPLT | CMPLT | 2024-05-13 13:53 | 2024-05-13 13:54 | User reported that the left handside controls were not working | (TC 05/13/2024 01:53:09 PM, Baghirzade, Nara) Found the footboard not connected to the connector on the bed. Reconnected and plugged to AC power. Found that all motion controls were locked; unlocked all controls. Checked all motion controls including both side panels and patient panels. Bed is functional.
65154 | 207076 | CM | CMPLT | CMPLT | 2024-05-13 13:57 | 2024-05-13 13:57 | Upgraded to V4.3 | (TC 05/13/2024 01:56:50 PM, Reid, Emmerson ) Upgraded to V4.3 and pushed APS IV PCA MAR 2022
| 207077 | RDS | CMPLT | CMPLT | 2024-05-13 14:19 | 2024-05-14 10:35 | Weekly checks in CCU | (TC 05/14/2024 10:34:21 AM, Scero, Tomlyn) weekly checks complete. 30J tests passed. Replaced electrodes on 1 defib and printer paper on another
1064422 | 207078 | CM | CMPLT | CMPLT | 2024-05-13 15:13 | 2024-05-13 15:25 | "Leads not detected" | (TC 05/13/2024 03:24:28 PM, Cortez, John) Tested with ECG simulator, no issues Keyboard cover ripped, replaced
32580 | 207079 | CM | CMPLT | CMPLT | 2024-05-13 16:02 | 2024-05-13 16:02 | Faulty cooling fan in the gantry | -Faulty cooling fan in the gantry - One of the OPCONT fan in the gantry was faulty -Removed OPCONT unit and replaced fan -Ran system test and it passed all tests.
1054669 | 207080 | CM | CMPLT | CMPLT | 2024-05-14 07:38 | 2024-10-02 10:07 | Sending pump for repair/exchange | (TC 05/14/2024 07:40:47 AM, Baghirzade, Nara) Repeatable system error when running a feed. Unable to resolve the error. Contacted cardinal health for RMA to exchange/repair under consignment. (TC 05/15/2024 06:49:00 AM, Baghirzade, Nara) SC 1003231111 (TC 05/28/2024 08:02:53 AM, Baghirzade, Nara) Received RMA. Shipped device for repair/exchange. KJV# 435529 (TC 09/10/2024 10:26:51 AM, Baghirzade, Nara) Sent follow up email regarding status of repair or if a replacement unit will be sent. (TC 10/02/2024 10:07:05 AM, Baghirzade, Nara) Received replacement pump. PO 7500206543 | Repeatable system error during feed
1062141 | 207081 | CM | CMPLT | CMPLT | 2024-05-14 07:43 | 2024-05-16 13:50 | Both power cords have bent prongs | (TC 05/14/2024 07:42:53 AM, Baghirzade, Nara) Main power cord and auxiliary power cord have bent prongs (all 3); was unable to straighten the ground prong. Requesting service from Stryker (TC 05/16/2024 01:47:24 PM, Baghirzade, Nara) Attached service report.
1062471 | 207082 | CM | CMPLT | CMPLT | 2024-05-14 07:46 | 2024-05-16 12:57 | Head portion of the bed not moving | (TC 05/14/2024 07:45:52 AM, Baghirzade, Nara) Head portion of the bed not moving; power cycled the bed and confirmed that there was nothing stuck under the bed. Requesting service from Stryker Duplicate WO closing this work order
1060032 | 207083 | CM | CMPLT | CMPLT | 2024-05-14 10:15 | 2024-05-14 10:17 | He sensor failed during patient test | (TC 05/14/2024 10:14:45 AM, Baghirzade, Nara) In the morning during calibration, the He sensor passed for DLCO. When Gaby had a patient test to do, the He sensor was giving a red light and not passing the DLCO cal. Opened ExpairTec, clicked on the fast TLCO tab, checked the helium tank for values (9.8%), adjusted the helium value to the tank value and ran the calibration. Took a picture of the old helium values in case it needs to be reverted to it. DLCO cal also passed. Gaby to confirm after next patient if its working.
1071053 | 207084 | ACC | CMPLT | CMPLT | 2024-05-14 11:50 | 2024-05-14 16:45 | Incoming inspection | (TC 05/14/2024 11:49:55 AM, Cortez, John) Unboxed and added to database
| 207085 | CM | CMPLT | CMPLT | 2024-05-14 12:40 | 2024-05-14 12:48 | "heart monitor not alarming in room 14" | (TC 05/14/2024 12:40:19 PM, Cortez, John) Tested monitor with sim and didnt come across any issues Spoke with nurse, who said that issue was at 6pm yesterday with o2 sat, which triggered no alarms They swapped sensors and X3s, issue resolved. They didnt know where original X3 went, but they may have sent O2 cables to biomed Checked for alarms in clinical audit trail and noticed that there were o2 sat alarms during throughout the timeframe when patient was in that room
| 207086 | CON | CMPLT | CMPLT | 2024-05-14 12:56 | 2024-09-25 09:47 | MSH MED ENG POCUS MWL DISCUSSION | (TC 05/14/2024 12:56:10 PM, Whittington, Scott) MSH MED ENG POCUS MWL DISCUSSION (TC 05/16/2024 01:52:53 PM, Whittington, Scott) MSH MED ENG POCUS MWL DISCUSSION (TC 05/16/2024 01:52:53 PM, Whittington, Scott) MSH MED ENG POCUS MWL DISCUSSION (TC 05/16/2024 01:52:53 PM, Whittington, Scott) MSH MED ENG POCUS MWL DISCUSSION | MSH MED ENG POCUS MWL DISCUSSION
1055643 | 207087 | CM | CMPLT | CMPLT | 2024-05-14 12:59 | 2024-05-16 12:56 | Side rail has physical damage on inner control panel | (TC 05/14/2024 12:59:34 PM, Cortez, John) Found with tag in ICU, noticed damage on inner rail, requested service form Hillrom (TC 05/16/2024 12:56:27 PM, Cortez, John) Repaired by Vince, service report attached
1055947 | 207088 | CM | CMPLT | CMPLT | 2024-05-14 13:06 | 2024-05-14 13:07 | Incoming inspection and functional check | (TC 05/14/2024 01:06:23 PM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. (TC 05/14/2024 01:06:52 PM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056522 | 207089 | CM | CMPLT | CMPLT | 2024-05-14 13:09 | 2024-05-14 13:09 | Battery replacement | (TC 05/14/2024 01:08:47 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055976 | 207090 | CM | CMPLT | CMPLT | 2024-05-14 13:09 | 2024-05-14 13:10 | Battery replacement | (TC 05/14/2024 01:09:25 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056441 | 207091 | CM | CMPLT | CMPLT | 2024-05-14 13:17 | 2024-05-14 13:18 | Repurposed for pulmonary function clinic | (TC 05/14/2024 01:16:37 PM, Baghirzade, Nara) Updated location and account on TMS; reprinted new asset tag. Completed functional testing; NIBP, temperature, and SpO2 modules are all functional. Checked and confirmed that the radical 7 has had a recent battery replacement/upgrade. Sent Gaby a message about the device being ready, will coordinate a time to drop it off and show Gaby and other PFT techs how to use the device.
1057758 | 207092 | CM | CMPLT | CMPLT | 2024-05-14 13:27 | 2024-05-14 13:29 | Restart required | (TC 05/14/2024 01:28:05 PM, Canjimavadivel, Kamallesh ) Focal point alerts prompted restart of overview, went to unit and rebooted. System ok
1071054 | 207093 | CM | CMPLT | CMPLT | 2024-05-14 13:50 | 2024-05-14 16:45 | Incoming inspection | (TC 05/14/2024 01:50:26 PM, Cortez, John) Unboxed and added to database
1071060 | 207094 | CM | CMPLT | CMPLT | 2024-05-14 13:58 | 2024-05-14 16:45 | Incoming inspection | (TC 05/14/2024 01:57:40 PM, Cortez, John) Unboxed and added to database
1046035 | 207095 | CM | CMPLT | CMPLT | 2024-05-14 13:59 | 2024-05-14 14:00 | Slight misalignment with wallstand | (TC 05/14/2024 01:59:17 PM, Fan, Tak) Auto position off. Home positions off. Recalibrated alignment.
1071061 | 207096 | CM | CMPLT | CMPLT | 2024-05-14 14:05 | 2024-05-14 16:44 | Incoming inspection | (TC 05/14/2024 02:05:04 PM, Cortez, John) Unboxed and added to database
| 207097 | CM | CMPLT | CMPLT | 2024-05-14 14:52 | 2024-05-14 14:53 | Relabelling and locating the safetynets on 11th floor | (TC 05/14/2024 02:52:09 PM, Baghirzade, Nara) Checked labelling against the serial numbers assigned on PSN. Found 11S PSN in 11N; relocated it back to 11S. A lot of the labelling was incorrect so relabelled the devices with the correct device numbers. Functional checks on all safetynets completed.
| 207098 | CM | CMPLT | CMPLT | 2024-05-14 14:54 | 2024-05-29 13:51 | Coordinating the Expair software install on computers for respirology | (TC 05/14/2024 02:53:42 PM, Baghirzade, Nara) Received asset numbers of computers in respirology from Steve. Went to south ambulatory pods to coordinate with the receptionist. Met with Gaby to confirm the scheduling as well. Tested the review software on my laptop; all ok. Will call IT for admin access to the computers. (TC 05/16/2024 11:37:13 AM, Baghirzade, Nara) Received admin access on all 6 computers in south ambulatory pods - respirology. Confirmed with Bria regarding access to exam rooms. (TC 05/17/2024 10:16:13 AM, Baghirzade, Nara) Completed software install and user configuration for all doctors in rooms 7, 9, 10, and 12. Expair is used as a review station only; no access to configuration, calibration, and back up. Users only have the option to review and view reports. Confirmed this for each doctor at each review station. Room 6 was locked -will coordinate access with Dr. Balter. Room 11 had the wrong asset number for the computer. Communicated with IT that I will require access again within 2 weeks time. Updated Steve and Dr. Blazer. (TC 05/28/2024 07:14:49 AM, Baghirzade, Nara) Coordinating with Steve, Dr. Balter, and south ambulatory pods -respirology- to install software. Dr. Balter decided not to have the software on his computer. Contacted IT for admin access for 5 computers. Confirmed the asset tags. (TC 05/29/2024 01:51:25 PM, Baghirzade, Nara) Completed the install of Expair and adjustments of all passwords to lowercase characters in the south ambulatory pod - respirology side. Updated Dr. Blazer, Steve, and all doctors that use the exam rooms.
1071073 | 207099 | ACC | CMPLT | CMPLT | 2024-05-14 14:55 | 2024-05-14 14:56 | Incoming inspection | (TC 05/14/2024 02:55:18 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071075 | 207100 | ACC | CMPLT | CMPLT | 2024-05-14 15:02 | 2024-05-14 15:03 | Incoming inspection | (TC 05/14/2024 03:02:06 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071077 | 207101 | CM | CMPLT | CMPLT | 2024-05-14 15:06 | 2024-05-14 09:56 | Incoming Inspection | Moved unit fron 6th floor storage to 12N. Unboxed and assembled the unit. Recorded info and assigned asset tag. Checked functionality. Tested ok.
1071080 | 207102 | ACC | CMPLT | CMPLT | 2024-05-14 15:13 | 2024-05-14 15:14 | Incoming inspection | (TC 05/14/2024 03:13:24 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071085 | 207103 | ACC | CMPLT | CMPLT | 2024-05-14 15:22 | 2024-05-14 15:24 | Incoming inspection | (TC 05/14/2024 03:23:54 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071084 | 207104 | CM | CMPLT | CMPLT | 2024-05-14 15:23 | 2024-05-14 09:55 | Incoming Inspection | Moved unit fron 6th floor storage to 12N. Unboxed and assembled the unit. Recorded info and assigned asset tag. Checked functionality. Tested ok.
1071089 | 207105 | ACC | CMPLT | CMPLT | 2024-05-14 15:31 | 2024-05-14 09:55 | Incoming Inspection | Moved unit fron 6th floor storage to 12N. Unboxed and assembled the unit. Recorded info and assigned asset tag Cut and discarfded boxes.
1071090 | 207106 | ACC | CMPLT | CMPLT | 2024-05-14 15:35 | 2024-05-14 15:36 | Incoming inspection | (TC 05/14/2024 03:35:03 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071091 | 207107 | ACC | CMPLT | CMPLT | 2024-05-14 15:43 | 2024-05-14 15:44 | Incoming inspection | (TC 05/14/2024 03:43:39 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071092 | 207108 | ACC | CMPLT | CMPLT | 2024-05-14 15:53 | 2024-05-14 15:54 | Incoming inspection | (TC 05/14/2024 03:53:39 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071093 | 207109 | CM | CMPLT | CMPLT | 2024-05-14 15:57 | 2024-05-13 14:00 | Incoming Inspection | Moved unit fron 6th floor storage to 12N. Unboxed and assembled the unit. Recorded info and assigned asset tag Cut and discarfded boxes. Checked functionlity. Tested ok.
1071094 | 207110 | ACC | CMPLT | CMPLT | 2024-05-14 16:00 | 2024-05-13 13:59 | Incoming Inspection | Moved unit fron 6th floor storage to 12N. Unboxed and assembled the unit. Recorded info and assigned asset tag Cut and discarfded boxes. Checked functionlity. Tested ok.
1071095 | 207111 | ACC | CMPLT | CMPLT | 2024-05-14 16:03 | 2024-05-14 16:04 | Incoming inspection | (TC 05/14/2024 04:03:22 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071096 | 207112 | ACC | CMPLT | CMPLT | 2024-05-14 16:12 | 2024-05-14 16:13 | Incoming inspection | (TC 05/14/2024 04:12:17 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071097 | 207113 | ACC | CMPLT | CMPLT | 2024-05-14 16:21 | 2024-05-14 16:22 | Incoming inspection | (TC 05/14/2024 04:21:35 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071098 | 207114 | ACC | CMPLT | CMPLT | 2024-05-14 16:29 | 2024-05-14 16:30 | Incoming inspection | (TC 05/14/2024 04:29:15 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071099 | 207115 | CM | CMPLT | CMPLT | 2024-05-14 16:39 | 2024-05-14 16:40 | Incoming inspection | (TC 05/14/2024 04:39:13 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071102 | 207116 | ACC | CMPLT | CMPLT | 2024-05-15 10:28 | 2024-05-15 10:29 | Incoming inspection | (TC 05/15/2024 10:28:20 AM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
| 207117 | PA | CMPLT | CMPLT | 2024-05-15 10:32 | 2024-05-16 12:14 | Ordering sterile water for Altrix cleaning NICU | (TC 05/15/2024 10:33:29 AM, Baghirzade, Nara) SC 1003231390 (TC 05/16/2024 12:13:36 PM, Baghirzade, Nara) Received. Order #: 4150277066
1071103 | 207118 | ACC | CMPLT | CMPLT | 2024-05-15 10:45 | 2024-05-15 10:46 | Incoming inspection | (TC 05/15/2024 10:45:43 AM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
| 207119 | CM | CMPLT | CMPLT | 2024-05-15 10:55 | 2024-05-21 15:38 | Battery order for Servos | (TC 05/15/2024 10:59:37 AM, Cortez, John) ordered x24 (TC 05/21/2024 03:38:22 PM, Cortez, John) batteries arrived
1071104 | 207120 | ACC | CMPLT | CMPLT | 2024-05-15 11:05 | 2024-05-15 11:06 | Incoming inspection | (TC 05/15/2024 11:06:00 AM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071105 | 207121 | ACC | CMPLT | CMPLT | 2024-05-15 11:13 | 2024-05-15 11:14 | Incoming inspection | (TC 05/15/2024 11:13:35 AM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1062421 | 207122 | CM | CMPLT | CMPLT | 2024-05-15 12:07 | 2024-05-15 12:08 | Left side rail, outer controls not working | (TC 05/15/2024 12:07:45 PM, Cortez, John) Confirmed issue. Power cycled bed and cleared error log Issue seems to be resolved
67970 | 207123 | CM | CMPLT | CMPLT | 2024-05-15 12:34 | 2024-05-15 12:36 | Wrong door and functional check | (TC 05/15/2024 12:34:37 PM, Baghirzade, Nara) Swapped the east door with a spare in biomed. Nurse requested a functional check as well. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 76%RH Incubator is fully functional.
1071107 | 207124 | ACC | CMPLT | CMPLT | 2024-05-15 12:40 | 2024-05-16 13:58 | Incoming Inspection | Moved unit and coordinated with BD to complete the Inspection.
69379 | 207125 | PM | CMPLT | CMPLT | 2024-05-15 12:45 | 2024-05-15 12:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/15/2024 12:44:59 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69361 | 207126 | PM | CMPLT | CMPLT | 2024-05-15 12:48 | 2024-05-15 12:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/15/2024 12:48:17 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1047866 | 207127 | PM | CMPLT | CMPLT | 2024-05-15 12:52 | 2024-05-15 12:54 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 05/15/2024 12:52:38 PM, Baghirzade, Nara) Completed full calibration of: -syringe size and position -force sensor -plunger position -syringe size sensor PM completed. See procedure document for PM results. PM schedule updated.
1054035 | 207128 | CM | CMPLT | NFF | 2024-05-15 13:05 | 2024-05-15 13:05 | Not powering on | (TC 05/15/2024 01:04:53 PM, Cortez, John) Swapped battery, charged original
69473 | 207129 | CM | CMPLT | CMPLT | 2024-05-15 13:11 | 2024-05-15 13:15 | Defective bottle side pressure sensor | (TC 05/15/2024 01:12:42 PM, Baghirzade, Nara) Downloaded error logs; noted errors for the bottle side pressure sensor. Ran analog sensor test and found that the bottle side pressure sensor stays at 4.9V. Replaced bottle side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
1061202 | 207130 | CM | CMPLT | NFF | 2024-05-15 13:20 | 2024-05-15 13:23 | Lights not bright enough | (TC 05/15/2024 01:22:41 PM, Reid, Emmerson ) Went up to check and nothing was wrong. (TC 05/15/2024 01:22:41 PM, Reid, Emmerson ) Went up to check and nothing was wrong.
1061209 | 207131 | CM | CMPLT | CMPLT | 2024-05-15 13:27 | 2024-05-15 13:29 | OR Light was lifting | (TC 05/15/2024 01:27:20 PM, Reid, Emmerson ) Light was lifting, I adjusted the tension and it is working now.
1056911 | 207132 | CM | CMPLT | CMPLT | 2024-05-15 13:31 | 2024-05-15 13:33 | Not connecting to the boom | (TC 05/15/2024 01:31:30 PM, Reid, Emmerson ) The LED on the ScaleOR was not turning on, but it was working. I connected it to AUX 2 and it worked. Not sure exactly what the issue was.
1071108 | 207133 | ACC | CMPLT | CMPLT | 2024-05-15 13:40 | 2024-05-15 13:41 | Incoming inspection | (TC 05/15/2024 01:40:32 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071109 | 207134 | ACC | CMPLT | CMPLT | 2024-05-15 13:51 | 2024-05-15 13:51 | Incoming inspection | (TC 05/15/2024 01:50:47 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
66490 | 207135 | CM | CMPLT | CMPLT | 2024-05-15 13:59 | 2024-05-15 14:01 | not holding suction levels | (TC 05/15/2024 01:59:35 PM, Rendon Lopez , Elizabeth) Cleaned, tested levels with and witout accessories. No fault found, vacuum is holding at different levels. Returned to L&D
1071111 | 207136 | ACC | CMPLT | CMPLT | 2024-05-15 14:01 | 2024-05-15 14:02 | Incoming inspection | (TC 05/15/2024 02:01:13 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
66701 | 207137 | CM | CMPLT | CMPLT | 2024-05-15 14:02 | 2024-05-15 14:05 | Monitor not printing | (TC 05/15/2024 02:01:56 PM, Rendon Lopez , Elizabeth) Cleaned recorder thermal head, replaced recorder paper. Ran calibration and demo mode for some time. Printer is working properly
1058962 | 207138 | CM | CMPLT | NFF | 2024-05-15 14:09 | 2024-05-15 14:10 | Not Working | (TC 05/15/2024 02:09:27 PM, Reid, Emmerson ) It was working
1071112 | 207139 | ACC | CMPLT | CMPLT | 2024-05-15 14:12 | 2024-05-15 14:13 | Incoming inspection | (TC 05/15/2024 02:11:42 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
| 207140 | CM | CMPLT | CMPLT | 2024-05-15 14:17 | 2024-05-15 14:18 | Holmium decommissioning meeting | (TC 05/15/2024 02:17:32 PM, Baghirzade, Nara) Holmium laser decommissioning meeting
1071114 | 207141 | ACC | CMPLT | CMPLT | 2024-05-15 14:20 | 2024-05-15 14:21 | Incoming inspection | (TC 05/15/2024 02:20:36 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071117 | 207142 | ACC | CMPLT | CMPLT | 2024-05-15 14:28 | 2024-05-15 14:29 | Incoming inspection | (TC 05/15/2024 02:28:26 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071118 | 207143 | ACC | CMPLT | CMPLT | 2024-05-15 14:42 | 2024-05-15 14:43 | Incoming inspection | (TC 05/15/2024 02:42:21 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1055171 | 207144 | CM | CMPLT | CMPLT | 2024-05-15 14:46 | 2024-05-22 13:50 | Says error and won't run medication | (TC 05/15/2024 03:40:11 PM, Rendon Lopez , Elizabeth) Downloaded error log, bottle PS fatal error seeral times. Replaced bottle PS, calibrated. Fails patient side occlusion during PM. Calibrated again, not resolved. Replaced patient PS, calibrated pressure sensors and rate, same issue (TC 05/22/2024 01:50:18 PM, Rendon Lopez , Elizabeth) Calibrated pressure several times. Reading is now in range. Completed check-in tests, all passed
1071142 | 207145 | ACC | CMPLT | CMPLT | 2024-05-15 14:52 | 2024-05-15 14:53 | Incoming inspection | (TC 05/15/2024 02:52:22 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071144 | 207146 | ACC | CMPLT | CMPLT | 2024-05-15 15:14 | 2024-05-15 15:15 | Incoming inspection | (TC 05/15/2024 03:14:28 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071241 | 207147 | ACC | CMPLT | CMPLT | 2024-05-15 15:40 | 2024-05-15 15:41 | Incoming inspection | (TC 05/15/2024 03:40:10 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071242 | 207148 | ACC | CMPLT | CMPLT | 2024-05-15 16:01 | 2024-05-15 16:02 | Incoming inspection | (TC 05/15/2024 04:01:09 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071243 | 207149 | ACC | CMPLT | CMPLT | 2024-05-15 16:14 | 2024-05-15 16:14 | Incoming inspection | (TC 05/15/2024 04:13:42 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071244 | 207150 | ACC | CMPLT | CMPLT | 2024-05-15 16:26 | 2024-05-15 16:27 | Incoming inspection | (TC 05/15/2024 04:26:10 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071245 | 207151 | ACC | CMPLT | CMPLT | 2024-05-15 16:33 | 2024-05-15 16:34 | Incoming inspection | (TC 05/15/2024 04:33:16 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071113 | 207152 | ACC | CMPLT | CMPLT | 2024-05-15 16:43 | 2024-05-15 16:43 | Incoming inspection | (TC 05/15/2024 04:43:01 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1071246 | 207153 | ACC | CMPLT | CMPLT | 2024-05-15 16:51 | 2024-05-15 16:52 | Incoming inspection | (TC 05/15/2024 04:51:28 PM, Cortez, John) Unboxed and added to database. Ran with test set, no issues.
1061215 | 207154 | CM | CMPLT | CMPLT | 2024-05-16 08:21 | 2024-05-16 08:23 | Monitor 3 not working | (TC 05/16/2024 08:21:20 AM, Reid, Emmerson ) Went to check it out while I was looking for the plastics microscope. The input was set but no signal was coming to this monitor. I went to check the input settings and it seemed to wake up the monitor and it was working.
1057870 | 207155 | CM | CMPLT | CMPLT | 2024-05-16 08:24 | 2024-05-16 08:27 | Roll stand broken | (TC 05/16/2024 08:24:58 AM, Reid, Emmerson ) Roll stand was not working properly. I was able to make some adjustments and fix it.
69487 | 207156 | PM | CMPLT | CMPLT | 2024-05-16 08:24 | 2024-05-16 08:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/15/2024 08:24:58 AM, DeGuzman, Leila) PM completed - all test passed. cleaned.
1066014 | 207157 | PM | CMPLT | CMPLT | 2024-05-16 08:29 | 2024-05-16 08:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/16/2024 08:29:51 AM, DeGuzman, Leila) PM completed - all test passed cleaned.
68653 | 207158 | PM | CMPLT | CMPLT | 2024-05-16 09:06 | 2024-05-16 09:07 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/16/2024 09:06:08 AM, DeGuzman, Leila) PM completed - all tests passed, cleaned.
1071248 | 207159 | ACC | CMPLT | CMPLT | 2024-05-16 09:33 | 2024-05-16 09:34 | Incoming Inspection | Retrieved unit from storage. Unboxed and assembled unit. Assigned asset and recorded in TMS. Checked functionallity. Tested ok. Delivered to 12N.
1061203 | 207160 | CM | CMPLT | CMPLT | 2024-05-16 10:37 | 2024-05-16 10:38 | Music not working | (TC 05/16/2024 10:37:14 AM, Reid, Emmerson ) Nurse PC wifi was having issues. Told them to call help desk.
1062471 | 207161 | CM | CMPLT | CMPLT | 2024-05-16 10:54 | 2024-05-16 11:24 | Bed headrest not moving up/down | (TC 05/16/2024 11:02:06 AM, Rendon Lopez , Elizabeth) Bed in 15M. Tested, confirmed issue. Notified Styker bed is in 15M, advised unit to not move the bed Closing WO same issue reported.
1047036 | 207162 | PM | CMPLT | CMPLT | 2024-05-16 10:59 | 2024-05-16 11:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/16/2024 10:59:26 AM, DeGuzman, Leila) PM completed - all tests passed. cleaned
66532 | 207163 | CM | CMPLT | CMPLT | 2024-05-16 11:06 | 2024-05-21 13:05 | Monitor not working | (TC 05/16/2024 11:07:07 AM, Rendon Lopez , Elizabeth) Confirmed issue, monitor not turning on. Removed from Triage and replaced with spare monitor. Issues connecting to Fetalink, tried to troubleshoot with no success, notified clinical team to place a IT ticket for fetalink (TC 05/16/2024 01:04:32 PM, Rendon Lopez , Elizabeth) Found issue with small power button board. Replaced board, cleaned, tested-OK. Cleaned recorder. (TC 05/21/2024 01:04:32 PM, Rendon Lopez , Elizabeth) Returned to Triage B, tested-OK
| 207164 | CM | CMPLT | CMPLT | 2024-05-16 11:32 | 2024-05-16 11:32 | Daily on call checks | (TC 05/16/2024 11:32:15 AM, Baghirzade, Nara) On-call daily checks completed.
1069399 | 207165 | CM | CMPLT | CMPLT | 2024-05-16 12:02 | 2024-05-16 12:05 | Skipping beats | (TC 05/16/2024 12:04:11 PM, Rendon Lopez , Elizabeth) Tested- No fault found. Returned to L&D
1071255 | 207166 | CM | CMPLT | CMPLT | 2024-05-16 13:03 | 2024-05-15 13:04 | Incoming Inspection | Unboxed and assembled. Assigned asset tag and recorded in TMS. Completed functional check and delivery accuracy. Tested ok. Delivered unit to 12N.
1071256 | 207167 | CM | CMPLT | CMPLT | 2024-05-16 13:15 | 2024-05-16 14:09 | Incoming Inspection | Unboxed and assembled. Assigned asset tag and recorded in TMS. Completed functional check and delivery accuracy. Tested ok. Delivered unit to 12N.
1071257 | 207168 | ACC | CMPLT | CMPLT | 2024-05-16 13:25 | 2024-05-15 13:25 | Incoming Inspection | Unboxed and assembled. Assigned asset tag and recorded in TMS. Completed functional check and delivery accuracy. Tested ok. Delivered unit to 12N.
1071260 | 207169 | ACC | CMPLT | CMPLT | 2024-05-16 13:29 | 2024-05-15 13:29 | Incoming Inspection | Unboxed and assembled. Assigned asset tag and recorded in TMS. Completed functional check and delivery accuracy. Tested ok. Delivered unit to 12N.
1061213 | 207170 | CM | CMPLT | CMPLT | 2024-05-16 13:31 | 2024-05-16 13:42 | OR ladder was removed | (TC 05/16/2024 01:31:22 PM, Reid, Emmerson ) Ladder appeared in biomed somehow. I investigated why and found no answers. Brought it back up and had it cleaned. I also looked into how many split legs are down in biomed right now for repair.
1071250 | 207171 | ACC | CMPLT | CMPLT | 2024-05-16 13:59 | 2024-05-16 14:00 | Asset tagging and drop off | (TC 05/16/2024 01:59:08 PM, Baghirzade, Nara) Asset tagged for inventory purposes. Dropped off with John Balcom.
1071251 | 207172 | ACC | CMPLT | CMPLT | 2024-05-16 14:00 | 2024-05-16 14:00 | Asset tagging and drop off | (TC 05/16/2024 02:00:04 PM, Baghirzade, Nara) Asset tagged for inventory purposes. Dropped off with John Balcom.
1071252 | 207173 | CM | CMPLT | CMPLT | 2024-05-16 14:03 | 2024-05-16 14:04 | Asset tagging and drop off | (TC 05/16/2024 02:03:25 PM, Baghirzade, Nara) Created asset on TMS for inventory purposes however did not put an actual asset sticker on the device as per Samir's recommendation. Dropped off with John Balcom
1071264 | 207174 | ACC | CMPLT | CMPLT | 2024-05-16 15:43 | 2024-05-16 15:43 | Incoming Inspection | Unboxed and assembled unit. Recorded in TMS and assigned asset tag. Checked accessories and functionality. Tested ok. Delivered to 12N.
| 207175 | CM | CMPLT | CMPLT | 2024-05-16 15:45 | 2024-05-16 15:46 | Suction tree issue | CAll re: suction tree not working. Went to check. Took unit to OR 12. Checked tubing connections and filter. Ok. Checked suction regulator. Found knob was damaged. Advised staff to discard the unit and order a new one.
63435 | 207176 | CM | CMPLT | NFF | 2024-05-16 16:33 | 2024-05-16 16:41 | "broken" | (TC 05/16/2024 04:33:23 PM, Cortez, John) Ran functional tests (accuracy, alarms) successfully Unable to replicate issue
1049362 | 207177 | CM | CMPLT | CMPLT | 2024-05-16 16:37 | 2024-05-17 16:01 | Contrast spill - Syringe not detected | Contrast spill - Syringe not detected Contrast was in side the front housing and piston A Injector was taken apart and cleaned Docking pin ribbon cable replaced. (TC 05/17/2024 03:57:12 PM, Debebe, Adam) Docking pin ribbon cable replaced. Tested injector functionality and put back together. Installed it back in ER CT room. Spare injector returned back to 5th floor.
63236 | 207178 | CM | CMPLT | NFF | 2024-05-16 16:43 | 2024-05-16 16:44 | "rotor malfunction" | (TC 05/16/2024 04:43:55 PM, Cortez, John) Visual check ok. Ran functional tests (accuracy, alarms) successfully Unable to replicate issue
68095 | 207179 | CM | CMPLT | NFF | 2024-05-16 16:45 | 2024-05-16 16:46 | "broken" | (TC 05/16/2024 04:45:28 PM, Cortez, John) Visual check ok. Ran functional tests (accuracy, alarms) successfully Unable to replicate issue
63207 | 207180 | CM | ACTIV | ISSUE | 2024-05-16 16:49 |  | "error showing downstream occlusion" | (TC 05/16/2024 04:49:00 PM, Cortez, John) Confirmed issue. System error 33 Will reach out to Cardinal Health (TC 07/22/2024 04:22:29 PM, Cortez, John) shipped to CH
| 207181 | CM | CMPLT | CMPLT | 2024-05-17 10:20 | 2024-05-17 10:20 | Daily on-call checks | (TC 05/17/2024 10:19:41 AM, Baghirzade, Nara) On-call daily checks completed.
1067975 | 207182 | CM | CMPLT | CMPLT | 2024-05-17 10:20 | 2024-05-17 12:54 | MSH ER PRISM NETWORK FAILED | (TC 05/17/2024 12:31:23 PM, Whittington, Scott) MSH ER PRISM NETWORK FAILED TECH ROAMED WITH THE ER PRISM AND ON RETURN WIRED NETWORK WOULD NOT CONNECT RESEATED NW CABLE AND REBOOTED OK SETUP WIRELESS AND CONNECTED VERIFIED AUTO SWITCHING BETWEEN WIRED AND WIRELESS MAC 34:98:B6:AC:8E:46 SN# 3K41265EA377F MODEL A6100-10000S | MSH ER PRISM NETWORK FAILED
| 207183 | CM | CMPLT | CMPLT | 2024-05-17 10:20 | 2024-05-17 10:21 | Zone A central station was down | (TC 05/17/2024 10:20:22 AM, Baghirzade, Nara) Zone A central display was off, there was still power to the PICIX and UPS is ok. The display adapter was loose; reconnected and display was working. Updated the team lead.
1063746 | 207184 | CM | CMPLT | CMPLT | 2024-05-17 10:22 | 2024-06-21 09:32 | Error 10 and 191; unable to run the altrix | (TC 05/17/2024 10:22:02 AM, Baghirzade, Nara) Unit was sent down as it was in an error state. Went in to service mode, cleared the RFU code, power-cycled the altrix and tried to run it at 25C. After a couple of minutes, it stops running and is in an error state; checked codes 10 and 191; all sensors are good with the exception to the pump. When running the altrix, can't hear that the pump is running and no water is running through the tubing. Will contact Stryker for support. (TC 05/24/2024 10:48:16 AM, Baghirzade, Nara) Did troubleshooting with a Stryker service tech; from the system state screen when running the heat, 2 flow sensors are not working. 2nd flow sensor is low and 3rd flow sensor is entirely defective. Will place an order for flow sensors. (TC 05/27/2024 01:45:28 PM, Baghirzade, Nara) Disassembled spare ED altrix to remove the flow sensors. Assembled the altrix back together. Disassembled the NICU altrix to get to the manifold. Replaced flow sensor 2 and flow sensor 3. (TC 05/28/2024 10:38:17 AM, Baghirzade, Nara) After assembling the device, the power was not turning on. Disassembled the device, checked the power supply board and found one of the connectors unplugged. Reconnected it and assembled the device. Ran the system state and found the flow sensors for port 2 and 3 is still not working and the ports are not lighting up. Concluding that the flow sensors that were taken off the ED altrix were also defective. Will wait until new sensors arrive from Stryker. (TC 06/20/2024 02:18:54 PM, Rhuland, Courtney) Replaced flow sensors 2&3 Flow sensor 3 working, flow sensor 2 not sensing any flow (TC 06/21/2024 09:32:17 AM, Rhuland, Courtney) Disassembled device and found the connector cable for flow sensor 2 to have a loose pin Reseated pin and all flow sensors are functional PM to be completed in seperate WO (TC 06/21/2024 12:48:50 PM, Rhuland, Courtney) After reassembly, flow sensors 1&2 were not working Replaced flow sensor cable with spare from old altrix All flow sensors showing and working properly
1061202 | 207185 | CM | ACTIV | ISSUE | 2024-05-17 10:44 |  | Aux 2 plate DVI connector issue | (TC 05/17/2024 10:45:14 AM, Reid, Emmerson ) Went with Jessica to try and attached the DVI connector to the plate, but it would not fit. Looks like we need longer jackscrews. We covered the opening with tap and left it. (TC 05/16/2024 10:44:42 AM, Huang, Jessica) got the call in OR 1 and the image from navigation system, the DVI connectors were not functional on both AUX plates. replaced the connector on the AUX 2 and tested it, working fine. informed the users (TC 05/17/2024 10:47:43 AM, Huang, Jessica) trying to replace the DVI connector on the AUX 1 and could not get it done. contacted Olympus to get the quote for the plate and was informed it was not available. trying to order some screws online (TC 06/21/2024 09:55:51 AM, Reid, Emmerson ) Went with Jessica and we were able to use the screws from the VGA connector on the DVI. I asked Rutvik to see if he could find and order these longer screws (TC 06/21/2024 10:39:32 AM, Huang, Jessica) worked with Emmerson, removed the screws from VGA and used them for the DVI connector.
1056665 | 207186 | CM | CMPLT | CMPLT | 2024-05-17 10:45 | 2024-05-21 10:55 | not showing the image on the monitors | please refer to WO 207216
1059778 | 207187 | FC | CMPLT | CMPLT | 2024-05-17 11:17 | 2024-05-17 11:20 | QC checklist | (TC 05/17/2024 11:19:48 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
500668 | 207188 | CM | CMPLT | CMPLT | 2024-05-17 12:28 | 2024-09-25 09:48 | MSH CEOU-3228(Rm3) | (TC 05/17/2024 12:27:42 PM, Whittington, Scott) MSH CEOUGE VOLUSON 4 IN ROOM 6 SOME RADS WANT TO USE THE 3D ; USED AE WITH -9, SAME AS IU22 SPARE | MSH CEOUGE VOLUSON 4 IN ROOM 6 SOME RADS WANT TO USE THE 3D ; USED AE WITH -9, SAME AS IU22 SPARE
1049608 | 207189 | CM | ACTIV | ISSUE | 2024-05-17 12:58 |  | MSH OPG SPP-HALL-D VOLUSON 10. ABRUPTLY SHUT OFF SEVERAL TIMES THIS PAST WEEK, ACCORDING TO VASSI |  | MSH OPG SPP-HALL-D VOLUSON 10. ABRUPTLY SHUT OFF SEVERAL TIMES THIS PAST WEEK, ACCORDING TO VASSI
1069363 | 207190 | CM | ACTIV | ISSUE | 2024-05-17 13:01 |  | MSH OPG SPP GE VOLUSON EXPERT 22 VASSI REPORTS THAT UP/DOWN OP PANEL CONTROL WORKING INTERMITTENTLY | (TC 06/03/2024 02:06:44 PM, Whittington, Scott) JUNE 3 MSH OPG SPP GE VOLUSON EXPERT 22 VASSI REPORTS THAT UP/DOWN OP PANEL CONTROL WORKING INTERMITTENTLY sn# F161650 REF# 1-551633119481 18006680732 | MSH OPG SPP GE VOLUSON EXPERT 22 VASSI REPORTS THAT UP/DOWN OP PANEL CONTROL WORKING INTERMITTENTLY sn# F161650 REF# 1-551633119481 18006680732
1054538 | 207191 | CM | CMPLT | CMPLT | 2024-05-17 14:46 | 2024-05-21 15:23 | "touch screen not accurate" | (TC 05/17/2024 02:46:17 PM, Cortez, John) Confirmed issue Touch calibration doesn't resolve issue. Accuracy is varies with each press Tried swapping display with spare. Touch accuracy is better, however, once in a while a press is very inaccurate (TC 05/21/2024 03:22:31 PM, Cortez, John) Leads tested ok with simulator, Touchscreen acceptable. Will return to service, but will monitor intermittent issues with touchscreen
1054985 | 207192 | CM | CMPLT | CMPLT | 2024-05-21 07:46 | 2024-06-04 13:08 | No "SD card" - sending for repair/replacement | (TC 05/21/2024 07:50:13 AM, Baghirzade, Nara) SC 1003233775 (TC 05/21/2024 01:04:37 PM, Baghirzade, Nara) Service tech from Masimo advised that I locate and re-seated the SD card for the Root, power cycle the Root, and see if it mitigates the issue. Re-seated the SD card and power cycled the root twice however after about 30 seconds, the Root alarms that there is no SD card. Reported my results to Masimo and requested a PO for repair/replacement (TC 05/28/2024 08:03:49 AM, Baghirzade, Nara) Received RMA. Shipped device for repair/exchange. KJV # 435528 (TC 06/04/2024 01:08:05 PM, Baghirzade, Nara) Received replacement Masimo. PO 7500206812
1054374 | 207193 | CM | CMPLT | NFF | 2024-05-21 08:23 | 2024-05-21 08:23 | Not sending vitals | (TC 05/21/2024 08:22:55 AM, Reid, Emmerson ) Checked the unit and there is no issue. Will see if more problems come up.
| 207194 | RDS | CMPLT | CMPLT | 2024-05-21 08:28 | 2024-05-21 08:30 | Weekly Checks For PACU | (TC 05/21/2024 08:29:45 AM, Reid, Emmerson ) Completed weekly checks.
65148 | 207195 | PM | CMPLT | CMPLT | 2024-05-21 08:34 | 2024-05-21 09:07 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/21/2024 09:06:42 AM, Reid, Emmerson ) PM Complete, see procedure for details.
1054374 | 207196 | CM | CMPLT | CMPLT | 2024-05-21 08:36 | 2024-05-21 08:43 | Upgrade to V4.3 | (TC 05/21/2024 08:43:22 AM, Reid, Emmerson ) Upgraded the pump to firmware V4.3
69202 | 207197 | PM | CMPLT | CMPLT | 2024-05-21 08:40 | 2024-05-21 09:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/21/2024 09:02:59 AM, DeGuzman, Leila) PM completed - all tests passed. cleaned
1071265 | 207198 | ACC | CMPLT | CMPLT | 2024-05-21 08:54 | 2024-05-16 08:54 | Incoming inspection | (TC 05/16/2024 08:54:18 AM, Rendon Lopez , Elizabeth) Incoming inspection
| 207199 | MEP | CMPLT | CMPLT | 2024-05-21 08:58 | 2024-07-10 10:55 | ED Final P3A | (TC 05/21/2024 08:30:00 AM, Canjimavadivel, Kamallesh ) AP installation validation with Brandon for Philips (TC 06/19/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Installed monitor in room P35 for T&O (TC 06/19/2024 02:30:00 AM, Canjimavadivel, Kamallesh ) Moved equipment for transition into space (TC 06/24/2024 02:56:00 PM, Canjimavadivel, Kamallesh ) Installed monitors in the rooms with arms. (TC 06/25/2024 03:21:14 PM, Canjimavadivel, Kamallesh ) Setting up equipment and updating TMS (TC 06/27/2024 02:59:02 PM, Canjimavadivel, Kamallesh ) Set up equipment and asset tags (TC 07/02/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Setting up space with all equipment, asset tags and doing final tests (TC 07/03/2024 09:59:29 AM, Canjimavadivel, Kamallesh ) Installed thermometers in the new (TC 07/10/2024 10:55:24 AM, Canjimavadivel, Kamallesh ) Finished setting up ED space prior to go-live
65170 | 207200 | CM | CMPLT | CMPLT | 2024-05-21 09:30 | 2024-05-21 09:31 | Upgraded to V4.3 | (TC 05/21/2024 09:30:39 AM, Reid, Emmerson ) Upgraded the pump to firmware V4.3
1071266 | 207201 | ACC | CMPLT | CMPLT | 2024-05-21 09:40 | 2024-05-16 09:40 | Incoming inspection | (TC 05/16/2024 09:40:04 AM, Rendon Lopez , Elizabeth) Incoming inspection
1071267 | 207202 | ACC | CMPLT | CMPLT | 2024-05-21 09:41 | 2024-05-16 09:41 | Incoming inspection | (TC 05/16/2024 09:40:46 AM, Rendon Lopez , Elizabeth) Incoming inspection
1071268 | 207203 | ACC | CMPLT | CMPLT | 2024-05-21 09:41 | 2024-05-16 09:41 | Incoming inspection | (TC 05/16/2024 09:41:25 AM, Rendon Lopez , Elizabeth) Incoming inspection
1071269 | 207204 | ACC | CMPLT | CMPLT | 2024-05-21 09:42 | 2024-05-17 09:42 | Incoming inspection | (TC 05/17/2024 09:42:18 AM, Rendon Lopez , Elizabeth) Incoming inspection
1071270 | 207205 | ACC | CMPLT | CMPLT | 2024-05-21 09:43 | 2024-05-17 09:43 | Incoming inspection | (TC 05/17/2024 09:43:07 AM, Rendon Lopez , Elizabeth) Incoming inspection
1071272 | 207206 | ACC | CMPLT | CMPLT | 2024-05-21 09:44 | 2024-05-17 09:44 | Incoming inspection | (TC 05/17/2024 09:44:18 AM, Rendon Lopez , Elizabeth) Incoming inspection
1071273 | 207207 | ACC | CMPLT | CMPLT | 2024-05-21 09:45 | 2024-05-17 09:45 | Incoming inspection | (TC 05/17/2024 09:45:06 AM, Rendon Lopez , Elizabeth) Incoming inspection
| 207208 | RDS | CMPLT | CMPLT | 2024-05-21 09:49 | 2024-05-21 14:40 | Weekly checks in CCU | (TC 05/21/2024 02:40:13 PM, Scero, Tomlyn) weekly checks complete. 30J tests passed
1056556 | 207209 | CM | CMPLT | CMPLT | 2024-05-21 10:25 | 2024-05-21 10:25 | Mother & Baby 17 Crash cart Fail ( connect electrodes) Error Message | (TC 05/13/2024 10:25:05 AM, Huang, Jessica) went to the unit to check the device, the electrode was connected and the green check mark on the display
| 207210 | MEP | ACTIV | ISSUE | 2024-05-21 10:25 |  | 3D Printing projects | (TC 05/21/2024 10:25:36 AM, Canjimavadivel, Kamallesh ) Modification to fridge lock design after feedback from test fit. Set up print and started. (TC 05/21/2024 10:25:36 AM, Canjimavadivel, Kamallesh ) Cleaning up left over residue from rhinoplasty mold. (TC 05/21/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Made modifications to fridge lock adapter design and sent for print (TC 05/22/2024 08:43:31 AM, Canjimavadivel, Kamallesh ) Secured lock to bracket and tested fit with fridge. Ready for install (TC 05/22/2024 01:30:49 PM, Canjimavadivel, Kamallesh ) Reached out to Ahmad from bone bank for feedback on tray we provided. He noticed slight curling which might impact the tray that will support. Modified tray design and sent for print (TC 05/22/2024 02:57:29 PM, Canjimavadivel, Kamallesh ) Prepped another nose mold and poured silicone (TC 05/23/2024 03:23:42 PM, Canjimavadivel, Kamallesh ) Fit fridge lock and installed (TC 05/23/2024 03:24:02 PM, Canjimavadivel, Kamallesh ) Removed nose from mold (TC 05/23/2024 03:24:36 PM, Canjimavadivel, Kamallesh ) Post processing tray for bone bank project. Brought to team and received feedback. Modified design with feedback (TC 05/27/2024 10:27:26 AM, Canjimavadivel, Kamallesh ) Set up 2 prints. Cleaned up supports from rhinoplasty models (TC 05/27/2024 02:27:03 PM, Canjimavadivel, Kamallesh ) Testing fit for vent knob kit (TC 05/27/2024 03:24:07 PM, Canjimavadivel, Kamallesh ) Updated vent knob design (TC 05/31/2024 03:13:01 PM, Canjimavadivel, Kamallesh ) Working on rhinoplasty model (TC 06/06/2024 08:33:42 AM, Canjimavadivel, Kamallesh ) Casted rhinoplasty model in silicone with cartilage segments. (TC 06/06/2024 01:47:33 PM, Canjimavadivel, Kamallesh ) Met with Emily to discuss modifications they would like for placenta stand (TC 06/06/2024 03:25:16 PM, Canjimavadivel, Kamallesh ) Post processing of the rhinoplasty model (TC 06/10/2024 03:16:47 PM, Canjimavadivel, Kamallesh ) Finished up placenta stand design and sent for print. Post processing the other components and started attaching with epoxy. (TC 06/11/2024 03:15:57 PM, Canjimavadivel, Kamallesh ) Applied cmultiple layers of clear coat. Designed and printed block stand pieces. (TC 06/12/2024 09:00:00 AM, Canjimavadivel, Kamallesh ) Finished up placenta stand tray (TC 06/12/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Going over workflow chart for anatomical models (TC 06/12/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Post processing of cartilage pieces and assembled rhinoplasty model. Poured silicone into mold and let it set overnight (TC 06/12/2024 03:23:47 PM, Canjimavadivel, Kamallesh ) Going over process flow chart (TC 06/13/2024 08:18:29 AM, Canjimavadivel, Kamallesh ) Removed rhinoplasty model from mold (TC 06/14/2024 08:33:16 AM, Canjimavadivel, Kamallesh ) Casting rhinoplasty model in silicone mold (TC 06/14/2024 01:34:08 PM, Canjimavadivel, Kamallesh ) Removed rhinoplasy model from mold and cleaned up model (TC 06/20/2024 03:06:38 PM, Canjimavadivel, Kamallesh ) Prepping rhinoplasty models (TC 06/25/2024 03:19:49 PM, Canjimavadivel, Kamallesh ) Going over process flow chart (TC 06/26/2024 01:58:13 PM, Canjimavadivel, Kamallesh ) Worked on proceess flow chart (TC 06/27/2024 02:59:39 PM, Canjimavadivel, Kamallesh ) Elucis demo (TC 07/17/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Finalized flow chart for surgical guides (TC 07/22/2024 10:40:18 AM, Canjimavadivel, Kamallesh ) Meet with Brandon from pharmacy to discuss potential project for fridge knob (TC 07/24/2024 03:05:06 PM, Canjimavadivel, Kamallesh ) Worked on risk doc (TC 07/31/2024 01:46:30 PM, Canjimavadivel, Kamallesh ) Working on risk doc with group (TC 07/31/2024 01:47:35 PM, Canjimavadivel, Kamallesh ) Working on rhinoplasty model (TC 07/30/2024 01:48:09 PM, Canjimavadivel, Kamallesh ) Working on rhinoplasty model (TC 08/01/2024 02:18:01 PM, Canjimavadivel, Kamallesh ) Rhinoplasty project (TC 08/02/2024 02:18:26 PM, Canjimavadivel, Kamallesh ) Rhinoplastry project (TC 08/06/2024 03:17:56 PM, Canjimavadivel, Kamallesh ) Working on rhinoplasty project (TC 08/07/2024 11:21:32 AM, Canjimavadivel, Kamallesh ) Discussion with EOS team (TC 08/07/2024 11:21:32 AM, Canjimavadivel, Kamallesh ) Working through risk doc (TC 08/07/2024 11:21:32 AM, Canjimavadivel, Kamallesh ) Working on rhinoplasty project (TC 08/09/2024 01:48:49 PM, Canjimavadivel, Kamallesh ) Working on rhinoplasty project (TC 08/12/2024 09:00:31 AM, Canjimavadivel, Kamallesh ) Prepping nose molds for rhinoplasty project (TC 08/12/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Preppring molds for rhinoplasty models (TC 08/13/2024 07:30:00 AM, Canjimavadivel, Kamallesh ) Finishing up pouring silicone for rhinoplasty models (TC 08/16/2024 11:04:07 AM, Canjimavadivel, Kamallesh ) Setting up printer for Ibrahim (TC 08/21/2024 02:40:03 PM, Canjimavadivel, Kamallesh ) Working on business case (TC 08/21/2024 02:40:03 PM, Canjimavadivel, Kamallesh ) Discussing details of vent knob project with Michelle from RT (TC 08/22/2024 11:16:49 AM, Canjimavadivel, Kamallesh ) Updated design for RT vent knob project and sent for print (TC 08/26/2024 08:28:05 AM, Canjimavadivel, Kamallesh ) Removing supports from vent knob. (TC 08/28/2024 10:37:07 AM, Canjimavadivel, Kamallesh ) Post processing of vent knob tools. Handed over 3 tools to RT (TC 08/27/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Business case (TC 09/26/2024 12:50:00 PM, Canjimavadivel, Kamallesh ) Working on business case (TC 10/03/2024 02:06:59 PM, Canjimavadivel, Kamallesh ) Modelled and printed prototype for clip requested by bone bank team (TC 10/31/2024 01:35:47 PM, Canjimavadivel, Kamallesh ) Designed prototype for transducer holder. Sent ver 1 for print (TC 11/05/2024 10:24:51 AM, Canjimavadivel, Kamallesh ) Revising tank knob model and sent for print (TC 11/26/2024 03:41:15 PM, Canjimavadivel, Kamallesh ) Business case discussion (TC 11/27/2024 02:15:12 PM, Canjimavadivel, Kamallesh ) Discuss outline for business case (TC 11/27/2024 02:15:12 PM, Canjimavadivel, Kamallesh ) Meet with Brandon from pharmacy for new flex lock bracket in 17M
1054976 | 207211 | CM | CMPLT | CMPLT | 2024-05-21 10:27 | 2024-05-21 10:27 | not working | (TC 05/13/2024 10:26:51 AM, Huang, Jessica) the power cord lose from the device so the monitor could not get charged charged the monitor, tested it, working fine
1064743 | 207212 | CM | CMPLT | CMPLT | 2024-05-21 10:28 | 2024-05-21 10:28 | not working | (TC 05/14/2024 10:28:13 AM, Huang, Jessica) replaced the SPO2 sensor and the BP cuff
541283 | 207213 | CM | CMPLT | CMPLT | 2024-05-21 10:29 | 2024-05-21 10:29 | not working | (TC 05/14/2024 10:29:22 AM, Huang, Jessica) passed calibration, problem not found
| 207214 | CM | CMPLT | CMPLT | 2024-05-21 10:31 | 2024-05-15 10:31 | checked all the AirSeal Insufflator in OR rooms
1070237 | 207215 | CM | CMPLT | CMPLT | 2024-05-21 10:34 | 2024-05-21 10:34 | low battery | (TC 05/17/2024 10:33:59 AM, Huang, Jessica) replaced batteries
1056665 | 207216 | CM | CMPLT | CMPLT | 2024-05-21 10:36 | 2024-05-21 10:36 | image not showing on the monitors | (TC 05/15/2024 10:35:49 AM, Huang, Jessica) got the call in OR 1, the image from microscope was not showing on the monitors. replaced the DVI cable and SDI cable, could not fix the issue. (TC 05/16/2024 10:37:49 AM, Huang, Jessica) took the microscope to OR 12 and troubleshooted on the device with Emmerson. found out the scaler is not functional. replaced it and informed OR team
| 207217 | CM | CMPLT | CMPLT | 2024-05-21 12:34 | 2024-06-04 10:25 | Ordering 2 chargers for ASCOM phones - will be charged back to 12N account | (TC 05/22/2024 11:27:07 AM, Samanian, Kamran) PO# 7500206811 (TC 06/04/2024 10:24:50 AM, Chukri, Samir) Received Parts .
69480 | 207218 | PM | CMPLT | CMPLT | 2024-05-21 13:07 | 2024-05-21 13:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/21/2024 01:07:38 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1054204 | 207219 | CM | CMPLT | CMPLT | 2024-05-21 13:08 | 2024-05-21 13:15 | Error code when probe cover put on | (TC 05/21/2024 01:15:09 PM, Canjimavadivel, Kamallesh ) Replaced probe Cleared error logs | E1.0 and E2.0 found in logs
68296 | 207220 | PM | CMPLT | CMPLT | 2024-05-21 13:09 | 2024-05-21 13:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/21/2024 01:08:52 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
68645 | 207221 | CM | CMPLT | CMPLT | 2024-05-21 13:11 | 2024-05-21 13:11 | Created by Failed PM WO # 202108 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/21/2024 01:10:56 PM, Baghirzade, Nara) Rate was calibrated. PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. | (TC 05/21/2024 01:09:52 PM, Baghirzade, Nara) Rate accuracy failed during PM. Rate was reading high. Will calibrate and PM.
69127 | 207222 | PM | CMPLT | CMPLT | 2024-05-21 13:12 | 2024-05-21 13:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/21/2024 01:12:21 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 207223 | CM | CMPLT | CMPLT | 2024-05-21 13:13 | 2024-05-21 13:13 | On-call daily checks | (TC 05/21/2024 01:13:07 PM, Baghirzade, Nara) On-call daily checks completed
67944 | 207224 | CM | CMPLT | CMPLT | 2024-05-21 13:14 | 2024-05-21 13:14 | Inside latch broken | (TC 05/21/2024 01:13:52 PM, Baghirzade, Nara) Supervised Laila in her repair of the inside latch for the door. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
1066462 | 207225 | FC | CMPLT | CMPLT | 2024-05-21 13:15 | 2024-05-21 13:15 | QC checklist | (TC 05/21/2024 01:14:49 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1050590 | 207226 | CM | CMPLT | CMPLT | 2024-05-21 13:17 | 2024-05-21 13:18 | Low battery - display is flickering | (TC 05/21/2024 01:17:02 PM, Baghirzade, Nara) Replaced 6 "D" batteries. Switched the scale on, display is no longer flickering. Scale is functional.
1058124 | 207227 | CM | CMPLT | CMPLT | 2024-05-21 13:27 | 2024-05-21 13:43 | Frozen screen | (TC 05/21/2024 01:43:01 PM, Canjimavadivel, Kamallesh ) Repalced battery and cleared logs. Tested ok
1062254 | 207228 | CM | CMPLT | CMPLT | 2024-05-21 13:53 | 2024-05-21 13:53 | User reported that the bed motion functions do not work | (TC 05/21/2024 01:53:13 PM, Baghirzade, Nara) "Limited motion functionality" error on display. Unplugged bed and switched off the battery; power cycled the bed. Switched battery on and plugged bed to AC; error was mitigated and all motion controls (patient and clinical side) are all functional.
1053734 | 207229 | CM | CMPLT | CMPLT | 2024-05-21 14:09 | 2024-05-21 14:09 | Functional check after use | (TC 05/21/2024 02:09:12 PM, Scero, Tomlyn) functional checks complete and passed
1058176 | 207230 | CM | CMPLT | CMPLT | 2024-05-21 14:33 | 2024-05-21 14:33 | Upgrade to V4.3 | (TC 05/21/2024 02:33:14 PM, Reid, Emmerson ) Upgraded pump to V4.3 and recored on the workbook.
1058176 | 207231 | PM | CMPLT | CMPLT | 2024-05-21 14:40 | 2024-05-21 14:52 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 05/21/2024 02:52:02 PM, Reid, Emmerson ) PM complete, see procedure for details
1058174 | 207232 | CM | CMPLT | CMPLT | 2024-05-21 15:04 | 2024-05-21 15:04 | Upgraded to V4.3 | (TC 05/21/2024 03:04:05 PM, Reid, Emmerson ) Upgraded to firmware V4.3 and marked off the workbook.
1066047 | 207233 | CM | CMPLT | CMPLT | 2024-05-21 15:24 | 2024-05-21 15:25 | No guardrail | (TC 05/21/2024 03:24:41 PM, DeGuzman, Leila) Performed check in tests - all passed. PCU plug in and shows guardrails
69081 | 207234 | CM | CMPLT | CMPLT | 2024-05-21 15:30 | 2024-05-21 15:30 | No guardrails | (TC 05/21/2024 03:29:55 PM, DeGuzman, Leila) Performed check in tests - all passed. PCU plug in and shows guardrails
1066041 | 207235 | CM | CMPLT | CMPLT | 2024-05-21 15:31 | 2024-05-21 15:31 | No guardrails | (TC 05/21/2024 03:31:24 PM, DeGuzman, Leila) Performed check in tests - all passed. PCU plug in and shows guardrails
68887 | 207236 | CM | CMPLT | CMPLT | 2024-05-21 15:33 | 2024-05-21 15:33 | No guardrails, flashing red | (TC 05/21/2024 03:33:10 PM, DeGuzman, Leila) PCU needs to plug in, connect channels, shows guardrails.
1054670 | 207237 | CM | ACTIV | ISSUE | 2024-05-21 15:58 |  | "not working" | (TC 05/21/2024 04:13:23 PM, Cortez, John) Peristaltic pump section dislodged from rest of outer casing Will reach out to Cardinal Health for replacement (TC 07/22/2024 04:21:34 PM, Cortez, John) shipped to CH
1058177 | 207238 | CM | CMPLT | CMPLT | 2024-05-22 08:26 | 2024-05-22 08:26 | Upgrade to V4.3 | (TC 05/22/2024 08:25:48 AM, Reid, Emmerson ) Upgraded firmware to V4.3
1058172 | 207239 | CM | CMPLT | CMPLT | 2024-05-22 08:46 | 2024-05-22 08:50 | Upgrade to V4.3 | (TC 05/22/2024 08:50:05 AM, Reid, Emmerson ) Upgraded to firmware V4.3 and added to workbook. Pushed APS Epidural Oct 2023
1058287 | 207240 | CM | CMPLT | CMPLT | 2024-05-22 08:47 | 2024-05-22 08:47 | System error | (TC 05/22/2024 08:47:01 AM, DeGuzman, Leila) Error code 120.4610 - PCU low battery needs to charge. Performed check-in tests passed. Ground Resistance = 249.0 mOhm Ground Leakage = 9.2 uA
68895 | 207241 | PM | CMPLT | CMPLT | 2024-05-22 09:09 | 2024-05-22 09:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/22/2024 09:09:16 AM, DeGuzman, Leila) PM completed. All tests passed. Ground Leakage: 254 mOhm Ground Resistance: 60.9 uA Battery has replaced (TC 05/27/2024 09:09:16 AM, DeGuzman, Leila) Battery conditioning: capacity is too low. Need to replaced the power board.
65123 | 207242 | CM | CMPLT | CMPLT | 2024-05-22 09:18 | 2024-05-22 09:18 | Upgraded to V4.3 | (TC 05/22/2024 09:18:13 AM, Reid, Emmerson ) Upgraded to firmware V4.3
1067340 | 207243 | CM | CMPLT | CMPLT | 2024-05-22 10:41 | 2024-05-22 10:46 | Expiratory cassette membrane needs replacement | (TC 05/22/2024 10:45:49 AM, Rendon Lopez , Elizabeth) Replaced membrane on expiratory cassette SN: 106678. pre-use checks and patient circuit tests, all passed. Returned to RT NICU
1067342 | 207244 | CM | CMPLT | CMPLT | 2024-05-22 10:52 | 2024-05-23 10:57 | Expiratory Cassette error | (TC 05/22/2024 10:57:35 AM, Rendon Lopez , Elizabeth) Ran pre-use checks an patient circuit test, all passed. Checked error logs, no errors found. Running system (TC 05/23/2024 10:57:00 AM, Rendon Lopez , Elizabeth) Ran overnight, no issues. Restarted, perfromed pre-use checks, all good. Returned to RT NICU
69615 | 207245 | CM | CMPLT | CMPLT | 2024-05-22 11:09 | 2024-05-22 11:09 | Missing red sticker | (TC 05/22/2024 11:08:45 AM, DeGuzman, Leila) Sent down, missing red sticker. Performed check-in, all tests passed.
| 207246 | CM | CMPLT | CMPLT | 2024-05-22 11:20 | 2024-05-22 11:20 | Daily on-call checks | (TC 05/22/2024 11:19:47 AM, Baghirzade, Nara) Daily on-call checks completed
68917 | 207247 | PM | CMPLT | CMPLT | 2024-05-22 11:24 | 2024-05-22 11:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/22/2024 11:24:02 AM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 151 mohms Leakage current: 64.4 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
1071295 | 207248 | ACC | CMPLT | CMPLT | 2024-05-22 11:24 | 2024-05-22 09:06 | Incoming Inspection | New bed for 12N. Moved bed into room. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 Ohm; 132 uA. Mounted Bed locator. Tested ok.
1065707 | 207249 | PM | CMPLT | CMPLT | 2024-05-22 11:28 | 2024-05-23 12:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/22/2024 11:27:57 AM, Baghirzade, Nara) PM passed and completed. Electrical Safety - Ground resistance: 125 mohms Leakage current: 62.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. Started battery conditioning. (TC 05/23/2024 12:11:15 PM, Baghirzade, Nara) Battery conditioning completed. Old battery capacity: 3400 mamphr New battery capacity: 3696 mamphr PM completed.
65888 | 207250 | CM | CMPLT | CMPLT | 2024-05-22 11:47 | 2024-05-22 11:47 | Check-in | (TC 05/13/2024 11:47:18 AM, DeGuzman, Leila) Performed check-in, all tests passed.
65899 | 207251 | CM | CMPLT | CMPLT | 2024-05-22 11:51 | 2024-05-13 11:51 | Check-in | (TC 05/13/2024 11:50:35 AM, DeGuzman, Leila) Performed Check-in, all tests passed
| 207252 | CM | CMPLT | CMPLT | 2024-05-22 12:58 | 2024-05-22 12:58 | Set up of the manikin in NICU procedure room | (TC 05/22/2024 12:57:49 PM, Baghirzade, Nara) Setting up the neo manikin in NICU procedure room with Kamran and service rep from Laerdal
1070239 | 207253 | CM | CMPLT | CMPLT | 2024-05-22 13:02 | 2024-05-21 13:03 | Low battery | (TC 05/21/2024 01:02:32 PM, Rendon Lopez , Elizabeth) Replaced batteries, tested-OK
| 207254 | PA | CMPLT | CMPLT | 2024-05-22 13:03 | 2024-05-28 13:16 | Ordering PCU power cords for biomed stock | (TC 05/22/2024 01:04:16 PM, Baghirzade, Nara) SC 1003235237 (TC 05/28/2024 01:16:30 PM, Baghirzade, Nara) PO 7500206928
1070240 | 207255 | CM | CMPLT | CMPLT | 2024-05-22 13:03 | 2024-05-21 13:03 | Low battery | (TC 05/21/2024 01:03:29 PM, Rendon Lopez , Elizabeth) Replaced batteries, tested-OK
65135 | 207256 | CM | CMPLT | CMPLT | 2024-05-22 14:27 | 2024-05-22 14:27 | Upgraded to V4.3 | (TC 05/22/2024 02:27:22 PM, Reid, Emmerson ) Upgraded to V4.3
1067992 | 207257 | CM | ACTIV | ISSUE | 2024-05-22 14:50 |  | MSH US PRISM REQUIRES TRACKBALL WITH PALM CONTROL 7 WIRELESS DONGLE | (TC 05/18/2024 02:49:49 PM, Whittington, Scott) MSH US PRISM REQUIRES TRACKBALL WITH PALM CONTROL 7 WIRELESS DONGLE | MSH US PRISM REQUIRES TRACKBALL WITH PALM CONTROL 7 WIRELESS DONGLE
1068220 | 207258 | CM | ACTIV | ISSUE | 2024-05-22 14:51 |  | MSH US PRISM REQUIRES WIRELESS DONGLE | (TC 05/18/2024 02:51:25 PM, Whittington, Scott) MSH US PRISM REQUIRES WIRELESS DONGLE | MSH US PRISM REQUIRES WIRELESS DONGLE
1054282 | 207259 | CM | CMPLT | CMPLT | 2024-05-22 15:08 | 2024-05-23 10:27 | Broken Probe To Be Replaced | (TC 05/23/2024 10:26:33 AM, Fan, Tak) 835681WX8 replaced by 1333211WX9 (TC 05/31/2024 04:09:09 PM, Fan, Tak) GE asksed for details for closing ticket.
502088 | 207260 | CM | CMPLT | CMPLT | 2024-05-22 15:17 | 2024-05-22 15:17 | Detector 2 Replacement | (TC 05/22/2024 03:17:24 PM, Paz, Ramon (Tino)) Inspected replacement detector. Uncrated the box and verified parts.
69210 | 207261 | CM | CMPLT | CMPLT | 2024-05-22 15:39 | 2024-05-23 15:39 | Error message | (TC 05/23/2024 03:39:07 PM, DeGuzman, Leila) Tag - Error message when programming infusion, (tubing already loaded properly) Device cleaned. Downloaded error log - found error code 240.4150 Performed analog sensor - bottle side sensor failed. Replaced bottle side pressure sensor: 0.9120 volts Performed check-in: all tests passed. Return to service.
1047035 | 207262 | CM | CMPLT | CMPLT | 2024-05-22 15:41 | 2024-05-23 14:07 | Channel error | (TC 05/23/2024 02:06:58 PM, DeGuzman, Leila) Tag - displaying channel error. Device cleaned. Download error log - found error code 241.4140 Performed analog sensor test - within range, passed. Performed check-in - all tests passed. Return to service.
1064418 | 207263 | CM | CMPLT | CMPLT | 2024-05-23 08:33 | 2024-05-22 08:33 | Port door broken and seals missing | (TC 05/22/2024 08:32:41 AM, Ferreira, Qwynn) replaced right port door replaced 3 seals tightened the hood handle Reviewed other components in case. done before knowing there was an incident regarding the seals previously.
68958 | 207264 | CM | CMPLT | CMPLT | 2024-05-23 08:35 | 2024-05-27 13:50 | Error code 800.8000 | (TC 05/23/2024 08:35:04 AM, Baghirzade, Nara) Downloaded the error log; noted constant error 800.8000 in the log. Checked service manual for the code and the only step is to contact BD. Emailed BD tech support. (TC 05/27/2024 01:50:22 PM, Baghirzade, Nara) Confirmed with BD that the error code does not impact the PCU and is there to catch any user related software faults. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 110 mohms Leakage current: 80.2 uA
1071301 | 207265 | CM | CMPLT | CMPLT | 2024-05-23 08:52 | 2024-05-22 13:04 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 Ohm; 131 uA. Mounted Bed locator. Tested ok.
1069503 | 207266 | CM | CMPLT | CMPLT | 2024-05-23 08:59 | 2024-05-22 09:25 | Scale needs calibration | (TC 05/22/2024 09:25:16 AM, Rendon Lopez , Elizabeth) Checked scale with weight, reading okay. Compared with other scale, same reading. Notified clinical staff to make sure when placing a blanket on the scale base doesn't touch the sides, as it could alter readings.
1071302 | 207267 | ACC | CMPLT | CMPLT | 2024-05-23 09:13 | 2024-05-22 09:13 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 40 Ohm; 132 uA. Mounted Bed locator. Tested ok.
1071303 | 207268 | ACC | CMPLT | CMPLT | 2024-05-23 09:22 | 2024-05-22 09:22 | Icoming Inspection | "New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 Ohm; 132 uA. Mounted Bed locator. Tested ok. "
1071304 | 207269 | ACC | CMPLT | CMPLT | 2024-05-23 09:34 | 2024-05-22 09:35 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 Ohm; 130 uA. Mounted Bed locator. Tested ok.
1071305 | 207270 | ACC | CMPLT | CMPLT | 2024-05-23 09:39 | 2024-05-22 13:44 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 40 Ohm; 132 uA. Mounted Bed locator. Tested ok.
1071306 | 207271 | ACC | CMPLT | CMPLT | 2024-05-23 09:48 | 2024-05-21 09:50 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 Ohm; 132 uA. Mounted Bed locator. Tested ok.
1071307 | 207272 | CM | CMPLT | CMPLT | 2024-05-23 09:53 | 2024-05-23 09:54 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 Ohm; 132 uA. Mounted Bed locator. Tested ok.
1071309 | 207273 | ACC | CMPLT | CMPLT | 2024-05-23 10:04 | 2024-05-23 15:53 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 40 mOhm; 132 uA. Mounted Bed locator. Tested ok.
1071310 | 207274 | ACC | CMPLT | CMPLT | 2024-05-23 10:07 | 2024-05-23 15:53 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 mOhm; 132 uA. Mounted Bed locator. Tested ok.
1071311 | 207275 | ACC | CMPLT | CMPLT | 2024-05-23 10:17 | 2024-05-23 15:52 | Incoming Inspection | "New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 40 mOhm; 134 uA. Mounted Bed locator. Tested ok. "
1071312 | 207276 | ACC | CMPLT | CMPLT | 2024-05-23 10:19 | 2024-05-23 15:52 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 40 mOhm; 133 uA. Mounted Bed locator. Tested ok.
1071313 | 207277 | ACC | CMPLT | CMPLT | 2024-05-23 10:27 | 2024-05-23 15:51 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 60 mOhm; 133 uA. Mounted Bed locator. Tested ok.
1071314 | 207278 | ACC | CMPLT | CMPLT | 2024-05-23 10:32 | 2024-05-23 15:51 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 mOhm; 133 uA. Mounted Bed locator. Tested ok.
1071315 | 207279 | ACC | CMPLT | CMPLT | 2024-05-23 10:37 | 2024-05-23 15:50 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 60 mOhm; 132 uA. Mounted Bed locator. Tested ok.
1071316 | 207280 | CM | CMPLT | CMPLT | 2024-05-23 10:41 | 2024-05-23 15:50 | Incoming Inspection | New bed for 12N. Moved bed into rooms. Recorded in TMS and assigned asset tag. Completed incoming inspection and EST. 50 mOhm; 133 uA. Mounted Bed locator. Tested ok.
1059745 | 207281 | CM | CMPLT | CMPLT | 2024-05-23 11:23 | 2024-05-23 11:28 | Will turn off shortly after turned on | (TC 05/23/2024 11:27:45 AM, Rendon Lopez , Elizabeth) Confirmed, no sound coming from probe. Replaced probe with spare one, tested-OK.
1055939 | 207282 | CM | CMPLT | CMPLT | 2024-05-23 11:33 | 2024-05-23 11:33 | No Temp or BP working | (TC 05/23/2024 11:32:52 AM, Ferreira, Qwynn) Swapped the BP tubing as previous one was leaking under higher pressure Temp prob was working and providing an accurate temp at 37C. If noticing temperature issues - replace temperature probe next time.
62779 | 207283 | CM | CMPLT | CMPLT | 2024-05-23 11:34 | 2024-05-23 11:34 | Wheel is broken | (TC 05/23/2024 11:34:21 AM, Ferreira, Qwynn) Lifted transport and replaced the wheel with a new one. Functional Test - OK