68977 | 207284 | PM | CMPLT | CMPLT | 2024-05-23 12:19 | 2024-05-23 12:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/23/2024 12:19:21 PM, Baghirzade, Nara) PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 201 mohms Leakage current: 69.2 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
1062032 | 207285 | CM | CMPLT | CMPLT | 2024-05-23 13:40 | 2024-05-29 10:53 | "makes noise when moving bed" | (TC 05/23/2024 01:39:45 PM, Cortez, John) Confirmed issue. Makes noise only when head section is moved. Unable to see anything wrong with actuator while patient on bed Asked ward clerk to swap bed (TC 05/23/2024 04:08:07 PM, Cortez, John) Checked bed again with patient out of bed, still didnt see anything wrong visually Put in request to Stryker (TC 05/29/2024 10:53:08 AM, Cortez, John) Mark repaired bed, service report attached
1059173 | 207286 | CM | CMPLT | CMPLT | 2024-05-23 13:46 | 2024-05-23 13:46 | Battery depleted | (TC 05/23/2024 01:46:05 PM, Reid, Emmerson ) Replaced the battery
1059141 | 207287 | CM | ACTIV | ISSUE | 2024-05-23 13:49 |  | Battery Issue | (TC 05/23/2024 01:48:48 PM, Reid, Emmerson ) Message came up that the battery should be replaced. The battery door was missing (broken off). I swapped the X3 with a spare and brought it down to biomed for repair.
1061203 | 207288 | CM | CMPLT | NFF | 2024-05-23 13:55 | 2024-05-23 13:56 | Light camera issue | (TC 05/23/2024 01:55:34 PM, Reid, Emmerson ) Went with Rutvik and looked at the light camera, but they fixed it before we got there. I proceded to show Rutvik how the cameras work and where they are kept.
68705 | 207289 | CM | CMPLT | CMPLT | 2024-05-23 14:01 | 2024-05-23 14:01 | User reported that the channels disconnect on their own | (TC 05/23/2024 02:01:28 PM, Baghirzade, Nara) Downloaded the error logs; no issues with the PCU however the latest dataset was not loaded. Left the PCU on, dataset loaded and then ran the PCU with the channels that it came down with; no issues or errors. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 170 mohms Leakage current: 73.6 uA Also did a functional check on all the channels that came down with the PCU, no issues with the LVPs.
1061446 | 207290 | CM | CMPLT | CMPLT | 2024-05-23 14:07 | 2024-05-23 16:10 | "bed alarm doesnt work" | (TC 05/23/2024 02:07:09 PM, Cortez, John) Confirmed issue. Patient seems to be centred, but bed says its unable to activate bed alarm Will wait until patient is out of bed (TC 05/23/2024 04:10:12 PM, Cortez, John) Tested again with patient out of bed. Bed alarm seems to be functional using myself and calibrated weights to test Nurse will contact me tomorrow if they still have issues
66377 | 207291 | CM | CMPLT | CMPLT | 2024-05-23 14:14 | 2024-05-23 14:15 | battery replaced
| 207292 | PA | CMPLT | CMPLT | 2024-05-23 14:18 | 2024-05-28 13:19 | Ordering BD parts for stock | (TC 05/23/2024 02:20:49 PM, Baghirzade, Nara) SC 1003236385 (TC 05/28/2024 01:19:18 PM, Baghirzade, Nara) PO 7500207027
1064645 | 207293 | CM | CMPLT | CMPLT | 2024-05-23 14:24 | 2024-05-23 14:24 | Functional check; no issue stated on the tag | (TC 05/23/2024 02:24:23 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
| 207294 | CM | CMPLT | CMPLT | 2024-05-23 14:26 | 2024-05-23 14:26 | On-call daily checks | (TC 05/23/2024 02:25:39 PM, Baghirzade, Nara) On-call daily checks completed
| 207295 | CM | CMPLT | CMPLT | 2024-05-23 14:27 | 2024-05-23 14:27 | Broken power cord for microsope in L&D triage | (TC 05/23/2024 02:27:21 PM, Baghirzade, Nara) Power cord was damaged and unusable. Replaced power cord, switched on microscope.
1071252 | 207296 | CM | CMPLT | CMPLT | 2024-05-23 14:30 | 2024-05-23 14:30 | Network information for the manikin | (TC 05/23/2024 02:29:37 PM, Baghirzade, Nara) Procedure Room Stack IP: 172.20.204.178 Tablet – 169.254.148.91 , jack : 16-0172D - Switch port 1/0/48 Laptop - 169.254.151.190 - jack : 16-0230D Switch port 2/0/47 Manikin - 169.254.59.163 – Jack 16-0170D - Switch port 1/0/47
69314 | 207297 | CM | CMPLT | CMPLT | 2024-05-23 14:34 | 2024-05-23 14:51 | Created by Failed PM WO # 205480 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/23/2024 02:51:04 PM, Rendon Lopez , Elizabeth) Replaced front case and bezel. Calibrated. PM completed, all tests passed | (TC 05/23/2024 02:33:08 PM, Rendon Lopez , Elizabeth) Door latch not staying up. and bezel ground wire damaged
1055594 | 207298 | CM | CMPLT | CMPLT | 2024-05-23 15:14 | 2024-05-23 15:15 | Move to 9th floor Psychiatry | (TC 05/23/2024 03:14:40 PM, DeGuzman, Leila) Performed functinality check - all tests passed. Place BP cuff, power cord, and SpO2 probe. Changed location to 9th floor Psychiatry.
1062936 | 207299 | CM | CMPLT | CMPLT | 2024-05-23 15:24 | 2024-05-23 15:25 | Move to 9th Floor Psychiatry | (TC 05/23/2024 03:24:34 PM, DeGuzman, Leila) Performed functinality check - all tests passed. Place BP cuff, power cord. Changed location to 9th floor Psychiatry.
1055917 | 207300 | CM | CMPLT | CMPLT | 2024-05-23 15:26 | 2024-05-23 15:27 | Move to 9th Floor Psychiatry | (TC 05/23/2024 03:26:37 PM, DeGuzman, Leila) Performed functinality check - all tests passed. Place BP cuff, power cord Changed location to 9th floor Psychiatry.
1055593 | 207301 | CM | CMPLT | CMPLT | 2024-05-23 15:29 | 2024-05-23 15:30 | Move to 9th Floor Psychiatry | (TC 05/23/2024 03:29:35 PM, DeGuzman, Leila) Performed functinality check - all tests passed. Place BP cuff, power cord Changed location to 9th floor Psychiatry.
1055619 | 207302 | CM | CMPLT | CMPLT | 2024-05-23 15:49 | 2024-05-23 15:51 | Move to 9th Floor Psychiatry, Battery replacement | (TC 05/23/2024 03:50:31 PM, DeGuzman, Leila) Battery has replaced and updated. Check function: passed. Move to 9th floor Psychiatry. Add SpO2 sensor
66922 | 207303 | CM | CMPLT | CMPLT | 2024-05-23 15:55 | 2024-05-23 15:55 | Move to 9th Floor Psychiatry, battery replacement | (TC 05/23/2024 03:54:46 PM, DeGuzman, Leila) Battery has replaced and updated. Check function: passed. Move to 9th floor Psychiatry.
1056505 | 207304 | CM | CMPLT | CMPLT | 2024-05-23 15:56 | 2024-05-23 15:59 | Move to 9th floor Psychiatry | (TC 05/23/2024 03:59:30 PM, DeGuzman, Leila) Check functions: tests passed.
1055976 | 207305 | CM | CMPLT | CMPLT | 2024-05-23 15:57 | 2024-05-23 15:58 | Move to 9th floor Psychiatry | (TC 05/23/2024 03:57:54 PM, DeGuzman, Leila) Check functions: tests passed.
1053347 | 207306 | CM | CMPLT | CMPLT | 2024-05-23 17:20 | 2024-05-23 17:20 | New tube installed. |  | New tube installed. HARP performed.
1064352 | 207307 | CM | CMPLT | CMPLT | 2024-05-23 17:27 | 2024-09-25 09:51 | MSH ICU SPARQ OOD A Unit shuts down |  | MSH ICU SPARQ OOD A Unit shuts down Called Philips and transferred logs. Ref#0123674323 JILLIAN ORDERED POWER SUPPLY, FEDEX BROUGHT IT TO TGH DOCK, BUT THEN TOOK IT BACK TO DEPOT DUE TO BE REDELIVERED MAY 24
1064420 | 207308 | CM | CMPLT | CMPLT | 2024-05-24 08:49 | 2024-05-24 08:49 | V2, V3 not working | (TC 05/24/2024 08:48:53 AM, Canjimavadivel, Kamallesh ) Confirmed issue, replaced chest lead set V1-V3. Tested okay with simulator.
| 207309 | CON | CMPLT | CMPLT | 2024-05-24 09:14 | 2024-09-25 09:51 | MSH DR RESEARCH WEIGHT LIMIT OF DETECTORS AND HOLDERS/COVERS | (TC 05/24/2024 09:13:47 AM, Whittington, Scott) MSH DR RESEARCH WEIGHT LIMIT OF DETECTORS AND HOLDERS/COVERS | MSH DR RESEARCH WEIGHT LIMIT OF DETECTORS AND HOLDERS/COVERS
1061413 | 207310 | CM | CMPLT | CMPLT | 2024-05-24 09:22 | 2024-05-28 11:06 | Damaged power cord | (TC 05/24/2024 09:22:29 AM, Baghirzade, Nara) Damaged main power cord. Contacted Stryker for service. (TC 05/28/2024 11:06:25 AM, Baghirzade, Nara) Attached service report
1060030 | 207311 | CM | CMPLT | CMPLT | 2024-05-24 09:23 | 2024-05-24 09:23 | Heater replacement | (TC 05/24/2024 09:23:08 AM, Baghirzade, Nara) Replaced the heater in the pneumotach; waited about 20 minutes for the heater to heat up. Measured the temperature around 31C using the IR thermometer. Updated Gaby
1067437 | 207312 | CM | CMPLT | CMPLT | 2024-05-24 09:32 | 2024-07-30 15:22 | Missing power supply, power cord, and defective mount | (TC 05/24/2024 09:44:51 AM, Baghirzade, Nara) As per Sri's request, inspected the VAC and found that the power supply adapter and power cord are missing. Removed the mounting mechanism to inspect why its getting stuck; found a large crack running lengthwise. Will need to order parts. Removed the VAC from 11 South (TC 05/28/2024 12:21:02 PM, Baghirzade, Nara) Followed up with Sri; was able to locate 2 power adapters and power cords. Only ordering mount. (TC 07/30/2024 03:21:34 PM, Baghirzade, Nara) Confirmed with Jessica that she had received and replaced the mount. Mount is functional.
69224 | 207313 | PM | CMPLT | CMPLT | 2024-05-24 09:47 | 2024-05-24 09:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 05/24/2024 09:47:31 AM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69207 | 207314 | FC | CMPLT | CMPLT | 2024-05-24 09:50 | 2024-05-24 09:51 | Functional test | (TC 05/24/2024 09:50:36 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1065892 | 207315 | FC | CMPLT | CMPLT | 2024-05-24 09:52 | 2024-05-24 09:52 | Functional test | (TC 05/24/2024 09:51:45 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1055181 | 207316 | FC | CMPLT | CMPLT | 2024-05-24 09:53 | 2024-05-24 09:53 | Functional test | (TC 05/24/2024 09:52:37 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
69167 | 207317 | CM | CMPLT | CMPLT | 2024-05-24 10:45 | 2024-05-24 10:45 | Calibration and functional test | (TC 05/24/2024 10:44:52 AM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
68859 | 207318 | CM | CMPLT | CMPLT | 2024-05-24 10:46 | 2024-05-24 10:48 | Power supply board replacement | (TC 05/24/2024 10:48:29 AM, DeGuzman, Leila) Replaced power supply board.
1063746 | 207319 | PA | CMPLT | CMPLT | 2024-05-24 10:50 | 2024-05-24 10:54 | Ordering flow sensors for repair | (TC 05/24/2024 10:54:12 AM, Baghirzade, Nara) SC 1003236888
1049065 | 207320 | CM | CMPLT | CMPLT | 2024-05-24 10:59 | 2024-06-14 10:43 | Created by Failed PM WO # 205514 : Planned Maintenance | (TC 06/14/2024 10:43:05 AM, Rendon Lopez , Elizabeth) Received replacement device and stand. Updated SN on TMS, Incoming inspection-OK | (TC 05/24/2024 10:57:45 AM, Rendon Lopez , Elizabeth) Found device in L&D OR with damaged hose and missing one stand caster. Contacted Stryker to receive replacement unit
| 207321 | CON | ACTIV | ISSUE | 2024-05-24 11:11 |  | MSH BI BIOPSY GUN RFP | (TC 05/24/2024 11:13:10 AM, Whittington, Scott) MSH BI BIOPSY GUN RFP | MSH BI BIOPSY GUN RFP
1050068 | 207322 | CM | CMPLT | CMPLT | 2024-05-24 14:10 | 2024-05-24 14:11 | SpO2 not working | (TC 05/24/2024 02:10:37 PM, Canjimavadivel, Kamallesh ) 2 extension cables were sent down in basket. 1 could not be repaired and was disposed. They other tested okay. Replaced BP cuff connector that was damaged.
1071336 | 207323 | CM | CMPLT | CMPLT | 2024-05-24 15:30 | 2024-05-24 09:53 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071337 | 207324 | ACC | CMPLT | CMPLT | 2024-05-24 15:32 | 2024-05-24 09:53 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071338 | 207325 | ACC | CMPLT | CMPLT | 2024-05-24 15:33 | 2024-05-24 09:52 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071339 | 207326 | CM | CMPLT | CMPLT | 2024-05-24 15:36 | 2024-05-30 09:52 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071340 | 207327 | ACC | CMPLT | CMPLT | 2024-05-24 15:41 | 2024-05-27 09:23 | Incoming | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071341 | 207328 | ACC | CMPLT | CMPLT | 2024-05-24 15:46 | 2024-05-27 09:15 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071342 | 207329 | ACC | CMPLT | CMPLT | 2024-05-24 15:52 | 2024-05-27 09:12 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071343 | 207330 | CM | CMPLT | CMPLT | 2024-05-24 15:53 | 2024-05-27 09:11 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071344 | 207331 | ACC | CMPLT | CMPLT | 2024-05-24 15:54 | 2024-05-27 09:09 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071345 | 207332 | ACC | CMPLT | CMPLT | 2024-05-24 15:54 | 2024-05-27 09:08 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071346 | 207333 | CM | CMPLT | CMPLT | 2024-05-24 15:55 | 2024-05-27 09:06 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071347 | 207334 | ACC | CMPLT | CMPLT | 2024-05-24 16:09 | 2024-05-27 09:05 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071348 | 207335 | CM | CMPLT | CMPLT | 2024-05-24 16:12 | 2024-05-27 09:04 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1071349 | 207336 | ACC | CMPLT | CMPLT | 2024-05-24 16:17 | 2024-05-27 09:03 | Incoming Inspection | Unboxed and assembled. Recorded in TMS and assigned asset tag. Checked functionality. Tested ok. Moved unit to 12N storage.
1062999 | 207337 | PM | CMPLT | CMPLT | 2024-05-24 16:30 | 2024-05-24 16:30 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/24/2024 04:30:05 PM, Cortez, John) Completed cleaning cycle
1046035 | 207338 | CM | CMPLT | CMPLT | 2024-05-24 20:20 | 2024-05-14 20:21 | MSH ER DRX EVO E5-02(ER) COLLIMATOR ISSUES, UNABLE TO MANUALLY CONTROL | (TC 05/12/2024 12:20:00 PM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) COLLIMATOR ISSUES, UNABLE TO MANUALLY CONTROL PROVIDED KEY FOR MANUAL OPERATION ROOM BUSY , PRIMARY PROBLEM WITH TABLE BUCKY AND ACSS TO BE RESOLVED | MSH ER DRX EVO E5-02(ER) COLLIMATOR ISSUES, UNABLE TO MANUALLY CONTROL
| 207339 | CM | CMPLT | CMPLT | 2024-05-27 09:40 | 2024-05-27 09:40 | Weekly checks in CCU | (TC 05/27/2024 09:40:16 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
62696 | 207340 | CM | CMPLT | CMPLT | 2024-05-27 10:29 | 2024-05-27 10:29 | Not charging | (TC 05/27/2024 10:29:04 AM, Canjimavadivel, Kamallesh ) Plugged in and saw charging indicator illuminate. Battery initially low but charging as expected. Battery diagnostics ok. Tested leads with simulator. Al lfunctions good. Will charge full before returning
1063377 | 207341 | CM | CMPLT | CMPLT | 2024-05-27 11:00 | 2024-05-27 11:00 | replaced batteries
1047183 | 207342 | CM | CMPLT | CMPLT | 2024-05-27 11:01 | 2024-05-27 11:01 | replaced batteries
1063375 | 207343 | CM | CMPLT | CMPLT | 2024-05-27 11:09 | 2024-05-27 11:09 | replaced batteries
1056665 | 207344 | CM | CMPLT | CMPLT | 2024-05-27 11:10 | 2024-05-27 11:10 | blurry image | (TC 05/22/2024 11:10:40 AM, Huang, Jessica) approached by the TL and the surgeon DR. O' Neill complained the image was blurry checked the device and it seemed functional informed TL
1056553 | 207345 | CM | CMPLT | CMPLT | 2024-05-27 11:13 | 2024-05-27 11:13 | replaced the battery | (TC 05/22/2024 11:12:56 AM, Huang, Jessica) replaced the battery and checked the device functional
| 207346 | CM | CMPLT | CMPLT | 2024-05-27 11:15 | 2024-05-27 11:15 | masimo monitor error message 16N client view 1 and 2 down
1053347 | 207347 | CM | CMPLT | CMPLT | 2024-05-27 13:03 | 2024-09-25 09:52 | MSH BI 1270 H1 3DIMENSIONS UPS UPGRADE | (TC 05/27/2024 01:03:32 PM, Whittington, Scott) MSH BI 1270 H1 3DIMENSIONS UPS UPGRADE (TC 06/05/2024 12:43:59 PM, Whittington, Scott) MSH BI 1270 H1 3DIMENSIONS UPS UPGRADE TWO OF THE ORIGINAL UPS DELIVERED WERE DOA REBUILDS 2 REPLACEMENTS RECEIVED AND PLACED IN H1 | MSH BI 1270 H1 3DIMENSIONS UPS UPGRADE
1053348 | 207348 | CM | CMPLT | CMPLT | 2024-05-27 13:06 | 2024-09-25 09:52 | MSH BI 1272 H3 3DIMENSIONS UPS UPGRADE | (TC 05/27/2024 01:05:48 PM, Whittington, Scott) MSH BI 1272 H3 3DIMENSIONS UPS UPGRADE (TC 06/06/2024 01:19:54 PM, Whittington, Scott) MSH BI 1272 H3 3DIMENSIONS UPS UPGRADE | MSH BI 1272 H3 3DIMENSIONS UPS UPGRADE
1053349 | 207349 | CM | CMPLT | CMPLT | 2024-05-27 13:09 | 2024-09-25 09:53 | MSH BI 1274 H2 3DIMENSIONS UPS UPGRADE | (TC 05/27/2024 01:09:22 PM, Whittington, Scott) MSH BI 1270 H2 3DIMENSIONS UPS UPGRADE (TC 06/06/2024 01:20:59 PM, Whittington, Scott) MSH BI 1274 H2 3DIMENSIONS UPS UPGRADE | MSH BI 1270 H2 3DIMENSIONS UPS UPGRADE
| 207350 | CM | CMPLT | CMPLT | 2024-05-27 13:47 | 2024-05-27 13:47 | Biomed huddle | (TC 05/27/2024 01:47:01 PM, Baghirzade, Nara) Biomed huddle
1051375 | 207351 | CM | CMPLT | CMPLT | 2024-05-27 13:48 | 2024-05-27 13:48 | Damaged power cord | (TC 05/27/2024 01:48:12 PM, Baghirzade, Nara) Replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 115 mohms Leakage current: 66.1 uA
| 207352 | RDS | CMPLT | CMPLT | 2024-05-27 14:29 | 2024-05-27 14:51 | Weekly check for ED
1049901 | 207353 | CM | CMPLT | CMPLT | 2024-05-27 15:28 | 2024-06-03 11:18 | MSH R61 SYSTEM STALLED IN NICU , TELESCOPE UP , NO POWER | (TC 05/27/2024 03:27:45 PM, Whittington, Scott) MSH R61 SYSTEM STALLED IN NICU , TELESCOPE UP , NO POWER REPLACED SPAG4659 CORD REEL FROM LOCAL STOCK ALLOWW TO CHARGE. SITTING AT 22% BATTERY CAPACITY WHEN AC POWER RESTORED ORDER REPLACEMENTS SPAG4659 CORD REEL (TC 05/31/2024 03:59:00 PM, Whittington, Scott) MSH R61 SYSTEM STALLED IN NICU , TELESCOPE UP , NO POWER NEW CORD REEL DOES NOT LOCK CABLE PARKED ON NICU 17, CART BATTERIES MAY HAVE BEEN COMPROMISED BY DRAINGING WILL COULD NOT MOVE SYSTEM, HE USED OVER RIDE BUTTON AND PUSHED TO DR ACOVE MSH5 FOUND SYSTEM UNABLE TO MOVE TUBEHEAD DIAGNOSTICS INDICATED AN INTERMITTENT RIGHT HAND HANDLE BAR SWITCH ABLE TO GET SYSTEM TO MOVE (TC 05/31/2024 08:43:00 PM, Whittington, Scott) MSH R61 SYSTEM STALLED IN NICU , TELESCOPE UP , NO POWER NEW CORD REEL DOES NOT LOCK CABLE PARKED ON NICU 17, CART BATTERIES MAY HAVE BEEN COMPROMISED BY DRAINGING WILL COULD NOT MOVE SYSTEM, HE USED OVER RIDE BUTTON AND PUSHED TO DR ACOVE MSH5 FOUND SYSTEM UNABLE TO MOVE TUBEHEAD DIAGNOSTICS INDICATED AN INTERMITTENT RIGHT HAND HANDLE BAR SWITCH ABLE TO GET SYSTEM TO MOVE (TC 05/27/2024 05:45:00 PM, Whittington, Scott) MSH R61 SYSTEM STALLED IN NICU , TELESCOPE UP , NO POWER (TC 06/03/2024 11:17:37 AM, Whittington, Scott) MSH R61 SYSTEM STALLED IN NICU , TELESCOPE UP , NO POWER NEW CORD REEL DOES NOT LOCK CABLE PARKED ON NICU 17, CART BATTERIES MAY HAVE BEEN COMPROMISED BY DRAINGING WILL COULD NOT MOVE SYSTEM, HE USED OVER RIDE BUTTON AND PUSHED TO DR ACOVE MSH5 FOUND SYSTEM UNABLE TO MOVE TUBEHEAD DIAGNOSTICS INDICATED AN INTERMITTENT RIGHT HAND HANDLE BAR SWITCH ABLE TO GET SYSTEM TO MOVE CAL HANDLE BAR JUNE 3 SYSTEM IN USE ALL MORNIG FOR PORTABLE BY LAK, TO BE RETURNED TO NICU | MSH R61 SYSTEM STALLED IN NICU , TELESCOPE UP , NO POWER NEW CORD REEL DOES NOT LOCK CABLE PARKED ON NICU 17, CART BATTERIES MAY HAVE BEEN COMPROMISED BY DRAINGING WILL COULD NOT MOVE SYSTEM, HE USED OVER RIDE BUTTON AND PUSHED TO DR ACOVE MSH5 FOUND SYSTEM UNABLE TO MOVE TUBEHEAD DIAGNOSTICS INDICATED AN INTERMITTENT RIGHT HAND HANDLE BAR SWITCH ABLE TO GET SYSTEM TO MOVE
1071370 | 207354 | ACC | CMPLT | CMPLT | 2024-05-27 19:51 | 2024-05-31 08:50 | MSH 5 US CANON PRISM PROBE PVT-745BTH SN# 99D241243 ACCEPTANCE | (TC 05/27/2024 07:54:12 PM, Whittington, Scott) MSH 5 US CANON PRISM PROBE PVT-745BTH SN# 99D241243 ACCEPTANCE (TC 05/31/2024 07:55:00 AM, Whittington, Scott) MSH 5 US CANON PRISM PROBE PVT-745BTH SN# 99D241243 ACCEPTANCE | MSH 5 US CANON PRISM PROBE PVT-745BTH SN# 99D241243 ACCEPTANCE
1067437 | 207355 | PA | CMPLT | CMPLT | 2024-05-28 07:32 | 2024-07-30 15:20 | Ordering mount mechanism | (TC 05/28/2024 07:35:14 AM, Baghirzade, Nara) SC 1003238843 (TC 07/30/2024 03:20:21 PM, Baghirzade, Nara) Received on June 5th. PO 7500207088
1054993 | 207356 | CM | CMPLT | CMPLT | 2024-05-28 08:54 | 2024-05-17 08:54 | Functional tests. Move to 12th Floor North | (TC 05/17/2024 08:53:46 AM, DeGuzman, Leila) Functional tests passed. Rutvik changed location to 12th Floor North.
1055933 | 207357 | CM | CMPLT | CMPLT | 2024-05-28 08:59 | 2024-05-17 08:59 | Functional tests. Move to 12th floor North | (TC 05/17/2024 08:58:52 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed location to 12th floor North.
1055903 | 207358 | CM | CMPLT | CMPLT | 2024-05-28 09:03 | 2024-05-28 09:03 | Funtional tests. Moved to 12th floor North | (TC 05/17/2024 09:02:43 AM, DeGuzman, Leila) Performed funtional tests, all passed. Rutvik changed location to 12th floor North.
1056209 | 207359 | CM | CMPLT | CMPLT | 2024-05-28 09:05 | 2024-05-17 09:06 | Functional tests, moved to 12th floor North | (TC 05/17/2024 09:05:36 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed location to 12th floor North.
1055910 | 207360 | CM | CMPLT | CMPLT | 2024-05-28 09:12 | 2024-05-17 09:13 | Funtional tests, moved to 12th floor North | (TC 05/17/2024 09:12:34 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed location to 12th floor North.
1055585 | 207361 | CM | CMPLT | CMPLT | 2024-05-28 09:16 | 2024-05-17 09:16 | Functional tests, moved to 12th floor North | (TC 05/17/2024 09:15:47 AM, DeGuzman, Leila) Performed functional tests, all passed Rutvik changed location to 12th floor North.
1056461 | 207362 | CM | CMPLT | CMPLT | 2024-05-28 09:19 | 2024-05-17 09:19 | Functional tests, moved to 12th floor North | (TC 05/17/2024 09:18:37 AM, DeGuzman, Leila) Performed functional tets, all passed. Rutvik changed location to 12th floor North.
1056432 | 207363 | CM | CMPLT | CMPLT | 2024-05-28 09:22 | 2024-05-17 09:22 | Funtional tests, moved to 12th floor North | (TC 05/17/2024 09:22:07 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed the location to 12th floor North.
1054997 | 207364 | CM | CMPLT | CMPLT | 2024-05-28 09:24 | 2024-05-17 09:26 | Functional tests, move to 12th floor North | (TC 05/17/2024 09:26:22 AM, DeGuzman, Leila) Performed functional tests, all passed. Battery not charging. Rutvik changed the location to 12th floor North. (TC 05/29/2024 09:26:22 AM, DeGuzman, Leila) Waiting for battery replacement
1055899 | 207365 | CM | CMPLT | CMPLT | 2024-05-28 09:30 | 2024-05-17 09:30 | Functional tests, moved to 12th floor North. | (TC 05/17/2024 09:29:59 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed the location to 12th floor North.
1055905 | 207366 | CM | CMPLT | CMPLT | 2024-05-28 09:32 | 2024-05-17 09:32 | Functional tests, moved to 12th floor North. | (TC 05/17/2024 09:31:53 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed the location to 12th floor North.
1055908 | 207367 | CM | CMPLT | CMPLT | 2024-05-28 09:34 | 2024-05-17 09:34 | Functional tests, moved to 12th floor North. | (TC 05/17/2024 09:34:26 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed the location to 12th floor North.
1055904 | 207368 | CM | CMPLT | CMPLT | 2024-05-28 09:36 | 2024-05-17 09:36 | Functional tests, moved to 12th floor North. | (TC 05/17/2024 09:36:10 AM, DeGuzman, Leila) Performed functional tests, all passed. Rutvik changed the location to 12th floor North.
1054975 | 207369 | CM | CMPLT | CMPLT | 2024-05-28 09:42 | 2024-05-27 09:42 | Functional tests, moved to 12th floor north | (TC 05/27/2024 09:42:00 AM, DeGuzman, Leila) Performed functional tests, all passed. Changed the location to 12th floor North.
| 207370 | CM | CMPLT | CMPLT | 2024-05-28 09:42 | 2024-06-20 08:11 | Power strip order to O.R requested by Steve and charge to R.T Anesthesia | (TC 05/28/2024 01:06:06 PM, Chukri, Samir) Search and ordered Power strip for O.R requested by Steve from R.T , charge the cost to R.T Anesthesia account . (TC 06/04/2024 08:20:15 AM, Chukri, Samir) Received 9 units from 10 (TC 06/20/2024 08:10:34 AM, Chukri, Samir) Received 10 power strip from Digikey and the remain holder 1 from Amazon . delivered to Emmerson.
1055900 | 207371 | CM | CMPLT | CMPLT | 2024-05-28 09:44 | 2024-05-27 09:44 | Functional tests, moved to 12th floor North. | (TC 05/27/2024 09:43:53 AM, DeGuzman, Leila) Performed functional tests, all passed. Changed the location to 12th floor North.
1056423 | 207372 | CM | CMPLT | CMPLT | 2024-05-28 09:46 | 2024-05-27 09:46 | Functional tests, moved to 12th floor North. | (TC 05/27/2024 09:46:29 AM, DeGuzman, Leila) Performed functional tests, all passed. Changed the location to 12th floor North.
1056430 | 207373 | CM | CMPLT | CMPLT | 2024-05-28 09:53 | 2024-05-27 09:53 | Functional tests, moved to 12th floor North. | (TC 05/27/2024 09:53:13 AM, DeGuzman, Leila) Performed functional tests, all passed. Changed the location to 12th floor North.
1055024 | 207374 | CM | CMPLT | CMPLT | 2024-05-28 10:39 | 2024-05-27 10:42 | Battery Replacement | (TC 05/27/2024 10:42:50 AM, DeGuzman, Leila) Replaced battery. Changed location to 12th floor North.
1055959 | 207375 | CM | CMPLT | CMPLT | 2024-05-28 10:47 | 2024-05-27 10:47 | Battery Replacement | (TC 05/27/2024 10:46:54 AM, DeGuzman, Leila) Replaced battery. Changed location to 12th floor North.
| 207376 | RDS | CMPLT | CMPLT | 2024-05-28 10:48 | 2024-05-28 10:49 | Weekly Checks For PACU | (TC 05/28/2024 10:49:10 AM, Reid, Emmerson ) Round complete, also fixed 4 SpO2 cables and gathered 4 CADD pumps for PM
1055609 | 207377 | CM | CMPLT | CMPLT | 2024-05-28 10:49 | 2024-05-27 10:49 | Battery Replacement | (TC 05/27/2024 10:49:16 AM, DeGuzman, Leila) Replaced battery. Changed location to 12th floor North.
1061202 | 207378 | CM | CMPLT | CMPLT | 2024-05-28 10:51 | 2024-05-28 10:51 | OR 1 wanted to move insuflator to boom | (TC 05/28/2024 10:51:20 AM, Reid, Emmerson ) Hose was already connected, they wanted confirmation it was okay. It was good.
1055964 | 207379 | CM | CMPLT | CMPLT | 2024-05-28 10:51 | 2024-05-27 10:52 | Battery Replacement | (TC 05/27/2024 10:51:31 AM, DeGuzman, Leila) Replaced battery. Changed location to 12th floor North.
1055958 | 207380 | CM | CMPLT | CMPLT | 2024-05-28 10:56 | 2024-05-27 10:56 | Battery Replacement | (TC 05/27/2024 10:56:29 AM, DeGuzman, Leila) Replaced the battery. Changed location to 12th floor North
1056490 | 207381 | CM | CMPLT | CMPLT | 2024-05-28 10:58 | 2024-05-27 10:59 | Battery Replacement | (TC 05/27/2024 10:58:35 AM, DeGuzman, Leila) Replaced the battery. Change location to 12th floor North.
1056512 | 207382 | CM | CMPLT | CMPLT | 2024-05-28 11:00 | 2024-05-27 11:00 | Battery Replacement | (TC 05/27/2024 11:00:13 AM, DeGuzman, Leila) Replaced the battery. Change location to 12th floor North
1055610 | 207383 | CM | CMPLT | CMPLT | 2024-05-28 11:02 | 2024-05-27 11:03 | Battery Replacement | (TC 05/27/2024 11:03:14 AM, DeGuzman, Leila) Replaced the battery. Changed location to 12th Floor North.
1055971 | 207384 | CM | CMPLT | CMPLT | 2024-05-28 11:05 | 2024-05-27 11:05 | Battery Replacement | (TC 05/27/2024 11:05:08 AM, DeGuzman, Leila) Replaced the battery. Change location to 12th floor North
1056218 | 207385 | CM | CMPLT | CMPLT | 2024-05-28 11:08 | 2024-05-27 11:08 | Battery Replacement | (TC 05/27/2024 11:08:22 AM, DeGuzman, Leila) Replaced the battery. Changed the location to 12th floor North.
1056488 | 207386 | CM | CMPLT | CMPLT | 2024-05-28 11:12 | 2024-05-27 11:12 | Battery Replacement | (TC 05/27/2024 11:12:20 AM, DeGuzman, Leila) Replaced the battery. Changed the location to 12th floor North.
1055992 | 207387 | CM | CMPLT | CMPLT | 2024-05-28 11:14 | 2024-05-27 11:14 | Battery Replacement | (TC 05/27/2024 11:14:09 AM, DeGuzman, Leila) Replaced the battery. Changed the location to 12th Floor North.
1061203 | 207388 | CM | CMPLT | CMPLT | 2024-05-28 11:29 | 2024-05-28 11:29 | Issue with Connecting to monitors | (TC 05/28/2024 11:28:36 AM, Reid, Emmerson ) They fixed it by strightening a pin on the DVI connector but wanted a spare so I gave them one.
1056437 | 207389 | CM | CMPLT | CMPLT | 2024-05-28 12:17 | 2024-05-28 12:17 | NIBP hose leaking | (TC 05/28/2024 12:17:14 PM, Baghirzade, Nara) NIBP hose was improperly connected to the Root; re-seated the connector. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 207390 | PA | CMPLT | CMPLT | 2024-05-28 14:01 | 2024-06-03 07:51 | Ordering flat network cables for NICU manikin | (TC 06/03/2024 07:51:19 AM, Baghirzade, Nara) Ordered through Amazon by Samir. Received order.
68895 | 207391 | CM | CMPLT | CMPLT | 2024-05-28 14:02 | 2024-05-30 08:07 | PCU power supply board replacement | (TC 05/28/2024 02:02:52 PM, DeGuzman, Leila) Replaced power board. (TC 05/30/2024 08:06:40 AM, DeGuzman, Leila) 05/29/2024: Run battery conditioning test First run: 3953 mamphr And run again.. (05/31/2024, De Guzman Leila) After the sencond run, battery capacity: 3943 mamphr
39506 | 207392 | CM | CMPLT | CMPLT | 2024-05-28 15:44 | 2024-05-27 15:44 | Error 3026 | Error 3026 replaced Power board on side B
1049362 | 207393 | CM | CMPLT | CMPLT | 2024-05-28 15:48 | 2024-05-24 15:47 | Contrast spill | Contrast spill Removed and cleaned piston tip on Side A Brought down spare injector - was faulty with error 3026 (repiared) Moved GenRad CT injector and installed it in ER for the weekend.
1048334 | 207394 | CM | CMPLT | CMPLT | 2024-05-29 07:22 | 2024-05-29 07:22 | SpO2 not working | (TC 05/29/2024 07:22:06 AM, Baghirzade, Nara) Damaged pin on SpO2 cable. Defective LNC cable. Replaced SpO2 cable and LNC cable. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 207395 | CM | CMPLT | CMPLT | 2024-05-29 07:32 | 2024-05-27 07:32 | Removed MX800 and X2 off stand; no longer needed in NICU | (TC 05/27/2024 07:31:45 AM, Baghirzade, Nara) Removed the research MX800 and X2 on a rolling stand as per Carol's request. Brought down to biomed, took MX800 and X2 off stand; MX800 is now the spare for NICU. X2 was given to Courtney for PM. Removed mounting plate off the rolling stand and stored the rolling stand in the MCR room (as per Kamran's request).
1051526 | 207396 | CM | CMPLT | CMPLT | 2024-05-29 09:07 | 2024-05-29 09:08 | Calibration Verification | (TC 05/29/2024 09:07:44 AM, Rhuland, Courtney) Temperature verified with 9600 Plus Calibration tester 36.0 C = 36.0 C 38.5 C = 38.5 C 41.0 C = 41.0 C Verification passed
1071378 | 207397 | ACC | CMPLT | CMPLT | 2024-05-29 09:11 | 2024-05-30 11:48 | Incoming inspection | (TC 05/30/2024 11:48:19 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071379 | 207398 | ACC | CMPLT | CMPLT | 2024-05-29 09:12 | 2024-05-30 11:49 | Incoming inspection | (TC 05/30/2024 11:48:52 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071380 | 207399 | ACC | CMPLT | CMPLT | 2024-05-29 09:14 | 2024-05-30 11:49 | Incoming inspection | (TC 05/30/2024 11:49:21 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071381 | 207400 | ACC | CMPLT | CMPLT | 2024-05-29 09:14 | 2024-05-30 11:50 | Incoming inspection | (TC 05/30/2024 11:49:35 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071382 | 207401 | ACC | CMPLT | CMPLT | 2024-05-29 09:15 | 2024-05-30 11:50 | Incoming inspection | (TC 05/30/2024 11:49:49 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071383 | 207402 | ACC | CMPLT | CMPLT | 2024-05-29 09:16 | 2024-05-30 11:50 | Incoming inspection | (TC 05/30/2024 11:50:01 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071384 | 207403 | ACC | CMPLT | CMPLT | 2024-05-29 09:17 | 2024-05-30 11:50 | Incoming inspection | (TC 05/30/2024 11:50:14 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071385 | 207404 | ACC | CMPLT | CMPLT | 2024-05-29 09:21 | 2024-05-30 11:51 | Incoming inspection | (TC 05/30/2024 11:51:23 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071386 | 207405 | ACC | CMPLT | CMPLT | 2024-05-29 09:22 | 2024-05-30 11:52 | Incoming inspection | (TC 05/30/2024 11:51:42 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071387 | 207406 | ACC | CMPLT | CMPLT | 2024-05-29 09:23 | 2024-05-30 11:52 | Incoming inspection | (TC 05/30/2024 11:51:56 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071388 | 207407 | ACC | CMPLT | CMPLT | 2024-05-29 09:24 | 2024-05-30 11:52 | Incoming inspection | (TC 05/30/2024 11:52:20 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071389 | 207408 | ACC | CMPLT | CMPLT | 2024-05-29 09:25 | 2024-05-30 11:53 | Incoming inspection | (TC 05/30/2024 11:52:33 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071390 | 207409 | ACC | CMPLT | CMPLT | 2024-05-29 09:26 | 2024-05-30 11:54 | Incoming inspection | (TC 05/30/2024 11:53:36 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071391 | 207410 | ACC | CMPLT | CMPLT | 2024-05-29 09:27 | 2024-05-30 11:54 | Incoming inspection | (TC 05/30/2024 11:53:50 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071392 | 207411 | ACC | CMPLT | CMPLT | 2024-05-29 09:28 | 2024-05-30 11:54 | Incoming inspection | (TC 05/30/2024 11:54:03 AM, Baghirzade, Nara) Unboxed and asset tagged.
1071393 | 207412 | ACC | CMPLT | CMPLT | 2024-05-29 09:28 | 2024-05-30 11:54 | Incoming inspection | (TC 05/30/2024 11:54:15 AM, Baghirzade, Nara) Unboxed and asset tagged.
| 207413 | TEDU | CMPLT | CMPLT | 2024-05-29 10:19 | 2024-05-28 10:19 | CMBES Conference | (TC 05/28/2024 10:18:44 AM, Scero, Tomlyn) Attended CESO/CMBES Conference from 730-330
1051527 | 207414 | CM | CMPLT | CMPLT | 2024-05-29 10:20 | 2024-05-29 10:20 | Calibration verification | (TC 05/29/2024 10:19:58 AM, Rhuland, Courtney) Verified accuracy with Rigel Uni-Sim 100/60 mmHg = 96/66 mmHg 120/80 mmHg = 118/82 mmHg 140/100 mmHg = 136/98 mmHg 180/140 mmHg = 174/148 mmHg Accuracy passed, within tolerance
1071395 | 207415 | ACC | CMPLT | CMPLT | 2024-05-29 10:26 | 2024-05-29 12:09 | Incoming inspection | (TC 05/29/2024 12:09:02 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1071396 | 207416 | CM | CMPLT | CMPLT | 2024-05-29 10:26 | 2024-05-29 12:07 | Incoming inspection | (TC 05/29/2024 12:06:47 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Removed power cord guard at the rear of the incubator.
1071397 | 207417 | ACC | CMPLT | CMPLT | 2024-05-29 10:27 | 2024-05-29 12:10 | Incoming inspection | (TC 05/29/2024 12:09:52 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1071403 | 207418 | ACC | CMPLT | CMPLT | 2024-05-29 10:29 | 2024-05-29 12:11 | Incoming inspection | (TC 05/29/2024 12:10:45 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1071404 | 207419 | ACC | CMPLT | CMPLT | 2024-05-29 10:30 | 2024-05-29 12:11 | Incoming inspection | (TC 05/29/2024 12:11:29 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1071406 | 207420 | ACC | CMPLT | CMPLT | 2024-05-29 10:30 | 2024-05-29 12:55 | Incoming inspection | (TC 05/29/2024 12:54:42 PM, Rhuland, Courtney) Incoming inspection completed.
1071408 | 207421 | ACC | CMPLT | CMPLT | 2024-05-29 10:31 | 2024-05-29 12:49 | Incoming inspection | (TC 05/29/2024 12:49:07 PM, Rhuland, Courtney) Incoming inspection completed.
1071409 | 207422 | ACC | CMPLT | CMPLT | 2024-05-29 10:32 | 2024-05-29 12:12 | Incoming inspection | (TC 05/29/2024 12:12:05 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1071410 | 207423 | ACC | CMPLT | CMPLT | 2024-05-29 10:32 | 2024-05-29 12:13 | Incoming inspection | (TC 05/29/2024 12:12:52 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed.
1071398 | 207424 | ACC | CMPLT | CMPLT | 2024-05-29 10:35 | 2024-05-29 12:02 | Incoming inspection | (TC 05/29/2024 12:02:17 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Removed power cord guard at the rear of the incubator.
1071399 | 207425 | ACC | CMPLT | CMPLT | 2024-05-29 10:35 | 2024-05-29 12:05 | Incoming inspection | (TC 05/29/2024 12:05:05 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Removed power cord guard at the rear of the incubator.
1071400 | 207426 | ACC | CMPLT | CMPLT | 2024-05-29 10:35 | 2024-05-29 12:06 | Incoming inspection | (TC 05/29/2024 12:05:42 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Removed power cord guard at the rear of the incubator.
1071401 | 207427 | ACC | CMPLT | CMPLT | 2024-05-29 10:35 | 2024-05-29 12:04 | Incoming inspection | (TC 05/29/2024 12:04:28 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Removed power cord guard at the rear of the incubator.
1071402 | 207428 | ACC | CMPLT | CMPLT | 2024-05-29 10:35 | 2024-05-29 12:47 | Incoming inspection | (TC 05/29/2024 12:46:45 PM, Rhuland, Courtney) Incoming inspection completed.
1071405 | 207429 | ACC | CMPLT | CMPLT | 2024-05-29 10:35 | 2024-05-29 12:08 | Incoming inspection | (TC 05/29/2024 12:07:32 PM, Baghirzade, Nara) Incoming inspection and asset tagging completed. Removed power cord guard at the rear of the incubator.
1071407 | 207430 | ACC | CMPLT | CMPLT | 2024-05-29 10:35 | 2024-05-29 12:48 | Incoming inspection | (TC 05/29/2024 12:47:54 PM, Rhuland, Courtney) Incoming inspection completed.
1062391 | 207431 | CM | CMPLT | CMPLT | 2024-05-29 10:36 | 2024-05-29 10:36 | "motion functions not working" | (TC 05/29/2024 10:35:37 AM, Cortez, John) Confirmed issue, error 3009-6-5-1503 present Power cycling bed seems to resolve issue All functions ok
| 207432 | CM | CMPLT | CMPLT | 2024-05-29 10:44 | 2024-06-04 14:18 | temp probe order, for stock | (TC 05/29/2024 10:51:48 AM, Cortez, John) placed order for 20x (TC 06/04/2024 02:18:27 PM, Cortez, John) arrived
1055946 | 207433 | CM | CMPLT | CMPLT | 2024-05-29 10:46 | 2024-05-29 10:46 | Not working | (TC 05/29/2024 10:45:55 AM, DeGuzman, Leila) Tag- Oxygen not working. Performed functional tests, all okay.
1056005 | 207434 | CM | CMPLT | CMPLT | 2024-05-29 10:48 | 2024-05-29 10:48 | Not working | (TC 05/29/2024 10:48:16 AM, DeGuzman, Leila) Came down attached to Root. Tag- Oxygen not working. Checked functions, all okay. Placed battery.
| 207435 | CM | CMPLT | CMPLT | 2024-05-29 11:00 | 2024-06-04 16:57 | Masimo safetynet errors, "serial number xxxxxx is alarming" | (TC 05/27/2024 11:00:19 AM, Cortez, John) Noticed errors on 10N station. Reached out to Naz, who explained that Masimo is aware. They resolved in short term by adding temporary "test" domain, and moved alerting SNs to this domain. (TC 06/04/2024 04:56:34 PM, Cortez, John) Resolved by Masimo
| 207436 | CM | CMPLT | CMPLT | 2024-05-29 11:17 | 2024-05-27 11:17 | Vital signs monitors in ED annex missing accessories | (TC 05/27/2024 11:17:07 AM, Cortez, John) Nurse provided list of missing accessories, in each room Supplied each monitor with stock or spares
65983 | 207437 | FC | CMPLT | CMPLT | 2024-05-29 11:59 | 2024-05-29 12:00 | No error reported; functional check | (TC 05/29/2024 11:59:31 AM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
65943 | 207438 | FC | CMPLT | CMPLT | 2024-05-29 12:00 | 2024-05-29 12:01 | No error stated; functional check | (TC 05/29/2024 12:00:34 PM, Baghirzade, Nara) Check-in tests passed and completed. IUI connectors and device cleaned.
1059778 | 207439 | FC | CMPLT | CMPLT | 2024-05-29 12:01 | 2024-05-29 12:01 | QC checklist | (TC 05/29/2024 12:01:08 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
69040 | 207440 | CM | CMPLT | CMPLT | 2024-05-29 13:25 | 2024-05-29 13:25 | Check-in | (TC 05/29/2024 01:25:26 PM, DeGuzman, Leila) Performed check-in tests - all passed. Cleaned the unit.
1055053 | 207441 | CM | CMPLT | CMPLT | 2024-05-29 14:13 | 2024-05-29 14:39 | Battery Replacement | (TC 05/29/2024 02:39:21 PM, DeGuzman, Leila) Came down attached to the root. Replaced the battery. Checked SpO2: passed.
1055005 | 207442 | CM | CMPLT | CMPLT | 2024-05-29 14:14 | 2024-05-29 14:23 | Temperature probe not working | (TC 05/29/2024 02:23:18 PM, DeGuzman, Leila) Tag- Temperature probe not working. Issue found: there's fluid in the probe well. Replaced temperature probe and well. Checked functions- all passed.
| 207443 | CM | CMPLT | CMPLT | 2024-05-29 14:17 | 2024-05-29 14:25 | Broken light otoscope | (TC 05/29/2024 02:25:28 PM, Rhuland, Courtney) Replaced burnt bulb in otoscope Functional check - OK
1069503 | 207444 | CM | CMPLT | NFF | 2024-05-29 14:31 | 2024-05-29 14:31 | Not reading accurately | (TC 05/29/2024 02:31:24 PM, Rhuland, Courtney) Brought test weights over and all measured exact Verified with other scale in department and read accurately No fault found
1054537 | 207445 | PM | CMPLT | CMPLT | 2024-05-29 16:30 | 2024-05-29 16:30 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 05/29/2024 04:29:51 PM, Cortez, John) Completed cleaning cycle Port 3 not being detected. There's damage to the port 3 hose, replaced with spare. Port detection issues seems to be resolved
68376 | 207446 | CM | CMPLT | CMPLT | 2024-05-30 07:56 | 2024-05-30 07:57 | Unable to use screen, very dark | (TC 05/30/2024 07:56:41 AM, Scero, Tomlyn) Called to OR15 because the AA couldn't read the screen on the AGM. The dimmness had been turned down to 10 so I managed to change it back to 100 for them. Restarted the self test and device was functioning fine.
1057873 | 207447 | CM | CMPLT | CMPLT | 2024-05-30 09:06 | 2024-07-08 11:43 | broken hose | (TC 05/30/2024 11:43:34 AM, Huang, Jessica) contacted the vendor for replacement unit. received it and sent it back to OR
1063793 | 207448 | CM | CMPLT | CMPLT | 2024-05-30 09:11 | 2024-07-08 12:01 | defective headlight | (TC 05/30/2024 12:01:12 PM, Huang, Jessica) picked up the light from OR desk and sent it for repair. received it back and sent it back to OR.
1057546 | 207449 | PM | CMPLT | CMPLT | 2024-05-30 11:25 | 2024-05-30 11:25 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 05/30/2024 11:24:49 AM, Scero, Tomlyn) Pm tests complete. No leaks. Battery replaced
1071432 | 207450 | ACC | CMPLT | CMPLT | 2024-05-30 11:39 | 2024-05-31 15:37 | Incoming Inspection | New unit. Unboxed and assembled the unit and rolling stand. Checked functionality and measured temp accuracy. Tested ok. EST - 170 mOhm; 19 uA. Delivered to OR.
| 207451 | TEDU | CMPLT | CMPLT | 2024-05-30 13:56 | 2024-05-30 13:56 | Women in STEM Luncheon - CMBES Conference | (TC 05/30/2024 01:56:07 PM, Scero, Tomlyn) Women in STEM Luncheon at the CMBES Conference
| 207452 | CM | CMPLT | CMPLT | 2024-05-30 14:34 | 2024-05-30 14:34 | Unboxing/asset tagging of phototherapy lights | (TC 05/30/2024 02:33:41 PM, Rhuland, Courtney) Unboxed, assembled and tagged phototherapy lights to be installed on omnibeds Meeting with NICU to determine location for mounting lights that will work.
69140 | 207453 | CM | CMPLT | CMPLT | 2024-05-31 08:25 | 2024-05-31 09:19 | Channel error | (TC 05/31/2024 09:18:44 AM, DeGuzman, Leila) Tag: Channel error message. Cleaned device. Downloaded error log, found error code: 241.4140.0 Performed Analog Sensor test: passed. Performed Check-in tests: all passed.
| 207454 | CM | CMPLT | CMPLT | 2024-05-31 09:11 | 2024-05-31 09:12 | NICU rounds | (TC 05/31/2024 09:11:56 AM, Rhuland, Courtney) NICU rounds completed
| 207455 | CM | CMPLT | CMPLT | 2024-05-31 09:13 | 2024-05-31 09:31 | RESUS rounds | (TC 05/31/2024 09:31:07 AM, Rhuland, Courtney) RESUS rounds completed
1060326 | 207456 | CM | CMPLT | CMPLT | 2024-05-31 10:26 | 2024-06-18 15:44 | Retire from Service - Sent for exchange | May 31 - Call re: spill happened in smoke evac. Went to check. Advised to clean and brought it down. Unit under consignment. Contacted vendor for RMA. No response. Cotnacted sales rep. Was informed no RMA required but was advised to issue no charge PO. Created SC. Awaiting approval. June 03 - Followed up. Received approval. ADvised MMC to issue PO. Scheduled pick up with KJV. Packed and shipped the unit. June 18 - Unit received. Unboxed and updated asset tag. Brought ESU down from OR and sync the unit. Checked functionality. Tested ok.
1055799 | 207457 | CM | CMPLT | CMPLT | 2024-05-31 10:28 | 2024-05-31 14:27 | Failed to check in, AC unplugged | (TC 05/31/2024 02:26:52 PM, Scero, Tomlyn) checked the defib on the crash cart. ensured the AC plug was properly connected. Repeated teh self test and 30J test.
1069387 | 207458 | CM | CMPLT | CMPLT | 2024-05-31 10:43 | 2024-06-07 08:42 | Effluent not draining/alarming blockage, kink | (TC 05/31/2024 11:17:28 AM, Canjimavadivel, Kamallesh ) Turned on device to see self tests. Missing handle for bag holder, reached out to Marlene for more information. Contacted vendor to initatie service call. (TC 06/06/2024 01:46:12 PM, Canjimavadivel, Kamallesh ) Worked with Milan from Baxter to go through testing and troubleshooting process. All functions passed, waiting for report (TC 06/07/2024 08:41:58 AM, Canjimavadivel, Kamallesh ) Attached service report, returned machine to unit
| 207459 | CM | CMPLT | CMPLT | 2024-05-31 11:07 | 2024-09-25 10:08 | Masimo Parts order | May 31 - Contacted vendor for parts info. Received. Created shopping cart. Awaiting approval. Received. Placed order via MMC. Sept 25 - Received parts. Soerted and Organized in the shelf.
1061333 | 207460 | FC | CMPLT | CMPLT | 2024-05-31 13:15 | 2024-05-31 13:15 | QC checklist | (TC 05/31/2024 01:14:45 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
| 207461 | RDS | CMPLT | CMPLT | 2024-05-31 13:16 | 2024-05-31 13:16 | Weekly Rounds: NICU & RESUS | (TC 05/31/2024 01:16:04 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
| 207462 | RDS | CMPLT | CMPLT | 2024-05-31 13:18 | 2024-05-31 13:18 | Weekly Rounds: 11 North | (TC 05/31/2024 01:17:58 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 207463 | RDS | CMPLT | CMPLT | 2024-05-31 13:20 | 2024-05-31 13:20 | Weekly Rounds: 11 South | (TC 05/31/2024 01:19:49 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
1059828 | 207464 | CM | CMPLT | NFF | 2024-05-31 13:21 | 2024-05-31 13:21 | User reported that there was water leaking | (TC 05/31/2024 01:21:24 PM, Baghirzade, Nara) Ran the device on system state first; no issues observed and all internal parts are functional. Ran the device on manual mode @ 25C for approximately 1 hour; no water leakage observed from the ports, hoses, and under the device (where the drain is located). Followed up with May. Followed up with Jen (person reporting this issue) that no fault was found with the device. Suspecting that the hose was loose and water leaked from there. Dropped off the altrix at the end of Zone B (near Zone C).
66799 | 207465 | PM | CMPLT | CMPLT | 2024-05-31 13:32 | 2024-05-31 14:10 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 05/31/2024 02:10:00 PM, Rhuland, Courtney) Worked with Leila to complete PM PM completed Refer to procedure for details
1061215 | 207466 | CM | CMPLT | CMPLT | 2024-05-31 13:38 | 2024-05-29 13:39 | Monitor Tilt Issue | Call re: monitor not staying in position. Went to check. Waited for access. Found Monitor sliding back when adjusting its position. Removed cover and adjusted tension. Tested ok.
1057009 | 207467 | CM | CMPLT | CMPLT | 2024-05-31 13:39 | 2024-05-31 14:02 | SpO2 parameter not available | (TC 05/30/2024 01:39:34 PM, Canjimavadivel, Kamallesh ) Confirmed issue at bedside, swapped with spare and brought down to biomed to assess (TC 05/31/2024 02:01:37 PM, Canjimavadivel, Kamallesh ) SpO2 parameter was deactivated, activated and SpO2 is working. Replaced MSL board as it was corroded
1061214 | 207468 | CM | CMPLT | CMPLT | 2024-05-31 13:41 | 2024-05-30 13:42 | Monitor Tilt issue | Call re: monitor not staying in position. Went to check. Waited for access. Found Monitor 2 sliding back when adjusting its position. Removed cover and adjusted tension. Tested ok.
1046196 | 207469 | CM | CMPLT | CMPLT | 2024-05-31 14:10 | 2024-05-31 14:10 | CT table side covers damaged and removed | (TC 05/29/2024 02:10:01 PM, Debebe, Adam) -CT table side telescopic covers damaged and removed. -Waited for the room to be available to put it back together as much as possible. -Unable to have access to the room due to long patient queues -Ordered parts with Siemens Will go back tomorrow to try to put it back together (TC 05/31/2024 02:10:29 PM, Debebe, Adam) Placed a call with Siemens as the screws are not holding anymore. Siemens will order and replace the parts. Siemens came in on June 05 with new telescopic table covers and replaced it.
68228 | 207470 | PM | CMPLT | CMPLT | 2024-05-31 14:16 | 2024-07-29 08:26 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 06/05/2024 02:07:37 PM, DeGuzman, Leila) Issue: Device turning off when performing NIBP tests. Possible cause power board issue. (TC 07/29/2024 08:25:33 AM, Baghirzade, Nara) Leila mentioned that X2 keeps shutting off when measuring NIBP. Replaced battery however the issue is still present. Confirmed that the device is not functional. Will retire the X2 as there are more to repurpose for NICU if required.
1056115 | 207471 | CM | CMPLT | CMPLT | 2024-05-31 15:26 | 2024-06-06 10:56 | Image transfer Issue | Call re: cannot send images to PACS. Went to check. Waited for room access. Checked unit and completed studies. Found multiple studies failed to send to PACS. Innvestigated further and they were having issues with sending studies since April. Checked all other Hyland devices as well and Foudn they all had same issue. Confirmed with KS that it is a server issue. Contacted Informatics. Awaiting response. MAy 30 - Informatics fixed issue on their end and advised to send all studies again. Went to endo and attempted to send studies. Failed. Tried in other rooms. Issue persisted. Noted error and sent it to informatics. Awaiting response. June 06 - Followed up and Confirmed that issue was on server end. Informatics will work with clinical staff to resolve it.
1056422 | 207472 | CM | CMPLT | CMPLT | 2024-05-31 15:55 | 2024-05-31 15:56 | "missing temp probe well" | (TC 05/31/2024 03:56:30 PM, Cortez, John) Added temp probe well Basket is bent. Attempted to straighten. Cleaned vent filter All parameters tested ok
1057008 | 207473 | CM | CMPLT | CMPLT | 2024-05-31 16:02 | 2024-05-30 16:03 | "SPO2 not working" | (TC 05/30/2024 04:02:35 PM, Cortez, John) Trunk issue. Replaced cable and informed nurse
| 207474 | CM | CMPLT | NFF | 2024-05-31 16:07 | 2024-05-31 16:08 | "SPO2 issue in ER Y52" | (TC 05/31/2024 04:07:35 PM, Cortez, John) NFF, all parameters tested ok However, there is a Msmt ext malf alert on banner, might need to replace FMS rack when room is free. Will forward info to Kam
1066921 | 207475 | FC | CMPLT | CMPLT | 2024-05-31 16:14 | 2024-05-30 16:14 | Check scale functionality | (TC 05/30/2024 04:14:01 PM, Cortez, John) Marlene has reported issue with bariatric beds in ICU Checked scale functionality of bed. Was able to zero and weigh calibrated weights successfully. Calibrated scale anyways. Accuracy ok afterwards (250 lb weights read 250.7 lbs)
1066918 | 207476 | FC | CMPLT | CMPLT | 2024-05-31 16:18 | 2024-05-30 16:18 | Scale functionality check | (TC 05/30/2024 04:17:48 PM, Cortez, John) Marlene has reported issue with bariatric beds in ICU Checked scale functionality of bed. Was able to zero and weigh calibrated weights successfully. Calibrated scale anyways. Accuracy ok afterwards (250 lb weights read 250.7 lbs)
50518 | 207477 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-23 14:28 | PM-MAJOR - Annual PM For: Vacuum Extractors, Obstetric | (TC 07/23/2024 02:27:48 PM, Rendon Lopez , Elizabeth) PM checks completed-OK
52230 | 207478 | PM | CANCL | CANCL | 2024-06-01 03:25 | 2024-07-17 07:52 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
53078 | 207479 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-05-22 10:45 | PM-MAJOR- Annual PM For: Pressure Meter, Digital | (TC 06/04/2024 10:45:29 AM, Chukri, Samir) Performed PM by Getinge tech. certificate calibration attached.
62779 | 207480 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-05 10:10 | PM-MAJOR - Annual PM For: Incubators, Infant, Transport | (TC 07/05/2024 10:05:12 AM, Ferreira, Qwynn) PM Passed Replaced Battery Replaced inner matress clip lock Functional Test Passed - Adjusted (R2)set point temperature - Increased 1C - Adjusted to primary from Secondary - Adjusted R3 to match setpoint - Temperature was tested at 36C - 36.1C Returned to unit (TC 07/05/2024 10:10:04 AM, Chukri, Samir) Worked with Qwynn
62966 | 207481 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-04 15:52 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/04/2024 03:52:30 PM, Rhuland, Courtney) PM completed Replaced 9V rechargeable battery Calibrated: Mains & VREF voltage Scale Functional check - OK Set 35C, 75% humidity Read 35C, 75.0% humidity Electrical safety - OK Leakage current - 80.4 uA Ground resistance - 64 mOhms
62972 | 207482 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62984 | 207483 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-23 10:36 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/23/2024 10:36:23 AM, Rhuland, Courtney) PM completed Cleaned fan filters Intensity: High - 39.3 uW/cm^2/nm Low - 15.6 uW/cm^2/nm Replaced broken timer Hours: 0 Chassis Leakage: 48.9 uA Ground Resistance: 186 mOhms
63024 | 207484 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-05 09:05 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/05/2024 09:05:18 AM, Rhuland, Courtney) PM complete Irradiance - 54 uW/cm^2/nm Hours - 6152 Leakage current - 13.4 uA Ground resistance - 84 mOhms
63027 | 207485 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-10-29 14:43 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/29/2024 02:43:24 PM, Alvarado, Karla) Replaced fan. Replaced power cord. Tests results: - Light Intensity small pad: 72.65 µW•cm-2•nm-1 - Hours: 5364.6 hours - Electrical safety results - Leakage current = 11.5 uA; GND Resistance = 174 mohms All tests passed.
63034 | 207486 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-05 08:42 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/05/2024 08:42:04 AM, Rhuland, Courtney) PM complete Irradiance - 55.6 uW/cm^2/nm Hours - 6680 Leakage current - 11.8 uA Ground resistance - 41 mOhms
63035 | 207487 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-05 08:51 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/05/2024 08:50:49 AM, Rhuland, Courtney) PM complete Irradiance - 54.7 uW/cm^2/nm Hours - 3851 Leakage current - 11.5 uA Ground resistance - 62 mOhms
63302 | 207488 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
63793 | 207489 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
63870 | 207490 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-17 09:26 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
63871 | 207491 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-10-21 13:10 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 10/08/2024 01:09:41 PM, Huang, Jessica) received the device from Mother and Baby. PM completed
63875 | 207492 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-20 10:10 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/20/2024 10:10:18 AM, DeGuzman, Leila) PM passed and completed. Calibrated Vref, Mains V, LV, and Scale. Switches and pedals were functional. Scale accuracy passed, 5000g. Humidity: 75% - 75% Temperature: 35 celsius - 35 celsius Electrical Safety passed; Ground resistace: 60 mOhm Leakage current: 15.9 uA Replaced procedure lights.
65129 | 207493 | PM | CMPLT | PMFAIL | 2024-06-01 03:25 | 2024-06-28 08:20 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/28/2024 08:19:33 AM, Reid, Emmerson ) Was sent down with a tag saying it had broken into pieces. The pole mount had come off likely from a drop but nothing had been damaged so I was able to put it together. I noticed the occlusion sensor seal is severely buldged and begginning to tear so I am going to send it out for repair.
65132 | 207494 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65133 | 207495 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65375 | 207496 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-10-21 13:19 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 10/02/2024 01:19:10 PM, Huang, Jessica) received the device from mother and baby, PM completed
65417 | 207497 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-18 12:03 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/18/2024 12:03:06 PM, Rhuland, Courtney) Replaced 2 discoloured porthole seals Replaced rechargeable 9V battery PM completed Calibrated: VREF & MAINS voltage Scale Functional check - OK Set 35C & 75% humidity Read 35C, 75.4% humidity Electrical safety - OK Leakage current - 74.4 uA Ground resistance - 188 mOhms
65853 | 207498 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
65859 | 207499 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
65883 | 207500 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-30 11:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/30/2024 11:54:36 AM, Baghirzade, Nara) Replaced: -rear case -front caase -mylar gasket -gasket for pressure sensor holder -latch module Full calibration and PM: Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65884 | 207501 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-03-05 09:59 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65892 | 207502 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-03-11 10:02 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65937 | 207503 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-03-05 11:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65938 | 207504 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-31 15:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/31/2024 03:45:35 PM, DeGuzman, Leila) PM completed.
65949 | 207505 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-30 11:51 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/30/2024 11:51:02 AM, Baghirzade, Nara) Replaced rear case. Replaced latch module. Full calibration and PM: Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65961 | 207506 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-06 12:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/01/2024 01:41:05 PM, Baghirzade, Nara) Replaced: -front case -rear case -mylar gasket. Assembled device. As soon as I connected to the PCU, got an error for a defective sensor. Will replace pressure sensor and complete the PM. (TC 08/06/2024 12:33:15 PM, Baghirzade, Nara) Replaced pressure sensor. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
65968 | 207507 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 13:04 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65977 | 207508 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-31 15:57 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/31/2024 03:56:45 PM, DeGuzman, Leila) PM completed.
66015 | 207509 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 08:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66016 | 207510 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 08:05 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66020 | 207511 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 08:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66023 | 207512 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 07:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66032 | 207513 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 07:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66035 | 207514 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 07:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66041 | 207515 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 07:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66044 | 207516 | PM | CMPLT | PMFAIL | 2024-06-01 03:25 | 2024-07-22 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/22/2024 11:17:23 AM, Rendon Lopez , Elizabeth) PM failed. Cracked front and rear case, waiting on part
66059 | 207517 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 07:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66402 | 207518 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
66464 | 207519 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-27 07:53 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced V-battery and air filter. Run battery check. 220 minutes, 16 Ah. Tested. Functional. SW2.51. Leakage: 89uA and 91 ohm.
66468 | 207520 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-05 13:57 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 34 Ah and 220 min. Tested. Functional. SW.2.51 Leakage: 90 uA and 95 ohm.
66471 | 207521 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-14 11:50 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has GS500. Replaced V-battery and air filter, GS500 filters. Run battery check. 230 minutes, 35 Ah. Tested. Functional. SW2.51 Leakage: 91 uA and 88 ohm.
66501 | 207522 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-12 09:04 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/12/2024 09:03:52 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
66513 | 207523 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-07 09:17 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/07/2024 09:16:53 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
66516 | 207524 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-12 09:07 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/12/2024 09:06:38 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
66526 | 207525 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-12 09:05 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/12/2024 09:05:22 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
66681 | 207526 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-10 14:48 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/10/2024 02:47:36 PM, Rhuland, Courtney) Replaced bent power cord Replaced missing corner grommet PM completed Calibrated: MAINS & LV voltage Scale Functional check - OK Set 35C & 75% humidity Read 35.1C & 74.9% humidity Electrical safety - OK Leakage curent - 87.7 uA Ground resistance - 181 mOhms (TC 07/11/2024 07:53:10 AM, Rhuland, Courtney) When removing humidifier reservoir, noticed humidifier was loose Replaced cracked red humidifier gasket & button Functional check - OK
66693 | 207527 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-13 08:28 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/13/2024 08:28:01 AM, Rhuland, Courtney) PM completed Replaced: Exam light 9V rechargeable battery Calibrated: MAINS & VREF voltage Scale Functional check - OK Set 35C, 75% humidity Read 35C, 76.3% humiditiy Electrical safety - OK Ground resistance - 135 mOhms Leakage current - 79.5 uA
66695 | 207528 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-05 14:14 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/05/2024 02:13:56 PM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66713 | 207529 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-24 13:02 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/24/2024 01:01:41 PM, Baghirzade, Nara) PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 57.6 uA Power cord resistance = 87 mohms
66714 | 207530 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-05 08:48 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/05/2024 08:48:17 AM, Rhuland, Courtney) PM complete Irradiance - 49.1 uW/cm^2/nm Hours - 6964 Leakage current - 13.4 uA Ground resistance - 0.71mOhms
66720 | 207531 | PM | CMPLT | PMFAIL | 2024-06-01 03:25 | 2024-09-05 14:25 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/05/2024 02:24:38 PM, Rendon Lopez , Elizabeth) Monitor fails recorder calibration, makes grinding noises and stucked.
66725 | 207532 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-05 14:24 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/05/2024 02:23:41 PM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66726 | 207533 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-18 10:58 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 06/18/2024 10:57:47 AM, DeGuzman, Leila) PM completed. All tests passed NIBP Tests Accuracy: 120/80 - 119/76 Leakage: 3.6 Linearity: 0.4 Valve: 0.9 SpO2: 100% US: good
66727 | 207534 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-25 15:55 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 07/25/2024 03:55:22 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Cleaned recorder, calibrated scale-OK
66731 | 207535 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-05 14:12 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/05/2024 02:12:09 PM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66762 | 207536 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66795 | 207537 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 14:26 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/01/2024 02:25:31 PM, DeGuzman, Leila) PM completed. See procedure for details.
66803 | 207538 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-25 10:13 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/25/2024 10:12:47 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
66805 | 207539 | PM | CNP | EIU | 2024-06-01 03:25 | 2024-07-30 14:43 | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66811 | 207540 | PM | CNP | ENL | 2024-06-01 03:25 | 2024-07-30 14:43 | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66817 | 207541 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-14 15:12 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/14/2024 03:12:01 PM, DeGuzman, Leila) PM passed and completed. Calibrated NIBP Battery replaced 2023. See procedure for details
66820 | 207542 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66821 | 207543 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-05 09:02 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 06/05/2024 09:01:57 AM, DeGuzman, Leila) PM completed. Go to procedures to see details.
66824 | 207544 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-23 12:04 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/23/2024 12:04:03 PM, DeGuzman, Leila) PM passed and completed. See procedure for details.
66833 | 207545 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-23 14:36 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/23/2024 02:35:43 PM, DeGuzman, Leila) PM passed and completed. See procedure for details.
66843 | 207546 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-10-10 11:12 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 10/10/2024 11:12:15 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
66847 | 207547 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-09 10:41 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/09/2024 10:37:42 AM, Alvarado, Karla) PM completed All tests Passed see procedure (TC 09/09/2024 10:40:55 AM, Rhuland, Courtney) Supervised Karla
66849 | 207548 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-11-06 14:12 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 11/06/2024 02:12:08 PM, Alvarado, Karla) Calibrated NBP PM completed All tests passed See procedure for details
66862 | 207549 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-25 10:33 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/25/2024 10:32:38 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
67757 | 207550 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67765 | 207551 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67872 | 207552 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-18 14:44 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/18/2024 02:44:30 PM, DeGuzman, Leila) PM completed. All tests passed.
67893 | 207553 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-20 10:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/20/2024 10:41:30 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Upgraded Software and library, replaced batteries
67895 | 207554 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-17 11:34 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/17/2024 11:34:16 AM, DeGuzman, Leila) PM completed. all tests passed.
67917 | 207555 | PM | CANCL | CANCL | 2024-06-01 03:25 | 2024-06-04 10:12 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
67920 | 207556 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-10 14:40 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/10/2024 02:40:08 PM, Rhuland, Courtney) Replaced broken hood from retired incubator Replaced broken porthole door Caibrated: Main and REF voltage Scale RH Functional check: Set 35C, 75% humidity Read 35C, 75.6% humidity Electrical safety: Leakage current - 83.5 uA Ground resistance - 92 mOhms
67951 | 207557 | PM | CMPLT | PMFAIL | 2024-06-01 03:25 | 2024-07-23 14:43 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/23/2024 02:43:23 PM, Rhuland, Courtney) During PM, noticed humidfier was loose When removing, chassis screwhold broke off Retiring unit and using for parts to repair other incubator with new chassis.
67957 | 207558 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-18 08:05 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/18/2024 08:04:37 AM, Rhuland, Courtney) Brought down from NICU storage room Repaired porthole door Repaired side door latch Calibrated: MAINS & LV voltage Scale Functional check - OK Set 35C, 75% humidity Read 35C, 74.6% humidity Electrical safety - OK Ground resistance - 81 mOhms Leakage current - 82.7 uA
67958 | 207559 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-17 13:31 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 06/17/2024 01:31:18 PM, Rhuland, Courtney) PM completed Calibrated: VREF & main voltage Scale Funcitonal check - OK Set 35C & 75% humidity Read 35C & 74.8% humidity Electrical safety - OK Ground resistance - 98 mOhms Leakage current - 86.3 uA
67971 | 207560 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-30 14:08 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/30/2024 02:08:16 PM, Rhuland, Courtney) PM completed Porthole doors repaired Calibrated: MAINS & LV voltage RH Scale Replaced: 9V rechargeable battery RH sensor During functional check, RH failed multiple times post calibration RH sensor found to be corroded - replaced Functional check - OK Set: 35C & 75% humidity Read: 35.1C & 75.2% humidity Electrical safety - OK Ground resistance - 53 mOhms Leakage current - 75 uA
68373 | 207561 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68615 | 207562 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 12:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68693 | 207563 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 13:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68709 | 207564 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 13:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68763 | 207565 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 13:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68787 | 207566 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 13:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68815 | 207567 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 13:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68927 | 207568 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 14:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68928 | 207569 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 14:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69048 | 207570 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69333 | 207571 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-09 12:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:57:10 PM, Patel, Rutvik) by BD
69344 | 207572 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69427 | 207573 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-06 10:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/06/2024 10:09:45 AM, Rendon Lopez , Elizabeth) Calibrated rate PM completed, all tests passed
69447 | 207574 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-09 12:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:51:23 PM, Patel, Rutvik) by BD
69622 | 207575 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 15:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
530502 | 207576 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-22 12:13 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 07/22/2024 12:13:13 PM, Rhuland, Courtney) PM Passed Disinfected reservoir Replaced reservoir gasket Replaced and lubricated O-ring seals Alarm signal test Passed Over Temperature alarm test Passed. Temperature Verification passed Electrical safety - OK Ground resistance - 40 mOhms Leakage current - 12.7 uA
65617 | 207577 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
32540 | 207578 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-08 15:39 | PM - Monthly | PM completed
32560 | 207579 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-02 15:45 | PM - Monthly | PM Major done Copied Scan PC Back up disk Slice count - 1107881 TOS phantom a(-998.1, 13.5) b(335.6, 14.1) c(134.4, 14.8) d(99.0, 14.7) e(-105.9, 14.2) f(-2.0, 15.0) -System passed all tests
32580 | 207580 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-03 16:37 | PM - Monthly | PM completed
501974 | 207581 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-03 15:01 | PM - Annual | (TC 07/03/2024 03:00:34 PM, Green, Paul) PM completed.
502366 | 207582 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-01 13:47 | PM-Annual | (TC 08/01/2024 01:46:58 PM, Fan, Tak) Gnd R = 98 mohm Max A = 0.3 / 156.8 / 356 uA ECG = 0.9 / 2.8 ISO = 7.8 M5Sc (458680YP6) = 0.5 / 1.6 uA 4Vc (302806YP5) = 0.5 / 6.2 uA
1046203 | 207583 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-30 13:49 | PM-Annual | (TC 07/30/2024 01:48:53 PM, Fan, Tak) PM completed. Gnd R = 152 mohm Max A = 0.2 / 162 / 315 uA ECG = 0.4 / 3.2 uA ISO = 12.6 uA M5Sc (461106YP7) = 0.6 / 1.7 uA
1046671 | 207584 | PM | ACTIV | READY | 2024-06-01 03:25 |  | PM-ELECTRICAL MECHANISM INSPECTION AND OPERATION TEST |  | DEPARTMENT CONTACT THE VENDOR TO PROVIDE PM AND MAINTENANCE
1047043 | 207585 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-03 13:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:34:27 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered
1047471 | 207586 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-26 10:52 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 1st year PM after 6th year PM last year. No GS500. It has PS500. Run battery check. 20 Ah and 220 min. Tested. Functional. SW.2.51 Leakage: 81 uA and 89 ohm.
1047473 | 207587 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-19 10:43 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 1st year PM after 6th year PM last year. It has no GS500. It has PS500. Run battery check. 220 minutes, 29 Ah. Tested. Functional. SW2.51. Leakage: 91 uA and 89 ohm.
1047833 | 207588 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-11 13:42 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/11/2024 01:42:27 PM, Rhuland, Courtney) PM completed See procedure for details
1047841 | 207589 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-28 13:56 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/28/2024 01:55:37 PM, Rhuland, Courtney) Replaced worn clutch Replaced plunger seal PM completed All tests passed Refer to procedure for details
1047842 | 207590 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-14 07:51 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/14/2024 07:51:09 AM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1047845 | 207591 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-17 09:29 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/17/2024 09:28:42 AM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1047853 | 207592 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-10-31 14:13 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 10/31/2024 02:12:43 PM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1047859 | 207593 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-14 14:24 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/14/2024 02:24:00 PM, Rhuland, Courtney) Clutch was worn and failed occlusion test Replaced clutch When running tests, loud knocking noise was present Replaced gear worm and issue resolved PM completed All tests passed Refer to procedure for details
1047861 | 207594 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-14 10:14 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/14/2024 10:14:11 AM, Rhuland, Courtney) Replaced cracked top case PM completed All tests passed Refer to procedure for details
1047864 | 207595 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-07 09:33 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/07/2024 09:32:34 AM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
1048606 | 207596 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM
1049047 | 207597 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | Planned Maintenance
1050247 | 207598 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1050264 | 207599 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-17 13:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/17/2024 01:31:25 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed
1050620 | 207600 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-30 13:31 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/30/2024 01:31:20 PM, DeGuzman, Leila) PM completed.
1051383 | 207601 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-23 09:06 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 07/23/2024 09:06:19 AM, Rhuland, Courtney) PM completed. All tests passed. Hours: 5182.9 Light Intensity: 48.4 uW/cm^2*nm Electrical Safety - OK Leakage: 11.6 uA Resistance: 139 mOhms
1051614 | 207602 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-11 10:55 | PM-MAJOR - Annual PM For: Testers | (TC 06/11/2024 10:55:11 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1051615 | 207603 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-05-22 10:19 | PM-MAJOR - Annual PM For: Testers | (TC 08/07/2024 10:19:57 AM, Chukri, Samir) certificate calibration by Peter from Gatinge, See attached
1054110 | 207604 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-25 15:57 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 07/25/2024 03:56:58 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Cleaned recorder, calibrated-OK
1054285 | 207605 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM For: Asperators
1054460 | 207606 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-04 10:21 | Lubrication cycle | Completed with PM
1054462 | 207607 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-11 10:35 | Lubrication cycle | Completed with PM
1054468 | 207608 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-15 14:49 | Lubrication cycle | Completed with PM
1054537 | 207609 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-10-30 15:23 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 06/07/2024 04:49:35 PM, Cortez, John) Ordered PM kit (TC 10/30/2024 03:22:40 PM, Cortez, John) Received spare PM kit from Stryker Replaced PM parts, Replaced 6x quick couplers Completed functional test 60 mohms, 120 uAs
1055111 | 207610 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-13 08:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055126 | 207611 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 10:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055748 | 207612 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1055749 | 207613 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-03-05 09:54 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1055770 | 207614 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
1057743 | 207615 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1058246 | 207616 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-16 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1059420 | 207617 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059425 | 207618 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1060913 | 207619 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-11 11:17 | Calibration Defibrillator Analyzer | (TC 06/11/2024 11:16:40 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1060914 | 207620 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | Calibration Defibrillator Analyzer
1060974 | 207621 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1062181 | 207622 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-11 14:51 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 09/11/2024 02:50:47 PM, Alvarado, Karla) 2498.4 hours Intensity light: - High: 34.5 uW/cm2/nm - Low:14.8 uW/cm2/nm Target light is functional GND Resistance: 0.172 ohms Leakage current: 47.3 uA PM completed. (TC 09/11/2024 02:51:27 PM, Baghirzade, Nara) Supervised
1063260 | 207623 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
1064401 | 207624 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064402 | 207625 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064403 | 207626 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064404 | 207627 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064489 | 207628 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-08-07 14:23 | PM-MAJOR- Annual PM for: Pulmo Vista | (TC 08/07/2024 02:22:35 PM, Canjimavadivel, Kamallesh ) PM perofrmed by vendor, see attached reports
1064493 | 207629 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-09-19 09:06 | ANNUAL PM - PM Major for: Warmers | WO 209847
1064502 | 207630 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-06-21 14:22 | PM - Annual | (TC 06/21/2024 02:21:48 PM, Green, Paul) This PM is performed by the vendor.
1064511 | 207631 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064512 | 207632 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064513 | 207633 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064514 | 207634 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064515 | 207635 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064516 | 207636 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM MAJOR - ANNUAL PM for: ERBE
1064974 | 207637 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-15 13:42 | Annual PM | (TC 07/15/2024 01:42:12 PM, Ferreira, Qwynn) PM Complete lead screws were cleaned and lubricated LPU was cleaned and jogged z-axis - OK x-axis - OK No loose cables or broken features Functional Check - OK
1064975 | 207638 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-15 13:46 | Annual PM | (TC 07/15/2024 01:45:53 PM, Ferreira, Qwynn) IPA concentration - OK Motor moving - OK PM Complete
1064976 | 207639 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | Annual PM
1066756 | 207640 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-17 07:49 | Lubrication cycle | Completed with PM
1066757 | 207641 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-11 14:05 | Lubrication cycle | Completed with PM
1067337 | 207642 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-11-11 09:23 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | Completed annual PM. Tested. Passed pre-use check. Functional.
1067338 | 207643 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
1067339 | 207644 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-11-22 10:02 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | performed annual PM. Tested. passed pre-use check. Functional.
1067340 | 207645 | PM | CMPLT | CMPLT | 2024-06-01 03:25 | 2024-07-05 09:30 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | performed annual PM. tested. Functional.
1067341 | 207646 | PM | ACTIV | ISSUE | 2024-06-01 03:25 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
1067342 | 207647 | PM | CMPLT | CMPLT | 2024-06-01 03:26 | 2024-11-18 11:02 | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | performed annual PM. Replaced exp cassette membrane. reset unit. tested. Passed pre-use check. Functional.
1067349 | 207648 | PM | CMPLT | CMPLT | 2024-06-01 03:26 | 2024-09-04 14:33 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 09/04/2024 02:33:18 PM, Rhuland, Courtney) PM completed Physical inspection - OK Test weight = 5Kg Measured 5.000Kg
1067604 | 207649 | PM | CMPLT | CMPLT | 2024-06-01 03:26 | 2024-07-16 09:27 | MSH ER JDMI i800 PM - Semi Annual | (TC 07/15/2024 07:22:16 AM, Whittington, Scott) MSH i800 PM - Semi Annual (TC 07/15/2024 07:22:16 AM, Whittington, Scott) MSH i800 PM - Semi Annual (TC 07/16/2024 09:27:02 AM, Whittington, Scott) MSH ER JDMI i800 PM - Semi Annual | MSH ER JDMI i800 PM - Semi Annual
1067979 | 207650 | PM | CMPLT | CMPLT | 2024-06-01 03:26 | 2024-07-16 09:26 | MSH 20 SPORTS CLINIC i800 PM - Semiannual | (TC 07/15/2024 07:25:27 AM, Whittington, Scott) MSH 20 SPORTS CLINIC i800 PM - Semiannual (TC 07/16/2024 09:26:20 AM, Whittington, Scott) MSH 20 SPORTS CLINIC i800 PM - Semiannual | MSH 20 SPORTS CLINIC i800 PM - Semiannual
1067992 | 207651 | PM | CMPLT | CMPLT | 2024-06-01 03:26 | 2024-07-16 09:23 | MSH US i800 PM - Semiannual | (TC 07/15/2024 07:28:02 AM, Whittington, Scott) MSH US i800 PM - Semiannual (TC 07/16/2024 09:22:50 AM, Dolinar, John) MSH US i800 PM - Semiannual | MSH US i800 PM - Semiannual
1068220 | 207652 | PM | CMPLT | CMPLT | 2024-06-01 03:26 | 2024-07-16 09:19 | MSH US i800 PM - Semiannual | (TC 07/15/2024 07:29:47 AM, Whittington, Scott) MSH US i800 PM - Semiannual (TC 07/16/2024 09:19:40 AM, Dolinar, John) MSH US i800 PM - Semiannual | MSH US i800 PM - Semiannual
1068221 | 207653 | PM | CMPLT | CMPLT | 2024-06-01 03:26 | 2024-07-16 09:19 | MSH US i800 PM - Semiannual | (TC 07/15/2024 07:31:12 AM, Whittington, Scott) MSH US i800 PM - Semiannual (TC 07/16/2024 09:13:33 AM, Dolinar, John) MSH US i800 PM - Semiannual | MSH US i800 PM - Semiannual
1070578 | 207654 | PM | ACTIV | ISSUE | 2024-06-01 03:26 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
1070579 | 207655 | PM | ACTIV | ISSUE | 2024-06-01 03:26 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
1070580 | 207656 | PM | ACTIV | ISSUE | 2024-06-01 03:26 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
1070581 | 207657 | PM | ACTIV | ISSUE | 2024-06-01 03:26 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
501583 | 207658 | CM | CMPLT | CMPLT | 2024-06-02 10:28 | 2024-06-03 12:54 | SPP 15-217 ROSE L&D Hallway (IU22) 501583 GREG RYAN REPORTS FAULTY MONITOR | (TC 06/03/2024 12:53:49 PM, Whittington, Scott) SPP 15-217 ROSE L&D Hallway (IU22) 501583 GREG RYAN REPORTS FAULTY MONITOR REPLACED MONITOR | SPP 15-217 ROSE L&D Hallway (IU22) 501583 GREG RYAN REPORTS FAULTY MONITOR
66857 | 207659 | CM | CMPLT | CMPLT | 2024-06-03 07:12 | 2024-06-03 07:33 | Missing battery | (TC 06/03/2024 07:32:39 AM, Rhuland, Courtney) Battery replaced Functional check - OK
| 207660 | RDS | CMPLT | CMPLT | 2024-06-03 08:51 | 2024-06-03 08:52 | Weekly check for ED
1056125 | 207661 | PM | CMPLT | CMPLT | 2024-06-03 09:55 | 2024-06-03 09:55 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 06/03/2024 09:55:02 AM, Canjimavadivel, Kamallesh ) Cleaned hoses and nozzles Cleaning cycle complete
| 207662 | CM | CMPLT | CMPLT | 2024-06-03 09:56 | 2024-06-03 09:57 | Battery change on scale | (TC 06/03/2024 09:56:32 AM, Rhuland, Courtney) Unit requested battery change on device Went up and replaced batteries but keypad was unresponsive Attempted to repair, but no change Advised unit to discuss with manager about purchasing new patient scale Patient scale was very old and had signs of rust
501583 | 207663 | CM | CMPLT | CMPLT | 2024-06-03 10:02 | 2024-06-11 08:42 | MSH 15 SPP 15-217 ROSE L&D Hallway (IU22) 501583 MAIN MONITOR INTERMITTENT, LOW BRIGHTNESS | (TC 06/03/2024 12:56:17 PM, Whittington, Scott) MSH 15 SPP 15-217 ROSE L&D Hallway (IU22) 501583 MAIN MONITOR INTERMITTENT, LOW BRIGHTNESS SITE# 52668740 REF# OBTAIN MONITOR FROM CEOU RM 12 SCAB TO ROSE, OK MOUNTED, DIFFICULTY HAVING IT RECOGNIZED EVENTUALLY RELIABLY STARTING UP WITH MAIN MONITOR RETURN TO L&D OBU FOR SPP (TC 06/03/2024 01:07:29 PM, Whittington, Scott) MSH 15 SPP 15-217 ROSE L&D Hallway (IU22) 501583 MAIN MONITOR INTERMITTENT, LOW BRIGHTNESS SITE# 52668740 REF# AFTER SEVERAL POWER UP/DOWN CYCLES. REVERTED TO BLANK MAIN MONITOR (TC 06/07/2024 01:07:29 PM, Whittington, Scott) MSH 15 SPP 15-217 ROSE L&D Hallway (IU22) 501583 MAIN MONITOR INTERMITTENT, LOW BRIGHTNESS SITE# 52668740 REF# AFTER SEVERAL POWER UP/DOWN CYCLES. REVERTED TO BLANK MAIN MONITOR JUNE 7 REPLACEMENT MONITOR DELIVERED, INSTALLED, RETURNED (TC 06/10/2024 11:19:56 AM, Whittington, Scott) MSH 15 SPP 15-217 ROSE L&D Hallway (IU22) 501583 MAIN MONITOR INTERMITTENT, LOW BRIGHTNESS PACK AND RETURN FAULTY MONITOR | MSH 15 SPP 15-217 ROSE L&D Hallway (IU22) 501583 MAIN MONITOR INTERMITTENT, LOW BRIGHTNESS SITE# 52668740 REF# JAMES 0123778873
1064400 | 207664 | CM | CMPLT | CMPLT | 2024-06-03 10:36 | 2024-06-05 13:31 | SpO2 issue | (TC 05/24/2024 10:43:56 AM, Rendon Lopez , Elizabeth) Monitor missing side panel. Removed device from OR4 AGM (TC 06/03/2024 10:44:58 AM, Rendon Lopez , Elizabeth) Tested SpO2-No issues. All functions working properly. Ordering side panel replacement (TC 06/05/2024 01:31:13 PM, Rendon Lopez , Elizabeth) Replaced side panel cover. Tested all functions-OK. Returned to L&D ORs
1057557 | 207665 | PM | CMPLT | CMPLT | 2024-06-03 11:13 | 2024-06-03 11:16 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 06/03/2024 11:16:15 AM, Scero, Tomlyn) PM tests complete and passed. Battery was replaced in 2022.
| 207666 | RDS | CMPLT | CMPLT | 2024-06-03 14:13 | 2024-06-04 14:24 | Weekly checks in CCU | (TC 06/04/2024 02:24:01 PM, Scero, Tomlyn) weekly checks complete
1047840 | 207667 | CM | CMPLT | CMPLT | 2024-06-03 14:50 | 2024-06-05 08:42 | Check clutch | (TC 06/03/2024 02:49:54 PM, Rhuland, Courtney) Disassembled and reassembled clutch assembly Replaced plunger gasket (TC 06/04/2024 02:43:55 PM, Rhuland, Courtney) Issue still present Replaced clutch Pressure cable broke when reassembling Will replace board with spare (TC 06/05/2024 08:41:31 AM, Rhuland, Courtney) Rebuilt device with spare screen All errors cleared PM to be completed in separate WO
1050072 | 207668 | CM | CMPLT | CMPLT | 2024-06-03 14:59 | 2024-06-05 11:04 | Damaged housing | (TC 06/03/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Disassembled unit to identify broken parts. Looked up part nurmbers in service manual (TC 06/04/2024 08:56:44 AM, Canjimavadivel, Kamallesh ) Reached out to Stevens for parts pricing, some items are on back order and one item is not set up in the system. Waiting for follow up email (TC 06/04/2024 03:19:19 PM, Canjimavadivel, Kamallesh ) Found parts from old unit that is to be retired. Started dismantling parts (TC 06/05/2024 11:04:02 AM, Canjimavadivel, Kamallesh ) Reassembled unit with replacement parts. Added cuff, all functions ok.
1065307 | 207669 | CM | CMPLT | CMPLT | 2024-06-03 16:27 | 2024-09-25 09:53 | MSH5 US RM 3 APLIO 5 SYSTEM NOT POWERING UP | (TC 07/08/2024 12:17:33 PM, Whittington, Scott) MSH5 US RM 3 APLIO 5 SYSTEM NOT POWERING UP CHECKED SYSTEM TO SEE IF IT IS POWER SUPPLY OR STRAINED AC CORD AD CORD OK,. SUSPECT POWER SUPPLY | MSH5 US RM 3 APLIO 5 SYSTEM NOT POWERING UP
| 207670 | CM | CMPLT | CMPLT | 2024-06-04 06:59 | 2024-06-03 14:00 | Biomed huddle | (TC 06/03/2024 02:00:00 PM, Baghirzade, Nara) Biomed huddle
1071252 | 207671 | CM | CMPLT | CMPLT | 2024-06-04 07:01 | 2024-06-03 13:00 | Cables for manikin | (TC 06/03/2024 01:00:00 PM, Baghirzade, Nara) Samir helped purchased QTY 4 cables for NICU manikin; cable connects manikin to the laptop. Purchased 2 6.6ft cables and 2 10ft. Placed 1 6.6ft cable inside the manikin case in John's office. Also let John know that we have a longer cable and spares in biomed if required.
1047839 | 207672 | CM | CMPLT | CMPLT | 2024-06-04 07:03 | 2024-06-03 10:03 | Pump did not alarm when feed was paused | (TC 06/03/2024 10:03:00 AM, Baghirzade, Nara) RN reported that the pump did not alarm when the feed was paused. Reviewed the alarm history; found all the alarms for near-empty syringe and feed complete were active. Ran a feed and paused it after a couple of minutes; set a timer and found that the medium priority (yellow indicator) alarm would sound after 2 minutes. Repeated the test and got the same results. In diagnostics menu, checked the alarm test and indicator test; bother are functional for any priority. Ran a continuous feed for a couple of hours, paused the feed multiple times and observed the medium priority alarm within 2 minutes after pausing/stopping the feed. No issues or errors observed. No fault found with the device. Device is fully functional.
63507 | 207673 | CM | CMPLT | CMPLT | 2024-06-04 07:14 | 2024-07-30 14:37 | Determining a decommissioning/disposal for the laser | (TC 06/03/2024 03:00:00 PM, Baghirzade, Nara) Working with Jessica, Andrew, and Reeba on decommissioning the holmium laser. Manufacturer quotes $2000 for the disposal including the pickup of the device. Contacted e-waste company that Sinai has a contract with, sent them the laser specifications such as dimensions and weight, and they have agreed to dispose/recycle it at no cost. Also contacted Nektor - MLSO for TGH - about their decommissioning/disposal plan and they contact the vendor for disposal. Reeba mentioned that she would like to donate the device however we need to find out if the license for the device is active before she can initiate donation work. (TC 07/30/2024 02:36:32 PM, Baghirzade, Nara) Reeba located a vendor that want the laser; adeel provided reeba with a form to complete. Laser was removed from hospital premises and Jessica removed all stickers/tags on the laser.
1063785 | 207674 | CM | CMPLT | CMPLT | 2024-06-04 08:39 | 2024-06-24 09:32 | Headlight out for repair | (TC 06/19/2024 02:35:05 PM, Reid, Emmerson ) Created no charge PO SC 1003254782 (TC 06/24/2024 09:32:29 AM, Reid, Emmerson ) Returned from repair and brought back to OR desk
1063786 | 207675 | CM | CMPLT | CMPLT | 2024-06-04 08:40 | 2024-06-24 09:33 | Headlight out for repair | (TC 06/19/2024 02:35:58 PM, Reid, Emmerson ) SC 1003254782 (TC 06/24/2024 09:33:13 AM, Reid, Emmerson ) Returned from repair and brought back to OR desk
1058153 | 207676 | CM | CMPLT | CMPLT | 2024-06-04 09:12 | 2024-06-04 09:12 | Upgrade to V4.3 | (TC 06/04/2024 09:12:21 AM, Reid, Emmerson ) Upgraded the unit to firmware V4.3 and reflected this in the workbook. Also push APS Epidural Oct 2023
1058178 | 207677 | CM | CMPLT | CMPLT | 2024-06-04 09:39 | 2024-06-04 09:39 | Upgrade to V4.3 | (TC 06/04/2024 09:38:47 AM, Reid, Emmerson ) Upgraded the unit to firmware V4.3 and reflected this in the workbook. Also pushed APS Epidural Oct 2023
1058171 | 207678 | CM | CMPLT | CMPLT | 2024-06-04 10:10 | 2024-06-04 10:10 | Upgraded to V4.3 | (TC 06/04/2024 10:09:59 AM, Reid, Emmerson ) Upgraded the unit to firmware V4.3 and reflected this in the workbook. Also pushed APS Epidural Oct 2023
67758 | 207679 | CM | CMPLT | CMPLT | 2024-06-04 10:36 | 2024-06-04 10:36 | Upgrade to V4.3 | (TC 06/04/2024 10:35:42 AM, Reid, Emmerson ) Upgraded the unit to firmware V4.3 and reflected this in the workbook. Also pushed APS Epidural Oct 2023
1064819 | 207680 | CM | CMPLT | CMPLT | 2024-06-04 10:44 | 2024-06-05 08:59 | SpO2 not picking up on monitor | (TC 06/04/2024 10:44:36 AM, Canjimavadivel, Kamallesh ) Went to room and confirmed issue. Isolated problem to X3. Swapped out with spare and brought down to shop (TC 06/05/2024 08:59:22 AM, Canjimavadivel, Kamallesh ) Tested again and SpO2 working normal, to be monitored for reoccurance. Returned to unit
1053491 | 207681 | CM | CMPLT | CMPLT | 2024-06-04 10:52 | 2024-07-25 15:00 | Screen frozen during use | (TC 06/04/2024 10:52:15 AM, Canjimavadivel, Kamallesh ) Downloaded logs to review (TC 07/24/2024 03:02:57 PM, Canjimavadivel, Kamallesh ) Set u pwith test lung to run bench test as per instruction from vendor. (TC 07/25/2024 02:59:53 PM, Canjimavadivel, Kamallesh ) Verified no additional issues with 24 hr bench test. Device will be returned to service. Updated vendor with status.
1064210 | 207682 | CM | ACTIV | ISSUE | 2024-06-04 10:55 |  | TE 28 + 20004 | (TC 06/04/2024 10:54:45 AM, Canjimavadivel, Kamallesh ) Spoke to stan and downloaded logs
1048333 | 207683 | CM | CMPLT | CMPLT | 2024-06-04 11:12 | 2024-06-04 13:18 | Defective battery | (TC 06/04/2024 01:18:15 PM, Rhuland, Courtney) Replaced battery Functional check - OK Brought to OPG
| 207684 | CM | CMPLT | CMPLT | 2024-06-04 12:22 | 2024-06-04 12:23 | Meeting with Janet | (TC 06/04/2024 12:22:34 PM, Baghirzade, Nara) Met with Janet to discuss the placement of exam lights, phototherapy lights, and gas tanks on the new omnibeds and incubators.
| 207685 | CM | CMPLT | CMPLT | 2024-06-04 12:24 | 2024-06-04 12:24 | Retiring old incubators and closing pending WOs on TMS | (TC 06/04/2024 12:23:47 PM, Baghirzade, Nara) Retired the following incubators: 67914 67956 67969 63873 67917 67928 67913 Closed any pending PM or repair WOs.
1054985 | 207686 | CM | CMPLT | CMPLT | 2024-06-04 13:09 | 2024-06-04 13:21 | Incoming inspection and set up for replacement Root | (TC 06/04/2024 01:20:37 PM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 207687 | CON | CMPLT | CMPLT | 2024-06-04 13:24 | 2024-08-01 11:07 | Vital sign monitor repurposing | (TC 06/04/2024 01:24:22 PM, Rhuland, Courtney) Relocated 6 VSMs from hospital to OPG SPP to replace out of support devices Removed old devices to be retired (TC 06/06/2024 10:39:34 AM, Rhuland, Courtney) walkthrough 60 murray to asess number of VSMs that will need to be replaced. (TC 06/20/2024 02:49:56 PM, Rhuland, Courtney) Retagged and exchanged VSMs at OPG 8th floor (TC 08/01/2024 11:06:54 AM, Rhuland, Courtney) Located VSM unused on 15th floor Loaded adult profile to device Brought to OPG and brought old device back to biomed
| 207688 | RDS | CMPLT | CMPLT | 2024-06-04 14:03 | 2024-06-04 14:03 | Weekly Rounds: 11 North | (TC 06/04/2024 02:02:44 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 207689 | RDS | CMPLT | CMPLT | 2024-06-04 14:05 | 2024-06-04 14:05 | Weekly Rounds: 11 South | (TC 06/04/2024 02:04:58 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 207690 | RDS | CMPLT | CMPLT | 2024-06-04 14:07 | 2024-06-04 14:07 | Weekly Rounds: NICU & RESUS | (TC 06/04/2024 02:07:18 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
| 207691 | CM | CMPLT | CMPLT | 2024-06-04 14:23 | 2024-07-03 10:20 | PIC+ patient interface cable | (TC 06/05/2024 01:48:48 PM, Scero, Tomlyn) placed the order (TC 06/17/2024 02:36:27 PM, Scero, Tomlyn) ETA given by medtronics, backordered until july 4/24 (TC 07/03/2024 10:19:45 AM, Scero, Tomlyn) order received
1049901 | 207692 | CM | CMPLT | CMPLT | 2024-06-04 15:26 | 2024-06-20 14:09 | Small battery not charging | (TC 06/04/2024 03:27:04 PM, Fan, Tak) Put spare in place. Charging on toaster, seems to be charging. Contacted Supervisors for replacement recommendations. (TC 06/20/2024 02:09:14 PM, Fan, Tak) Batt ordered and delivered.
1066919 | 207693 | CM | CMPLT | CMPLT | 2024-06-04 16:41 | 2024-06-04 16:42 | Scale uncalibrated message | (TC 06/04/2024 04:41:07 PM, Cortez, John) Calibrated scale, and verified accuracy
| 207694 | CM | CMPLT | CMPLT | 2024-06-04 17:05 | 2024-06-28 15:15 | misc biomed - june 2024 | (TC 06/03/2024 05:05:18 PM, Cortez, John) Checked ECG carts to confirm Wifi card model. Confirmed to be CF20G, or option D21 (TC 06/04/2024 05:07:37 PM, Cortez, John) Tightened thermometer holder in ICU room 10 (TC 06/05/2024 04:44:33 PM, Cortez, John) Steve from RT requested Expair to be insalled on RT laptop for remote use. REached out to IT for admin rights, and connected with Jacky from IT to install VPN on laptop Waiting on physician's user ids (TC 06/07/2024 05:01:27 PM, Cortez, John) Installed Expair on RT laptop but still need to figure out network setup. Also waiting on IT to setup VPN. (TC 06/18/2024 01:45:46 PM, Cortez, John) ICU called, patient monitor in room 10 wasn't turning on. Power cord was just loose on the boom (TC 06/18/2024 04:13:39 PM, Cortez, John) Continued working to setup RT laptop for remote viewing Contacted Chris who suggested to copy directory from current workstations that use Expair Logins and network seems to hav ebeen setup properly. Next step is VPN access (TC 06/21/2024 04:01:42 PM, Cortez, John) Amanda from ICU reached out saying the battery for defib Pod A wasn't charging. Stopped by and saw that outlet's ground fault protection tripped. Hit switch on outlet and power returned. (TC 06/28/2024 03:14:42 PM, Cortez, John) Went to PFT with Adeel and Steve, to look into request from Gabby to get scheduling software (efficcess) on network We need credentials to access full version. Adeel will reach out (TC 06/26/2024 03:24:40 PM, Cortez, John) systems meeting
1062661 | 207695 | CM | CMPLT | CMPLT | 2024-06-05 08:57 | 2024-07-03 08:47 | send the device to Conmed for PM | (TC 06/12/2024 08:44:00 AM, Huang, Jessica) requested the loaner and received it, swapped the device on the moible cart and sent it for PM (TC 07/03/2024 08:46:20 AM, Huang, Jessica) sent the quote to Reeba for approval. received the device back and installed it in OR 9. (TC 07/03/2024 08:47:27 AM, Patel, Rutvik) working with Jessica to install the device in OR 9
32510 | 207696 | CM | CMPLT | CMPLT | 2024-06-05 09:00 | 2024-06-05 09:01 | MSH DR1 DETECTOR GRIDS/GUIDES HAVE BEEN LUBRICATED BY PERSONS UNKNOWN. REQUIRES THOROUGH OIL REMOVAL | (TC 06/04/2024 09:00:00 PM, Whittington, Scott) MSH DR1 DETECTOR GRIDS/GUIDES HAVE BEEN LUBRICATED BY PERSONS UNKNOWN. REQUIRES THOROUGH OIL REMOVAL OIL REMOVED | MSH DR1 DETECTOR GRIDS/GUIDES HAVE BEEN LUBRICATED BY PERSONS UNKNOWN. REQUIRES THOROUGH OIL REMOVAL
32520 | 207697 | CM | CMPLT | CMPLT | 2024-06-05 09:05 | 2024-06-05 09:13 | msh dr2 hdd full/ oil removal followup | (TC 06/05/2024 09:04:39 AM, Whittington, Scott) msh dr2 hdd full/ oil removal followup | msh dr2 hdd full/ oil removal followup
1061212 | 207698 | CM | CMPLT | CMPLT | 2024-06-05 10:08 | 2024-06-05 11:07 | Recording Issue | June 04 - Call re: cannot record the procedure. Went to check. Waited for access. Found system was not able to record procedure. In recording menu, when selecting video preview; the screen displays “system ideal” message. When attempting to start recording, error message “Capture failed” appears. Rebooted nCare. Did not work. Rebooted VMC. still did not work. Contacted Olympus rep. He contacted another rep. Attempted to resolve the issue but no result. Was advised to reach out to Olympus US support. June 05 - Contacted Olypus US and opened service ticket. Connected with a rep and was advised to Remove the nCare unit and brought it down to biomed. Arranged for the monitor and network access. Updated IP for the tech suppoer rep to remote in. Issue resolved. Updated unit to old IP and reinstalled it back in the OR cabinet. Checked functionality. Tested ok.
66829 | 207699 | PM | CMPLT | CMPLT | 2024-06-05 10:37 | 2024-06-05 11:10 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 06/05/2024 11:10:12 AM, DeGuzman, Leila) PM completed. Go to procedures to see details.
| 207700 | PA | CMPLT | CMPLT | 2024-06-05 11:18 | 2024-06-05 11:24 | Masimo accessories biomed stock | (TC 06/05/2024 11:24:03 AM, Canjimavadivel, Kamallesh ) Placed order for stock accessories
66127 | 207701 | CM | CMPLT | CMPLT | 2024-06-05 13:25 | 2024-06-05 13:25 | Broken bracket | (TC 06/05/2024 01:25:08 PM, Rendon Lopez , Elizabeth) Removed device from NICU RT. Dvice is being replaced by new model. Retired from service and disposed
67945 | 207702 | CM | CMPLT | CMPLT | 2024-06-05 13:40 | 2024-06-05 13:40 | Air probe failure | (TC 06/05/2024 01:40:16 PM, Rhuland, Courtney) Problem seems to be intermittent - 2 previous work orders for same issue but unable to replicate Replaced air probe Replaced scale that could not be recognized Functional check - OK
69628 | 207703 | CM | CMPLT | CMPLT | 2024-06-05 14:30 | 2024-06-06 10:58 | Channel error | (TC 06/05/2024 03:35:09 PM, DeGuzman, Leila) Device came down, issue: channel error. Cleaned device, Downloaded error and event log. Error code found: 240.4150.0 Performed check-in test: all passed. Performed analog sensor: failed sensor in fluid side. Sensor to be change... (TC 06/06/2024 10:57:55 AM, DeGuzman, Leila) Replaced the bottle-side pressure sensor. Performed analog pressure test: passed/within range. Performed check-in tests: all passed.
68669 | 207704 | CM | CMPLT | CMPLT | 2024-06-05 14:34 | 2024-06-06 09:32 | Broken | (TC 06/06/2024 09:32:20 AM, DeGuzman, Leila) Came down says broken. Cleaned device. Downloaded error log, error code: 240.4150.0 Replaced the bottle-side pressure sensor. Performed analog sensor: passed/within range. Performed check-in: all test passed.
1061210 | 207705 | CM | CMPLT | CMPLT | 2024-06-05 15:14 | 2024-06-05 15:14 | Monitor image flipped | (TC 06/05/2024 03:13:43 PM, Reid, Emmerson ) Was able to flip the image back in the settings. Working now. Also assembled 2 hoses for nitrogen with teflon tape and retrieved the biomed cart i left up in OR
1063794 | 207706 | CM | CMPLT | CMPLT | 2024-06-05 16:01 | 2024-07-11 14:23 | Buzzing noise | Jun 05 - Call re: not working properly. Went to check. Spokw with staff and was informed the light is making excessive noise. BOrught it down to test it. Found two small screws on the headlight very loose that were vibrating when turning it on. Tightned it. The cap of the headlight still made buzzing noise. Consulted JH. WAs advised to contact the surgi one. Contacted vendor. Awaiting resposne. Jun 06 - Followed up. Vendor rep dropped by to meet and offered to take the headlight. Created SC. Awaiting apporval. Received. Issued PO vi MMC. Jun 19 - Followed up on the status. Was informed the tech support is still working on it. Jul 11 - Received unit. Went to OR and borrowed battery. Checked functionality Tested ok. Returned to OR.
1050075 | 207707 | CM | CMPLT | CMPLT | 2024-06-06 09:24 | 2024-07-18 13:56 | Screen distorted | (TC 06/06/2024 01:43:59 PM, Canjimavadivel, Kamallesh ) Started disassembling unit to replace screen (TC 06/14/2024 10:20:55 AM, Canjimavadivel, Kamallesh ) Attempted to swap screen from older device that is to be retired, screens are not the same. REquested quote from vendor. (TC 06/18/2024 08:45:38 AM, Canjimavadivel, Kamallesh ) Followed up with vendor on quote. Placed order for display (TC 07/18/2024 01:56:05 PM, Canjimavadivel, Kamallesh ) Installed new display and tested ok.
1058274 | 207708 | PM | CMPLT | CMPLT | 2024-06-06 09:51 | 2024-06-06 09:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 06/06/2024 09:51:09 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69584 | 207709 | CM | CMPLT | CMPLT | 2024-06-06 10:26 | 2024-06-06 10:26 | Channel error | (TC 06/06/2024 10:26:01 AM, Rendon Lopez , Elizabeth) Downloaded error log, bottle pressure sensor fatal error. Replaced bottle PS, and worn hinge plate. Calibrated, check in tests, all passed.
1071401 | 207710 | CM | CMPLT | CMPLT | 2024-06-06 10:31 | 2024-06-06 10:31 | No issue stated | (TC 06/06/2024 10:31:21 AM, Rhuland, Courtney) No issue stated on tag Ran incubator for 45 minutes and error code 7 appeared - Thermistor reference voltage is out of range Performed system calibration Error 15 appeared - Safety Relay 1 test failure Reseated calibration plug and error cleared Ran for 2 hours and no issues Functional check - OK
1061209 | 207711 | CM | CMPLT | CMPLT | 2024-06-06 10:32 | 2024-06-06 10:33 | Touch Screen Frozen | Call re: touch screen not working. Went to check. Waited for access. Found touch screen not responding. Rebooted VMC. Issue persisted. Disconnect and reconnected all cables. Rebooted VMC. Issue resolved. Tested ok.
1050572 | 207712 | CM | CMPLT | CMPLT | 2024-06-06 11:30 | 2024-06-06 13:27 | Not working | call re: not working. Went to check. Waited for access Found the monitor displaying a grey square but not the actual image. Checked all connections and Verified cable integrity. Rebooted the unit. Arranged fro a scope to test. Issue persisted. Upon investigation, found scope guide input was swithced to a different input. Rectified it. Checked functionality. Tested ok.
65904 | 207713 | CM | CMPLT | CMPLT | 2024-06-06 11:32 | 2024-06-06 13:54 | Occlusion | (TC 06/06/2024 01:53:31 PM, DeGuzman, Leila) Device came down: reads occlusion. Problem not found. Cleaned device. Performed check-in tests: all passed Returned to service.
1055115 | 207714 | CM | CMPLT | CMPLT | 2024-06-06 11:35 | 2024-06-07 08:34 | Not charging | (TC 06/06/2024 02:10:09 PM, DeGuzman, Leila) Reported issue was alarming low battery - not charging. Cleaned device. Turn on device, very low battery. Charging device and performing battery conditioning test. (TC 06/07/2024 08:33:39 AM, DeGuzman, Leila) Battery capacity: 3765 mamphr Performed check-in tests: All passed. Returned to service.
1071398 | 207715 | CM | CMPLT | CMPLT | 2024-06-06 12:44 | 2024-06-25 13:09 | Error code 82 | (TC 06/06/2024 12:46:59 PM, Rhuland, Courtney) Error code 82 present when incubator is turned on Power cycled as instructed by service manual but error still present Put unit into service mode through hardware and error log read "No Hash File" Contacted Getinge and waiting to hear from service tech (TC 06/06/2024 02:00:29 PM, Rhuland, Courtney) Spoke with service tech and he suggested it could be the coin battery Replaced battery but error persisted Board replacement to come (no charge, under warranty) (TC 06/25/2024 01:08:55 PM, Rhuland, Courtney) Replaced board pins came loose from connector - reseated Calibrated RH Functional check - OK
1047843 | 207716 | CM | CMPLT | CMPLT | 2024-06-06 12:56 | 2024-06-06 13:23 | System failure - required biomed password | (TC 06/06/2024 01:23:19 PM, Rhuland, Courtney) Primary audible alarm BGND test alarm present Cleared error Checked connection to speaker Ran multiple test infusions and could not replicate error
1046203 | 207717 | CM | CMPLT | CMPLT | 2024-06-06 15:09 | 2024-06-06 14:08 | Display cuts out when moving monitor. | (TC 06/06/2024 02:08:00 PM, Fan, Tak) Replaced HDMI cable.
501071 | 207718 | CM | CMPLT | CMPLT | 2024-06-06 15:21 | 2024-06-07 08:38 | MSH CEOU RM 12 Monitor is black | (TC 06/06/2024 03:21:33 PM, Fan, Tak) Removed AVIO from spare system for Jillian. Reinstalled. (TC 06/07/2024 08:38:26 AM, Chau, Ken) replaced monitor, tested okay (TC 06/11/2024 12:07:00 PM, Whittington, Scott) MAY 31 MSH CEOU ROOM 12 MONITOR INSTALLED REPLACEMENT MONITOR HAD SPLASH SCREEN , BUT I/P FAUTY? DOA | MSH CEOU RM 12 Monitor is black
1062421 | 207719 | CM | CMPLT | NFF | 2024-06-06 16:12 | 2024-06-06 16:12 | "Can't move up or down" | (TC 06/06/2024 04:12:21 PM, Cortez, John) Up and down movements were just locked out. Spoke with ward clerk and support services clerk about feature.
69291 | 207720 | CM | CMPLT | CMPLT | 2024-06-07 08:50 | 2024-06-07 11:23 | Channel error | (TC 06/07/2024 11:23:27 AM, DeGuzman, Leila) Device came down with an issue channel error message. Cleaned device. Downloaded Error log: error code - 240.4150.0 Performed analog pressure sensor test: failed Replaced pressure sensor. Performed check-in tests: all passed. All good. returned to service.
1066036 | 207721 | CM | CMPLT | CMPLT | 2024-06-07 08:51 | 2024-06-07 11:18 | Channel knocked off onto the table. | (TC 06/07/2024 11:18:28 AM, DeGuzman, Leila) Device came down with the issue channel knocked off onto the floor, and need to be checked to make sure not broken. Performed check-in tests: all passed
1050474 | 207722 | CM | CMPLT | CMPLT | 2024-06-07 09:39 | 2024-08-29 09:09 | Paper tray latch not working | Jun 6 - Call re: not working. Schedule and went to check. Found printer paper tray not able to hold its position when inserted and bounces back open. Removed back cover and inspected unit. Found paper tray latch not able to hold lock position. SPokw with nurse and Temporarily taped it for their use. Will connected with vendor for support. Jun 7 - Contacted vendor. Was informed that someone from tech support will be in touch. Awaiting response. June 13 - Followed up. Was informed that Olympus still services the unit. Scheduled and removed printer from service. Contacted Olympus for RMA. Awaiting. June 18 - Followed up. Was informed no RMA required. Created SC and requeted approval. Received. Issued PO via MMC. Packed unit. Scheduled and shipped unit via KJV. Awaiting delivery. Jul 11 - Followed up. Tech support indicated that they did not receive the unit. Confirmed with the carrier that it was delivered to Olympus on June 20 - Investigated further with Olympus found the unit with Olympus CS. Was informed that they will send repair estimate soon. Aug 29 - Received repair quote of >$3000. Was advised not to move forward with the repair. Requested olympus to ship the unit back. Inspection fee will be charged. Sept 25 - Unit received. Unboxed. Removed stickers. Disposed unit.
1061212 | 207723 | CM | CMPLT | CMPLT | 2024-06-07 10:24 | 2024-06-07 10:26 | Coriomatrix error | Call re: video placement issue. Went to check. Waited for access. Found error "routing not possible, missing video matrix". Venrified video system connections. Reebooted VMC. Issue persisted. Found RED LED on coriomatrix. Rebooted coriomatrix and VMC again. Issue resolved. Checked functionality. Tested ok.
1070248 | 207724 | CM | CMPLT | CMPLT | 2024-06-07 11:06 | 2024-06-07 11:06 | Batteries need replacing | (TC 06/07/2024 11:05:56 AM, Rhuland, Courtney) Purchased batteries from Rexall and replaced batteries in scale Verified that Wifi was disabled to help with batteries draining quickly
66823 | 207725 | PM | CMPLT | CMPLT | 2024-06-07 11:37 | 2024-06-07 14:05 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 06/07/2024 02:04:48 PM, DeGuzman, Leila) PM completed. All tests passed Go to procedures to see details.
1053348 | 207726 | CM | CMPLT | CMPLT | 2024-06-07 13:51 | 2024-09-25 09:54 | MSH BI 1272 H3 ERROR GEN IMAGE CAPTURE 25:53 | (TC 06/07/2024 01:50:56 PM, Whittington, Scott) MSH BI 1272 H3 ERROR GEN IMAGE CAPTURE 25:53 SO YOUNG REPORTED TO MIE (TC 06/07/2024 10:50:00 AM, Whittington, Scott) JUNE 6 MSH BI 1272 H3 ERROR GEN IMAGE CAPTURE 25:53 SO YOUNG REPORTED TO MIE (TC 06/10/2024 10:50:00 AM, Whittington, Scott) JUNE 6 MSH BI 1272 H3 ERROR GEN IMAGE CAPTURE 25:53 SO YOUNG REPORTED TO MIE | MSH BI 1272 H3 ERROR GEN IMAGE CAPTURE 25:53 SO YOUNG REPORTED TO MIE
501069 | 207727 | CM | CMPLT | CMPLT | 2024-06-07 13:56 | 2024-09-25 09:54 | MSH CEOU RM 11 BOOT UP ERROR | (TC 06/27/2024 03:35:10 PM, Whittington, Scott) MSH CEOU RM 11 BOOT UP ERROR HAD NOT BEEN EXPERIENCED IN 2 WEEKS | MSH CEOU RM 11 BOOT UP ERROR
1061209 | 207728 | CM | CMPLT | CMPLT | 2024-06-07 15:13 | 2024-06-07 15:14 | Bluetooth speaker Issue | Call re; Speaker not working. Went to check. Waited for access. Found bluetooth not connecting, Rebooted VMC. Checked connectivity. Issue resolved. Tested ok.
1065409 | 207729 | CM | CMPLT | CMPLT | 2024-06-07 15:15 | 2024-06-07 15:15 | Parameter config issue | Request from nurse to change the parameter display screen. Went to check. Confirgured dispaly as requested. Tested ok.
| 207730 | CM | CMPLT | CMPLT | 2024-06-07 15:17 | 2024-07-11 14:20 | Elmed Insulation tester Issue | Jun 7 - Request from MDRD to troubleshoot Elmed insulation tester that they use to check probe insulation integrity. Consulted SC. Found IC was placed incorrectly. Rectified it. Found speaker not working due to depleted battery. Ordered new batteries via SAP. Awaiting delivery. Jul 11 - Replaced battery. Rectified IC position. Checked functionality. The continuity test works however Found speaker malfunction. Spoke with Garry Bassi and was advised that he will reach out to the 3rd party company who will repair the unit. Returned to MDRD.
32540 | 207731 | CM | CMPLT | CMPLT | 2024-06-07 15:23 | 2024-06-07 15:23 | Table vertical movement error | -Table vertical movement disabled -Issue was due to a faulty GTS battert -Replaced battery and configured table values -Table 0 adjustment done -Tested system functionality and returned for use.
1061458 | 207732 | CM | CMPLT | CMPLT | 2024-06-07 16:54 | 2024-06-21 14:19 | "motion functions work for a short time, then stop" | (TC 06/07/2024 04:54:18 PM, Cortez, John) Potential battery issue Prompt says battery disconnected even with battery switch in on position. Tried to swap batteries but ordered batteries aren't the same size. Will need to order again Provided unit with spare bed and informed Barb (TC 06/17/2024 10:17:37 AM, Cortez, John) Orderes batteries (TC 06/21/2024 02:19:20 PM, Cortez, John) Batteries arrived. Replaced batteries and charged overnight
66846 | 207733 | PM | CMPLT | CMPLT | 2024-06-10 08:37 | 2024-06-10 09:22 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 06/10/2024 09:22:17 AM, DeGuzman, Leila) PM completed, all tests passed. Go to procedure to see details.
67948 | 207734 | CM | CMPLT | CMPLT | 2024-06-10 09:02 | 2024-06-10 09:02 | Broken door hinge | (TC 06/10/2024 09:02:06 AM, Rhuland, Courtney) Repaired broken door hinge Replaced exam light Functional check - OK
| 207735 | RDS | CMPLT | CMPLT | 2024-06-10 09:12 | 2024-06-10 09:12 | Weekly Checks For PACU | (TC 06/10/2024 09:12:20 AM, Reid, Emmerson ) Rounds complete (TC 06/10/2024 09:12:20 AM, Reid, Emmerson ) Rounds complete
| 207736 | RDS | CMPLT | CMPLT | 2024-06-10 09:13 | 2024-06-10 09:14 | Weekly Checks For OR | (TC 06/10/2024 09:13:57 AM, Reid, Emmerson ) Rounds complete (TC 06/10/2024 09:13:57 AM, Reid, Emmerson ) Rounds complete
1054158 | 207737 | CM | CMPLT | CMPLT | 2024-06-10 09:19 | 2024-07-15 14:38 | Not working | (TC 06/07/2024 09:18:56 AM, Rendon Lopez , Elizabeth) Sending out for repair (TC 06/18/2024 11:53:54 AM, Rendon Lopez , Elizabeth) Incoming inspection, missing screw cover, returned to spectramedx under warranty (TC 07/15/2024 02:37:51 PM, Rendon Lopez , Elizabeth) Repair completed. Incoming inspection-OK
1058534 | 207738 | CM | CMPLT | CMPLT | 2024-06-10 09:21 | 2024-09-25 11:21 | Bladder Scanner charger not working | (TC 06/10/2024 09:21:17 AM, Reid, Emmerson ) Checked the voltage and found the the issue was in the cable. I noticed the cable was also damaged and likely cause the issue just passed the transformer. I reached out to CHS tech support to see if warranty replacement was possible and if not to purchase a replacement cable. (TC 06/24/2024 08:51:07 AM, Reid, Emmerson ) Submitted NC PO request and completed the service Request form SC 1003257915 (TC 06/26/2024 01:24:23 PM, Reid, Emmerson ) Sent out with KJV (TC 09/25/2024 11:21:25 AM, Reid, Emmerson ) Came back from repair after needed much follow up and several months. Connected it back with the bladder scanner.
1047030 | 207739 | CM | CMPLT | CMPLT | 2024-06-10 09:21 | 2024-06-18 11:40 | Not working | (TC 06/07/2024 09:21:20 AM, Rendon Lopez , Elizabeth) Sending out for repair (TC 06/18/2024 11:40:12 AM, Rendon Lopez , Elizabeth) incoming inspection-OK
1057502 | 207740 | CM | CMPLT | CMPLT | 2024-06-10 09:22 | 2024-06-18 11:30 | Background noise | (TC 06/07/2024 09:22:22 AM, Rendon Lopez , Elizabeth) Loose crystals. Sending out for repair (TC 06/18/2024 11:30:21 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1056130 | 207741 | CM | CMPLT | CMPLT | 2024-06-10 09:24 | 2024-06-17 11:13 | No FHR Sound | (TC 06/07/2024 09:23:52 AM, Rendon Lopez , Elizabeth) Sending out for repair (TC 06/17/2024 11:13:23 AM, Rendon Lopez , Elizabeth) Broken beyond repair. Retire from service.
1060134 | 207742 | CM | CMPLT | CMPLT | 2024-06-10 09:25 | 2024-06-18 12:01 | Beeping noises | (TC 06/07/2024 09:25:35 AM, Rendon Lopez , Elizabeth) Sending out for repair (TC 06/18/2024 12:01:05 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1063183 | 207743 | CM | CMPLT | CMPLT | 2024-06-10 09:27 | 2024-06-18 11:42 | Beeping noises | (TC 06/07/2024 09:28:18 AM, Rendon Lopez , Elizabeth) Sending out for repair (TC 06/18/2024 11:42:02 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1067117 | 207744 | CM | CMPLT | CMPLT | 2024-06-10 09:29 | 2024-06-18 11:26 | Not detected | (TC 06/07/2024 09:30:08 AM, Rendon Lopez , Elizabeth) Sending out for repair under warranty (TC 06/18/2024 11:26:20 AM, Rendon Lopez , Elizabeth) Repair not covered under warranty because of physical damage. proceeded with flat rate repair. Incoming inspection-OK
1051433 | 207745 | CM | CMPLT | CMPLT | 2024-06-10 09:31 | 2024-06-18 12:05 | Not reading Ctx | (TC 06/07/2024 09:32:19 AM, Rendon Lopez , Elizabeth) Sending out for repair (TC 06/18/2024 12:04:33 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1051391 | 207746 | CM | CMPLT | CMPLT | 2024-06-10 09:34 | 2024-06-18 11:34 | Not working | (TC 06/07/2024 09:34:52 AM, Rendon Lopez , Elizabeth) Sending out for repair (TC 06/18/2024 11:33:37 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1055530 | 207747 | CM | CMPLT | CMPLT | 2024-06-10 10:06 | 2024-06-10 10:06 | Broken temperature probe | (TC 06/10/2024 10:05:41 AM, Rhuland, Courtney) Replaced temperature probe Functional check - OK
1061214 | 207748 | CM | CMPLT | CMPLT | 2024-06-10 10:45 | 2024-06-10 10:46 | PACs PC not showing on Monitors | (TC 06/10/2024 10:45:45 AM, Reid, Emmerson ) Switched the display mode to douplicate from only on 1, also showed Amanda how this works. Also repaired a tourniquet cuff by replacing the o ring and tested it.
| 207749 | CM | CMPLT | CMPLT | 2024-06-10 11:01 | 2024-06-28 10:57 | order for WA otoscope bulbs (3100), biomed stock | (TC 06/10/2024 11:02:34 AM, Cortez, John) placed order for 1 bx (TC 06/28/2024 10:56:34 AM, Cortez, John) arrived
| 207750 | RDS | CMPLT | CMPLT | 2024-06-10 11:15 | 2024-06-10 13:45 | Weekly checks in CCU | (TC 06/10/2024 01:45:25 PM, Scero, Tomlyn) weekly checks complete. 30J tests passed.
| 207751 | CM | CMPLT | CMPLT | 2024-06-10 14:00 | 2024-07-04 13:29 | Purchasing a new External Pacemaker Analyzer | (TC 06/11/2024 02:02:30 PM, Scero, Tomlyn) Emailed Medset for a quote and ordered on ereq (TC 06/28/2024 01:29:33 PM, Scero, Tomlyn) tester arrived. unboxed and asset tagged/ (TC 07/05/2024 08:55:39 AM, Chukri, Samir) completed the work order.
68571 | 207752 | CM | CMPLT | CMPLT | 2024-06-10 15:25 | 2024-06-10 15:25 | Needs to be repaired | (TC 06/10/2024 03:25:20 PM, Rendon Lopez , Elizabeth) Sent down from 15M Antenatal. Device missing battery, replaced battery, tested, lots of background noise. Unable to repair, notified manager. Retired.
1061203 | 207753 | CM | CMPLT | CMPLT | 2024-06-11 08:19 | 2024-06-11 08:25 | PACs PC monitor not working | (TC 06/11/2024 08:24:34 AM, Reid, Emmerson ) Went up to check even though this is a helpdesk issue. Monitor was not displaying any input. When I removed the secondary HDMI the image reapeared. The display settings were set to extend. I fixed the settings and it is working now.
1048600 | 207754 | CM | CMPLT | CMPLT | 2024-06-11 09:12 | 2024-06-11 09:12 | Upgraded to V4.3 | (TC 06/11/2024 09:12:11 AM, Reid, Emmerson ) Upgraded the firmware to V4.3 and push the palliative library
63931 | 207755 | PM | CMPLT | CMPLT | 2024-06-11 09:13 | 2024-06-11 09:14 | PM-MAJOR-Annual -PM for : EEG simulator | (TC 06/11/2024 09:13:55 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1048602 | 207756 | CM | CMPLT | CMPLT | 2024-06-11 09:14 | 2024-06-11 09:14 | Upgraded to V4.3 | (TC 06/11/2024 09:14:15 AM, Reid, Emmerson ) Upgraded the firmware to V4.3 and pushed the palliative library
68148 | 207757 | PM | CMPLT | CMPLT | 2024-06-11 09:29 | 2024-06-11 09:30 | PM-MAJOR - Annual PM For: Pulse CO-Oximeter | (TC 06/11/2024 09:29:41 AM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached.
1061210 | 207758 | CM | CMPLT | CMPLT | 2024-06-11 09:40 | 2024-06-10 12:41 | Vitals not showing up | CAll re: Vitals not showing up on wall monitor. Went to check. Waited for access. Found gas machine cavle connection loose. Removed, reconnected and secured the cable. Issue resolved. Tested ok.
65159 | 207759 | CM | CMPLT | CMPLT | 2024-06-11 09:44 | 2024-06-11 09:44 | Upgraded to V4.3 | (TC 06/11/2024 09:43:55 AM, Reid, Emmerson ) Upgraded the firmware to V4.3 and pushed the APS IV PCA MAR 2022 library
65159 | 207760 | PM | CMPLT | CMPLT | 2024-06-11 09:45 | 2024-06-11 10:05 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/11/2024 10:05:13 AM, Reid, Emmerson ) PM complete, see procedure for details.
| 207761 | RDS | CMPLT | CMPLT | 2024-06-11 09:56 | 2024-06-11 10:00 | Weekly check for ED
1070801 | 207762 | CM | CMPLT | CMPLT | 2024-06-11 09:57 | 2024-06-11 09:57 | Phototherapy light install | (TC 06/11/2024 09:56:50 AM, Rhuland, Courtney) Adjusted exam light, oxygen tank holder and phototherapy light Configured setup so lights were not too close to heaters confirming with NICU staff about configuration before continuing with the entire fleet
| 207763 | RDS | CMPLT | CMPLT | 2024-06-11 09:57 | 2024-07-03 12:00 | Weekly check for ICU
1051663 | 207764 | CM | CMPLT | CMPLT | 2024-06-11 10:18 | 2024-06-14 14:08 | PCM wants doppler mounted on stand | (TC 06/10/2024 10:18:26 AM, Rendon Lopez , Elizabeth) Requested stand ordering information from STAT Medical (TC 06/11/2024 10:18:26 AM, Rendon Lopez , Elizabeth) Ordered new stand (TC 06/14/2024 02:07:42 PM, Rendon Lopez , Elizabeth) Received stand, assembled, mounted doppler. Tested-Ok. Returned to 15M
1067323 | 207765 | CM | CMPLT | CMPLT | 2024-06-11 10:51 | 2024-07-15 13:49 | Turns off when plugged in | (TC 06/11/2024 10:52:29 AM, Rendon Lopez , Elizabeth) Confirmed problem, kept to send out for repair under warranty (TC 07/15/2024 01:48:31 PM, Rendon Lopez , Elizabeth) Unable to repair. Retired from service
67741 | 207766 | CM | CMPLT | CMPLT | 2024-06-11 10:56 | 2024-06-11 11:08 | Frayed probe wire | (TC 06/11/2024 11:07:33 AM, DeGuzman, Leila) Probe connecting wire frayed. Replaced the probe. Cleaned device and battery connection Checked functions, all good. Returned to service.
1049964 | 207767 | CM | CMPLT | CMPLT | 2024-06-11 10:59 | 2024-06-11 11:12 | Broken probe | (TC 06/11/2024 11:11:55 AM, DeGuzman, Leila) Replaced the probe. Cleaned device and battery connection Checked functions, all good. Returned to service.
1049961 | 207768 | CM | CMPLT | CMPLT | 2024-06-11 11:07 | 2024-06-11 11:13 | Broken probe | (TC 06/11/2024 11:12:48 AM, DeGuzman, Leila) Replaced the probe. Cleaned device and battery connection Checked functions, all good. Returned to service.
66491 | 207769 | CM | CMPLT | CMPLT | 2024-06-11 11:11 | 2024-06-11 11:11 | Machine says there is suction, but no suction when tested with hand | (TC 06/11/2024 11:10:44 AM, Rendon Lopez , Elizabeth) Tested vacuum, adjusted canister accessories, all functions working and values in range. NFF, returned to L&D
| 207770 | CM | CMPLT | CMPLT | 2024-06-11 11:23 | 2024-07-02 14:25 | Computer issue - Date and Time | Computer not able to save date and time format. Contacted HD. They remoted in and adjusted settings but did not resolve the issue. Was informed that they need to look into it. Jun 19 - Followed up. Was infomed that they are still looking into this issue. Jun 21 - Followed up. Was informed by the tech that they will need to file an approval for the date format change. Awaiting approval. Jul 02 - Followed up. Was informed IS cannot change the date format as it is standardized across SHS. Case closed.
1064167 | 207771 | PM | CMPLT | CMPLT | 2024-06-11 12:17 | 2024-06-11 12:17 | PM MAJOR | (TC 06/11/2024 12:16:54 PM, Chukri, Samir) Performed calibration by Getinge tech, calibration certificate attached
1051809 | 207772 | CM | CMPLT | CMPLT | 2024-06-11 13:10 | 2024-06-11 13:10 | Not accurate reading | (TC 06/11/2024 01:09:54 PM, DeGuzman, Leila) Checked funtions: Found error code: E05 - heater error Front panel buttons not working. Swap with other spare Old SN: A1953736X Checked Function: All okay. Return to service.
66554 | 207773 | CM | CMPLT | CMPLT | 2024-06-11 14:09 | 2024-06-11 11:51 | Buzzing sound when plug-in | (TC 06/11/2024 11:51:11 AM, DeGuzman, Leila) No buzzing sound occured. Checked Functions: all okay. Go to procedures to see details.
1057950 | 207774 | CM | CMPLT | CMPLT | 2024-06-11 14:17 | 2024-06-21 09:26 | Remote Issue/Arm board Broken | Jun 10 - Call Re: table not moving. Went to check. Patient on the table. Waited for access. Foudn remote not responding. Rebooted. Issue resolved. Also found armboard broken. Contacted vendor for purchase info. Jun 11 - Followed up. Info received. Created SC and sent for approval. Received. Issued PO via MMC. Awaiting delivery. Jun 21 - Delivery received. Unboxed and delivered to OR. Cut and Disposed boxes.
1049362 | 207775 | CM | CMPLT | CMPLT | 2024-06-11 15:40 | 2024-06-11 15:40 | Faulty Injector Head Cable | Faulty injector main cable Replaced head main cable that runs from the control room base to the scan room. Pulled the cable through the conduit connected and tested for functionality.
1053347 | 207776 | CM | CMPLT | CMPLT | 2024-06-12 07:30 | 2024-06-11 07:30 | MSH BI H1 - Noise During Tilt | (TC 06/10/2024 07:29:41 AM, Chau, Ken) Heard loose screw in system, located 1 loose part, contacted holigic and was informed system is not safe to use as the chain link part is iff -put system off line -error also associated was with the paddle slider (TC 06/11/2024 07:29:41 AM, Chau, Ken) Worked with hologic to tightened chain link Hologic replaced slider and fixed issue Removed other loose screw that was associated with the board
66524 | 207777 | CM | CMPLT | CMPLT | 2024-06-12 11:11 | 2024-06-12 11:11 | Recorder not printing | (TC 06/06/2024 11:11:13 AM, Rendon Lopez , Elizabeth) Removed from room replaced with spare (TC 06/07/2024 11:11:13 AM, Rendon Lopez , Elizabeth) Replaced recorder with spare repaired recorder. Tested all functions, OK (TC 06/12/2024 11:11:13 AM, Rendon Lopez , Elizabeth) Returned to L&D room 14
1055617 | 207778 | CM | CANCL | DUP | 2024-06-12 11:36 | 2024-06-13 09:31 | Machine not turning on, missing BP cuff | (TC 06/13/2024 09:23:14 AM, DeGuzman, Leila) Reported device in not turning on, missing NIBP cuff. Place NIBP cuff. Check funtions and battery: all okay.
1056447 | 207779 | CM | CMPLT | CMPLT | 2024-06-12 11:44 | 2024-06-13 09:21 | Reported device in not turning on, missing NIBP cuff. | (TC 06/13/2024 09:20:41 AM, DeGuzman, Leila) Reported device in not turning on, missing NIBP cuff. Place NIBP cuff. Check funtions and battery: all okay.
1055617 | 207780 | CM | CMPLT | CMPLT | 2024-06-12 11:45 | 2024-06-13 09:45 | Came down with the Root | (TC 06/13/2024 09:45:06 AM, DeGuzman, Leila) Patient cable was frayed. Replaced the patient cable. Checked function; Okay.
1056508 | 207781 | CM | CMPLT | CMPLT | 2024-06-12 11:46 | 2024-06-13 09:38 | Came down with the Root | (TC 06/13/2024 09:38:18 AM, DeGuzman, Leila) Replaced battery. Place SpO2 set.
32510 | 207782 | CM | CMPLT | CMPLT | 2024-06-12 13:53 | 2024-09-25 09:54 | MSH DR1 EXPOSURE ISSUES | (TC 06/12/2024 01:52:48 PM, Whittington, Scott) MSH DR1 EXPOSURE ISSUES | MSH DR1 EXPOSURE ISSUES
1061211 | 207783 | CM | CMPLT | CMPLT | 2024-06-12 14:05 | 2024-06-12 14:16 | Monitor turning Blue | Call re: blue screen. Went to check. Waited for access. Found surgical tower connected to incorrect port. Rectified it. Checked functionality.Tested ok.
67987 | 207784 | RQ | ACTIV | ISSUE | 2024-06-12 14:06 |  | OR repair - not working | (TC 06/12/2024 02:06:55 PM, Chukri, Samir) Received and checked found power supply need repair order parts from Digikey . vendor qoute for power board exchange is $1500 pluse Bulb $1000. we try to repair the power board in house and we have a bulb $800 | picked up by Jessica
1046060 | 207785 | RQ | ACTIV | ISSUE | 2024-06-12 14:20 |  | OR Repair - No Light | (TC 06/12/2024 02:21:09 PM, Chukri, Samir) Samir working on it ... looking for parts try to repair in house, Reeba request to retired this unit. | Not working no light picked up by Jessica Samir working on it ... looking for parts try to repair in house, Reeba request to retired this unit. (high cost to repair it by Vendor)
67987 | 207786 | RQ | CANCL | CANCL | 2024-06-12 14:27 | 2024-06-13 09:01 | not working light
1056665 | 207787 | CM | CMPLT | CMPLT | 2024-06-12 14:29 | 2024-06-12 14:32 | Cannot connect to monitor | Call re: cannot connect microscope to wall monitor. Went to check. Waited for access. Found the BNC cable connector was forced onto the boom connector and its center pins was bent. Rectified it. Found staff connected the DVI cable onto the boom instead of microscope. Rectified it. Rebooted the microscope and SCALEOR. Issue resolved. Tested ok.
1056444 | 207788 | CM | CMPLT | CMPLT | 2024-06-13 09:33 | 2024-06-13 09:33 | Missing O2 cord, and BP cuff | (TC 06/13/2024 09:32:34 AM, DeGuzman, Leila) Reported O2 cord and missing NIBP cuff. Place NIBP cuff. Check funtions: all okay.
64921 | 207789 | RQ | CMPLT | CMPLT | 2024-06-13 10:19 | 2024-06-17 13:40 | OR Repair - just test | (TC 06/17/2024 01:40:04 PM, Chukri, Samir) completed. | OR test
1063649 | 207790 | CM | CMPLT | CMPLT | 2024-06-13 11:56 | 2024-06-13 11:57 | MSH BI RAD REVIEW WORK STATION FIRST INSIDE DOOR, WOULD NOT BOOT | (TC 06/13/2024 11:56:25 AM, Whittington, Scott) MSH BI RAD REVIEW WORK STATION FIRST INSIDE DOOR, WOULD NOT BOOT CHECKED AND RESEATED CONNECTIONS SEVERAL ATTEMPTS, SYSTEM BOOTED USER NAME mammo PW mamo | MSH BI RAD REVIEW WORK STATION FIRST INSIDE DOOR, WOULD NOT BOOT
1067990 | 207791 | CM | CMPLT | CMPLT | 2024-06-13 12:39 | 2024-06-13 12:39 | MSH ER VIVD IQ UNABLE TO INITIATE PROBES | (TC 06/13/2024 12:38:46 PM, Whittington, Scott) MSH ER VIVD IQ UNABLE TO INITIATE PROBES The main probe input wasn’t fully inserted & locked down. Corrected Ok now | MSH ER VIVD IQ UNABLE TO INITIATE PROBES
1060974 | 207792 | RQ | CMPLT | CMPLT | 2024-06-13 13:20 | 2024-06-17 12:54 | OR Repair - ForceTriade E300 | (TC 06/13/2024 01:27:54 PM, Reid, Emmerson ) Testing if this shows up (TC 06/17/2024 12:54:22 PM, Reid, Emmerson ) Reached out to the vendor. He said this was a scanning issue. I cleaned the glass on the scanner and sent back to the unit. | ESU E300 error. Reported by Amanda from OR13.
65543 | 207793 | CM | CMPLT | CMPLT | 2024-06-13 13:33 | 2024-06-13 13:26 | BP Leak | Call re: BP not working. Went to check. waited for access. FOund NIBP hose old and warn out and connector jammed with virox. Arranged for a new NIBP hose and replaced it. Checked functionality. Tested ok. Returned to service.
1053010 | 207794 | CM | CMPLT | CMPLT | 2024-06-13 15:42 | 2024-07-11 14:13 | System froze during post-stress | (TC 06/13/2024 03:42:02 PM, Fan, Tak) Checked system and spoke to tech. System froze during 1.5 min clip acquisition. Tech waited 5 minutes then restarted the machine. Lost the clip. Could not find clip. System appears OK. Will follow up about normal clip sizes. (TC 07/11/2024 12:13:00 PM, Fan, Tak) Followed up with Rashmi. Apparently it happened multiple times on different machines with the same user. Happens on the continuous capture clip on a post-stress exam. Crashing ruins the stress test since scanning window is 90 sec post stress. Clip length can be up to 90 sec. (TC 07/11/2024 02:13:00 PM, Fan, Tak) Followed up with Kevin from GE apps. Apparently it can happen on the post-stress continuous capture if the user Freezes and unfreezes frequently. But he suggested getting apps training for new staff. | User restarted system, image lost.
32580 | 207795 | CM | CMPLT | CMPLT | 2024-06-13 20:40 | 2024-06-13 20:40 | Gantyr Rotation Error | Gantyr Rotation Error Error has been happening for the last few days - first thing in the morning requiring to restart the system Part Servo Amp was on back order Received the part and replaced it System tested and no error
1058187 | 207796 | CM | CMPLT | CMPLT | 2024-06-14 09:10 | 2024-06-14 09:10 | Upgraded to V4.3 | (TC 06/14/2024 09:09:36 AM, Reid, Emmerson ) Upgraded firmware to V4.3 and pushed APS IV PCA Mar 2022 library
65135 | 207797 | CM | CANCL | DUP | 2024-06-14 09:42 | 2024-06-14 09:42 | Upgrade to V4.3 | (TC 06/14/2024 09:42:07 AM, Reid, Emmerson ) Upgraded firmware to V4.3 and pushed APS IV PCA Mar 2022 library.
65132 | 207798 | CM | CMPLT | CMPLT | 2024-06-14 09:47 | 2024-06-14 09:47 | Upgraded to V4.3 | (TC 06/14/2024 09:47:16 AM, Reid, Emmerson ) Upgraded firmware to V4.3 and pushed APS IV PCA Mar 2022 library
1054767 | 207799 | CM | CMPLT | CMPLT | 2024-06-14 10:25 | 2024-06-14 13:32 | BP hose is broken, temp probe not reading | (TC 06/14/2024 01:31:34 PM, Canjimavadivel, Kamallesh ) Swapped out dual hose connector, replaced temp probe. All functions pass
1071504 | 207800 | CM | CMPLT | CMPLT | 2024-06-14 11:05 | 2024-06-13 11:05 | Testing function, adding to TMS | (TC 06/13/2024 11:04:58 AM, Rendon Lopez , Elizabeth) 15M Manager requested to check. All functions-OK, added to TMS
1057499 | 207801 | CM | CMPLT | CMPLT | 2024-06-14 11:16 | 2024-06-07 11:16 | Not reading contractions | (TC 06/07/2024 11:15:48 AM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
1064955 | 207802 | CM | CMPLT | CMPLT | 2024-06-14 11:17 | 2024-06-07 11:17 | Background noise | (TC 06/07/2024 11:16:47 AM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
1069316 | 207803 | CM | CMPLT | CMPLT | 2024-06-14 11:17 | 2024-06-07 11:17 | No working | (TC 06/07/2024 11:17:29 AM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
1048003 | 207804 | CM | CMPLT | CMPLT | 2024-06-14 11:18 | 2024-06-07 11:19 | Doesn't work | (TC 06/07/2024 11:18:46 AM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
1057483 | 207805 | CM | CMPLT | CMPLT | 2024-06-14 11:19 | 2024-06-07 11:19 | Doesn't work | (TC 06/07/2024 11:19:29 AM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
1069781 | 207806 | CM | CMPLT | CMPLT | 2024-06-14 11:20 | 2024-06-07 11:20 | Doesn't work | (TC 06/07/2024 11:20:16 AM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
1051390 | 207807 | CM | CMPLT | CMPLT | 2024-06-14 11:21 | 2024-06-07 11:21 | Doesn't work | (TC 06/07/2024 11:21:06 AM, Rendon Lopez , Elizabeth) Tested, NFF. Returned to L&D
1063331 | 207808 | CM | CMPLT | CMPLT | 2024-06-14 12:54 | 2024-09-25 09:55 | MSH OPG SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK | (TC 06/14/2024 12:54:35 PM, Whittington, Scott) MSH OPG SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK | MSH OPG SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK
1063808 | 207809 | CM | CMPLT | CMPLT | 2024-06-14 12:58 | 2024-08-06 12:37 | MSH SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK | (TC 06/14/2024 12:58:33 PM, Whittington, Scott) MSH SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK (TC 08/06/2024 12:35:27 PM, Whittington, Scott) MSH SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK FIREWALL HAS BEEN RECONFIGURED TO ACCEPT THE IP OK | MSH SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK
1057302 | 207810 | CM | CMPLT | CMPLT | 2024-06-14 13:00 | 2024-09-25 09:56 | MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK | (TC 06/14/2024 12:59:50 PM, Whittington, Scott) MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK (TC 08/21/2024 01:03:40 PM, Whittington, Scott) MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK (TC 08/21/2024 01:04:22 PM, Whittington, Scott) MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK (TC 08/21/2024 05:04:00 PM, Whittington, Scott) MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK | MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK
1050904 | 207811 | CM | CMPLT | CMPLT | 2024-06-14 13:01 | 2024-09-25 09:56 | MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK | (TC 06/14/2024 01:00:40 PM, Whittington, Scott) MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK | MSH OGP SPP Ultrasound GE VOLUSON 10 SPP>HSC DATA LINK
1071464 | 207812 | CM | CMPLT | CMPLT | 2024-06-14 14:15 | 2024-06-10 14:15 | Battery malfunction | (TC 06/10/2024 02:14:39 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071465 | 207813 | CM | CMPLT | CMPLT | 2024-06-14 14:16 | 2024-06-10 14:16 | Not working | (TC 06/10/2024 02:15:52 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071467 | 207814 | CM | CMPLT | CMPLT | 2024-06-14 14:17 | 2024-06-10 14:17 | Doesn't work | (TC 06/10/2024 02:16:44 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1046431 | 207815 | CM | CMPLT | CMPLT | 2024-06-14 14:17 | 2024-06-10 14:18 | Doesn't work | (TC 06/10/2024 02:17:33 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071477 | 207816 | CM | CMPLT | CMPLT | 2024-06-14 14:18 | 2024-06-10 14:18 | Doesn't work | (TC 06/10/2024 02:18:27 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071478 | 207817 | CM | CMPLT | CMPLT | 2024-06-14 14:19 | 2024-06-10 14:19 | Doesn't work | (TC 06/10/2024 02:19:05 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071479 | 207818 | CM | CMPLT | CMPLT | 2024-06-14 14:20 | 2024-06-10 14:20 | Doesn't work | (TC 06/10/2024 02:19:48 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071481 | 207819 | CM | CMPLT | CMPLT | 2024-06-14 14:20 | 2024-06-10 14:21 | Doesn't work | (TC 06/10/2024 02:20:38 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071482 | 207820 | CM | CMPLT | CMPLT | 2024-06-14 14:21 | 2024-06-10 14:22 | Doesn't work | (TC 06/10/2024 02:21:32 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071483 | 207821 | CM | CMPLT | CMPLT | 2024-06-14 14:23 | 2024-06-10 14:23 | Doesn't work | (TC 06/10/2024 02:23:13 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1071484 | 207822 | CM | CMPLT | CMPLT | 2024-06-14 14:26 | 2024-06-10 14:26 | Doesn't work | (TC 06/10/2024 02:25:48 PM, Rendon Lopez , Elizabeth) Replaced battery with used one. Tested-OK.
1046426 | 207823 | CM | CMPLT | CMPLT | 2024-06-14 14:26 | 2024-06-11 14:27 | Doesn.t work | (TC 06/11/2024 02:26:36 PM, Rendon Lopez , Elizabeth) Cleaned contacts, charged. Tested-OK
1071480 | 207824 | CM | CMPLT | CMPLT | 2024-06-14 14:28 | 2024-06-11 14:28 | Doesn't work | (TC 06/11/2024 02:27:41 PM, Rendon Lopez , Elizabeth) Added to TMS. Cleaned contacts, charged. Tested-OK
67733 | 207825 | CM | CMPLT | CMPLT | 2024-06-14 14:28 | 2024-06-11 14:29 | Doesn't work | (TC 06/11/2024 02:28:35 PM, Rendon Lopez , Elizabeth) Cleaned contacts, charged. Tested-OK
67725 | 207826 | CM | CMPLT | CMPLT | 2024-06-14 14:29 | 2024-06-11 14:30 | Doesn't work | (TC 06/11/2024 02:29:31 PM, Rendon Lopez , Elizabeth) Cleaned contacts, charged. Tested-OK
| 207827 | CM | CMPLT | CMPLT | 2024-06-14 15:52 | 2024-11-12 10:34 | OR Inventory | OR Inventory (TC 07/02/2024 03:53:27 PM, Patel, Rutvik) Updated TMS record on OR Assets (TC 07/23/2024 08:43:01 AM, Patel, Rutvik) Completed OR invenstory. Awaiting OR Team approval to divide the assets. (TC 07/24/2024 12:43:00 PM, Patel, Rutvik) Completed OR invenstory. Awaiting OR Team approval to divide the assets. (TC 08/14/2024 01:56:02 PM, Patel, Rutvik) Updated TMS Record and OR Assets (TC 11/11/2024 03:33:00 PM, Patel, Rutvik) Printed stickers and tagged all OR equipment.
| 207828 | CM | CMPLT | CMPLT | 2024-06-14 15:53 | 2024-06-14 16:28 | 12N Move | Worked in 12 N for moving and organizing beds. Cut and disposed boxeds and garbabge.
1071511 | 207829 | CM | CMPLT | CMPLT | 2024-06-17 07:37 | 2024-06-17 07:44 | Incoming inspection | (TC 06/17/2024 07:44:14 AM, Rhuland, Courtney) Incoming inspection completed Electrical safety - OK Asset tagged (TC 06/17/2024 10:53:01 AM, Whittington, Scott) MSH L&D SONOSITE DEMO LX
1071512 | 207830 | CM | CMPLT | CMPLT | 2024-06-17 07:38 | 2024-06-17 07:42 | Incoming inspection | (TC 06/17/2024 07:42:16 AM, Rhuland, Courtney) Incoming inspection completed Electrical safety - OK Asset tagged (TC 06/17/2024 10:36:56 AM, Whittington, Scott) MSH OR SONOSITE DEMO PX
1064452 | 207831 | CM | CMPLT | PMFAIL | 2024-06-17 08:11 | 2024-06-17 08:14 | Retired | Retired Device not working.
| 207832 | RDS | CMPLT | CMPLT | 2024-06-17 09:05 | 2024-06-17 13:21 | Weekly checks in CCU | (TC 06/17/2024 01:21:28 PM, Scero, Tomlyn) Checks complete. restarted CCU overview 1049218 because i was getting a high GDI warning. plugged in defib 1055799 because I got an email that said the AC power wasn't connected. 30J tests passed.
1047862 | 207833 | PM | CMPLT | CMPLT | 2024-06-17 09:33 | 2024-06-17 10:04 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 06/17/2024 10:03:35 AM, Rhuland, Courtney) PM completed All tests passed Refer to procedure for details
67877 | 207834 | PM | CMPLT | CMPLT | 2024-06-17 09:37 | 2024-06-17 10:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/17/2024 10:40:24 AM, DeGuzman, Leila) PM completed. All tests passed.
1059519 | 207835 | RQ | CMPLT | NFF | 2024-06-17 10:22 | 2024-06-17 10:22 | OR Repair - NIM not turning on | (TC 06/17/2024 10:22:02 AM, Reid, Emmerson ) Power cord was disconnected. I reconnected it and it turned on. Tested with simulator and it was working. | Reported that it was not turning on and possibly missing a cable. Power cord was disconnected. I reconnected it and it turned on. Tested with simulator and it was working. Returned to the OR desk
65416 | 207836 | CM | CMPLT | CMPLT | 2024-06-17 10:31 | 2024-06-17 10:32 | Fan making noise | (TC 06/17/2024 10:31:31 AM, Rhuland, Courtney) Fan making squeaking sound when turned on Replaced fan and rubber seal Noise stopped Calibrated RH Functional check - OK
| 207837 | CM | CMPLT | CMPLT | 2024-06-17 10:38 | 2024-06-28 12:20 | Transport Battery
1057940 | 207838 | RQ | CMPLT | CMPLT | 2024-06-17 10:50 | 2024-07-11 08:22 | OR Repair - Broken Leg Segment | (TC 06/17/2024 10:59:35 AM, Reid, Emmerson ) Emailed the vendor to get a quote for the broken part (TC 07/11/2024 08:21:52 AM, Reid, Emmerson ) Recieved the part and repaired the leg segment.
67890 | 207839 | PM | CMPLT | CMPLT | 2024-06-17 11:12 | 2024-06-17 14:33 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/17/2024 02:32:39 PM, DeGuzman, Leila) PM completed. All tests passed.
1056911 | 207840 | CM | CMPLT | CMPLT | 2024-06-17 11:19 | 2024-06-24 08:34 | Lamp 2 Over 500 hours | (TC 06/17/2024 11:19:36 AM, Reid, Emmerson ) Checked the lamp this morning and it needs to be replaced. I don't have the lamp on hand so I reached out to the valleymed for a replacement (TC 06/24/2024 08:33:34 AM, Reid, Emmerson ) Recieved the lamp. Replaced the lamp and reset the timer.
69318 | 207841 | CM | CMPLT | CMPLT | 2024-06-17 11:40 | 2024-06-17 11:42 | Not working | (TC 06/17/2024 11:42:00 AM, Rendon Lopez , Elizabeth) Downloaded error log, Bottle pressure sensor fatal errors. Replaced Bottle PS, calibrated. Check in tests, all passed.
67890 | 207842 | CM | CMPLT | NFF | 2024-06-17 13:18 | 2024-06-17 13:18 | Occlusion | (TC 06/17/2024 01:18:12 PM, Rendon Lopez , Elizabeth) Ran on same settings as previous patient, no issues found. Updated SW and reloaded library
| 207843 | CM | CMPLT | CMPLT | 2024-06-17 13:39 | 2024-07-08 09:50 | 3D-Formlabs Resin and Tank Order
501583 | 207844 | CM | CMPLT | CMPLT | 2024-06-17 14:05 | 2024-09-25 09:57 | MSH 15 SPP 15-217 ROSE L&D MAIN MONITOR INTERMITTENT AGAIN, JUST REPLACED LAST WEEK | (TC 06/18/2024 12:47:41 PM, Whittington, Scott) MSH 15 SPP 15-217 ROSE L&D MAIN MONITOR INTERMITTENT AGAIN, JUST REPLACED LAST WEEK SITE 52668740 REF# ALEX 0123830939 (TC 06/24/2024 03:08:16 PM, Whittington, Scott) MSH 15 SPP 15-217 ROSE L&D MAIN MONITOR INTERMITTENT AGAIN, JUST REPLACED LAST WEEK SITE 52668740 REF# ALEX 0123830939 | MSH 15 SPP 15-217 ROSE L&D MAIN MONITOR INTERMITTENT AGAIN, JUST REPLACED LAST WEEK SITE 52668740 REF# ALEX 0123830939
1071516 | 207845 | ACC | CMPLT | CMPLT | 2024-06-17 15:18 | 2024-06-17 15:21 | Incoming Inspection | June 17 - Old unit found in biomed shop. No asset tag or history and details was recorded in TMS. Assigned asset tag and recorded in TMS. Cleaned, Vacuumed and Inpsected unit. Checked functionality. Tested ok. Installed unit in IVF as thier's was broeken and sent for repair.
1062999 | 207846 | CM | CMPLT | CMPLT | 2024-06-17 16:48 | 2024-06-17 16:49 | "remove from service notice" | (TC 06/17/2024 04:48:33 PM, Cortez, John) Confirmed issue, Error 17 in logs "The MC did not receive the CAN HB from the FMC", container was also empty Cleared log. Ran cleaning cycle with no issues/errors
1063784 | 207847 | RQ | CMPLT | CMPLT | 2024-06-18 07:25 | 2024-08-12 08:45 | OR. Repair - Light not working | (TC 06/20/2024 08:42:47 AM, Reid, Emmerson ) Tested with a battery and let it run. The fan did not activate and it ended up shitting down and overheating after a few minutes. Reached out to the vendor for an RMA (TC 06/26/2024 01:39:18 PM, Reid, Emmerson ) Submitted NC PO SC 1003259479 (TC 06/26/2024 01:39:18 PM, Reid, Emmerson ) Submitted NC PO SC 1003259479 (TC 06/26/2024 01:39:18 PM, Reid, Emmerson ) Submitted NC PO SC 1003259479 (TC 06/26/2024 01:46:09 PM, Reid, Emmerson ) Sent out with KJV (TC 08/12/2024 08:44:38 AM, Reid, Emmerson ) Recieved and returned to the OR | Picked up by Samir 7:17 Jun 18 24 , give it to Emmerson.
1061213 | 207848 | PA | CMPLT | CMPLT | 2024-06-18 09:10 | 2024-06-18 09:12 | DVI adaptors from Amazon | (TC 06/18/2024 09:11:54 AM, Reid, Emmerson ) Found and ordered
1061209 | 207849 | CM | CMPLT | CMPLT | 2024-06-18 09:33 | 2024-06-18 09:34 | Monitor 3 not showing | Call re: monitor 3 not working. Went to check. Waited for access. Found monitor 3 blank screen with screen control lock off and input switched. Corrected that. Locked screen control. Checked functionality. Tested ok.
1054227 | 207850 | CM | CMPLT | CMPLT | 2024-06-18 11:36 | 2024-09-25 09:57 | MSH OPG SPP>HSC NETWORK | (TC 06/18/2024 11:36:00 AM, Whittington, Scott) MSH OPG SPP>HSC NETWORK (TC 06/19/2024 12:43:01 PM, Whittington, Scott) MSH OPG SPP>HSC NETWORK | MSH OPG SPP>HSC NETWORK
1049608 | 207851 | CM | CMPLT | CMPLT | 2024-06-18 11:39 | 2024-09-25 09:58 | MSH OPG SPP>HSC NETWORK | (TC 06/18/2024 11:38:48 AM, Whittington, Scott) MSH OPG SPP>HSC NETWORK | MSH OPG SPP>HSC NETWORK
502452 | 207852 | CM | CMPLT | CMPLT | 2024-06-18 11:42 | 2024-09-25 09:59 | MSH OPG SPP>HSC NETWORK | (TC 06/18/2024 11:41:36 AM, Whittington, Scott) MSH OPG SPP>HSC NETWORK | MSH OPG SPP>HSC NETWORK
502453 | 207853 | CM | CMPLT | CMPLT | 2024-06-18 11:43 | 2024-09-25 09:59 | MSH OPG SPP>HSC NETWORK | (TC 06/18/2024 11:43:21 AM, Whittington, Scott) MSH OPG SPP>HSC NETWORK | MSH OPG SPP>HSC NETWORK
1069363 | 207854 | CM | CMPLT | CMPLT | 2024-06-18 11:44 | 2024-09-25 10:00 | MSH OPG SPP>HSC NETWORK | (TC 06/18/2024 11:44:20 AM, Whittington, Scott) MSH OPG SPP>HSC NETWORK | MSH OPG SPP>HSC NETWORK
1049618 | 207855 | CM | CMPLT | CMPLT | 2024-06-18 11:45 | 2024-09-25 10:00 | MSH OPG SPP>HSC NETWORK | (TC 06/18/2024 11:45:15 AM, Whittington, Scott) MSH OPG SPP>HSC NETWORK | MSH OPG SPP>HSC NETWORK
1049641 | 207856 | CM | CMPLT | CMPLT | 2024-06-18 11:55 | 2024-07-15 14:36 | innacurate readings | (TC 06/18/2024 11:56:52 AM, Rendon Lopez , Elizabeth) Confirmed issue, reading not changing accurately with different weights applied. Sending out for repair under warranty. (TC 07/15/2024 02:36:25 PM, Rendon Lopez , Elizabeth) Unable to repair. Retired from service
1048500 | 207857 | CM | CMPLT | CMPLT | 2024-06-18 12:28 | 2024-09-25 10:01 | MSH US EPIQ ROAMER OPERATOR CONTROL PANEL CYLINDER HAS SLOW DRIFT DOWN | (TC 06/18/2024 12:27:46 PM, Whittington, Scott) MSH US EPIQ ROAMER OPERATOR CONTROL PANEL CYLINDER HAS SLOW DRIFT DOWN (TC 07/16/2024 02:18:11 PM, Dolinar, John) (TC 06/18/2024 12:27:46 PM, Whittington, Scott) MSH US EPIQ ROAMER OPERATOR CONTROL PANEL CYLINDER HAS SLOW DRIFT DOWN cylinders are available when system returns to floor | MSH US EPIQ ROAMER OPERATOR CONTROL PANEL CYLINDER HAS SLOW DRIFT DOWN SITE 75025255 REF# ALEX 0123830921
1071527 | 207858 | ACC | CMPLT | CMPLT | 2024-06-18 12:54 | 2024-09-25 10:02 | MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE | (TC 06/19/2024 02:23:36 PM, Whittington, Scott) MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE SETTING UP CONTROL PANEL TO DUPLICATE 250 DUNDAS APLIO 300 LAYOUT (TC 06/20/2024 03:40:45 PM, Whittington, Scott) MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE SETTING UP CONTROL PANEL TO DUPLICATE 250 DUNDAS APLIO 300 LAYOUT (TC 07/03/2024 02:12:29 PM, Whittington, Scott) MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE From: Abe, Sheila <Sheila.Abe@sinaihealth.ca> Sent: Wednesday, July 3, 2024 8:14 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Neish, Angela - Sinai Health System <Angela.Neish@sinaihealth.ca>; Wang, Wendy - Sinai Health System <Wendy.Wang@sinaihealth.ca> Subject: [External] Ultrasound machine for the medically fragile program Hi Scott, I hope all is well with you. We wanted to inform you that we will be using a new procedure room for our medically fragile program. It will not be located in the OFC room; instead, it will be on the same floor L & D- OR room 1. Please let me know if there is anything we need to do with the ultrasound machine. Thanks, Sheila Sheila Abe, RN Clinic Manager, Mount Sinai Fertility 250 Dundas Street West, 7th floor Toronto, ON. M5T 2Z5 Sheila.abe@sinaihealth.ca From: Whittington, Scott [mailto:Scott.Whittington@uhn.ca] Sent: July 03, 2024 11:17 AM To: Abe, Sheila <Sheila.Abe@sinaihealth.ca> Cc: Neish, Angela <Angela.Neish@sinaihealth.ca>; Wang, Wendy <Wendy.Wang@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] Ultrasound machine for the medically fragile program CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Sheila I will move the system up to MSH 15 L&D OBU OR 1. There is usually space within OBU to park equipment . I will test it in an available network jack and verify it is still connecting with eIVF. Thanks Scott 416-402-2462 (TC 08/21/2024 04:52:49 PM, Whittington, Scott) MSH 15 IVF L&D APLIO 500 TEST WITH SHEILA, DR JONES SYSTEM IS ON DHCP, DID NOT SEE TEST PT 1071527 (TC 08/23/2024 12:28:20 PM, Whittington, Scott) MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE MIRZA PROVIDED STATIC IP 172.20.158.5 (TC 09/03/2024 04:31:46 PM, Whittington, Scott) MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE TEST WITH SHEILA, GOMA, MELANIE (TC 09/03/2024 04:32:37 PM, Dolinar, John) MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE | MSH 15 IVF L&D APLIO 500 SPRESSO ACCEPTANCE
1059828 | 207859 | PM | CMPLT | CMPLT | 2024-06-18 13:03 | 2024-06-18 13:04 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 06/18/2024 01:03:32 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1054961 | 207860 | CM | CMPLT | NFF | 2024-06-18 13:51 | 2024-06-17 13:51 | "not turning on" | (TC 06/17/2024 01:51:03 PM, Cortez, John) Battery not charged. Charged battery overnight, battery health appears ok All parameters tested ok
64992 | 207861 | PM | CMPLT | CMPLT | 2024-06-18 13:54 | 2024-06-18 14:00 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/18/2024 01:59:47 PM, DeGuzman, Leila) PM completed. All tests passed.
1056425 | 207862 | CM | CMPLT | NFF | 2024-06-18 14:19 | 2024-06-18 14:20 | "BP not measuring" | (TC 06/18/2024 02:19:39 PM, Cortez, John) Battery depleted Charged battery overnight, All parameters tested ok. Calibrated BP anyway, as it was off slightly, by 3 mmhg. Verified BP again
1049903 | 207863 | CM | CMPLT | CMPLT | 2024-06-18 15:37 | 2024-06-18 15:37 | Battery charger not working | (TC 06/18/2024 03:36:51 PM, Fan, Tak) Pins OK. Checked battery in both slots, wouldn't charge. Other battery was OK. Replaced the battery.
1061206 | 207864 | CM | CMPLT | CMPLT | 2024-06-18 15:46 | 2024-06-18 15:47 | Monitor Tilting Issue | Call re: monitor not staying in position. Went to check. Found loose tension. Removed side cover and adjusted it., Checked functionality. Tested ok.
1048537 | 207865 | CM | CMPLT | CMPLT | 2024-06-18 16:29 | 2024-06-18 16:28 | Battery needs replacement. | Fixed.
1056125 | 207866 | PM | CMPLT | CMPLT | 2024-06-19 07:31 | 2024-06-20 15:03 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 06/18/2024 10:00:00 AM, Canjimavadivel, Kamallesh ) Cleaned hoses and started cleaning cycle. Error code alarmed (TC 06/20/2024 03:02:56 PM, Canjimavadivel, Kamallesh ) Completed with repair WO
1056125 | 207867 | CM | CMPLT | CMPLT | 2024-06-19 07:32 | 2024-06-20 14:57 | Error code 10 RFU | (TC 06/18/2024 03:00:00 AM, Canjimavadivel, Kamallesh ) Checked service manual for error code, related to difference it temp sensor reading. Checked technical state and found manifold sensor reading different from rest of the sensors. Opened up device to replace sensor. (TC 06/19/2024 03:24:12 PM, Canjimavadivel, Kamallesh ) Installed 2 new temp probes and and repalced other 2 from old device (TC 06/20/2024 02:57:02 PM, Canjimavadivel, Kamallesh ) Reassembled device and ran cleaning cycle. Alarm no longer present and all temp values were within range. Ran cut-off test and verifications tests, all pass. Returned evice to unit.
| 207868 | RDS | CMPLT | CMPLT | 2024-06-19 07:38 | 2024-06-18 09:30 | Weekly check for ED
1057945 | 207869 | RQ | CMPLT | CMPLT | 2024-06-19 09:09 | 2024-09-17 12:45 | OR Repair - Surgical Table parts order | (TC 06/19/2024 09:19:20 AM, Reid, Emmerson ) Placed an order for the parts listed. (TC 07/30/2024 01:35:54 PM, Reid, Emmerson ) Recieved the two headboards (TC 08/01/2024 10:47:25 AM, Patel, Rutvik) Repalced and swithced over OR table mounting components to the new head boards. (TC 08/14/2024 08:13:30 AM, Reid, Emmerson ) Recieved and repaired 1-of-3 widening segments (TC 08/16/2024 10:30:05 AM, Reid, Emmerson ) Recieved 2-of-3 Widening plates and repaired the second widening segment. (TC 09/17/2024 12:44:45 PM, Reid, Emmerson ) All parts recieved and the parts have been repaired. | Placing an order for TS7000 table segment parts that have broken
1053348 | 207870 | CM | CMPLT | CMPLT | 2024-06-19 09:43 | 2024-09-25 10:02 | MSH H3 - Generator 25:56 | (TC 06/19/2024 09:43:02 AM, Chau, Ken) H3 25:56 generator error. Got sytem running by adjsuting paddles and calibrating but error typically is caused by tube or generator, Hologic to look into logs to see which part should be replaced (TC 06/21/2024 03:28:44 PM, Chau, Ken) Up down error -Hologic to perform calibration to resolve issue
1049901 | 207871 | CM | CMPLT | CMPLT | 2024-06-19 09:58 | 2024-06-19 09:58 | MSH R61 - "Battery Charger Not Working" | (TC 06/19/2024 09:57:48 AM, Chau, Ken) Faulty large detector battery -removed from service and replaced
| 207872 | CM | CMPLT | CMPLT | 2024-06-19 10:22 | 2024-06-21 14:20 | Order for Stryker procuity batteries, biomed stock | (TC 06/19/2024 01:58:42 PM, Cortez, John) Ordered batteries (TC 06/21/2024 02:20:26 PM, Cortez, John) batteries arrived
1062119 | 207873 | PM | CMPLT | CMPLT | 2024-06-19 10:50 | 2024-06-20 11:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 06/20/2024 11:11:28 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Upgraded Software and library, replaced batteries
1061215 | 207874 | CM | CMPLT | CMPLT | 2024-06-19 14:30 | 2024-06-19 14:41 | Suction Not working | Call re: suction not working. Went to check. Waited for access. Checked suction and tubing connection. Found filter was blocking the pressure. Advised to replace the filter. Checked functionality. Tested ok.
1061422 | 207875 | CM | ACTIV | ISSUE | 2024-06-19 14:34 |  | "controls not working" | (TC 06/19/2024 02:35:31 PM, Cortez, John) Confirmed issue, Power cycling bed, seems to resolve controls for footboard display and left side rail, but right side rail still not functioning. Control panel for right side rail also has different layout for some reason Will reach out to stryker when 10N swaps out bed (TC 06/25/2024 10:58:40 AM, Cortez, John) Placed service request to stryker
1061211 | 207876 | CM | CMPLT | CMPLT | 2024-06-19 14:44 | 2024-06-19 14:51 | Monitor Tilting Issue | Call re: monitor not staying in position. Went to check. Found monitor 3 had loose tension and was tilting back verticle when adjusting it at an angle. Removed side cover and adjusted it., Checked functionality. Tested ok.
1059536 | 207877 | CM | CMPLT | CMPLT | 2024-06-19 15:21 | 2024-09-25 10:03 | MSH OB ANAES 15th Floor - Murray Side L&D EDGE II PROBE SN# 04XNQ7 XTAL | (TC 06/20/2024 03:41:43 PM, Whittington, Scott) MSH OB ANAES 15th Floor - Murray Side L&D EDGE II PROBE SN# 04XNQ7 XTAL (TC 06/27/2024 03:39:17 PM, Whittington, Scott) MSH OB ANAES 15th Floor - Murray Side L&D EDGE II PROBE SN# 04XNQ7 XTAL (TC 06/26/2024 03:39:17 PM, Whittington, Scott) MSH OB ANAES 15th Floor - Murray Side L&D EDGE II PROBE SN# 04XNQ7 XTAL | MSH OB ANAES 15th Floor - Murray Side L&D EDGE II PROBE SN# 04XNQ7 XTAL
1061108 | 207878 | CM | CMPLT | CMPLT | 2024-06-19 15:24 | 2024-07-02 14:24 | Head part does not elevate | Jun 19 - Was informed by SC that someone left this bed for repair since Jan 2024. Move unit close to hallway. Removed mattress. FOund the unit could not turn on and its batteries were completely depleted. Plugged in the unit to charge battery. Inspected unit and checked functionality. All functions are working fine. Will have to check the battery later. Contacted BN. Was informed that the stryker will be coming in to replace battery and PM the bed within a couple of months. Jun 20 - Checked functionality after recharging the battery. Found the head part not elevating when the bed goes below certain level. Contacted Stryker for assistance. Awaiting response. Jun 21 - Followed up. Shceduled call with Stryker. Worked with the rep over the phone to troubleshoot the issue. Could not resolve it. Was informed that they will check if the bed is under contract and get back to me. Jun 22 - Followed up. Got confirmation that the bed is covered under contract. Scheduled and called in the rep for service. Rep came in but did not have the part available for replacement. Palced the order. Will schedule repair once part is delveired. Jul 02 - Part delivered. Schedueld and Rep came in for service. Replaced part. Checked functionality. Tested ok. Delivered unit back to 9 South.
1055924 | 207879 | CM | CMPLT | CMPLT | 2024-06-19 15:27 | 2024-06-19 15:27 | Temp probe not working | (TC 06/19/2024 03:27:14 PM, DeGuzman, Leila) Came down temp probe not working. Replaced temp probe and well Checked functions: all okay
1057051 | 207880 | CM | CMPLT | CMPLT | 2024-06-19 15:31 | 2024-06-19 15:32 | Temp probe not working | (TC 06/19/2024 03:31:35 PM, DeGuzman, Leila) Came down temp probe not working. Checked functions: Temp probe reading well. All functions are okay.
1055983 | 207881 | CM | CMPLT | CMPLT | 2024-06-19 15:36 | 2024-06-19 15:36 | Battery Replacement | (TC 06/19/2024 03:36:26 PM, DeGuzman, Leila) Came down attached to the Masimo Root. Checked function ad battery. Functions are good. Battery replaced.
1057052 | 207882 | CM | CMPLT | CMPLT | 2024-06-19 15:40 | 2024-06-19 15:40 | Battery Replacement | (TC 06/19/2024 03:39:59 PM, DeGuzman, Leila) Came down attached to the Masimo Root. Checked functions all okay. Checked battery and replaced.
1070799 | 207883 | CM | CMPLT | CMPLT | 2024-06-20 08:39 | 2024-06-20 08:39 | Scale not accurate | (TC 06/20/2024 08:39:04 AM, Rhuland, Courtney) Scale measured 850g off from test weight Calibrated scale Weight after calibration was accurate Functional check - OK
1070799 | 207884 | CM | CMPLT | CMPLT | 2024-06-20 09:09 | 2024-06-20 09:11 | Installing phototherapy light | (TC 06/20/2024 09:11:17 AM, Rhuland, Courtney) Installed phototherapy light and moved exam light and oxygen tank holder Set up a "show and tell" with NICU staff for this afternoon to get feedback on configuration and warn about melting lights
65492 | 207885 | CM | CMPLT | CMPLT | 2024-06-20 11:28 | 2024-06-10 11:28 | Missing battery | (TC 06/10/2024 11:28:07 AM, Cortez, John) Replaced battery
1059517 | 207886 | CM | CMPLT | CMPLT | 2024-06-20 13:35 | 2024-08-21 09:47 | defective PI box | (TC 08/21/2024 09:47:25 AM, Huang, Jessica) picked the device from OR desk the PI box was not connecting to the console and it was giving the error. contacted medtronic and troubleshooting the device. the issue could not get fixed. required the PO and got approval from Reeba to send the PI box for repair. received it back in Aug 2024, tested it, working fine. sent the device back to OR
1049901 | 207887 | CM | CMPLT | CMPLT | 2024-06-20 14:11 | 2024-06-24 08:47 | Battery not connecting | (TC 06/20/2024 02:11:10 PM, Fan, Tak) Battery taped down to det. The latch was stuck and preventing battery attachment. Replaced with spare. Ordered new latch. (TC 06/24/2024 08:47:30 AM, Fan, Tak) Replaced latch with new one.
| 207888 | CM | CMPLT | CMPLT | 2024-06-20 14:54 | 2024-06-20 14:54 | on call (OR 13 Pyxis Dispensing Cabinet) | (TC 06/14/2024 02:54:15 PM, Huang, Jessica) received a call from the OR desk and was informed to be back to OR 13.
1066115 | 207889 | FC | CMPLT | CMPLT | 2024-06-20 15:18 | 2024-06-20 15:19 | QC check | (TC 06/20/2024 03:18:34 PM, Canjimavadivel, Kamallesh ) QC check pass
1061217 | 207890 | CM | CMPLT | CMPLT | 2024-06-20 16:23 | 2024-06-20 16:28 | Surgical Light Handle Issue | Request from Reeba to remove red cover on surgical light handles from 3 different ORs. Went to check. Found two ORs in the middle fo the prep and was advised to remove it at the end of the day. Found OR 6 surgical light red cover stuck. Also found somm dents and scratches on the push button. Had to remove the entire handle. Replace it with a handle from OR 12. Brought back the damaged handle. Forcefully removed red cover. Dismantle the handle and rectified push button issue. Filed red cover and push button for smooth insertion. Checked functionality. Tested ok. Installed it in OR 12.
| 207891 | CM | CMPLT | CMPLT | 2024-06-21 13:40 | 2024-09-27 14:09 | Altrix temperature probe cables for biomed stock
1061214 | 207892 | CM | CMPLT | CMPLT | 2024-06-21 14:12 | 2024-06-21 14:14 | Wall monitor does not display PACS computer | Call re: Nurse cant set up Wall monitor to display PACS computer. Went to check. Waited for access. Found DP cable loose. Fixed it. Adjusted display configuration in PACS computer from exteded to duplicate. Issue resolved. Checked functionality. Tested ok.
1063746 | 207893 | CM | CMPLT | CMPLT | 2024-06-21 14:34 | 2024-06-24 13:31 | Error 10 | (TC 06/21/2024 02:34:31 PM, Rhuland, Courtney) During cleaning, error code 10 was thrown Manifold temperature found to be out of range Ordered spare cables and will test other cables (TC 06/24/2024 01:30:45 PM, Rhuland, Courtney) Used Dremmel to polish spare temperature probe Replaced temperature probe All temperatures (manifold, returns, and safeties) were within range of each other No errors present
1050633 | 207894 | CM | CMPLT | CMPLT | 2024-06-21 20:41 | 2024-09-25 10:04 | MSH 15 L&D TE7C CORD CAP END CABLE JACKET TORN | (TC 06/24/2024 04:29:31 PM, Whittington, Scott) MSH 15 L&D TE7C CORD CAP END CABLE JACKET TORN | MSH 15 L&D TE7C CORD CAP END CABLE JACKET TORN
| 207895 | RDS | CMPLT | CMPLT | 2024-06-24 08:36 | 2024-06-24 08:45 | Weekly Checks For PACU | (TC 06/24/2024 08:45:06 AM, Reid, Emmerson ) Weekly rounds complete.
| 207896 | RDS | CMPLT | CMPLT | 2024-06-24 09:36 | 2024-06-26 12:54 | weekly checks in CCU | (TC 06/26/2024 12:53:37 PM, Scero, Tomlyn) weekly checks complete. plugged in and 30J test passed.
1059578 | 207897 | CM | CMPLT | CMPLT | 2024-06-24 10:33 | 2024-06-24 10:33 | Reading inaccurate | (TC 06/24/2024 10:33:27 AM, Reid, Emmerson ) Calibrated on the temp checker, pass first attempt. Sending back
| 207898 | CM | CMPLT | CMPLT | 2024-06-24 10:36 | 2024-06-24 10:36 | installed the power bar in the OR rooms
1048528 | 207899 | CM | CMPLT | CMPLT | 2024-06-24 12:55 | 2024-06-24 13:43 | Channel error | (TC 06/24/2024 01:43:04 PM, DeGuzman, Leila) Issue found: Module not calibrated. -performed calibration -performed check-in All tests passed. Needs software update.
66711 | 207900 | CM | CMPLT | CMPLT | 2024-06-24 13:48 | 2024-06-24 13:48 | Broken porthole door | (TC 06/24/2024 01:47:53 PM, Rhuland, Courtney) Repaired porthole door Calibrated RH Repaired backwards flip door Repaired loose exam light Functional check - OK
1061202 | 207901 | CM | CMPLT | CMPLT | 2024-06-24 13:48 | 2024-06-24 13:52 | Boom arms cannot move | Call from OR asking assistace to move arms. Went to check. Spoke with surgeon and was informed that he cannot adjust the monor position where he wanted. Changed the angle of the moving arm and showed staff how to adjust and explained movement is limited up to 270* angle.
64064 | 207902 | PM | CMPLT | CMPLT | 2024-06-24 15:32 | 2024-06-24 15:32 | PM MAJOR-Uninterruptable Power Supply, Battery Replacement | (TC 06/24/2024 03:32:00 PM, Cortez, John) Switch restarted over the weekend Battery self-test failed Replaced battery
64761 | 207903 | PM | CMPLT | CMPLT | 2024-06-24 15:38 | 2024-06-24 15:39 | PM MAJOR-Uninterruptable Power Supply, Battery Replacement | (TC 06/24/2024 03:38:36 PM, Cortez, John) Replaced battery
1054985 | 207904 | CM | CMPLT | CMPLT | 2024-06-24 15:41 | 2024-07-17 14:18 | shipped the device for warranty repair | (TC 07/17/2024 02:18:01 PM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
64061 | 207905 | CM | CMPLT | CMPLT | 2024-06-24 16:20 | 2024-06-24 16:24 | Switch in closet restarted over weekend | (TC 06/24/2024 04:24:00 PM, Cortez, John) Replaced battery
1049903 | 207906 | CM | CMPLT | CMPLT | 2024-06-24 16:23 | 2024-09-25 10:04 | MSH R63 DETECTOR BATTERY WILL NOT CHARGE | (TC 06/24/2024 04:22:45 PM, Whittington, Scott) MSH R63 DETECTOR BATTERY WILL NOT CHARGE REPLACE BATTETRY | MSH R63 DETECTOR BATTERY WILL NOT CHARGE
| 207907 | CM | CMPLT | CMPLT | 2024-06-24 16:35 | 2024-09-25 10:05 | MSH ANGIO MEDI REQUESTED PARTS RETURN FOR ULTIMAX | (TC 06/24/2024 04:34:49 PM, Whittington, Scott) MSH ANGIO MEDI REQUESTED PARTS RETURN FOR ULTIMAX | MSH ANGIO MEDI REQUESTED PARTS RETURN FOR ULTIMAX
1057829 | 207908 | CM | CMPLT | CMPLT | 2024-06-25 08:01 | 2024-06-25 08:01 | Not working | (TC 06/25/2024 08:00:40 AM, Canjimavadivel, Kamallesh ) Needed to restart central to fix recorder, test print passed after restart
62696 | 207909 | CM | CMPLT | CMPLT | 2024-06-25 08:40 | 2024-06-25 08:41 | V2 not working | (TC 06/25/2024 08:41:03 AM, Canjimavadivel, Kamallesh ) Confirmed issue with V1-V3 chest lead set. Replaced chest lead set, tested ok with simulator
1061211 | 207910 | CM | CMPLT | CMPLT | 2024-06-25 08:47 | 2024-06-25 08:47 | Light handle loose | (TC 06/25/2024 08:47:19 AM, Reid, Emmerson ) Was called yesterday to check the handle this morning. Went up to check and the handles seemed fine for all the lights. I checked the monitors as well and tightened one of them. Everything is working fine.
1049902 | 207911 | CM | CMPLT | CMPLT | 2024-06-25 08:51 | 2024-06-24 15:51 | Returned to MSH from PMH. | (TC 06/24/2024 03:51:00 PM, Fan, Tak) Reconfigured and returned.
1060002 | 207912 | CM | CMPLT | NFF | 2024-06-25 09:14 | 2024-06-25 09:15 | Only 1 lead working at a time | (TC 06/25/2024 09:14:37 AM, Canjimavadivel, Kamallesh ) Tested with simulator, all leads working. Unable to replicate reported issue.
1062152 | 207913 | CM | CMPLT | CMPLT | 2024-06-25 09:17 | 2024-06-25 09:17 | Not turning on | (TC 06/25/2024 09:17:15 AM, Canjimavadivel, Kamallesh ) Repalced batteries, tested ok
| 207914 | CM | CMPLT | CMPLT | 2024-06-25 10:07 | 2024-08-06 16:38 | Order for test cable- Zoll IBP adapter | (TC 06/25/2024 10:07:20 AM, Cortez, John) Ordered 1 cable for now (TC 08/06/2024 04:37:44 PM, Cortez, John) Arrived
1056421 | 207915 | CM | CMPLT | NFF | 2024-06-25 10:33 | 2024-06-25 10:34 | "BP not working" | (TC 06/25/2024 10:33:56 AM, Cortez, John) All parameters tested ok. NFF Calibrated BP anyway
1054694 | 207916 | CM | CMPLT | CMPLT | 2024-06-25 10:44 | 2024-06-25 10:44 | "prongs on power plug missing" | (TC 06/25/2024 10:43:53 AM, Cortez, John) Replaced prong attachment
1067990 | 207917 | CM | CMPLT | CMPLT | 2024-06-25 12:34 | 2024-06-25 13:05 | MSH ER VIVID IQ WILL NOT BOOT, GOES TO BIOS PASSWORD REQUEST Renm@c5 is the PW, did not respond | (TC 06/25/2024 01:04:19 PM, Whittington, Scott) MSH ER VIVID IQ WILL NOT BOOT, GOES TO BIOS PASSWORD REQUEST Renm@c5 is the PW, did not respond SITE 416596VIQ REF# ZACHARY 1-551890310425 18006680732 Tak reseated the SSD and the system has booted up. I will bring the system back down to ER for clinical use. GE will make a follow-up service visit to check the system out. (TC 06/25/2024 01:04:31 PM, Fan, Tak) BIOS indicated all Sata slots were "empty". Looks like no operating system found so always boots to BIOS. Reseated the SSD and the system has booted up. Checked BIOS again, and SSD was identified. | MSH ER VIVID IQ WILL NOT BOOT, GOES TO BIOS PASSWORD REQUEST Renm@c5 is the PW, did not respond SITE 416596VIQ REF# ZACHARY 1-551890310425 18006680732 Tak reseated the SSD and the system has booted up. I will bring the system back down to ER for clinical use. GE will make a follow-up service visit to check the system out.
1064843 | 207918 | CM | CMPLT | NFF | 2024-06-25 13:22 | 2024-06-25 13:28 | "reads asystole despite getting pulse reading from SPO2" | (TC 06/25/2024 01:27:39 PM, Cortez, John) Tested parameters with patient simulator All parameters tested ok Issue was possibly a leads problem
| 207919 | CON | ACTIV | ISSUE | 2024-06-25 13:28 |  | MSH WOMENS 7 INFANT HEALTH MEETING RE ULTRASOUND PURCHASE | (TC 06/25/2024 01:28:00 PM, Whittington, Scott) MSH WOMENS 7 INFANT HEALTH MEETING RE ULTRASOUND PURCHASE | MSH WOMENS 7 INFANT HEALTH MEETING RE ULTRASOUND PURCHASE
65194 | 207920 | CM | CMPLT | CMPLT | 2024-06-25 13:58 | 2024-06-25 13:58 | Asked to check the Microscope | (TC 06/25/2024 01:58:02 PM, Reid, Emmerson ) Muareen asked if I could check the microscope because they wanted to use it tomorrow. I checked and was not able to get the video displayed, everything else was working. I called Steve from Osis and he told me to bypass the recorder becuase it was having a hardware issue and could not be repaired. I did this and was able to run a cable to the camera that could connect to the boom. I let Maureen know about this and that I would be available if they needed help tomorrow.
67950 | 207921 | CM | CMPLT | CMPLT | 2024-06-25 14:43 | 2024-06-25 14:43 | Broken parts | (TC 06/25/2024 02:43:14 PM, Rhuland, Courtney) Repaired side wall latch Repaired porthole hinge Functional check - OK
1071407 | 207922 | CM | CMPLT | CMPLT | 2024-06-26 07:55 | 2024-06-26 08:57 | System failure 25 | (TC 06/26/2024 08:57:13 AM, Rhuland, Courtney) Reseated fan to clear error Replaced missing corner grommet Calibrated RH Functional check - OK
1058584 | 207923 | CM | CMPLT | CMPLT | 2024-06-26 08:12 | 2024-06-26 08:12 | Not working | (TC 06/26/2024 08:12:19 AM, Reid, Emmerson ) Was sent down saying not working. I ran on the genius checker and it passed first try. Working now.
1059938 | 207924 | CM | CMPLT | CMPLT | 2024-06-26 08:22 | 2024-06-26 08:23 | Analysis error | (TC 06/26/2024 08:22:40 AM, Canjimavadivel, Kamallesh ) Tested with simulator, was able to get reading and anaylize various waveforms. Could not replicate issue.
1059596 | 207925 | CM | CMPLT | CMPLT | 2024-06-26 08:25 | 2024-06-26 08:25 | Does not give a number | (TC 06/26/2024 08:24:51 AM, Reid, Emmerson ) Ran on the genius checker and it passed calibration first try. Checked on my hand if it would read and it did. Working now.
1059570 | 207926 | CM | CMPLT | CMPLT | 2024-06-26 08:33 | 2024-06-26 08:34 | Not turning on dead batteries | (TC 06/26/2024 08:33:31 AM, Reid, Emmerson ) Batteries were dead so I replaed them and ran on the genius checker. Passed first attempt. Working now.
| 207927 | CM | CMPLT | CMPLT | 2024-06-26 12:55 | 2024-07-23 11:39 | warranty exchange for SPO2 sensors and cables for OR (RT) | (TC 06/26/2024 01:07:35 PM, Scero, Tomlyn) Collected all broken SPO2 sensors and cables given to me from the OR. Scanned the serial numbers into an excel worksheet. Contacted the Masimo tech support and gave the list of broken sensors to be covered by warranty. Created a no charge PO to give to Masimo. (TC 06/26/2024 02:37:15 PM, Scero, Tomlyn) PO received from medbuy. PO 7500208541 (TC 06/27/2024 10:45:48 AM, Scero, Tomlyn) Packaged sensors and requested shipping from KJV. waybill # 441207 (TC 07/23/2024 11:38:49 AM, Scero, Tomlyn) Warranty replacements received. 7 sensors and 3 trunk cables Stock taken to the OR
1070842 | 207928 | CM | CMPLT | CMPLT | 2024-06-26 14:42 | 2024-06-26 14:46 | Phototherapy light installation | (TC 06/26/2024 02:45:39 PM, Rhuland, Courtney) Installed photherapy light Moved O2 tank holder down Verified scale with test weight
1070737 | 207929 | CM | CMPLT | CMPLT | 2024-06-26 14:42 | 2024-06-26 14:43 | Phototherapy light installation | (TC 06/26/2024 02:42:38 PM, Rhuland, Courtney) Installed photherapy light Moved O2 tank holder down Verified scale with test weight
1070740 | 207930 | CM | CMPLT | CMPLT | 2024-06-26 14:42 | 2024-06-26 14:44 | Phototherapy light installation | (TC 06/26/2024 02:44:23 PM, Rhuland, Courtney) Installed photherapy light Moved O2 tank holder down Verified scale with test weight
1064583 | 207931 | CM | CMPLT | CMPLT | 2024-06-27 08:31 | 2024-06-27 08:31 | Not plugged in | (TC 06/27/2024 08:31:28 AM, Reid, Emmerson ) Defib was returned by L&D and never put back properly. Was left unplugged and pads were opened. I plugged it in, replaced the pads, and did 30J test. Working now.
1061209 | 207932 | CM | ACTIV | ISSUE | 2024-06-27 08:59 |  | Boom cover fell off | (TC 06/27/2024 08:59:42 AM, Reid, Emmerson ) Amanada called to let me know the side cover for the boom came off and one of the connecting pins broke. I replaced it with one from OR12 and looking to order a replacement or repair the part.
1071571 | 207933 | ACC | CMPLT | CMPLT | 2024-06-27 09:02 | 2024-07-04 07:42 | Incoming inspection | Config added and monitor installed in purple zone
1071572 | 207934 | CM | CMPLT | CMPLT | 2024-06-27 09:08 | 2024-07-04 07:42 | Incoming inspection | Config added and monitor installed in purple zone
1056466 | 207935 | CM | CMPLT | CMPLT | 2024-06-27 10:36 | 2024-06-27 15:08 | Oximetry not working / RTC battery low | (TC 06/27/2024 03:07:36 PM, DeGuzman, Leila) Issue found: RTC battery low. Device needs to send out for service Checked all functions passed for: SpO2 NIBP Temp
1059591 | 207936 | CM | CMPLT | CMPLT | 2024-06-27 11:12 | 2024-06-27 11:14 | Not Calibrating | (TC 06/27/2024 11:13:46 AM, Reid, Emmerson ) Tried testing many times over multiple days and it is still not calibrating. Unit is not under warranty and not fixable. Retiring this asset.
1058580 | 207937 | CM | CMPLT | CMPLT | 2024-06-27 11:16 | 2024-06-27 11:16 | Not calibrating | (TC 06/27/2024 11:16:24 AM, Reid, Emmerson ) Tried testing many times over multiple days and it is still not calibrating. Unit is not under warranty and not fixable. Retiring this asset.
1058574 | 207938 | CM | CMPLT | CMPLT | 2024-06-27 11:17 | 2024-06-27 11:18 | Not calibrating | (TC 06/27/2024 11:17:31 AM, Reid, Emmerson ) Tried testing many times over multiple days and it is still not calibrating. Unit is not under warranty and not fixable. Retiring this asset.
1071305 | 207939 | CM | CMPLT | CMPLT | 2024-06-27 11:31 | 2024-06-27 11:34 | Bed Error | Found Bed error during on call Daily check. Error 3009-3-12-701. Checked service manual. Found "Display backlight enable error". Went to unit. Disconnected bed and turned battery off. Performed reboot. Error resolved. checked functionality. Tested ok.
1071313 | 207940 | CM | CMPLT | CMPLT | 2024-06-27 11:36 | 2024-06-27 11:39 | Bed Error - On call check | Found Bed error during on call Daily check. Error 3009-1-5-1304. Checked service manual. Found "CAN message timeout - SDO retry". Went to unit. Disconnected bed and turned battery off. Performed reboot. Error resolved. checked functionality. Tested ok.
1071315 | 207941 | CM | CMPLT | CMPLT | 2024-06-27 11:40 | 2024-06-27 11:40 | Bed error - On call check | Found Bed error during on call Daily check. Error 3009-3-5-1370. Checked service manual. Found "Check display controller PCB Ethernet comm error". Went to unit. Disconnected bed and turned battery off. Performed reboot. Error resolved. checked functionality. Tested ok.
| 207942 | PA | CMPLT | CMPLT | 2024-06-27 12:46 | 2024-06-27 12:47 | Replacement LNCS DCI Sp02 Finger Sensor | (TC 06/27/2024 12:46:31 PM, Reid, Emmerson ) Gave a new finger sensor for PAU to replace a broken one
1056524 | 207943 | CM | CMPLT | CMPLT | 2024-06-27 15:14 | 2024-06-27 15:15 | Not working | (TC 06/27/2024 03:14:52 PM, DeGuzman, Leila) Replaced battery. Checked functions, all working.
65698 | 207944 | CM | CMPLT | CMPLT | 2024-06-27 15:28 | 2024-06-27 15:31 | Repair | Call re: Not working. Went to check and brought it down to biomed. Cleaned unit. Inpsected unit and Checked functionality. No issue found. Tested ok. Delivered back to OR.
63072 | 207945 | CM | CMPLT | CMPLT | 2024-06-27 15:53 | 2024-06-27 15:53 | Battery Replacement and Basic Repairs | (TC 06/27/2024 03:53:26 PM, Ferreira, Qwynn) Replaced the battery 3D printed a door knob Added set screw to IV Pump mount Functional Test - OK (TC 06/27/2024 03:54:26 PM, Chukri, Samir) Worked with Qwynn
65129 | 207946 | CM | CMPLT | CMPLT | 2024-06-28 08:13 | 2024-06-28 08:13 | Upgraded to V4.3 | (TC 06/28/2024 08:13:11 AM, Reid, Emmerson ) Upgraded the pump to V4.3 and p[ush APS IV PCA MAR 2022 library.
65166 | 207947 | CM | CMPLT | CMPLT | 2024-06-28 08:27 | 2024-06-28 08:27 | Upgraded to V4.3 | (TC 06/28/2024 08:27:10 AM, Reid, Emmerson ) Upgraded to V4.3 and push APS IV PCA MAR 2022 library
1057108 | 207948 | CM | CMPLT | CMPLT | 2024-06-28 08:30 | 2024-06-28 10:31 | O2 not working | (TC 06/28/2024 10:31:26 AM, DeGuzman, Leila) Issue found: Red LNC-10 patient cable from the radical 7 (asset#: 1057109) attached not working. Changed the cable Checked functions all passed for: NIBP Temp SpO2
| 207949 | OC | CMPLT | CMPLT | 2024-06-28 09:00 | 2024-07-03 09:42 | ON Call June 28-July 5 | (TC 06/28/2024 08:59:45 AM, Scero, Tomlyn) on call daily check report complete (TC 07/03/2024 09:42:22 AM, Scero, Tomlyn) daily on call checks (TC 07/02/2024 09:42:39 AM, Scero, Tomlyn) daily on call checks
1055803 | 207950 | CM | CMPLT | CMPLT | 2024-06-28 09:47 | 2024-06-28 09:48 | Plug in AC | (TC 06/28/2024 09:47:25 AM, Scero, Tomlyn) Plug in AC alert from daily checks. When we located the unit it was plugged in. Completed the 30J test and passed. (TC 06/28/2024 09:48:17 AM, Reid, Emmerson ) Assisted locating unit
1057109 | 207951 | CM | CMPLT | CMPLT | 2024-06-28 10:31 | 2024-06-28 11:45 | Battery Replacement | (TC 06/28/2024 11:45:04 AM, DeGuzman, Leila) Found the battery hasn't been replaced and the battery looks bulging. Replaced the battery and the patient cable. Checked functions: all okay.
1055907 | 207952 | CM | CMPLT | CMPLT | 2024-06-28 10:34 | 2024-06-28 10:34 | "display will not turn on, no readings" | (TC 06/28/2024 10:34:12 AM, Cortez, John) Battery needs to be charged, charged overnight Added spare SPO2 sensor, Tightened handle, Cleaned filter All parameters tested ok
501066 | 207953 | CM | CMPLT | CMPLT | 2024-06-28 12:05 | 2024-07-04 15:27 | MSH CEOU RM 4 error 227 | (TC 06/27/2024 12:05:52 PM, Whittington, Scott) MSH CEOU RM 4 error 227 SITE# 553547 REF# 0123876052 (TC 07/04/2024 03:26:04 PM, Whittington, Scott) JILLIAN REPLACED HDD AND RELOADED SW | MSH CEOU RM 4 error 227 SITE# 553547 REF# 0123876052
501071 | 207954 | CM | CMPLT | CMPLT | 2024-06-28 12:09 | 2024-09-25 10:06 | MSH CEOU-3237(Rm12) | (TC 06/28/2024 12:28:24 PM, Whittington, Scott) MSH CEOU-3237(Rm12) POSSIBLE ERROR 005 TRANSDUCER SITE 553550 REF# KEITH 0123883760 | MSH CEOU-3237(Rm12) POSSIBLE ERROR 005 TRANSDUCER SITE 553550 REF# KEITH 0123883760
| 207955 | CM | CMPLT | CMPLT | 2024-06-28 14:52 | 2024-08-13 10:28 | Ordering batteries for Philips MP5 monitor, biomed stock | (TC 06/28/2024 02:58:45 PM, Cortez, John) ordered 2x batteries (TC 08/13/2024 10:27:59 AM, Cortez, John) arrived
1046787 | 207956 | CM | CMPLT | CMPLT | 2024-06-28 15:03 | 2024-08-15 15:48 | mattress ripped | (TC 06/28/2024 03:03:39 PM, Cortez, John) Swapped bed with spare, Will request service (TC 08/15/2024 03:48:16 PM, Cortez, John) Repaired by Vince (hillrom)
1054479 | 207957 | CM | CMPLT | NFF | 2024-06-28 15:06 | 2024-06-28 15:08 | "odd noise when moving bed with intellidrive engaged, right/foot side of bed" | (TC 06/28/2024 03:06:06 PM, Cortez, John) Tried moving with/without intellidrive, Unable to replicate issue. Marlene will reach out if issue returns
502453 | 207958 | CM | ACTIV | ISSUE | 2024-06-28 15:13 |  | MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl | (TC 06/28/2024 03:13:44 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl VASSI & KATIE REPORTED TROUBLE WITH SYSTEM IVOLVING TRACKBALL AND MEASUREMENST ETC FOUND THE SYSTEM WITH NO CURSOR ON SCREEN CHECKED AND CLEANED TRKBL PT & FAMILY IN ROOM ADVISED OB TO CHECK OPERATION BEFORE I LEFT,OB CONTINUED WITH CASE WHEN TRKBL VISIBLE CALLED BACK IN SHORT TIME TOLD NO MEASUREMENTS ABLE TO MAKE MEASUREMENTS WORK (TC 07/02/2024 03:13:44 PM, Whittington, Scott) JUNE 28 MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl VASSI & KATIE REPORTED TROUBLE WITH SYSTEM IVOLVING TRACKBALL AND MEASUREMENST ETC FOUND THE SYSTEM WITH NO CURSOR ON SCREEN CHECKED AND CLEANED TRKBL PT & FAMILY IN ROOM ADVISED OB TO CHECK OPERATION BEFORE I LEFT,OB CONTINUED WITH CASE WHEN TRKBL VISIBLE CALLED BACK IN SHORT TIME TOLD NO MEASUREMENTS ABLE TO MAKE MEASUREMENTS WORK JULY 2 ERRORS CONTINUED NOTIFIED GE Please consider this a confirmation of your purchase order number : and a copy of your terms and conditions. Order Number: 800646686 No firm ETA this is back ordered Part Ordered: KTZ304071 Quantity: 1 Price: 0$ Return Auth. (RMA): 50026857 Returnable: Yes (TC 08/08/2024 04:19:27 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl OP PANEL RECEIVED AUGUST 8 (TC 08/09/2024 01:19:00 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl OP PANEL RECEIVED AUGUST 8 (TC 08/08/2024 04:19:27 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl OP PANEL RECEIVED AUGUST 8 STORED IN CEOU CONSULT ROOM NEXT TO ROOM 6 (TC 08/09/2024 01:19:00 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl OP PANEL RECEIVED AUGUST 8 (TC 08/08/2024 04:19:27 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl OP PANEL RECEIVED AUGUST 8 STORED IN CEOU CONSULT ROOM NEXT TO ROOM 6 (TC 08/09/2024 01:19:00 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl OP PANEL RECEIVED AUGUST 8 (TC 08/08/2024 04:19:27 PM, Whittington, Scott) MSH OPG D BETA VOLUSON mesurements/trkbl/op pnl OP PANEL RECEIVED AUGUST 8 STORED IN CEOU CONSULT ROOM NEXT TO ROOM 6 | msh opg spp d room beta voluson mesurements/trkbl/op pnl REF# 1-552026368223 PN# KTZ304071 DESCRIPTION: RTU110 CONSOLE EC320 TECH APPROVAL 212035838 PARTS ORDER REF# 800646686 CURRENTLY NO PARTS IN STOCK JULY 10 TANYA NEW PN# RTU CONSOLE PN# KTZ304229 REF# 800651568 RMA# 50026881
500882 | 207959 | CM | CMPLT | CMPLT | 2024-06-28 15:19 | 2024-07-08 12:06 | MSH RPCU xg probe plt 704 streaks | (TC 06/28/2024 03:19:01 PM, Whittington, Scott) msh rpcu xg probe plt 704 streaks (TC 07/03/2024 01:19:00 PM, Whittington, Scott) JUNE 28 msh rpcu xg probe plt 704 streaks (TC 07/03/2024 01:19:00 PM, Whittington, Scott) JUNE 28 msh rpcu xg probe plt 704 streaks (TC 07/08/2024 12:14:33 PM, Whittington, Scott) msh rpcu xg probe plt 704 streaks REPLAQCED BY FORMER PMH MIE 1066350 | msh rpcu xg probe plt 704 streaks REPLACED BY FORMER PMH MIE 1066350
26102 | 207960 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26114 | 207961 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26126 | 207962 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26135 | 207963 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26794 | 207964 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-09-12 08:06 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 09/12/2024 08:05:30 AM, Rhuland, Courtney) PM completed All tests passed See procedure doc for details
50997 | 207965 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Lifts, Patient Transfer
51571 | 207966 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52058 | 207967 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52092 | 207968 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS
52466 | 207969 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
62501 | 207970 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
63727 | 207971 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Testers
64493 | 207972 | PM | CMPLT | UNL | 2024-07-01 03:15 | 2024-08-12 12:51 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 08/12/2024 12:51:16 PM, Baghirzade, Nara) Looked for the aspirator in NICU and RESUS however was not able to locate it during rounds throughout July and 1st week in August. Checked in RT workshop, NICU RT, and RESUS RT. Will reopen the WO if the device is found,
65134 | 207973 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65136 | 207974 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65851 | 207975 | PM | CNP | ENL | 2024-07-01 03:15 | 2024-09-05 11:23 | PM-MAJOR - Annual PM For: Aspirators, Emergency
65893 | 207976 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-03-22 11:06 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65925 | 207977 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-07-31 07:18 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/31/2024 07:18:03 AM, Baghirzade, Nara) Replaced: -rear case -front case -mylar gasket -latch module Full calibration and check-in: Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
65945 | 207978 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-01 13:11 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65981 | 207979 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-01 12:55 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
66001 | 207980 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-01 08:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66031 | 207981 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-01 07:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66037 | 207982 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-01 07:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66466 | 207983 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-09-04 13:09 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Replaced neonatal flow sensor and cable. Replaced expiratory valve. Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 16 Ah and 220 min. Tested. Functional. SW.2.51 Leakage: 88 uA and 91 ohm.
66502 | 207984 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-08 09:16 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/08/2024 09:16:00 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
66532 | 207985 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-14 09:23 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/14/2024 09:22:56 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66700 | 207986 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-08 09:15 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/08/2024 09:14:53 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
66707 | 207987 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-09-06 11:40 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/06/2024 11:40:21 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
66708 | 207988 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66718 | 207989 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-19 10:50 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/19/2024 10:49:58 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66722 | 207990 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66730 | 207991 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-06 14:37 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/06/2024 02:36:51 PM, Rendon Lopez , Elizabeth) PM completed, cleaned and calibrated recorder. Tested, all tests passed.
66763 | 207992 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
66780 | 207993 | PM | CNP | ENL | 2024-07-01 03:15 | 2024-07-30 14:42 | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67759 | 207994 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67874 | 207995 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-07-30 11:26 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/30/2024 11:26:02 AM, Rendon Lopez , Elizabeth) Upgraded Software, downloaded library. PM completed, all tests passed
67887 | 207996 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-07-30 12:05 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/30/2024 12:05:06 PM, Rendon Lopez , Elizabeth) Upgraded Software, downloaded library. PM completed, all tests passed
67929 | 207997 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-06 14:51 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/06/2024 02:51:09 PM, Rhuland, Courtney) Replaced: porthole door 3 porthole seals Calibrated: Mains & VREF voltage Scale Functional check - OK Scale - OK Set 35C & 75% humidity Read 35C & 76.4% humidity Electrical safety - OK Ground resistance - 74 mOhms Leakage current - 84.9 uA
67964 | 207998 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-29 13:07 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/29/2024 01:06:39 PM, Rhuland, Courtney) PM completed Replaced: Rechargeable 9V battery 2 side door hinges Calibrated: VREF & MAINS voltage Scale RH Functional check - OK Set 35C, 75% humidity Read 35C, 75.5% humidity Electrical safety - OK Ground resistance - 116 mOhms Leakage current - 92.6 uA
68371 | 207999 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68377 | 208000 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68543 | 208001 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68612 | 208002 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 12:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68663 | 208003 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68734 | 208004 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 13:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68798 | 208005 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 13:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68808 | 208006 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-07-03 14:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 02:00:31 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered Transferred network configuration & data set
68812 | 208007 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68842 | 208008 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68862 | 208009 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-07-03 13:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:53:14 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered Transferred network configuration & data set
68912 | 208010 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69003 | 208011 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 14:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69008 | 208012 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 14:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69040 | 208013 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-07-17 11:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/17/2024 11:52:03 AM, DeGuzman, Leila) PM passed and completed.
69059 | 208014 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-01 10:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/01/2024 10:26:22 AM, DeGuzman, Leila) Cleaned device. Updated software Calibrated rate accuracy. PM passed and completed.
69123 | 208015 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 14:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69145 | 208016 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69149 | 208017 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69163 | 208018 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 15:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69313 | 208019 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-09 12:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:53:56 PM, Patel, Rutvik) by BD
69339 | 208020 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-09 12:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:48:07 PM, Patel, Rutvik) by BD
69421 | 208021 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69453 | 208022 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-07-03 13:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:42:03 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered
69481 | 208023 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-15 11:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/15/2024 11:12:17 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed
69546 | 208024 | PM | ACTIV | ISSUE | 2024-07-01 03:15 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69560 | 208025 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-11-21 10:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/21/2024 10:46:15 AM, Cortez, John) Updated software PM completed successfully, results logged in ASM
69663 | 208026 | PM | CANCL | DUP | 2024-07-01 03:15 | 2024-09-10 07:30 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69674 | 208027 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-16 15:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
32540 | 208028 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-12 15:40 | PM - Monthly | PM cancelled due to TGH Emerg CT service
32560 | 208029 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-07 16:33 | PM - Monthly | Cancelled due to TGH Emerg CT service
32580 | 208030 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-08 16:36 | PM - Monthly | Cancelled due to TGH Emerg CT service
501440 | 208031 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-02 21:35 | PM - Semi-Annual | PM done.
501440 | 208032 | PM | CMPLT | CMPLT | 2024-07-01 03:15 | 2024-08-09 20:33 | HARP - Semi-Annual | HARP
502044 | 208033 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-26 13:55 | PM - Semi-Annual | (TC 08/26/2024 01:55:26 PM, Green, Paul) Replaced gantry and head desiccant. Greased and replaced head o-ring seal. Topped up liquid alcohol for cooling system (half full).
1046440 | 208034 | PM | ACTIV | READY | 2024-07-01 03:16 |  | Annual PM
1046600 | 208035 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-08 12:06 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 08/08/2024 12:05:59 PM, Rhuland, Courtney) Replaced battery and vacuumed
1046601 | 208036 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-08 12:08 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 08/08/2024 12:08:17 PM, Rhuland, Courtney) Replaced battery and vacuumed
1046602 | 208037 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-08 12:09 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 08/08/2024 12:08:59 PM, Rhuland, Courtney) Replaced battery and vacuumed
1046685 | 208038 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Insufflators
1047040 | 208039 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-11-06 13:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/06/2024 01:39:46 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1047098 | 208040 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-13 07:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1047831 | 208041 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-16 10:51 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult | (TC 08/15/2024 10:50:20 AM, Scero, Tomlyn) 6 year PM parts replaced. Battery replaced. Pm tests started (TC 08/16/2024 10:51:08 AM, Scero, Tomlyn) All PM tests complete and passed. IPM-L tests results sheet attached
1047832 | 208042 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
1048486 | 208043 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Insufflators
1048545 | 208044 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Insufflators
1050045 | 208045 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-13 07:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1050257 | 208046 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-13 07:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1051088 | 208047 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1051165 | 208048 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-19 09:01 | PM-MAJOR - Annual PM For: Bladder Scanner | (TC 08/27/2024 09:01:00 AM, Huang, Jessica) PM completed
1051737 | 208049 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1054814 | 208050 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-07-11 14:34 | PM-MAJOR - Annual PM For: Ventilators | (TC 07/11/2024 02:34:11 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 26674 hrs Membrane capacity: 35%
1054815 | 208051 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-07-23 16:31 | PM-MAJOR - Annual PM For: Ventilators | (TC 07/23/2024 04:30:46 PM, Cortez, John) PM parts replaced Pre-op check passed Total operating time: 20574 hrs Total vent time: 11314 hrs Membrane capacity: 61%
1055112 | 208052 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-10 09:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055123 | 208053 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-16 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1056041 | 208054 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1056401 | 208055 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-08 12:10 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 08/08/2024 12:10:24 PM, Rhuland, Courtney) Replaced battery and vacuumed
1057512 | 208056 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057568 | 208057 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-19 12:44 | PM MAJOR - ANNUAL PM for: Electrosurgical Unit | (TC 08/19/2024 12:44:16 PM, Reid, Emmerson ) OR called about an issue with the unit. PM was due so I completed the PM and found no issue. Will install the unit back into the room it came from.
1057836 | 208058 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1058247 | 208059 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-16 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1058285 | 208060 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-16 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1059059 | 208061 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059416 | 208062 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059417 | 208063 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059418 | 208064 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059422 | 208065 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059423 | 208066 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059426 | 208067 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059428 | 208068 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059429 | 208069 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059430 | 208070 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059431 | 208071 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059433 | 208072 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059434 | 208073 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059435 | 208074 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059436 | 208075 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059437 | 208076 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059438 | 208077 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059439 | 208078 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059442 | 208079 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059444 | 208080 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059446 | 208081 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059448 | 208082 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059449 | 208083 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059450 | 208084 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059452 | 208085 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059453 | 208086 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059455 | 208087 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059456 | 208088 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059457 | 208089 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059458 | 208090 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059459 | 208091 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059460 | 208092 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059462 | 208093 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059463 | 208094 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059464 | 208095 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059544 | 208096 | PM | CANCL | CANCL | 2024-07-01 03:16 | 2024-08-16 09:57 | PM-MAJOR - Annual PM For: Central Monitor
1059877 | 208097 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1059944 | 208098 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-09 09:24 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
1060332 | 208099 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1061105 | 208100 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 10:11 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Tested ok.
1061106 | 208101 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 11:44 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061107 | 208102 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-28 09:44 | PM-MAJOR- for 9S Bed Elecric | PM comleted by stryker. Checked funtionality. Tested ok.
1061108 | 208103 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-28 10:03 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061109 | 208104 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-13 16:18 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061110 | 208105 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 10:38 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061111 | 208106 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-13 16:20 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061112 | 208107 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-28 09:46 | PM-MAJOR- for 9S Bed Elecric | PM completed by stryker. Checekd functionality. Tested ok.
1061113 | 208108 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-13 16:21 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061114 | 208109 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 09:29 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Tested ok.
1061115 | 208110 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 11:15 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061116 | 208111 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-13 16:23 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061117 | 208112 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-13 16:24 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061118 | 208113 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 10:33 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Battery replaced. Tested ok.
1061119 | 208114 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 10:47 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061120 | 208115 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-13 16:26 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Checked functionality. Tested ok.
1061121 | 208116 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-27 10:17 | PM-MAJOR- for 9S Bed Elecric | PM Completed by Stryker. Tested ok.
1061154 | 208117 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1061192 | 208118 | PM | CANCL | CANCL | 2024-07-01 03:16 | 2024-09-03 10:13 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | Device decommission and used for parts. Still waitinf for replacement from Stryker. PM kit to be used for different device
1061364 | 208119 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-06 21:56 | HARP - Semi Annual | PM was done.
1061364 | 208120 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-09 21:54 | PM - Semi Annual | PM was done.
1064323 | 208121 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | PM-MAJOR - ANNUAL PM for: Insufflators | shipping the device to ConMed for PM (TC 10/29/2024 01:46:42 PM, Patel, Rutvik) Repair estimate recevied from . Sent it for approval Received. Advised MMC to update the PO and send it to vendor.
1064645 | 208122 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-09 11:21 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1064646 | 208123 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-08-09 11:17 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
1064648 | 208124 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-07-31 07:35 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/31/2024 07:35:03 AM, Baghirzade, Nara) Replaced: -rear case -left handle Full calibration and check-in: Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1067589 | 208125 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-09-25 11:03 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 09/25/2024 11:03:28 AM, Baghirzade, Nara) PM completed. Lithium ion battery - ok and functional - power failure alarm and retention of settings is functional. All switches and controllers are functional. Warmer is functional. SpO2 is functional. Resuscitation system is functional.
1067591 | 208126 | PM | CMPLT | CMPLT | 2024-07-01 03:16 | 2024-07-25 08:23 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 07/25/2024 08:22:51 AM, Baghirzade, Nara) PM completed. Lithium ion battery - ok and functional - power failure alarm and retention of settings is functional. All switches and controllers are functional. Warmer is functional. SpO2 is functional. Resuscitation system is functional. Placed warmer back to NICU 18th fl. storage.
1070151 | 208127 | PM | ACTIV | ISSUE | 2024-07-01 03:16 |  | Quaterly PM for: 3D Pritner
| 208128 | CM | CMPLT | CMPLT | 2024-07-02 08:21 | 2024-07-02 08:32 | Sterile water for Altrix | (TC 07/02/2024 08:32:06 AM, Rhuland, Courtney) Ordered water for altrix cleaning
| 208129 | RDS | CMPLT | CMPLT | 2024-07-02 08:29 | 2024-07-02 08:30 | Weekly Checks For PACU | (TC 07/02/2024 08:29:43 AM, Reid, Emmerson ) PACU rounds complete.
1056466 | 208130 | CM | CMPLT | CMPLT | 2024-07-02 09:46 | 2024-07-25 14:24 | Battery is not holding charge. Shipping it to Masimo to exchange. | Masimo replaced the unit for battery failure with another one, sn.2000174350. Programmed. tested. functional.
1048528 | 208131 | CM | CMPLT | CMPLT | 2024-07-02 10:20 | 2024-07-02 10:20 | Calibration and check-in | (TC 07/02/2024 10:20:07 AM, Rhuland, Courtney) System calibration required after software upgrade Calibration and check-in completed All tests passed Refer to ASM for details
68813 | 208132 | CM | CMPLT | CMPLT | 2024-07-02 10:45 | 2024-07-02 10:47 | Device not charging | (TC 07/02/2024 10:47:08 AM, Rhuland, Courtney) Power cord replaced Functional check - OK
1070763 | 208133 | CM | CMPLT | CMPLT | 2024-07-02 12:54 | 2024-07-02 12:54 | Fan always on high speed | (TC 07/02/2024 12:54:01 PM, Rhuland, Courtney) Fan always on high speed error present after running for a few minutes Replaced heat sink temperature sensor from stock and will replace with warranty part once received Error cleared Verified scale accuracy Installed phototherapy light Functional check - OK
1054368 | 208134 | CM | CMPLT | CMPLT | 2024-07-02 13:22 | 2024-09-05 13:12 | Battery replacement | Battery was replaced with ED go-live
1071595 | 208135 | ACC | CMPLT | CMPLT | 2024-07-02 13:36 | 2024-07-04 10:53 | Incoming inspection | (TC 07/04/2024 10:52:48 AM, Scero, Tomlyn) Incoming inspection complete. asset tagged and IFU uploaded.
63187 | 208136 | CM | CMPLT | CMPLT | 2024-07-02 13:45 | 2024-07-02 13:50 | Broken housing | (TC 07/02/2024 01:49:45 PM, Rhuland, Courtney) Went to check light and appears to have been dropped Plastic housing around light is cracked and coming apart No repairs can be done, reaching out to vendors for quotes for replacement.
1061212 | 208137 | CM | CMPLT | CMPLT | 2024-07-02 14:27 | 2024-07-02 14:29 | Monitor Issue | CAll re: Monitor issue. Went to check. Waited for access. Was infoemd that surgeon experienced video lag on the monbitor. They also mentioned about monitors getting artifacts when turniing cautery on. Checked connections and cables. Cautery was inbuilt in Robot. Separated all cables from each other to eliminate interference. Waited for some time to see if the issue occurs again. Could not duplicate the issue. Tested ok.
1046275 | 208138 | CM | CMPLT | CMPLT | 2024-07-02 15:19 | 2024-07-23 14:49 | MSH IV2 SCANROOM 2 GOMA REPORTED OPERATOR PANEL ISSUES FREEZE & CALIPER | (TC 07/02/2024 03:19:22 PM, Whittington, Scott) MSH IV2 SCANROOM 2 GOMA REPORTED OPERATOR PANEL ISSUES FREEZE & CALIPER spoke to louiza, rooms are very busy today | MSH IV2 SCANROOM 2 GOMA REPORTED OPERATOR PANEL ISSUES FREEZE & CALIPER
| 208139 | CM | CMPLT | CMPLT | 2024-07-02 15:25 | 2024-09-25 10:07 | MSH US PROBE 681 VTE IMAGE QUALITY LABELLED TWH PROBE MIE 500747 SN# 99D1846691 | (TC 07/02/2024 03:25:42 PM, Whittington, Scott) MSH US PROBE 681 VTE IMAGE QUALITY LABELLED TWH PROBE MIE 500747 SN# 99D1846691 FOUND IMAGE ARTIFACT NOT XTAL, BUT MEMBRANE RELATED, IN NEAR FIELD RECOMMENDED CONTINUED USE | MSH US PROBE 681 VTE IMAGE QUALITY LABELLED TWH PROBE MIE 500747 SN# 99D1846691
1067975 | 208140 | CM | CMPLT | CMPLT | 2024-07-02 15:28 | 2024-09-25 10:07 | MSH US I800 IN ER PVT781 IMAGE ARTIFACT PROBE#51 CURRENTLY IN MDRD | (TC 07/02/2024 03:28:51 PM, Whittington, Scott) MSH US I800 IN ER PVT781 IMAGE ARTIFACT CURREENTLY BEING CLEANED (TC 07/02/2024 03:29:00 PM, Chau, Ken) MSH US I800 IN ER PVT781 IMAGE ARTIFACT CURREENTLY BEING CLEANED | MSH US I800 IN ER PVT781 IMAGE ARTIFACT PROBE#51 CURRENTLY IN MDRD CURREENTLY BEING CLEANED
32560 | 208141 | CM | CMPLT | CMPLT | 2024-07-02 15:49 | 2024-07-02 15:49 | Blue screen on Scan PC - | Blue screen on Scan PC System won't turn on Replaced System disk CT turned on with no issues. -Back up disk will be copied during PM - today
66692 | 208142 | CM | CMPLT | NFF | 2024-07-03 07:47 | 2024-07-03 07:47 | "Scale not working" | (TC 07/03/2024 07:47:27 AM, Rhuland, Courtney) Tested and verified scale was working and accurate 5kg test weight measured 5000g Functional check - OK No fault found
1070839 | 208143 | CM | CMPLT | CMPLT | 2024-07-03 09:51 | 2024-07-03 09:52 | Phototherapy light install | (TC 07/03/2024 09:51:41 AM, Rhuland, Courtney) Installed phototherapy light Verified accuracy of scale
1070798 | 208144 | CM | CMPLT | CMPLT | 2024-07-03 09:54 | 2024-07-03 09:54 | Phototherapy light install | (TC 07/03/2024 09:53:48 AM, Rhuland, Courtney) Installed phototherapy light Verified accuracy of scale
1070761 | 208145 | CM | CMPLT | CMPLT | 2024-07-03 09:55 | 2024-07-03 09:55 | Phototherapy light install | (TC 07/03/2024 09:54:52 AM, Rhuland, Courtney) Installed phototherapy light Verified accuracy of scale
1049892 | 208146 | PM | CMPLT | CMPLT | 2024-07-03 10:20 | 2024-07-03 10:23 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 07/03/2024 10:23:09 AM, Canjimavadivel, Kamallesh ) Replaced battery
66512 | 208147 | CM | CMPLT | CMPLT | 2024-07-03 10:24 | 2024-07-04 14:13 | Recorder not printing | (TC 07/02/2024 01:54:31 PM, Rendon Lopez , Elizabeth) Confirmed issue, removed from room (TC 07/03/2024 02:13:02 PM, Rendon Lopez , Elizabeth) Testing printer, faded print and stocked during calibration. Replaced printer head, cleaned and greased motor. Replaced back, calibrated, tested-Ok (TC 07/04/2024 02:13:02 PM, Rendon Lopez , Elizabeth) Returned to L&D room 18
1049903 | 208148 | CM | CMPLT | CMPLT | 2024-07-03 11:14 | 2024-07-02 11:14 | Noisy drive | (TC 07/02/2024 11:14:07 AM, Fan, Tak) Replaced drive motors and caster wheels. And wear pads (follow up from PM). (TC 07/03/2024 11:14:52 AM, Fan, Tak) Drive is unbalanced. Replaced strain gauges, calibrated drive, adjusted the handlebar brackets. (TC 07/25/2024 03:40:00 PM, Fan, Tak) Drives too fast/sensitive. Tried to adjust the strain gauges and handle bar height. The bar might be warped. or drive board maybe faulty. (TC 07/29/2024 03:13:22 PM, Whittington, Scott) MSH R63 NOISY DRIVE (TC 07/29/2024 03:13:58 PM, Dolinar, John) MSH R63 NOISY DRIVE LUBRICATE STRAIN GAUGE MECHANICAL MECHANISM, TEST PROBLEM STILL THERE REPAIR WIRING TO DRIVE ENCODER OK (TC 07/29/2024 04:56:20 PM, Fan, Tak) MSH R63 NOISY DRIVE | MSH R63 NOISY DRIVE
1051031 | 208149 | CM | CMPLT | CMPLT | 2024-07-03 11:17 | 2024-07-02 11:18 | No wired network | (TC 07/02/2024 11:17:32 AM, Fan, Tak) Checked connection and cables. Confirmed network cable issue. Replaced cable and tested, OK. Reported back to Faith Zhu.
| 208150 | CM | CMPLT | CMPLT | 2024-07-03 11:29 | 2024-07-03 11:30 | Spare fetal monitor part | (TC 07/03/2024 11:30:20 AM, Rendon Lopez , Elizabeth) Ordering spare parts for fetal monitors repairs Received
65974 | 208151 | PM | CMPLT | CMPLT | 2024-07-03 13:32 | 2024-07-03 13:32 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/03/2024 01:32:11 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered Confirmed calibration
65895 | 208152 | PM | CMPLT | CMPLT | 2024-07-03 13:39 | 2024-07-03 13:39 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/03/2024 01:39:07 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered
65901 | 208153 | PM | CMPLT | CMPLT | 2024-07-03 13:41 | 2024-07-03 13:41 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 07/03/2024 01:40:40 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered Confirmed calibration
68788 | 208154 | PM | CMPLT | CMPLT | 2024-07-03 13:57 | 2024-07-03 13:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/03/2024 01:57:00 PM, Rhuland, Courtney) PM completed by BD Unboxed and PM stickered Transferred network configuration & data set
1058131 | 208155 | CM | CMPLT | CMPLT | 2024-07-03 14:01 | 2024-07-03 14:01 | Not turning on even with new battteries | (TC 07/03/2024 02:00:42 PM, DeGuzman, Leila) Put new batteries, turned on. Checked errors: E5.0 - heater not working. Changed the probe Checked functions all okay. returned to service.
502044 | 208156 | CM | CMPLT | CMPLT | 2024-07-03 14:58 | 2024-08-23 15:10 | Laptop battery replacment. | (TC 07/03/2024 02:58:03 PM, Green, Paul) Removed laptop from mount. Battery is wrong model. We need a Dell battery. Emailed Digirad and let them know. They will send a replacment. (TC 08/23/2024 03:09:46 PM, Green, Paul) From: Michael Herz <mherz@ttgimagingsolutions.com> Sent: August 23, 2024 12:24 PM To: Call Center <CallCenter@ttgimagingsolutions.com> Cc: Green, Paul <Paul.Green@uhn.ca> Subject: [External] CA148470 SN:E036 at MSH *parts order* Please order the following items, P1 to the address below. PN117894, WIN7, SEE below to confirm 117883 117801 Shipping Address: Mount Sinai Hospital Attn: Barry Ivo 6th floor Nuc Med 600 University Ave Toronto, ON M5G 1X5 Michael Herz, CBET-R Cell 260-388-9409 Michael Herz / Senior Field Tech Support Engineer TTG Imaging Solutions, LLC T: (800) 856-1671 E: mherz@ttgimagingsolutions.com 8205 B&G Court, Stokesdale, NC, 27357 ¦ View map www.ttgimagingsolutions.com From: Green, Paul <Paul.Green@uhn.ca> Sent: Friday, August 23, 2024 10:07 AM To: Michael Herz <mherz@ttgimagingsolutions.com> Subject: E036 at MSH [EXTERNAL EMAIL] This email was sent from outside of TTG Imaging Solutions. Please use caution opening attachments or clicking links. Shipping Address: Mount Sinai Hospital Attn: Barry Ivo 6th floor Nuc Med 600 University Ave Toronto, ON M5G 1X5 From: Green, Paul <Paul.Green@uhn.ca> Sent: August 23, 2024 10:27 AM To: Michael Herz <mherz@ttgimagingsolutions.com> Cc: Green, Paul <Paul.Green@uhn.ca> Subject: FW: [External] ERGO Dell Latitude E6500 Issues - keyboard, overheating and GPU Hi Michael, This is another laptop with performance issues and keyboard not functioning. This is for Ergo E036 with Dell Latitude E6500 which has a contract. We also require desiccants for this system as well. Regards, Paul
1054019 | 208157 | CM | CMPLT | CMPLT | 2024-07-03 15:21 | 2024-08-28 13:26 | "only reads zero" | (TC 07/02/2024 03:20:51 PM, Cortez, John) Probe failed Caliscan test Reached out to CHS rep (TC 07/05/2024 03:43:36 PM, Cortez, John) Shipped out console+probe Requested loaner (TC 08/28/2024 01:25:53 PM, Cortez, John) Repaired unit arrived. Assembled, Functional test passed
1061211 | 208158 | CM | CMPLT | CMPLT | 2024-07-03 15:36 | 2024-07-03 15:40 | Surgical Light Issue | Call re: Light not able to turn on through Ipad. Went to check. Waited for access. Foudn one of the light had the wifi off. Turned it on. Needed to access other lights. Also found wifi not connected on IPad. Tried many times to connect the wifi but did not work. Consulted JH as well. Patient on the table and was asked to come back later. Went back later. Removed IPad from consol and rebooted the device. Issue resolved. Checked functionality. Tested ok.
63872 | 208159 | CM | CMPLT | CMPLT | 2024-07-04 08:50 | 2024-07-04 08:51 | Scale not working | (TC 07/04/2024 08:51:17 AM, Rhuland, Courtney) Scale not measuring Calibrated and verified scale accuracy Test passed Functional check - OK
1054960 | 208160 | CM | CMPLT | CMPLT | 2024-07-04 09:07 | 2024-07-03 09:07 | "missing power cord" | (TC 07/03/2024 09:07:22 AM, Cortez, John) Added spare cord and secured with cable tie Added/charged adult cuff All parameters tested ok
1054161 | 208161 | CM | CMPLT | NFF | 2024-07-04 09:22 | 2024-07-03 09:22 | Not working | (TC 07/03/2024 09:22:28 AM, Rendon Lopez , Elizabeth) Tested, NFF
1054420 | 208162 | CM | CMPLT | CMPLT | 2024-07-04 09:27 | 2024-08-15 10:40 | Not working | (TC 07/03/2024 09:27:25 AM, Rendon Lopez , Elizabeth) Confirmed issue, not detecting contractions. Kept to send out for repair (TC 08/15/2024 10:39:34 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1071620 | 208163 | CM | CMPLT | CMPLT | 2024-07-04 09:28 | 2024-08-15 14:33 | Interferance | (TC 07/03/2024 09:28:42 AM, Rendon Lopez , Elizabeth) Added device to TMS. Confirmed issue, keep to send out for repair. (TC 08/15/2024 02:33:16 PM, Rendon Lopez , Elizabeth) Incoming Inspection-OK
1064952 | 208164 | CM | CMPLT | CMPLT | 2024-07-04 09:52 | 2024-08-16 10:50 | Connection prong broken | (TC 07/03/2024 09:52:19 AM, Rendon Lopez , Elizabeth) Confirmed issue, damaged cable. keep to send out for repair (TC 08/16/2024 10:49:57 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1071585 | 208165 | ACC | CMPLT | CMPLT | 2024-07-04 09:54 | 2024-07-03 09:54 | Incoming inspection | (TC 07/02/2024 09:53:58 AM, Rendon Lopez , Elizabeth) Inspected, added to TMS (TC 07/03/2024 09:53:58 AM, Rendon Lopez , Elizabeth) Delivered to Anesthesia team
1071584 | 208166 | ACC | CMPLT | CMPLT | 2024-07-04 09:55 | 2024-07-03 10:00 | Incoming inspection | (TC 07/02/2024 10:00:17 AM, Rendon Lopez , Elizabeth) Inspected, added to TMS (TC 07/03/2024 10:00:17 AM, Rendon Lopez , Elizabeth) Delivered to Anesthesia
1061211 | 208167 | CM | CMPLT | CMPLT | 2024-07-04 10:27 | 2024-07-04 10:33 | Monitor Issue | Two calls 1) call re: monitor not disaplying image. Went to check. When I went, was told that issue is resolved. 2) call again re: monitor turning blue every now and then. Went to check. Found they were using stryker tower. Wated for some time. Observed all olympus monitors going blue for a moment coming back to original display image. Chekced all connections. Stood by for half an hour to confirm the issue. However, found no issue with stryker tower monitor. Removed old SDI cable and replaced and routed new SDI cable connecting Stryker tower and Olympus boom. Waited around for observation. Issue seemed to be resoved. Checked functionality. Tested ok.
1056125 | 208168 | PM | CMPLT | CMPLT | 2024-07-04 10:30 | 2024-07-04 10:34 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/04/2024 10:33:21 AM, Canjimavadivel, Kamallesh ) Cleaned hoses and nozzles Cleaning cycle complete
62669 | 208169 | CM | CMPLT | CMPLT | 2024-07-04 13:10 | 2024-06-10 13:10 | Otoscope bulb out | (TC 06/10/2024 01:09:56 PM, Cortez, John) Added 3100 bulb
1054537 | 208170 | PM | CMPLT | CMPLT | 2024-07-04 13:40 | 2024-07-04 13:40 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/04/2024 01:40:08 PM, Cortez, John) Completed cleaning cycle
67871 | 208171 | PM | CMPLT | CMPLT | 2024-07-04 14:15 | 2024-07-04 15:45 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/04/2024 03:45:07 PM, Rendon Lopez , Elizabeth) Upgraded Software, downloaded library. PM completed, all tests passed
63389 | 208172 | CM | CMPLT | CMPLT | 2024-07-04 14:35 | 2024-07-04 14:42 | Not working | (TC 07/04/2024 02:42:24 PM, Rhuland, Courtney) Light was not turning on Power cord on back of device was loose Reseated cord - light powered on Focusing sleeve was missing - replaced with spare sleeve Checked other lights for similar issue
1063000 | 208173 | PM | CMPLT | CMPLT | 2024-07-04 16:18 | 2024-07-04 16:19 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/04/2024 04:18:37 PM, Cortez, John) Completed cleaning cycle
1070760 | 208174 | CM | CMPLT | CMPLT | 2024-07-05 08:21 | 2024-07-05 08:26 | Phototherapy light install | (TC 07/05/2024 08:26:03 AM, Rhuland, Courtney) Installed phototherapy light Moved O2 tank down Verified accuracy of scale
1061207 | 208175 | CM | CMPLT | CMPLT | 2024-07-05 09:01 | 2024-07-04 09:03 | Monitor Handle and Tilting Issue | Call re: not able to adjust Monitor position. Went to check. Waited for access. Found Monitor was sliding back when adjusting it at an angle. Removed cover and adjusted tension. Tested ok. Also found another monitor with loose handle. Removed back cover and tightened it. Tested ok.
1067338 | 208176 | CM | CMPLT | CMPLT | 2024-07-05 09:32 | 2024-07-05 09:32 | O2 cell/sensor failed during pre-use check | Tested unit for 5 days. Functional. passed pre-use check. No error.
1058169 | 208177 | CM | ACTIV | ISSUE | 2024-07-05 10:06 |  | Created by Failed PM WO # 206816 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic |  | (TC 07/05/2024 10:04:42 AM, Reid, Emmerson ) Not working, need to send out to repiar. Had an issue while upgrading and the firmware is failing.
1058169 | 208178 | CM | ACTIV | ISSUE | 2024-07-05 10:07 |  | CADD Firmware erorr - Sending for repair
65154 | 208179 | CM | ACTIV | ISSUE | 2024-07-05 10:08 |  | Created by Failed PM WO # 202080 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic |  | (TC 05/13/2024 01:58:06 PM, Reid, Emmerson ) Unit needs to be sent out for repair.
65154 | 208180 | CM | ACTIV | ISSUE | 2024-07-05 10:09 |  | CADD occlusion sensor seal broken - sending out for repair
65129 | 208181 | CM | ACTIV | ISSUE | 2024-07-05 10:10 |  | Created by Failed PM WO # 207493 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic |  | (TC 06/28/2024 08:19:33 AM, Reid, Emmerson ) Was sent down with a tag saying it had broken into pieces. The pole mount had come off likely from a drop but nothing had been damaged so I was able to put it together. I noticed the occlusion sensor seal is severely buldged and begginning to tear so I am going to send it out for repair.
65129 | 208182 | CM | ACTIV | ISSUE | 2024-07-05 10:10 |  | CADD occlusion sensor seal broken - sending out for repair
67184 | 208183 | PM | CMPLT | CMPLT | 2024-07-05 11:47 | 2024-07-05 11:47 | PM-MAJOR - Annual PM For: Universal SpO2 Finger Simulator | (TC 07/05/2024 11:47:17 AM, Chukri, Samir) Certificate of calibration by Clare instrument received . attached certificate.
1063746 | 208184 | PM | CMPLT | CMPLT | 2024-07-05 12:02 | 2024-07-05 12:03 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/05/2024 12:02:40 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1059828 | 208185 | PM | CMPLT | CMPLT | 2024-07-05 12:28 | 2024-07-05 12:28 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/05/2024 12:28:24 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
| 208186 | CM | CMPLT | CMPLT | 2024-07-05 13:31 | 2024-07-05 13:34 | White Balance Issue | Call re: monitor not able to white balance. Scheduled and went to check. Waited for clinical staff. Worked with them to set uo the scope and check. Found they were using Hologic scope with Olympus lightsource. White balance failed. Advised the staff to use Olympus scope. Set it up. Passed. Checked both scopes again. Hologic scope failed to white balance. Determined the issue was with the scope and not the video system. Tested ok.
1071301 | 208187 | CM | CMPLT | CMPLT | 2024-07-05 13:35 | 2024-07-08 13:57 | Bed connector issue | Jul 05 - Call re: bed not turning on. Patient was in the room. Scheduled it after patient discharge. Went to check. FOund bed not turning ON. Rebooted the bed. Did not work. FOund foot board sitiing loose on the bed connector. Removed it. Found bed connector pins were damaged (possibly due to attempting to attach foot board on the bed with force). Confirmed bed under contract. Took pictures and Contacted stryker and requested to send the rep in for repair. Jul 08 - Followed up. Coordinated with staff and Scheduled repair with the rep. Worked with the rep to repair the bed. Issue resolved. Checked functionality. Requested the rep to provide clinical training ont eh operation of the bed. Tested ok. Refer to report attached.
1067353 | 208188 | CM | CMPLT | CMPLT | 2024-07-05 13:48 | 2024-08-29 08:24 | Roll stand arm broken | Jul 05 - call re: scope guide stand broken. Scheduled and went to check. Waited for access to the room. Found roll stand arm screw loose. Attempted to fix it but found that screw holding the scope guide in place was tripped and threads were damaged. The entore arm will need to be replaced. Took pictures and contacted vendor for purchase info. Jul 11 - Followed up. info received. Created SC. Requested approval from manager Met with manager to explain the issue. Approval received. Placed order via MMC. Awaiting delivery. Aug 8 - Followed up multiple times to get an ETA from Olympus. No acknowledgement or response from them. Escalated issue to sales rep. Still no response. Email 4th time. Awaiting response. Aug 28 - Received part. Coordinated ad scheduled replacement after hours. Replaced arm. Secured cables. Checked functionality. Tested ok.
501071 | 208189 | CM | CMPLT | CMPLT | 2024-07-05 14:49 | 2024-07-10 10:25 | MSH CEOU RM # 12 - 'Black Lower and Upper Screen' | (TC 07/04/2024 02:49:23 PM, Chau, Ken) Created case for FSE from Philips to look at logs (TC 07/05/2024 02:49:48 PM, Chau, Ken) Issue caused by C5-1 probe, been ordered as reported by Philips logs Performed EMF tested in rooms that is creating artifacts and all seems to be at a very low level and not the cause of artifacts (TC 07/10/2024 10:24:43 AM, Chau, Ken) Error 005 again occured during boot up without probes -replaced the probe with SN#F0BK5V -system was working after reboot but the error 005 seems to be intermittent and Philips has ordered NAIM board to be replaced
| 208190 | CM | CMPLT | CMPLT | 2024-07-05 15:47 | 2024-07-09 10:56 | misc biomed - July 2024 | (TC 07/05/2024 03:46:39 PM, Cortez, John) Focal point meeting (TC 07/04/2024 03:47:08 PM, Cortez, John) Bladderscan battery charger sent down from ICU, not charging Confirmed issue with power supply (TC 07/09/2024 10:50:56 AM, Cortez, John) on-call check (TC 07/09/2024 10:55:49 AM, Cortez, John) Masimo client down 10N. Restarted remotely (TC 07/10/2024 09:41:35 AM, Cortez, John) on-call daily checks (TC 07/16/2024 04:02:07 PM, Cortez, John) PFT called about minor leak on 4th floor office Suggested them to unplug PC underneath and move to side (TC 07/17/2024 04:04:33 PM, Cortez, John) Checked PFT again in the morning to see if ceiling leak got worse. It's ok. 16th floor CPET ok as well (TC 07/17/2024 04:06:28 PM, Cortez, John) Installations in Pod C room 25-28 patched Pod C central (TC 07/18/2024 03:56:05 PM, Cortez, John) systems meeting (TC 07/26/2024 09:29:21 AM, Cortez, John) Helped swap in new SCDs in ICU (TC 07/26/2024 09:29:58 AM, Cortez, John) ICU rooms 25-28 go-live meeting (TC 07/25/2024 09:29:21 AM, Cortez, John) Helped swap in new SCDs in ICU (TC 07/25/2024 09:29:58 AM, Cortez, John) ICU rooms 25-28 go-live meeting (TC 07/24/2024 09:35:41 AM, Cortez, John) pm meeting (TC 07/29/2024 04:04:32 PM, Cortez, John) CE meeting (TC 07/29/2024 04:06:11 PM, Cortez, John) Move batteries/supplies for Stryker bed PMs to 12th floor (TC 07/29/2024 04:07:43 PM, Cortez, John) Tightened panels on Stryker boom (ICU room 28) as per Amanda's request. Noticed that panels don't use actual screws, but fasteners that can get loose easily Checked all rooms in Pod C and tightened loose fasterners. Spoke briefly with Marlene about it
66230 | 208191 | CM | CMPLT | CMPLT | 2024-07-05 16:15 | 2024-07-05 16:16 | Moved to 10 south | (TC 07/05/2024 04:15:30 PM, Cortez, John) Swapped with defective unit, asset# 1051446
1050704 | 208192 | CM | CMPLT | CMPLT | 2024-07-05 16:50 | 2024-07-05 16:50 | Moved to 10S | (TC 07/05/2024 04:50:09 PM, Cortez, John) checked functionality replaced #1051446
1070759 | 208193 | CM | CMPLT | CMPLT | 2024-07-08 08:12 | 2024-07-08 08:13 | Phototherapy light install | (TC 07/08/2024 08:12:51 AM, Rhuland, Courtney) Installed phototherapy light Verified scale accuracy
1056665 | 208194 | CM | CMPLT | CMPLT | 2024-07-08 09:09 | 2024-07-08 09:09 | images were not showing on the monitors | (TC 06/13/2024 09:09:41 AM, Huang, Jessica) received the call in OR 1 and the images from microscope were not showing on the monitors. checked the connections and restarted the microscope. working fine
1056665 | 208195 | CM | CMPLT | CMPLT | 2024-07-08 09:12 | 2024-07-08 09:12 | no video input | (TC 07/03/2024 11:31:30 AM, Huang, Jessica) the device was used in OR 1 and was showing "no video input" on the screen. turned on the camera button and restarted the device. working fine.
| 208196 | RDS | CMPLT | CMPLT | 2024-07-08 09:57 | 2024-07-08 11:27 | Weekly checks in CCU | (TC 07/08/2024 11:26:50 AM, Scero, Tomlyn) Weekly checks complete. 30 J test passed. 1 defib wasn't plugged in and the other the pads weren't connected.
1067933 | 208197 | PM | CMPLT | CMPLT | 2024-07-08 10:35 | 2024-07-08 10:39 | PM-MAJOR - Annual PM: For SimMan | (TC 07/08/2024 10:38:33 AM, Canjimavadivel, Kamallesh ) PM was performed and completed by vendor.
1061211 | 208198 | CM | CMPLT | CMPLT | 2024-07-08 11:33 | 2024-07-08 11:33 | ipad was not working | (TC 07/03/2024 11:33:15 AM, Huang, Jessica) removed the ipad from the wall and power cycle it. working fine
1070241 | 208199 | CM | CMPLT | CMPLT | 2024-07-08 11:36 | 2024-07-08 11:36 | replaced batteries
1070234 | 208200 | CM | CMPLT | CMPLT | 2024-07-08 11:37 | 2024-07-08 11:37 | replaced batteries
| 208201 | CM | CMPLT | CMPLT | 2024-07-08 11:37 | 2024-07-08 11:37 | biomed conference
1070238 | 208202 | CM | CMPLT | CMPLT | 2024-07-08 11:42 | 2024-07-08 11:42 | replaced batteries
| 208203 | CM | CMPLT | CMPLT | 2024-07-08 11:46 | 2024-07-08 11:46 | working on the masimo monitor error message
1061209 | 208204 | CM | CMPLT | CMPLT | 2024-07-08 11:51 | 2024-07-08 11:51 | monitor 2 image was upside down | (TC 06/03/2024 11:50:57 AM, Huang, Jessica) was informed the image on the monitor 2 was upside down and checked it in the room. working fine
1070235 | 208205 | CM | CMPLT | CMPLT | 2024-07-08 11:54 | 2024-07-08 11:54 | checked the setting
1070239 | 208206 | CM | CMPLT | CMPLT | 2024-07-08 11:55 | 2024-07-08 11:55 | checked the setting
1070241 | 208207 | CM | CMPLT | CMPLT | 2024-07-08 11:57 | 2024-07-08 11:57 | checked the setting | (TC 06/03/2024 11:56:48 AM, Huang, Jessica) checked the setting
1070238 | 208208 | CM | CMPLT | CMPLT | 2024-07-08 11:57 | 2024-07-08 11:57 | checked the setting
1061219 | 208209 | CM | CMPLT | CMPLT | 2024-07-08 11:59 | 2024-07-08 11:59 | no issue found | (TC 06/03/2024 11:59:17 AM, Huang, Jessica) received the call from OR desk and saying OR 18 needed help. went to OR 18 and it was empty.
| 208210 | CM | CMPLT | CMPLT | 2024-07-08 12:04 | 2024-07-08 12:04 | installed the monkey bar on the bed
1061213 | 208211 | CM | CMPLT | CMPLT | 2024-07-08 12:07 | 2024-07-08 12:07 | wall monitor installation
1061212 | 208212 | CM | CMPLT | CMPLT | 2024-07-08 12:09 | 2024-07-08 12:09 | could not display the image on the wall monitor | (TC 06/12/2024 12:09:19 PM, Huang, Jessica) the wall monitor could not display the image. working with Oscar and Derrick to fix the issue.
1048492 | 208213 | CM | CMPLT | CMPLT | 2024-07-08 12:13 | 2024-07-08 12:13 | lost image on the monitor | (TC 06/13/2024 12:12:57 PM, Huang, Jessica) the tower was being used in OR 10. the screen was showing blue and came back. contacted the sales rep from Stryker. Checked all the conections and informed OR team.
32520 | 208214 | CM | ACTIV | ISSUE | 2024-07-08 12:32 |  | MSH DR2 DETECTOR COVER CLAMSHELL |  | MSH DR2 DETECTOR COVER CLAMSHELL
| 208215 | CM | CMPLT | CMPLT | 2024-07-08 13:43 | 2024-07-23 11:13 | BIS host cable replacement | (TC 07/08/2024 01:43:02 PM, Scero, Tomlyn) BIS x host cable order from Philips (TC 07/09/2024 02:02:58 PM, Scero, Tomlyn) PO 7500209036 received (TC 07/23/2024 11:13:05 AM, Scero, Tomlyn) cabled received
1071295 | 208216 | CM | CMPLT | CMPLT | 2024-07-08 13:45 | 2024-07-08 13:47 | Not working | call re: not working. Went to check. Found bed in the hall way and foot board screen not working. Moved bed close to outlet and turned it on. Found food board was not inserted propely. Checked connectors. Found foot board connector was not sitiing smoothly on the bed. Adjsted the tilting angle and installed foot board. Checked functionality. Tested ok. Moved bed into a room. Tested ok.
| 208217 | CM | CMPLT | CMPLT | 2024-07-08 13:55 | 2024-07-23 11:21 | SPO2 sensors | (TC 07/08/2024 02:19:22 PM, Scero, Tomlyn) ordered sensors and cables (TC 07/09/2024 02:02:18 PM, Scero, Tomlyn) PO 7500209034 received (TC 07/23/2024 11:21:11 AM, Scero, Tomlyn) Sensors received and taken to OR for the RTs to try.
1048500 | 208218 | CM | CMPLT | CMPLT | 2024-07-08 14:54 | 2024-07-11 14:30 | msh us epiq monitor site 75025255 | (TC 07/09/2024 03:29:03 PM, Whittington, Scott) JULY 8 msh us epiq monitor site 75025255 REF# 0123928752 DIAGNOSE FAULTY MONITOR' ORDER REPLACEMENT JUNE 9 RECEIVE & INSTALL REPLACEMENT MONITOR (TC 07/11/2024 02:30:12 PM, Whittington, Scott) msh us epiq monitor site 75025255 REF# 0123928752 FAULTUY MONITOR RETURNED | msh us epiq monitor site 75025255 REF# 0123928752
1054725 | 208219 | CM | CMPLT | CMPLT | 2024-07-08 14:57 | 2024-07-16 13:11 | MSH IRUS EPIQ site 81256332 C5-1 sn F045KN | (TC 07/09/2024 03:30:44 PM, Whittington, Scott) JULY 8 MSH IRUS EPIQ site 81256332 C5-1 sn F045KN site 81256332 REF# 0123928766 REPLACEMENT ORDERED | MSH IRUS EPIQ site 81256332 C5-1 sn F045KN site 81256332 REF# 0123928766 msh 5 irus epiq c5-1 probe acc MIE 1054931 WO 208289 JUL 16 CL
1061202 | 208220 | CM | ACTIV | ISSUE | 2024-07-08 15:15 |  | Suction not very strong | (TC 07/08/2024 03:15:01 PM, Reid, Emmerson ) Checked the tree and it was working. Changed the filter and will put a building service request to fix the suction
| 208221 | CM | CMPLT | CMPLT | 2024-07-08 15:23 | 2024-07-08 15:23 | broken matteress on 11 south | (TC 06/20/2024 09:06:05 AM, Huang, Jessica) received the call from 11 south regarding the broken mattress. went to the room to check it and contacted the vendor. moved it to 12th floor. assisting Mark and sent it back to the unit.
1063361 | 208222 | CM | CMPLT | CMPLT | 2024-07-09 08:11 | 2024-07-08 08:11 | Wont turn on | (TC 07/09/2024 08:10:59 AM, Ferreira, Qwynn) Device would not turn on Battery was not installed. Received battery from unit and charged. Device is working and all functions - OK (TC 07/08/2024 08:10:59 AM, Ferreira, Qwynn) Device would not turn on Battery was not installed. Received battery from unit and charged. Device is working and all functions - OK
63801 | 208223 | CM | CMPLT | NFF | 2024-07-09 10:55 | 2024-07-09 10:55 | "not printing" | (TC 07/09/2024 10:54:52 AM, Cortez, John) Printed normally. NFF
1070843 | 208224 | CM | CMPLT | CMPLT | 2024-07-09 11:22 | 2024-07-09 11:22 | Phototherapy installation | (TC 07/09/2024 11:22:27 AM, Rhuland, Courtney) Installed phototherapy light Verified accuracy of scale
1070738 | 208225 | CM | CMPLT | CMPLT | 2024-07-09 11:25 | 2024-07-09 11:26 | Phototherapy light install | (TC 07/09/2024 11:25:36 AM, Rhuland, Courtney) Installed phototherapy light Verified accuracy of scale
| 208226 | PA | CMPLT | CMPLT | 2024-07-09 11:45 | 2024-07-24 08:36 | Servo batteries | (TC 07/09/2024 11:48:35 AM, Canjimavadivel, Kamallesh ) Placed order for batteries (TC 07/24/2024 08:35:43 AM, Canjimavadivel, Kamallesh ) All batteries received, gave to Nigel to install
1071645 | 208227 | CM | CMPLT | CMPLT | 2024-07-09 11:58 | 2024-07-09 15:21 | Incoming inspection | (TC 07/09/2024 03:20:31 PM, DeGuzman, Leila) Incoming inspection completed. Ground Resistance: 73 mOhm Chassis Leakage Current: 141.9 uAmp
67955 | 208228 | CM | CMPLT | CMPLT | 2024-07-09 12:33 | 2024-07-09 12:34 | Fan failure | (TC 07/09/2024 12:34:09 PM, Rhuland, Courtney) Reseated fan bushing and fan Error cleared Replaced 2 side door hinges Functional check - OK
1054692 | 208229 | CM | ACTIV | ISSUE | 2024-07-09 14:39 |  | "door doesn't stay latched" | (TC 07/09/2024 02:39:35 PM, Cortez, John) Outer housing coming apart near peristaltic pump. Reached out to Cardinal Health for RMA (TC 07/22/2024 04:20:50 PM, Cortez, John) shipped to CH
| 208230 | CM | CMPLT | CMPLT | 2024-07-09 14:45 | 2024-07-09 15:57 | Medical Device Evaluation with Human Factors | Attended Medical Device Evaluation study with Brian Leung at UHN Healthcare Human Factors
1071650 | 208231 | ACC | CMPLT | CMPLT | 2024-07-10 08:07 | 2024-07-10 10:58 | Incoming inspection | (TC 07/10/2024 10:58:02 AM, Canjimavadivel, Kamallesh ) Leftover from COVID-19 stock. Repurposed for ED. Replaced batteries and attached bracket to holder. Installed in yellow zone.
1071651 | 208232 | ACC | CMPLT | CMPLT | 2024-07-10 08:09 | 2024-07-10 11:03 | Incoming inspection | (TC 07/10/2024 11:03:16 AM, Canjimavadivel, Kamallesh ) Leftover from COVID-19 stock. Repurposed for ED. Replaced batteries and attached bracket to holder. Installed in yellow zone.
1071652 | 208233 | ACC | CMPLT | CMPLT | 2024-07-10 08:16 | 2024-07-10 11:04 | Incoming Inspection | (TC 07/10/2024 11:03:52 AM, Canjimavadivel, Kamallesh ) Leftover from COVID-19 stock. Repurposed for ED. Replaced batteries and attached bracket to holder. Installed in yellow zone.
65933 | 208234 | CM | CMPLT | CMPLT | 2024-07-10 08:17 | 2024-09-30 11:26 | Channel error | (TC 07/10/2024 08:34:12 AM, Rhuland, Courtney) Channel error 5-13-1531 present Not in service manual Called BD and found to be a software issue and logic board needs to be replaced Placed order for board (TC 09/30/2024 11:25:53 AM, Alvarado, Karla) Replaced logic board. Performed at ASM: Flash Syringe, Calibration, and Check- in test. All test passed.
1071653 | 208235 | ACC | CMPLT | CMPLT | 2024-07-10 08:18 | 2024-07-10 11:04 | Incoming Inspection | (TC 07/10/2024 11:04:19 AM, Canjimavadivel, Kamallesh ) Leftover from COVID-19 stock. Repurposed for ED. Replaced batteries and attached bracket to holder. Installed in yellow zone.
| 208236 | CM | CMPLT | CMPLT | 2024-07-10 08:23 | 2024-07-10 08:33 | Medfusion parts for biomed stock | (TC 07/10/2024 08:33:01 AM, Rhuland, Courtney) Placed order for medfuison parts
501584 | 208237 | CM | CMPLT | CMPLT | 2024-07-10 08:26 | 2024-07-10 08:30 | MSH OPG SPP D2 IU22 NOT SENDING TO PACS | (TC 07/10/2024 08:26:25 AM, Whittington, Scott) MSH OPG SPP D2 IU22 NOT SENDING TO PACS PACS ERROR HAD OCCURED , WHICH PROHIBITED SENDING TO JDMIPACS VASSI ATTEMPTED TO SEND TO DOWNTIME, BUT UNABLE FOUND THAT DOWNTIME WAS REACHABLE FROM THE SYSTEM, BUT THE WINDOW THAT IS USED FOR RESEND DOD NOT CONTAIN DOWNTIME, ONLY JDMIPACS DELETED DOWNTIME SETTINT, AND RE-ENTERED DOWNTIME NOW APPEARS AS A DESTINATION ON RESEND WINDOW STUDY SENT TO DOWNTIME, VASSI TO CONTACT 7773 TO HAVE THE STUDY CROSSED OVER TO CORAL (TC 07/10/2024 08:26:25 AM, Whittington, Scott) MSH OPG SPP D2 IU22 NOT SENDING TO PACS PACS ERROR HAD OCCURED , WHICH PROHIBITED SENDING TO JDMIPACS VASSI ATTEMPTED TO SEND TO DOWNTIME, BUT UNABLE FOUND THAT DOWNTIME WAS REACHABLE FROM THE SYSTEM, BUT THE WINDOW THAT IS USED FOR RESEND DOD NOT CONTAIN DOWNTIME, ONLY JDMIPACS DELETED DOWNTIME SETTINT, AND RE-ENTERED DOWNTIME NOW APPEARS AS A DESTINATION ON RESEND WINDOW STUDY SENT TO DOWNTIME, VASSI TO CONTACT 7773 TO HAVE THE STUDY CROSSED OVER TO CORAL | MSH OPG SPP D2 IU22 NOT SENDING TO PACS
1059881 | 208238 | CM | CMPLT | CMPLT | 2024-07-10 08:33 | 2024-09-25 10:08 | MSH-5-RPCU ON XG MIE 1063193 HAS CRYSTAL FAULTS PLT1204BX SN# 99B2115416 | (TC 07/10/2024 09:58:15 AM, Whittington, Scott) MSH-5-RPCU ON XG MIE 1063193 HAS CRYSTAL FAULTS PLT1204BX SN# 99B2115416 | MSH-5-RPCU ON XG MIE 1063193 HAS CRYSTAL FAULTS PLT1204BX SN# 99B2115416
1051739 | 208239 | CM | CMPLT | CMPLT | 2024-07-10 08:36 | 2024-09-25 10:09 | MSH-5-RPCU PLT-704AT PLACED ON MIE 1063193 PLT-704AT SN# 99B1953302 | (TC 07/10/2024 09:59:45 AM, Whittington, Scott) MSH-5-RPCU PLT-704AT PLACED ON MIE 1063193 PLT-704AT SN# 99B1953302 | MSH-5-RPCU PLT-704AT PLACED ON MIE 1063193 PLT-704AT SN# 99B1953302
69165 | 208240 | CM | CMPLT | CMPLT | 2024-07-10 10:11 | 2024-07-10 10:12 | Channel Error | (TC 07/10/2024 10:11:42 AM, DeGuzman, Leila) Downloaded error code: 242.4030 Cleaned device. Ran device, no error showing. Performed check-in, all tests passed. return to service
| 208241 | CON | CMPLT | CMPLT | 2024-07-10 11:27 | 2024-09-25 10:11 | MSH CEOU INTERDERENCE ISSUES WITH ARTIFACT IN ROOMS 7,8,9,10,11 | (TC 06/19/2024 11:31:30 AM, Whittington, Scott) MSH CEOU INTERERENCE ISSUES WITH ARTIFACT IN ROOMS 7,8,9,10,11 (TC 07/31/2024 08:12:00 AM, Whittington, Scott) MSH CEOU INTERFERENCE ISSUESREMOVE TEMPORARY POWER CABLE FROM CEILINGS (TC 06/19/2024 11:31:00 PM, Whittington, Scott) MSH CEOU INTERDERENCE ISSUES WITH ARTIFACT IN ROOMS 7,8,9,10,11 (TC 07/31/2024 08:12:00 AM, Whittington, Scott) MSH CEOU INTERFERENCE ISSUESREMOVE TEMPORARY POWER CABLE FROM CEILINGS | MSH CEOU INTERDERENCE ISSUES WITH ARTIFACT IN ROOMS 7,8,9,10,11
32510 | 208242 | CM | CMPLT | CMPLT | 2024-07-10 11:29 | 2024-09-25 10:13 | MSH DR1 EXPOSURE ISSUES | (TC 07/02/2024 11:30:30 AM, Whittington, Scott) MSH DR1 EXPOSURE ISSUES (TC 07/02/2024 11:30:30 AM, Whittington, Scott) MSH DR1 EXPOSURE ISSUES (TC 07/02/2024 11:31:00 PM, Whittington, Scott) MSH DR1 EXPOSURE ISSUES | MSH DR1 EXPOSURE ISSUES
1063708 | 208243 | CM | CMPLT | CMPLT | 2024-07-10 13:34 | 2024-07-11 08:12 | Scan engine and SW upgrade | (TC 07/10/2024 11:34:00 AM, Fan, Tak) Scan engine and SW upgraded by vendor. Reconfigured system. Submitted helpdesk MAC whitelisting request. (TC 07/11/2024 08:12:02 AM, Fan, Tak) Wired network setup. Tested, OK. Wifi doesn't work though. Followed up with Mirza.
67761 | 208244 | CM | CMPLT | CMPLT | 2024-07-10 13:42 | 2024-07-10 13:42 | Upgraded to V4.3 | (TC 07/10/2024 01:42:09 PM, Reid, Emmerson ) Upgraded to V4.3 and pushed APS Epidural Oct 2023
67761 | 208245 | PM | CMPLT | PMFAIL | 2024-07-10 13:46 | 2024-07-10 15:15 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/10/2024 03:14:32 PM, DeGuzman, Leila) Performed PM. Failed accuracy test. See procedure for details.
65117 | 208246 | CM | CMPLT | CMPLT | 2024-07-10 14:01 | 2024-07-10 14:01 | Upgraded to V4.3 | (TC 07/10/2024 02:01:09 PM, Reid, Emmerson ) Upgraded to V4.3 and pushed APS IV PCA Mar 2022
1058189 | 208247 | CM | CMPLT | CMPLT | 2024-07-10 14:05 | 2024-07-10 14:06 | Upgraded to V4.3 | (TC 07/10/2024 02:05:38 PM, Reid, Emmerson ) Upgraded to V4.3 and pushed APS IV PCA Mar 2022
67761 | 208248 | CM | ACTIV | ISSUE | 2024-07-10 15:17 |  | Created by Failed PM WO # 208245 : PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 08/20/2024 10:36:57 AM, Reid, Emmerson ) Ran the 20ml infusion twice and it over-infused both times. Needs to be sent for repair. | (TC 07/10/2024 03:14:32 PM, DeGuzman, Leila) Performed PM. Failed accuracy test. See procedure for details.
1066929 | 208249 | CM | CMPLT | CMPLT | 2024-07-10 18:15 | 2024-07-10 18:16 | No display | Call re: not working. Went to check. Waited for access. Found no display on screen. Checked connections. Found monitor input was switched. Rectified it. Checked functionality. Advisede staff on how to select appropriate imput. Tested ok.
32520 | 208250 | CM | CMPLT | CMPLT | 2024-07-10 20:34 | 2024-07-24 14:10 | MSH DR2 OGP MOVEMENT DET LONG, TABLESIDE CONTROLLER | (TC 07/06/2024 08:34:53 PM, Whittington, Scott) MSH DR2 OGP MOVEMENT DET LONG, TABLESIDE CONTROLLER DET LONG DRIVE SYSTEM REPAIRED TABLESIDE CONTROLLER REPAIRED (TC 07/06/2024 10:55:00 PM, Whittington, Scott) MSH DR2 OGP MOVEMENT DET LONG, TABLESIDE CONTROLLER DET LONG DRIVE SYSTEM REPAIRED TABLESIDE CONTROLLER REPAIRED | MSH DR2 OGP MOVEMENT DET LONG, TABLESIDE CONTROLLER
1058191 | 208251 | CM | CMPLT | CMPLT | 2024-07-11 11:07 | 2024-07-11 11:37 | Keep prompting occlusion | (TC 07/11/2024 11:37:08 AM, DeGuzman, Leila) Run device, no prompting occlussion.
1060119 | 208252 | CM | CMPLT | NFF | 2024-07-11 11:38 | 2024-08-20 14:14 | Bed Repair | (TC 08/20/2024 02:13:50 PM, Ferreira, Qwynn) Mark cleared error and said no problem was found
1071398 | 208253 | CM | CMPLT | CMPLT | 2024-07-11 13:54 | 2024-08-02 13:40 | System error 25/will not hold date&time | (TC 07/11/2024 02:11:37 PM, Rhuland, Courtney) Reseated fan to clear system error 25 When powered on, message appeared to change coin battery Date and time would not hold Replaced coin battery and time and date were still incorrect Reached out to getinge for tech support and waiting on reply (TC 08/02/2024 01:40:12 PM, Rhuland, Courtney) Replaced SBC board Replaced coin battery Set date and time Calibrated RH Functional check - OK
1061432 | 208254 | CM | CMPLT | CMPLT | 2024-07-11 14:23 | 2024-07-11 14:23 | Bent power cord | (TC 07/11/2024 02:23:26 PM, Rhuland, Courtney) Replaced power cord that was bent Functional check - OK
1071406 | 208255 | CM | CMPLT | CMPLT | 2024-07-11 14:26 | 2024-07-25 10:41 | Broken light shade | (TC 07/11/2024 02:30:53 PM, Rhuland, Courtney) Light shade had been snapped off Removed broken inner piece and received quote Placed order for new light shade (TC 07/25/2024 10:40:46 AM, Rhuland, Courtney) Replaced light shade Tested light Found incubator and reattached light
| 208256 | RDS | CMPLT | CMPLT | 2024-07-11 14:51 | 2024-07-11 14:56 | Weekly check for ICU
| 208257 | MEP | CMPLT | CMPLT | 2024-07-12 09:32 | 2024-09-25 10:45 | MSH WIRELESS SURVEY JULY 2024 | (TC 07/12/2024 09:32:54 AM, Whittington, Scott) MSH WIRELESS SURVEY JULY 2024 (TC 07/11/2024 06:22:00 PM, Whittington, Scott) MSH WIRELESS SURVEY | MSH WIRELESS SURVEY JULY 2024 From: Nicholson, Andrew <andrew.nicholson@sinaihealth.ca> Sent: July 9, 2024 10:52 AM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Abdou, Makar - Sinai Health System <Makar.Abdou@sinaihealth.ca> Cc: Wong, Erin - Sinai Health System <Erin.Wong@sinaihealth.ca>; Justin Fiege <Justin.Fiege@dxc.com>; 'mirza.riaz@dxc.com' <mirza.riaz@dxc.com>; Samanian, Kamran - Sinai Health System <Kamran.Samanian@sinaihealth.ca>; Tucci, Anthony - Sinai Health System <Anthony.Tucci@sinaihealth.ca>; Noseworthy, Joanne - Sinai Health System <Joanne.Noseworthy@sinaihealth.ca>; Pejcic, Sonja <Sonja.Pejcic@uhn.ca>; Rice, Murray <Murray.Rice@uhn.ca> Subject: Re: [External] Mount Sinai Wifi upgrade Hi Scott, It would be great if you could get something out to Makar this week. Regards, Andrew ________________________________________ From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: July 9, 2024 10:36 AM To: Abdou, Makar; Nicholson, Andrew Cc: Wong, Erin; Justin Fiege; 'mirza.riaz@dxc.com'; Samanian, Kamran; Tucci, Anthony; Noseworthy, Joanne; Pejcic, Sonja - UHN - University Health Network; Rice, Murray - UHN - University Health Network Subject: Re: [External] Mount Sinai Wifi upgrade CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Andrew We can pull details from our database and forward to you. What is your timeline? Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Nicholson, Andrew <andrew.nicholson@sinaihealth.ca> Sent: Tuesday, July 9, 2024 9:27:15 AM To: Abdou, Makar - Sinai Health System <Makar.Abdou@sinaihealth.ca>; Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Wong, Erin - Sinai Health System <Erin.Wong@sinaihealth.ca>; Justin Fiege <Justin.Fiege@dxc.com>; 'mirza.riaz@dxc.com' <mirza.riaz@dxc.com>; Samanian, Kamran - Sinai Health System <Kamran.Samanian@sinaihealth.ca>; Tucci, Anthony - Sinai Health System <Anthony.Tucci@sinaihealth.ca>; Noseworthy, Joanne - Sinai Health System <Joanne.Noseworthy@sinaihealth.ca> Subject: Re: [External] Mount Sinai Wifi upgrade Thanks Makar. Hi Scott, we are looking for some technical WIFI specs to ensure devices are supported in our WIFI end state. I believe these devices are under your teams management- are you able to provide the info to Makar? If not, who can provide this info? Regards, Andrew ________________________________________ From: Abdou, Makar Sent: July 9, 2024 8:37 AM To: Whittington, Scott - UHN - University Health Network Cc: Nicholson, Andrew; Wong, Erin; Justin Fiege; 'mirza.riaz@dxc.com'; Samanian, Kamran; Tucci, Anthony; Noseworthy, Joanne Subject: RE: [External] Mount Sinai Wifi upgrade Hi Scott, Thanks for sharing the insights, So by any chance do you have list of assets or inventory for the medical imaging devices your group is using? Thanks, Makar Abdou Senior Project Manager, Project Management Office Information Services T 416-586-4800 x2694 A patient care, teaching and research centre affiliated with University of Toronto. From: Whittington, Scott [mailto:Scott.Whittington@uhn.ca] Sent: July 08, 2024 3:37 PM To: Abdou, Makar <Makar.Abdou@sinaihealth.ca> Cc: Nicholson, Andrew <andrew.nicholson@sinaihealth.ca>; Wong, Erin <Erin.Wong@sinaihealth.ca>; Justin Fiege <Justin.Fiege@dxc.com>; 'mirza.riaz@dxc.com' <mirza.riaz@dxc.com>; Samanian, Kamran <Kamran.Samanian@sinaihealth.ca>; Tucci, Anthony <Anthony.Tucci@sinaihealth.ca>; Noseworthy, Joanne <Joanne.Noseworthy@sinaihealth.ca> Subject: RE: [External] Mount Sinai Wifi upgrade CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Makor Any medical imaging devices my group works with at MSH and are on the MSH network are MSH owned devices, such as ultrasounds, and mobile xrays. Is that the info you are looking for? Thanks, Scott 416-402-2462 From: Abdou, Makar <Makar.Abdou@sinaihealth.ca> Sent: July 8, 2024 3:08 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Nicholson, Andrew - Sinai Health System <andrew.nicholson@sinaihealth.ca>; Wong, Erin - Sinai Health System <Erin.Wong@sinaihealth.ca>; Justin Fiege <Justin.Fiege@dxc.com>; 'mirza.riaz@dxc.com' <mirza.riaz@dxc.com>; Samanian, Kamran - Sinai Health System <Kamran.Samanian@sinaihealth.ca>; Tucci, Anthony - Sinai Health System <Anthony.Tucci@sinaihealth.ca> Subject: [External] Mount Sinai Wifi upgrade Hi Scott, Hope you are doing great, we are going through some Wifi upgrades at Mount Sinai, So liked to connect with you to get a list of UHN Wifi connected devices. We are using the below fields to collect the connectivity details, this is very important to ensure a seamless upgrade. Line No. Campus Dept Device Type Make Model SSID WIFI Standard WPA WPA2 AAA AES PSK TKIP MSHCHAPV2 TLS Please let me know if you will be providing the information, or kindly guide us to the contacts for who can help from UHN, If you have any further questions please let me know, Thanks, Makar Abdou Senior Project Manager, Project Management Office Information Services T 416-586-4800 x2694 A patient care, teaching and research centre affiliated with University of Toronto
1055953 | 208258 | CM | CMPLT | CMPLT | 2024-07-12 13:14 | 2024-07-12 13:56 | Temperature probe not working | (TC 07/12/2024 01:55:36 PM, DeGuzman, Leila) Checked device, issue found: Module Error (80) when taking temperature. Checked other functions PASS for: NIBP SpO2 Device need to send out for serivce.
1062554 | 208259 | CM | ACTIV | ISSUE | 2024-07-12 14:10 |  | Screeching noise when head/foot section lifted | (TC 07/12/2024 02:09:45 PM, Cortez, John) Cleared section width alarm, but there's a screeching noise when head section is lifted, Scale also needs to be calibrated. Asked ward clerk to swap bed. Calibrated scale and reached out to Stryker (TC 07/13/2024 11:22:30 AM, Cortez, John) WO 11392256
| 208260 | OC | CMPLT | CMPLT | 2024-07-12 14:38 | 2024-07-12 14:39 | On-call - July 5-11 2024
1046542 | 208261 | CM | CMPLT | CMPLT | 2024-07-12 14:47 | 2024-07-12 14:47 | Not working | (TC 07/12/2024 02:47:22 PM, DeGuzman, Leila) Issue found: low battery. Replaced batteries and cleaned device Checked functions, all okay. Returned to service.
1059778 | 208262 | FC | CMPLT | CMPLT | 2024-07-12 14:48 | 2024-07-12 14:48 | QC Check | Cleaned and completed QC check list and reset parameters to patient default.
1063310 | 208263 | CM | CMPLT | CMPLT | 2024-07-12 15:02 | 2024-07-12 15:42 | Beam Issue | Call re: Beam issue. Went to check with ER. Spoke with suergeon and was informed that the laser had double beam issue. Moved laser into a different room and tested the laser with different couplers and micromanipulator. Found beam spread was a bit wide but could not confirm the double beam issue. Checked functionality with surgeon and all tests worked fine. Tested ok.
1051703 | 208264 | CM | CMPLT | CMPLT | 2024-07-12 15:04 | 2024-07-12 15:04 | Not working, retire from service | (TC 07/12/2024 03:03:52 PM, DeGuzman, Leila) Checked device, replace new battery. still not working. Retired from service.
1058117 | 208265 | CM | CMPLT | CMPLT | 2024-07-12 15:11 | 2024-07-19 08:56 | Urgent Medical Device Correction | (TC 07/19/2024 08:56:24 AM, Reid, Emmerson ) Completed the remediation with with Baxter Tech. See attached Service Report.
1054468 | 208266 | CM | CMPLT | CMPLT | 2024-07-12 15:27 | 2024-07-15 14:46 | Created by Failed PM WO # 202259 : PM-MAJOR - Annual PM For: Infusion Pump, High Flow, Warmer | (TC 07/15/2024 02:45:34 PM, Canjimavadivel, Kamallesh ) Opened up device and found switch for door clamp was stuck in depressed position. Disassembled and cleaned out all components of air detector clamp assembly. Reassambled and carried out remainder of PM. All tests pass, returned device to unit. | (TC 07/12/2024 03:25:39 PM, Canjimavadivel, Kamallesh ) Started PM and found damaged component. Contacted tech support and was able to get required part number and service manual. Received quote for part.
32510 | 208267 | CM | CMPLT | CMPLT | 2024-07-14 10:51 | 2024-07-14 10:53 | MSH DR1 DETECTOR DIFFICULT TO CAL | (TC 07/14/2024 10:51:01 AM, Whittington, Scott) MSH DR1 DETECTOR DIFFICULT TO CAL TROUBLE SHOOT, DIAGNOSE, CAL | MSH DR1 DETECTOR DIFFICULT TO CAL
| 208268 | PA | CANCL | CANCL | 2024-07-15 08:00 | 2024-07-18 07:43 | Top key for H-1200 | Parts order not needed | Stock parts
| 208269 | CM | CMPLT | CMPLT | 2024-07-15 08:23 | 2024-07-13 20:37 | On Call Call Back | Call from NICU re: Staff assist (nurse call) alarming. Confirmed with SC that . Called them back and informed them that biomed do not look after nurse call and advised to connect with IT for assistance. PRovided ext.
| 208270 | CM | CMPLT | CMPLT | 2024-07-15 11:04 | 2024-07-22 09:54 | Order for Servo-U pm kits, biomed stock | (TC 07/15/2024 11:12:41 AM, Cortez, John) ordered x3 pm kits (TC 07/22/2024 09:54:23 AM, Cortez, John) arrived
1061211 | 208271 | CM | CMPLT | CMPLT | 2024-07-15 13:27 | 2024-07-15 13:38 | Display Interference | CAll re: display flickering. Went to check. Waited for access.. Spokw with surgeon and was informed that all displays flickers when activating monopolar. Checked the wiring and arrangements. Found all wires from scope and esu clamped together to OR table for stability. Advised to isolate all cables. Stuck arounf for a while to ensure issue is resolved. Tested ok.
| 208272 | CM | CMPLT | CMPLT | 2024-07-15 13:41 | 2024-07-15 13:41 | Bed alerts Issue | Foudn Bed alerts on 3 bed in 3 different rooms in 12N. Went to unit. Disconnected and rebooted all bed. Checked functionality. Tested ok.
1071301 | 208273 | CM | CMPLT | CMPLT | 2024-07-15 13:43 | 2024-07-15 13:43 | Bed alarm Issue | Call re: bed causing nurse call alarm to go off. Went to check. Checked all connections. Reset the bed. checked functionality. Tested ok.
1062316 | 208274 | CM | CMPLT | CMPLT | 2024-07-15 13:47 | 2024-07-26 16:01 | Bed alert - Battery Issue | Jul 15 - Received bed alert from Lifenet SW. Went to check. Found error "switch on battery" Disconnected and rebooted the bed. Error persisted. Called stryker rep for assistance. Followed instructions and attempted to troubleshoot over the phone. Issue still not resolved. Took pictures and sent it to stryker and requested to schedule repair under contract. Jul 17 - Followed up. Coordinated with tech to come in for repair. The tech came it but found the battery he brought does not fit the bed. Was informed that he will acwuire correct batteries and schedule the visit again. Jul 29 - The rep came in. WOrked with him to resolve the issue. Battery replaced. refer to report attahced.
501105 | 208275 | CM | CMPLT | CMPLT | 2024-07-15 15:16 | 2024-09-25 10:45 | MSH CEOU iu22- Artifact | (TC 07/15/2024 03:16:17 PM, Chau, Ken) Artifact on scan head when in high pen and gain increased Removing ethernet cable, noise/artifact reduced signifcantly moved machine to other location and no issue move a different IU22 into the same room using same outlet, issue on both machine Ran on UPS and no issue Issue seems to reside within the outlets (TC 07/16/2024 07:23:03 AM, Chau, Ken) touch screen cal in rm 10 completed room creating artifact and comes with ethernet -touch base with IT and they created WO to have ethernet moved from 1 bay to another and to work with facilities -philips heard of this issue from another hospital, IT did something but they do not recall exactly what, they will contact other hospital and see waht the resolution was (TC 07/23/2024 02:08:28 PM, Chau, Ken) Working with IT to move cables around to see if it will resolve issue but cannot remove issue -IT to move network cable with Engineering team to hopefully resolve issue caused by network jack -demonstrated with IT team that everyitme plugging and unplugging the network jack the artifact will come and go. Doesnt matter which machine is plugged into that will cause issue -if multiple machine plugged into that room with network connected will cause artifacts -reached out to Philips as they recall this issue happend at a different hospital and their network team did something to resolve issue, they will try and reach out to other hospital to see what the resolution was Configured all the rooms to have pixel spacing enabled (TC 07/23/2024 03:47:06 PM, Dolinar, John) MSH CEOU INTERFERENCE ARTIFACT ISSUE ASSISTING KEN (TC 07/24/2024 07:46:16 AM, Chau, Ken) Installed netgear switch and no more artifacts -asked CEOU manager to purchase additional 4 switches until network team resolve issue (TC 07/25/2024 03:08:39 PM, Chau, Ken) Tested with network team and moved cable from network stacks... issue still there, they will replace to shielded cables (TC 07/30/2024 02:18:10 PM, Chau, Ken) Installed rooms with some switches and did not notice artifacts (TC 07/31/2024 09:29:00 AM, Chau, Ken) MSH CEOU iu22- Artifact followup
1062544 | 208276 | CM | CMPLT | CMPLT | 2024-07-15 16:04 | 2024-07-15 16:04 | "scale uncalibrated" | (TC 07/15/2024 04:03:55 PM, Cortez, John) Calibrated scale, Accuracy check ok Confirmed functiionality with actual patient
| 208277 | CM | CMPLT | CMPLT | 2024-07-15 17:13 | 2024-07-15 17:13 | Site coverage
| 208278 | CM | CMPLT | CMPLT | 2024-07-16 07:20 | 2024-07-18 10:06 | Sterile water for NICU
1067912 | 208279 | CM | CMPLT | CMPLT | 2024-07-16 08:49 | 2024-07-16 08:49 | Missing power cord | (TC 07/16/2024 08:49:22 AM, Canjimavadivel, Kamallesh ) Installed power cord and secured to roll stand with zip-tie
69613 | 208280 | CM | CMPLT | CMPLT | 2024-07-16 08:53 | 2024-07-16 08:56 | Channel Error | (TC 07/16/2024 08:56:00 AM, DeGuzman, Leila) No error found. Performed check-in tests: all tests passed.
1065707 | 208281 | CM | CMPLT | CMPLT | 2024-07-16 09:26 | 2024-07-16 09:26 | Channel error | (TC 07/16/2024 09:26:06 AM, DeGuzman, Leila) Replaced power cord. Performed check-in tests: all passed Ground resistance: 139 mOhm Leakage current: 36.1 uAmp Returned to service.
1067604 | 208282 | CM | ACTIV | ISSUE | 2024-07-16 09:29 |  | MSH ER JDMI US i800 BACKUP USB MISSING | (TC 07/16/2024 09:28:39 AM, Whittington, Scott) MSH ER JDMI US i800 BACKUP USB MISSING | MSH ER JDMI US i800 BACKUP USB MISSING
500847 | 208283 | CM | CMPLT | CMPLT | 2024-07-16 09:39 | 2024-07-24 14:09 | msh 5 us aplio xg front wheel locks not releasing | (TC 07/16/2024 09:38:56 AM, Whittington, Scott) msh 5 us aplio xg front wheel locks not releasing (TC 07/16/2024 09:52:27 AM, Whittington, Scott) msh 5 us aplio xg front wheel locks not releasing (TC 07/16/2024 09:52:27 AM, Dolinar, John) msh 5 us aplio xg front wheel locks not releasing | msh 5 us aplio xg front wheel locks not releasing
1055162 | 208284 | CM | CMPLT | CMPLT | 2024-07-16 09:54 | 2024-07-16 10:10 | Channel Error | (TC 07/16/2024 10:09:41 AM, DeGuzman, Leila) User problem: cannot enter a volume. Run in basic infusion, can enter Rate and VTBI. Performed Check-in tests: all passed.
| 208285 | CM | CMPLT | CMPLT | 2024-07-16 10:07 | 2024-07-16 10:12 | Blood pressure check | (TC 07/16/2024 10:11:42 AM, Rhuland, Courtney) Physician reported issue with BP cuff Checked all cuffs and devices No issues found Told users to replace cuff if issue persisted
1066001 | 208286 | CM | CMPLT | CMPLT | 2024-07-16 11:00 | 2024-07-16 11:00 | Channel Error | (TC 07/16/2024 11:00:18 AM, DeGuzman, Leila) Channel error: 200.5040 Flashed pump for software update. Performed Check-in tests: all passed.
69220 | 208287 | CM | CMPLT | CMPLT | 2024-07-16 12:06 | 2024-07-16 12:07 | Channel Error | (TC 07/16/2024 12:06:32 PM, DeGuzman, Leila) Downloaded error log: 200.5040 No error showed up. Performed check-in tests: all passed
1061999 | 208288 | CM | CMPLT | NFF | 2024-07-16 12:44 | 2024-08-20 14:14 | Bed has no function - Error 3009-6-5-1503 | (TC 07/16/2024 12:43:52 PM, Ferreira, Qwynn) Inspected the bed - no obvious signs of wear or broken/loose cables or connections. Reset bed - Did not work Reached out to marak for updates. (TC 08/20/2024 02:14:28 PM, Ferreira, Qwynn) Bed error codes were reset but no problem was found
1054931 | 208289 | ACC | CMPLT | CMPLT | 2024-07-16 13:06 | 2024-07-16 13:13 | MSH IRUS PROBE C5-1 SN# F0GD55 ACCEPTANCE | (TC 07/16/2024 01:06:38 PM, Whittington, Scott) MSH IRUS PROBE C5-1 SN# F0GD55 ACCEPTANCE JULY 16 2024 SN# F0GD55 REPLACES SN# F045KN | MSH IRUS PROBE C5-1 SN# F0GD55 ACCEPTANCE SYSTEM MSH IRUS EPIQ site 81256332 PROBE FAULT C5-1 sn F045KN SYSTEM MIE1054725 WO 208219
66723 | 208290 | CM | CMPLT | CMPLT | 2024-07-16 15:42 | 2024-07-16 15:43 | Recorder not printing | (TC 07/16/2024 03:42:44 PM, Rendon Lopez , Elizabeth) Room 325 not printing. Cleaned recorder thermal head, calibrated, printed strips, working properly.
67959 | 208291 | CM | CMPLT | NFF | 2024-07-17 09:03 | 2024-07-17 09:03 | Skin temp probe not correlating with baby temperature | (TC 07/17/2024 09:02:54 AM, Rhuland, Courtney) Connected baby temp probe Set baby temp to 36C Incubator analyzer, baby probe and baby temp all read 36C No fault found Functional check - OK
1056420 | 208292 | CM | CMPLT | CMPLT | 2024-07-17 09:41 | 2024-07-16 09:52 | "1)plug doesn't go in wall 2)spo2 not working" | (TC 07/16/2024 09:52:20 AM, Cortez, John) Power cord tested ok Spo2 waveform cuts out intermittently, replaced Cleaned filter, added radical battery All parameters tested ok
1051031 | 208293 | CM | CMPLT | CMPLT | 2024-07-17 10:07 | 2024-07-17 15:03 | NICU Echopac PC slowing down | (TC 07/16/2024 04:07:00 PM, Fan, Tak) Typically happens from many studies being processed on echopac but software seems relatively light when checked. Seemed normal as well. Will follow up again with Faith. (TC 07/17/2024 03:03:13 PM, Fan, Tak) System in use. Issue was the watermark is not what it seems and needs fine tuning depending on the computer space.
1070739 | 208294 | CM | CMPLT | CMPLT | 2024-07-17 10:13 | 2024-07-17 10:14 | Phototherapy light installation | (TC 07/17/2024 10:13:47 AM, Rhuland, Courtney) Installed phototherapy light on unit Verified accuracy of scale
67885 | 208295 | CM | CMPLT | CMPLT | 2024-07-17 10:15 | 2024-07-17 10:15 | Won't run epidural bolus, keeps occluding | (TC 07/17/2024 10:15:20 AM, Rendon Lopez , Elizabeth) Upgraded software and library. Running on user mode. (TC 07/17/2024 08:48:01 AM, Rendon Lopez , Elizabeth) Tested, NFF. Performed checks, all passed.
69384 | 208296 | PM | CMPLT | CMPLT | 2024-07-17 10:37 | 2024-07-17 11:31 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/17/2024 11:31:09 AM, Rendon Lopez , Elizabeth) Replaced left IUI. PM completed, all tests passed
69509 | 208297 | CM | CMPLT | CMPLT | 2024-07-17 11:08 | 2024-07-17 11:08 | Channel error | (TC 07/17/2024 11:08:00 AM, DeGuzman, Leila) Downloaded error log: found error: Bottle side pressure sensor 240.4150 Analog sensor tests reads 4.99 volts Replaced bottle-side sensor. Perfomed check-in tests: all passed. Returned to service.
69347 | 208298 | CM | CMPLT | CMPLT | 2024-07-17 11:32 | 2024-07-17 11:33 | Channel error | (TC 07/17/2024 11:33:07 AM, Rendon Lopez , Elizabeth) Downloaded error log. Bottle pressure sensor fatal error. Replaced bottle PS, calibrated, check-in tests, all passed.
66705 | 208299 | CM | CMPLT | CMPLT | 2024-07-17 13:39 | 2024-07-17 13:39 | Broken foot door | (TC 07/17/2024 01:39:20 PM, Rhuland, Courtney) Door came down to biomed Repaired broken hinge Went up to NICU to locate omnibed Found in RESUS Reattached door
1057483 | 208300 | CM | CMPLT | CMPLT | 2024-07-17 13:43 | 2024-08-15 14:37 | Not working | (TC 07/17/2024 01:43:08 PM, Rendon Lopez , Elizabeth) Measured values lower than normal. Keeping to send out for repair (TC 08/15/2024 02:37:27 PM, Rendon Lopez , Elizabeth) PO 7500209930
1057494 | 208301 | CM | CMPLT | CMPLT | 2024-07-17 13:45 | 2024-07-17 13:52 | Not working | (TC 07/17/2024 01:51:59 PM, Rendon Lopez , Elizabeth) Tested, NFF
1047061 | 208302 | CM | CMPLT | CMPLT | 2024-07-17 13:57 | 2024-08-15 14:30 | Not working | (TC 07/17/2024 02:00:33 PM, Rendon Lopez , Elizabeth) Readings lower than normal. Keep to send out for repair (TC 08/15/2024 02:29:44 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
66511 | 208303 | CM | CMPLT | CMPLT | 2024-07-17 14:07 | 2024-07-18 16:07 | Monitor in OR4 not working with wireless station | (TC 07/16/2024 02:44:59 PM, Rendon Lopez , Elizabeth) Removed monitor from OR4, using spare (TC 07/16/2024 02:44:59 PM, Rendon Lopez , Elizabeth) Testing monitor with smae wireless station/transducers, working, no issues. Monitor working with wired transducers as well. Tested for some time, no issues. Cleaned recorder. (TC 07/18/2024 04:06:48 PM, Rendon Lopez , Elizabeth) Returned monitor to OR 4
69317 | 208304 | CM | CMPLT | PMFAIL | 2024-07-17 14:46 | 2024-07-17 15:20 | Channel Error | (TC 07/17/2024 03:20:05 PM, DeGuzman, Leila) Downloaded error log, found bottle-side sensor high failure. Check in analog sensor tests, confirmed its high voltage when not pressed. Replaced with new sensor, new sensor value is 0.9792V. Tried to performed Patient occlusion, failed: 13.285 psi Tried to calibrate, failed because of high voltage value of the sensor.
69316 | 208305 | CM | CMPLT | PMFAIL | 2024-07-17 14:46 | 2024-07-17 15:31 | Channel Error | (TC 07/17/2024 03:30:36 PM, DeGuzman, Leila) Downloaded error log, found bottle-side sensor high failure. Check in analog sensor tests, confirmed its high voltage when not pressed. Replaced with new sensor, new sensor value is 0.978V. Tried to performed Patient occlusion, failed: 13.463 psi Tried to calibrate, failed because of high voltage value of the sensor.
1063784 | 208306 | CM | CANCL | CANCL | 2024-07-17 14:58 | 2024-08-08 14:07 | Does not turn ON
| 208307 | CM | CMPLT | CMPLT | 2024-07-17 15:18 | 2024-07-16 15:20 | Suction Tree Issue | Call re: not working. Went to check and brought it down to biomed. FOund suction regulator sitting loose on the carousal. Fixed it. Found suction hose connector cracked and loose. Fixed it and tightned with teflon tape. Found suction tubing leaking due to loose connection points at multiple locations. Fixed all. Checked functionality. Tested ok.
1064512 | 208308 | CM | CMPLT | CMPLT | 2024-07-17 15:28 | 2024-07-17 15:42 | APC Issue | Call re: not working. Went to check. Was infomed to come back when the room is available. Went back. Spoke with staff and was informed that the unit does not fire when activating the foot switch. Checked the fiber and functionality with the unit. Found intermittent issue. Took out new fiber and checked functionality with two more different unit. Worked fine. Advised to use new fiber during the procedure. Tested ok.
| 208309 | PA | CMPLT | CMPLT | 2024-07-18 07:48 | 2024-07-26 08:32 | Li-ion battery for stock | (TC 07/18/2024 07:56:35 AM, Canjimavadivel, Kamallesh ) Placed order for batteries (TC 07/26/2024 08:32:03 AM, Canjimavadivel, Kamallesh ) Received batteries, placed on shelf
1057108 | 208310 | CM | CMPLT | CMPLT | 2024-07-18 08:01 | 2024-07-18 08:01 | not working | unit came with Philips cuff and hose. Replaced them. tested. functional.
1055953 | 208311 | CM | CMPLT | CMPLT | 2024-07-18 08:03 | 2024-07-18 08:03 | temperature module error | replaced temperature probe. tested. functional.
62696 | 208312 | CM | CMPLT | NFF | 2024-07-18 08:28 | 2024-07-18 08:35 | Not turning on even when plugged in | (TC 07/18/2024 08:35:19 AM, Canjimavadivel, Kamallesh ) Turned on ECG cart both plugged in and without, booted up with no issues and worked with simulator. Was not able to replicate reported issue.
66842 | 208313 | PM | CMPLT | CMPLT | 2024-07-18 11:25 | 2024-07-18 11:54 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/18/2024 11:53:31 AM, DeGuzman, Leila) PM completed. See procedure for details.
| 208314 | RQ | ACTIV | READY | 2024-07-18 11:47 |  | THIS IS A TEST, DISREGARD THIS WEB REQUEST - Ibrahim S
66810 | 208315 | PM | CMPLT | CMPLT | 2024-07-18 13:05 | 2024-07-18 13:30 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/18/2024 01:30:10 PM, DeGuzman, Leila) PM completed. See procedure for details.
1067437 | 208316 | FC | CMPLT | CMPLT | 2024-07-18 13:19 | 2024-07-16 13:20 | QC checklist | (TC 07/16/2024 01:19:33 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1055916 | 208317 | CM | CMPLT | CMPLT | 2024-07-18 13:20 | 2024-07-16 13:20 | Not measuring BP | (TC 07/16/2024 01:20:22 PM, Baghirzade, Nara) NIBP tubing was incorrectly seated and had a loose connection; placed it on the connector properly. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1061333 | 208318 | FC | CMPLT | CMPLT | 2024-07-18 13:22 | 2024-07-16 13:22 | QC checklist | (TC 07/16/2024 01:22:06 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
1055591 | 208319 | CM | CMPLT | CMPLT | 2024-07-18 13:24 | 2024-07-17 13:24 | Patient admit button was no longer available on the display for SafetyNet | (TC 07/17/2024 01:24:00 PM, Baghirzade, Nara) Rebooted the Masimo multiple times; button still not available. Looked through both the service and user manual but no information on this issue. SafetyNet config was reloaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Patient admit button is now available on the device. Took the Masimo to 11 South and measured own SpO2, confirmed that the numerics were being displayed on the 11S safetynet monitor.
65996 | 208320 | CM | CMPLT | CMPLT | 2024-07-18 13:27 | 2024-07-17 13:27 | System error; unable to bypass the system error | (TC 07/17/2024 01:27:21 PM, Baghirzade, Nara) Checked the service manual and the recommendation was to replace the logic board. Replaced the logic board. Completed software upgrade. Transferred network configuration and dataset file. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 161 mohms Leakage current: 78.5 uA
1063361 | 208321 | CM | CMPLT | CMPLT | 2024-07-18 13:36 | 2024-09-12 15:43 | Loss of ultrasound data | (TC 07/18/2024 01:41:46 PM, Ferreira, Qwynn) Tested the device and error message was present Reset the scanner and cleared some memory to see if any improvement No damage on cables or pin connections Reached out to company for repair of teh probe/warranty (TC 07/24/2024 08:30:38 AM, Ferreira, Qwynn) Shipping the ultrasound and probe to CHS with no charge po (TC 09/12/2024 03:31:52 PM, Patel, Rutvik) Unti received. Unboxed and assembled it. Charged battery. Checked functionality and tested measurement accuracy with phantom. All measurement within range. (125 ml, 124ml, 124ml). Tested ok. Delivered back to 12S.
68895 | 208322 | CM | CMPLT | CMPLT | 2024-07-18 13:36 | 2024-07-17 13:36 | Software to be reloaded again | (TC 07/17/2024 01:36:07 PM, Baghirzade, Nara) Software upgrade completed. Transferred network configuration and dataset files.
1055581 | 208323 | CM | CMPLT | CMPLT | 2024-07-18 13:39 | 2024-07-18 13:39 | No issue stated on biomed tag | (TC 07/18/2024 01:38:50 PM, Baghirzade, Nara) No issue found with the device. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 208324 | CM | CMPLT | CMPLT | 2024-07-18 13:40 | 2024-07-17 13:40 | NICU meeting | (TC 07/17/2024 01:39:50 PM, Baghirzade, Nara) NICU meeting regarding post flooding steps for zone C.
66851 | 208325 | PM | CMPLT | CMPLT | 2024-07-18 13:41 | 2024-07-18 14:00 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/18/2024 02:00:10 PM, DeGuzman, Leila) PM completed. See procedures for details.
1054986 | 208326 | CM | CMPLT | CMPLT | 2024-07-18 13:41 | 2024-07-18 13:41 | Patient admit button not on display; no numerics shown on PSN central monitor | (TC 07/18/2024 01:41:19 PM, Baghirzade, Nara) No patient admit button on the display despite multiple reboots of the Root. Looked through the service and user manual however no information found for this issue. SafetyNet config was re-loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Patient admit button is available on the Root. Took the Masimo to 11 North and measured own SpO2, confirmed that the numerics were being displayed on the 11N safetynet monitor.
1055914 | 208327 | CM | CMPLT | CMPLT | 2024-07-18 13:44 | 2024-07-18 13:44 | Not measuring BP | (TC 07/18/2024 01:43:53 PM, Baghirzade, Nara) Replaced the NIBP cuff and hose; Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 208328 | CM | CMPLT | CMPLT | 2024-07-18 13:45 | 2024-08-06 14:16 | Found wound VAC on biomed repair shelf | (TC 07/18/2024 01:45:13 PM, Baghirzade, Nara) Contacted Julie and she requested that I call KCI to organize the pick up as finding the owner would be time consuming. Called KCI and found out that the VAC was requested by 11 South however the VAC needs to go back to KCI. Organized the pick up from biomed. Confirmation # 13337460, SN: VFVS00806. Followed up with Brittany and Mykayla to let them know that rental vacs need to go back from their unit and not from biomed. Mykayla acknowledged and will let the staff know. (TC 07/29/2024 08:43:25 AM, Baghirzade, Nara) VAC was not picked up by KCI; sent a follow up email. (TC 08/06/2024 02:15:39 PM, Baghirzade, Nara) Wound VAC was picked up by 3M staff on Friday, August 2nd. No pickup/drop off receipt as they have a new system. Signed electronically on his phone for pickup acknowledgement.
67156 | 208329 | CM | CMPLT | CMPLT | 2024-07-18 14:04 | 2024-07-18 14:04 | Temp repurpose for diabetes clinic | (TC 07/18/2024 02:04:22 PM, Rhuland, Courtney) Temporarily (~1 week) brought to 60 Murray diabetes clinic while waiting for new BP devices Performed functional check & NIBP test All tests passed Brought to clinic
| 208330 | CM | CMPLT | CMPLT | 2024-07-18 14:16 | 2024-07-18 14:16 | X3 battery charge back | (TC 07/18/2024 02:16:09 PM, Canjimavadivel, Kamallesh ) 3 batteries were installed during go-live. charge back to unit
67208 | 208331 | PM | CMPLT | CMPLT | 2024-07-18 14:17 | 2024-07-18 15:01 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 07/18/2024 03:01:25 PM, DeGuzman, Leila) PM completed. See procedure for details.
| 208332 | CM | CMPLT | CMPLT | 2024-07-18 14:38 | 2024-07-23 07:13 | Batteries for incubators
500668 | 208333 | RQ | CMPLT | CMPLT | 2024-07-18 15:07 | 2024-07-18 15:23 | testtest | (TC 07/18/2024 03:23:06 PM, Fan, Tak) Test acknowledgement. | testing testing 123 test
1069392 | 208334 | CM | CMPLT | CMPLT | 2024-07-18 15:44 | 2024-07-18 15:44 | "unable to clear air from system" | (TC 07/18/2024 03:44:10 PM, Cortez, John) Ran a test treatment, but unable to repeat issue. Self-tests passed
| 208335 | CM | CMPLT | CMPLT | 2024-07-19 08:28 | 2024-07-19 08:30 | NICU rounds | (TC 07/19/2024 08:29:34 AM, Rhuland, Courtney) NICU rounds completed
| 208336 | CM | CMPLT | CMPLT | 2024-07-19 08:30 | 2024-07-19 08:34 | RESUS rounds | (TC 07/19/2024 08:34:13 AM, Rhuland, Courtney) RESSUS rounds completed
1050617 | 208337 | CM | CMPLT | CMPLT | 2024-07-19 08:47 | 2024-07-19 08:47 | Channel Error | (TC 07/19/2024 08:47:19 AM, DeGuzman, Leila) Performed Check-in tests: all passed Cleaned device.
1057549 | 208338 | CM | CMPLT | CMPLT | 2024-07-19 09:35 | 2024-07-18 09:35 | Not working, switching off | (TC 07/18/2024 09:35:22 AM, Rendon Lopez , Elizabeth) Monitor from OR 4. Tested in AGM in OR1, no fault found, returnde to OR 4, working properly, all tests passed.
66700 | 208339 | CM | CMPLT | CMPLT | 2024-07-19 09:37 | 2024-07-18 09:46 | Recorder ink issues | (TC 07/18/2024 09:45:51 AM, Rendon Lopez , Elizabeth) Cleaned thermal head print, calibrated recorder, tested-OK
1059828 | 208340 | PM | CMPLT | CMPLT | 2024-07-19 09:41 | 2024-07-19 11:01 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/19/2024 11:00:43 AM, DeGuzman, Leila) Drained, disinfected, rinsed and refilled with sterile water
66718 | 208341 | CM | CMPLT | CMPLT | 2024-07-19 09:47 | 2024-07-18 09:47 | Recorder ink issues | (TC 07/18/2024 09:47:05 AM, Rendon Lopez , Elizabeth) Cleaned thermal head print, calibrated recorder, tested-OK
1054111 | 208342 | CM | CMPLT | CMPLT | 2024-07-19 09:48 | 2024-07-18 09:48 | Recorder ink issues | (TC 07/18/2024 09:48:06 AM, Rendon Lopez , Elizabeth) Cleaned thermal head print, calibrated recorder, tested-OK
66523 | 208343 | CM | CMPLT | CMPLT | 2024-07-19 10:56 | 2024-07-19 10:56 | Transducer input broken | (TC 07/19/2024 10:55:40 AM, Rendon Lopez , Elizabeth) Fetal transducer connector piece broken off inside inlet, removed piece. Tested all inputs-OK. Cleaned recorder, tested all functions-OK
1063746 | 208344 | PM | CMPLT | CMPLT | 2024-07-19 11:03 | 2024-07-19 13:57 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/19/2024 01:56:53 PM, DeGuzman, Leila) Drained, disinfected, rinsed and refilled with sterile water
1056125 | 208345 | PM | CMPLT | CMPLT | 2024-07-19 11:08 | 2024-07-19 11:08 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 07/19/2024 11:08:04 AM, Canjimavadivel, Kamallesh ) Cleaned hoses and nozzles Cleaning cycle complete
1065076 | 208346 | PM | CMPLT | CMPLT | 2024-07-19 13:24 | 2024-07-19 13:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/19/2024 01:23:44 PM, Baghirzade, Nara) Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
69501 | 208347 | CM | CMPLT | CMPLT | 2024-07-19 13:26 | 2024-07-19 13:26 | LVP came down from Oncology with complaints that it did not detect air bubbles | (TC 07/19/2024 01:25:35 PM, Baghirzade, Nara) Downloaded error logs; no errors noted. Completed AIL test on ASM - no issues with the AIL test. Completed analog test successfully. Rate was reading high during rate accuracy test. Rate was calibrated. Check-in tests passed and completed. IUI connectors and device cleaned. | LVP came down from Oncology with complaints that it did not detect air bubbles. LVP came down with another LVP (69380) and PCU (68619)
69380 | 208348 | CM | CMPLT | CMPLT | 2024-07-19 13:28 | 2024-07-19 13:28 | LVP came down from Oncology with complaints that it did not detect air bubbles | (TC 07/19/2024 01:27:42 PM, Baghirzade, Nara) LVP came down from Oncology with complaints that it did not detect air bubbles. LVP came down with another LVP (69380) and PCU (68619) Completed AIL test - no issues found. Check-in tests passed and completed. IUI connectors and device cleaned.
66708 | 208349 | CM | CMPLT | CMPLT | 2024-07-19 13:45 | 2024-07-19 13:45 | recorder ink issues | (TC 07/19/2024 01:44:58 PM, Rendon Lopez , Elizabeth) Cleaned recorder thermal head. Calibrated, tested-OK
66728 | 208350 | CM | CMPLT | CMPLT | 2024-07-19 13:46 | 2024-07-19 13:46 | Reorder ink issues | (TC 07/19/2024 01:46:06 PM, Rendon Lopez , Elizabeth) Cleaned recorder thermal head. Calibrated, tested-OK
63437 | 208351 | CM | CMPLT | CMPLT | 2024-07-19 13:55 | 2024-07-19 13:55 | User reported that it had error message (not specified) | (TC 07/19/2024 01:54:48 PM, Baghirzade, Nara) Ran the pump on various rates to completion 3 times; no issues or errors observed. The only thing to note was that the battery was completely depleted and needed to be charge, which is likely connected to the error that the user observed. Pump functional, returned to 11 N
1050614 | 208352 | CM | CMPLT | CMPLT | 2024-07-19 14:00 | 2024-07-19 14:03 | Beeps red when turn on | (TC 07/19/2024 02:02:39 PM, DeGuzman, Leila) No issue found. Performed check-in tests: all passed. Return to service.
1056386 | 208353 | RQ | PENDG | REQST | 2024-07-19 14:08 |  | this is a web request test. disregard this as it is not a real request. |  | test test 12345
69104 | 208354 | PM | CMPLT | CMPLT | 2024-07-19 14:42 | 2024-07-19 14:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/19/2024 02:43:00 PM, DeGuzman, Leila) PM passed and completed.
1062936 | 208355 | CM | CMPLT | CMPLT | 2024-07-19 14:44 | 2024-07-19 14:45 | Does nto turn on | Call re: not working. Went to check and brought it down to biomed. Foudn battery completely drained and power cord connection loose. Charged it for a while. Clenaed unit. Chceked battery diagnostic and functionality. No issue found. Tested ok. Returned to 9 south.
1061214 | 208356 | CM | CMPLT | CMPLT | 2024-07-19 14:46 | 2024-07-19 14:48 | Monitor tilting | Call re; Monitor tilting issue. Went to check. Waited for access. Founf monitor tilting back to vertical positon when adjusting it to an angle. Removed cover and adjusted tension. FOund monito handle loose. Fixed it. Checked functionality. Tested ok.
32520 | 208357 | CM | CMPLT | CMPLT | 2024-07-19 17:33 | 2024-07-22 13:33 | MSH DR2 System not auto positioning. | (TC 07/19/2024 05:37:36 PM, Dolinar, John) Detector x axis movement faulty. (TC 07/19/2024 05:37:55 PM, Heron, Chris) Detector x axis movement faulty. (TC 07/17/2024 10:06:10 AM, Heron, Chris) MSH DR2 DETEC TOR LONG. X INTERMITTENT (TC 07/22/2024 10:07:12 AM, Heron, Chris) MSH DR2 DETECTOR LONG. X INTERMITTENT (TC 07/18/2024 10:07:12 AM, Heron, Chris) MSH DR2 DETECTOR LONG. X INTERMITTENT (TC 07/22/2024 01:49:39 PM, Whittington, Scott) MSH DR2 Detector x axis movement faulty. DIAGNOSE DRIVE AS INTERMITTENT LOCATE AND TEST DRIVE FROM LOCAL STOCK INSTALL , TEST OK | MSH DR2 Detector x axis movement faulty.
1061333 | 208358 | FC | CMPLT | CMPLT | 2024-07-22 07:24 | 2024-07-19 14:24 | QC checklist | (TC 07/19/2024 02:24:00 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
| 208359 | RDS | CMPLT | CMPLT | 2024-07-22 07:34 | 2024-07-19 14:37 | Weekly Rounds: 11 North | (TC 07/19/2024 02:37:00 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 208360 | RDS | CMPLT | CMPLT | 2024-07-22 07:40 | 2024-07-19 14:39 | Weekly Rounds: 11 South | (TC 07/19/2024 02:39:00 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 208361 | RDS | CMPLT | CMPLT | 2024-07-22 07:42 | 2024-07-19 14:41 | Weekly Rounds: NICU & RESUS | (TC 07/19/2024 02:41:00 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
| 208362 | CM | ACTIV | ISSUE | 2024-07-22 09:41 |  | OR stirrups repair
1050633 | 208363 | CM | CMPLT | CMPLT | 2024-07-22 10:09 | 2024-07-22 10:10 | MSH 15 L&D TE7C not powering up | (TC 07/17/2024 10:09:24 AM, Heron, Chris) MSH 15 L&D TE7C not powering up REPLACED AC CORD CAP | MSH 15 L&D TE7C not powering up
1051376 | 208364 | CM | CMPLT | CMPLT | 2024-07-22 10:10 | 2024-07-22 10:10 | Damaged power cable | (TC 07/22/2024 10:10:21 AM, DeGuzman, Leila) Replaced power cable. Device okay.
1067975 | 208365 | CM | CMPLT | CMPLT | 2024-07-22 10:15 | 2024-07-22 10:15 | MSH ER JDMI PRISM NOT BOOTING | (TC 07/22/2024 10:15:16 AM, Heron, Chris) MSH ER JDMI PRISM NOT BOOTING REPLACE POWER CORD (TC 07/18/2024 10:15:16 AM, Heron, Chris) MSH ER JDMI PRISM NOT BOOTING REPLACE POWER CORD | MSH ER JDMI PRISM NOT BOOTING
68599 | 208366 | CM | CMPLT | CMPLT | 2024-07-22 10:15 | 2024-07-22 10:15 | check-in | (TC 07/22/2024 10:15:19 AM, DeGuzman, Leila) Came down no tag. Downloaded error log: no issue found. Performed check-in tests, all passed. return to service
69430 | 208367 | CM | CMPLT | CMPLT | 2024-07-22 10:18 | 2024-07-22 10:18 | check-in | (TC 07/22/2024 10:18:01 AM, DeGuzman, Leila) Came down no tag. Downloaded error log: no issue found. Performed check-in tests, all passed. return to service
| 208368 | PA | CMPLT | CMPLT | 2024-07-22 11:13 | 2024-07-30 08:46 | Ordering latch assembly LVP | (TC 07/22/2024 11:14:02 AM, Baghirzade, Nara) SC 1003275854 (TC 07/30/2024 08:45:37 AM, Baghirzade, Nara) PO 7500209559
66044 | 208369 | CM | CMPLT | CMPLT | 2024-07-22 11:18 | 2024-11-13 13:53 | Created by Failed PM WO # 207516 : PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/11/2024 11:23:57 AM, Baghirzade, Nara) Replaced rear case, front case, and logic board. (TC 11/13/2024 08:59:04 AM, Baghirzade, Nara) After replacing the rear case, front case, and logic board, the PCU would not switch on (beeps but screen stays blank). Disassembled device to ensure that all connectors were properly seated and assembled PCU again; issue still persists. Disassembled the device and checked both fuses; fuses are ok. Re-seated both boards and all conectors; assembled device however issue still persists. Swapped out the power supply board with a spare from an old unit; device is switching on but has a "settings unrestorable" error. Will complete the software upgrade, dataset transfer, and network config transfer. (TC 11/13/2024 01:53:27 PM, Baghirzade, Nara) Assembled device after repair. Completed software upgrade, main boot block upgrade, processor upgrade, and keyboard upgrade. Network configuration and dataset transferred. PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 134 mohms Leakage current: 67.1 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated. | (TC 07/22/2024 11:17:23 AM, Rendon Lopez , Elizabeth) PM failed. Cracked front and rear case, waiting on part
68908 | 208370 | PM | CMPLT | CMPLT | 2024-07-22 11:21 | 2024-07-22 11:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/22/2024 11:40:03 AM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
68908 | 208371 | CM | CMPLT | CMPLT | 2024-07-22 11:22 | 2024-07-22 11:26 | Side broken, channel fell off without operator touching it | (TC 07/22/2024 11:26:23 AM, Rendon Lopez , Elizabeth) Checked channel latch and IUI, no issues found. tested with different pumps-OK
1059494 | 208372 | CM | CMPLT | CMPLT | 2024-07-22 11:28 | 2024-07-22 11:30 | Retire | Retired. Device not working.
69604 | 208373 | PM | CMPLT | CMPLT | 2024-07-22 11:55 | 2024-07-22 12:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/22/2024 12:01:55 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed.
1057863 | 208374 | CM | CMPLT | NFF | 2024-07-22 12:31 | 2024-07-22 12:38 | Overheating and shutting off | (TC 07/22/2024 12:38:12 PM, Rhuland, Courtney) Replaced circulating fluid Ran device for a few hours - did not overheat/shut down Verified temp with temperature check No fault found
| 208375 | CM | CMPLT | CMPLT | 2024-07-22 14:02 | 2024-07-22 14:02 | Off-site PFT laptop | (TC 07/22/2024 02:02:19 PM, Baghirzade, Nara) Steve dropped off the off-site PFT laptop as Dr. Lapinsky was having connectivity issues with Expair. Contacted Jacky Zhang from IT to set up an admin account for me for palo alto; after access was granted, I was able to log in using his instructions. First attempt to log in to Expair failed, was displaying a "cannot connect to the medisoft server", expair would only operate on local mode. Let the laptop sit for a few minutes and then tried again, this time it worked. Contacted Jacky to find out what might have gone wrong, he suspects firewall or that the server is not officially on an allow list however waiting a couple of minutes after the laptop boots up works. Let Steve know and he set up a meeting with Dr. Lapinksy. Met with Dr. and went over the log in steps; the only incorrect step he made was that he did not click on the palo alto icon before inputting his login. He logged in and immediately tried to login to Expair; same error was presented so I asked him to wait a minute and try again. After he did it again, it was working. He took the laptop home and I let him know to contact me if there were any issues with logging in. Did not hear back from him or Steve so will assume that Expair is functional.
68619 | 208376 | CM | CMPLT | CMPLT | 2024-07-22 14:29 | 2024-07-22 14:29 | Functional check | (TC 07/22/2024 02:29:27 PM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 173 mohms Leakage current: 79.2 uA No issues found with the device.
69317 | 208377 | CM | CMPLT | CMPLT | 2024-07-22 14:32 | 2024-07-22 14:32 | Replaced patient-side pressure sensor and full calibration | (TC 07/22/2024 02:31:37 PM, Baghirzade, Nara) Received LVP from Laila after she had replaced the bottle-side pressure sensor; patient occlusion was failing and reading high. Completed an analog sensor test and voltages were within range however when performing a pressure calibration (factory default), it was alerting that the downstream pressure sensors were defective. Replaced patient-side pressure sensor and ran the pressure calibration again; completed successfully. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
69316 | 208378 | CM | CMPLT | CMPLT | 2024-07-22 14:35 | 2024-07-22 14:35 | Replaced patient-side pressure sensor and full calibration | (TC 07/22/2024 02:35:17 PM, Baghirzade, Nara) Received LVP from Laila after she had replaced the bottle-side pressure sensor; patient occlusion was failing and reading high. Completed an analog sensor test and voltages were within range however when performing a pressure calibration (factory default), it was alerting that the downstream pressure sensors were defective. Replaced patient-side pressure sensor and ran the pressure calibration again; completed successfully. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
| 208379 | CM | CMPLT | CMPLT | 2024-07-22 14:36 | 2024-07-22 14:36 | Biomed PM meeting | (TC 07/22/2024 02:36:24 PM, Baghirzade, Nara) Biomed meeting
1059778 | 208380 | FC | CMPLT | CMPLT | 2024-07-22 14:38 | 2024-07-22 14:38 | QC checklist | (TC 07/22/2024 02:37:46 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South
1064423 | 208381 | CM | CMPLT | CMPLT | 2024-07-22 14:38 | 2024-07-22 14:38 | Light is orange instead of blue | (TC 07/22/2024 02:38:16 PM, Rhuland, Courtney) Replaced intensity rocker Functional check - OK
67930 | 208382 | CM | CMPLT | CMPLT | 2024-07-22 14:40 | 2024-07-22 14:41 | User reported system error | (TC 07/22/2024 02:41:04 PM, Baghirzade, Nara) No error observed during boot up/ found no issues with the incubator. Functional tests passed and completed. Scale is functional. Tilt is functional. Calibrated RH as it was out of range. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.8%RH Incubator is fully functional.
1046275 | 208383 | CM | ACTIV | ISSUE | 2024-07-22 14:51 |  | MSH CFRH 7flr APLIO 300 OPERATOR PANEL FREEZE & CALIPER BUTTONSINTERMITTENT, REQUIRES MEMBRANE SHEET; SCAN RANGE POT TCS PANEL LOWER RIGHT INTERMITTE | (TC 07/22/2024 02:50:57 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 OPERATOR PANEL FREEZE & CALIPER BUTTONSINTERMITTENT, REQUIRES MEMBRANE SHEET; SCAN RANGE POT TCS PANEL LOWER RIGHT INTERMITTENT (TC 07/22/2024 02:50:57 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 OPERATOR PANEL FREEZE & CALIPER BUTTONSINTERMITTENT, REQUIRES MEMBRANE SHEET; SCAN RANGE POT TCS PANEL LOWER RIGHT INTERMITTENT (TC 07/23/2024 02:50:57 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 OPERATOR PANEL FREEZE & CALIPER BUTTONSINTERMITTENT, REQUIRES MEMBRANE SHEET; SCAN RANGE POT TCS PANEL LOWER RIGHT INTERMITTENT | MSH CFRH 7flr APLIO 300 OPERATOR PANEL FREEZE & CALIPER BUTTONSINTERMITTENT, REQUIRES MEMBRANE SHEET; SCAN RANGE POT TCS PANEL LOWER RIGHT INTERMITTENT SITE ON46333US REF# PN# PRICE#
1069292 | 208384 | CM | CMPLT | CMPLT | 2024-07-22 15:03 | 2024-07-24 14:06 | MSH US ER G7 PROBE SN# FSD2375657 TROPHON PROBE # 51 ARTIFACT | (TC 07/24/2024 02:06:41 PM, Whittington, Scott) MSH US ER G7 PROBE SN# FSD2375657 TROPHON PROBE # 51 ARTIFACT | MSH US ER G7 PROBE SN# FSD2375657 TROPHON PROBE # 51 ARTIFACT
1069292 | 208385 | ACC | CMPLT | CMPLT | 2024-07-22 15:04 | 2024-07-24 14:07 | MSH US ER G7 PROBE PVT 781 VTE SN# ACCEPTANCE | (TC 07/24/2024 02:07:37 PM, Whittington, Scott) JULY 24 2024 SN# FSD2452683 REPLACED : MSH US ER G7 PROBE SN# FSD2375657 TROPHON PROBE # 51 ARTIFACT | MSH US ER G7 PROBE PVT 781 VTE SN# ACCEPTANCE JULY 24 2024 SN# FSD2452683 REPLACED : MSH US ER G7 PROBE SN# FSD2375657 TROPHON PROBE # 51 ARTIFACT
1056430 | 208386 | CM | CMPLT | CMPLT | 2024-07-22 15:38 | 2024-07-22 15:42 | Repair | Call re: not working. Went to check and brought it down to biomed. Found monitor displaced from the mounting stand. Fixed it. Found temp probe cover broken and stuck inside probe well. Removed it and cleaned it. Battery low. Charged unit for a while. Checked functionality. Tested ok. Delivered back to 12N.
1049784 | 208387 | PM | CMPLT | CMPLT | 2024-07-23 09:51 | 2024-07-23 09:51 | PM MAJOR-Monitor, Physiologic, Central Station | (TC 07/23/2024 09:50:47 AM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch Vacuumed PC PM complete
| 208388 | RDS | CMPLT | CMPLT | 2024-07-23 09:54 | 2024-07-23 14:54 | Weekly checks in CCU | (TC 07/23/2024 02:54:14 PM, Scero, Tomlyn) weekly checks complete. 30J tests passed.
1049786 | 208389 | PM | CMPLT | CMPLT | 2024-07-23 09:54 | 2024-07-23 09:59 | PM MAJOR-Monitors, Physiologic, Central Station, Main | (TC 07/23/2024 09:59:00 AM, Canjimavadivel, Kamallesh ) Rebooted PC to complete patch Vacuumed PC PM complete
1062550 | 208390 | CM | ACTIV | PV | 2024-07-23 13:17 |  | Communication lost to the main controller | (TC 07/23/2024 01:17:55 PM, Baghirzade, Nara) Patient was on the bed; unable to use any of the motion controls. Tried to shutdown/power cycle the bed however was unable to shut it down. Bed was missing asset tag; printed and placed asset tag. Contacted Stryker for support. | Motion controls not working; unable to shutdown/power cycle the bed
1053391 | 208391 | CM | CMPLT | CMPLT | 2024-07-23 13:28 | 2024-07-23 13:29 | Before/after patient use | (TC 07/23/2024 01:28:45 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 49 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
66712 | 208392 | CM | CMPLT | CMPLT | 2024-07-23 13:31 | 2024-07-23 13:31 | No issue stated on yellow biomed tag | (TC 07/23/2024 01:31:07 PM, Baghirzade, Nara) No fault found. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74%RH Incubator is fully functional.
1066460 | 208393 | FC | CMPLT | CMPLT | 2024-07-23 13:32 | 2024-07-23 13:32 | QC checklist | (TC 07/23/2024 01:32:19 PM, Baghirzade, Nara) QC checklist completed as per Sarabeth's request. Returned to 11 North
69433 | 208394 | CM | CMPLT | CMPLT | 2024-07-23 13:33 | 2024-07-23 13:34 | No issue stated on biomed tag | (TC 07/23/2024 01:33:31 PM, Baghirzade, Nara) Downloaded error log; did not find any significant errors. Completed analog sensors test; no issues found. During check-in, the rate accuracy failed (reading slightly high). Calibrated rate. Check-in tests completed. IUI connectors and device cleaned. (TC 07/23/2024 01:33:31 PM, Baghirzade, Nara) Latch assembly was broken. Replaced latch assembly. Check-in tests completed. IUI connectors and device cleaned.
69130 | 208395 | CM | CMPLT | CMPLT | 2024-07-23 13:37 | 2024-07-23 13:38 | User stated "error" when connecting to PCU | (TC 07/23/2024 01:37:35 PM, Baghirzade, Nara) Downloaded error log; did not find any significant errors. Completed analog sensors test; no issues found. During check-in, the rate accuracy failed (reading slightly high). Calibrated rate. Check-in tests completed. IUI connectors and device cleaned.
1058259 | 208396 | CM | CMPLT | CMPLT | 2024-07-23 13:38 | 2024-07-23 13:39 | No issue reported; functional check | (TC 07/23/2024 01:38:31 PM, Baghirzade, Nara) Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 110 mohms Leakage current: 65.4 uA
1047847 | 208397 | CM | CMPLT | CMPLT | 2024-07-23 14:00 | 2024-08-12 08:15 | Device won't swtich on | (TC 07/23/2024 01:59:36 PM, Baghirzade, Nara) Originally handled/disassembled/assembled by Courtney, she reported that she was unable to switch on the device. Received the device, replaced the battery and the data transmission cable; greased the lead screw and checked that all other parts/components were in place. Assembled the device back, was able to switch it on. Plunger moves back and forth smoothly. Will complete full calibration and PM. (TC 07/24/2024 10:55:54 AM, Baghirzade, Nara) Completed full calibration. When running the pump, it was giving the coarse travel/check clutch and plunger error however this would only occur when the plunger is fully extended. Put the plunger midway and the alarm was not showing up/occlusion alarm was passing. Will need to disassemble and check or replace the clutch. (TC 07/24/2024 10:57:57 AM, Baghirzade, Nara) Disassembled the device fully. Replaced the clutch. Checked all other components, did not see any other issues. Assembled the device back. Completed full calibration. When running the pump, the same issue is present where it's coarse travel when plunger is fully extended but when it is midway or closer to the end of infusion, there are no coarse travel errors and occlusion alarm is functional. Loud knocking/whirring noises heard when plunger is fully extended and running. No noises when plunger is midway or closer to the end of the infusion. Will have to disassemble and check the plunger. (TC 07/26/2024 11:09:47 AM, Baghirzade, Nara) Disassembled the device. Re-seated the metal bracket that has the motor on it. Replaced plunger rod seal. Checked that the plunger moves smoothly back and forth. Assembled the device back. Completed full calibration. During the occlusion alarm check, there was a coarse travel/check plunger error again. Replaced the plungerhead. Asssembled the plungerhead and when switching on the pump, it immediately has the "force sensor test failure" error. Will replace the force sensor. (TC 08/12/2024 08:14:43 AM, Baghirzade, Nara) After replacing the force sensor, the plunger was not moving smoothly and was creating a lot of noise. Disassembled the device; swapped out the entire plunger assembly with a spare. Greased the lead screw. Checked to make sure all parts are secure and not broken. Assembled the device back, ran a functional for about an hour, and no issues or errors found. Plunger is moving smoothly and quietly.
| 208398 | CM | CMPLT | CMPLT | 2024-07-23 14:04 | 2024-07-23 14:04 | Meeting with Andrew | (TC 07/23/2024 02:04:25 PM, Baghirzade, Nara) Meeting with Andrew regarding EEG and Lasers.
1070242 | 208399 | CM | CMPLT | CMPLT | 2024-07-23 14:06 | 2024-07-23 14:06 | Replaced batteries | (TC 07/23/2024 02:05:51 PM, Baghirzade, Nara) Replaced 6 D-batteries. Scale is functional. Returned to PNAC.
67951 | 208400 | CM | CANCL | CANCL | 2024-07-23 14:45 | 2024-07-23 14:46 | Created by Failed PM WO # 207557 : PM-MAJOR - Annual PM For: Incubators, Infant, Mobile |  | (TC 07/23/2024 02:43:23 PM, Rhuland, Courtney) During PM, noticed humidfier was loose When removing, chassis screwhold broke off Retiring unit and using for parts to repair other incubator with new chassis.
1057528 | 208401 | CM | CMPLT | CMPLT | 2024-07-23 14:56 | 2024-07-25 14:27 | SPO2 sensor not reading on monitor | (TC 07/23/2024 03:02:56 PM, Scero, Tomlyn) Confirmed with nurse that the SPO2 wasn't reading from the patient on the monitor. I gave her the CCU spare2 monitor and the values were showing. (TC 07/25/2024 02:27:28 PM, Scero, Tomlyn) Device was tested with an SPO2 simulator. No issues were present. spo2 waveform was detected right away.
| 208402 | RDS | CMPLT | CMPLT | 2024-07-23 15:12 | 2024-07-23 15:18 | Weekly check for ED
| 208403 | RDS | CMPLT | CMPLT | 2024-07-23 15:13 | 2024-07-31 13:33 | Weekly check for ICU
1050725 | 208404 | CM | CMPLT | CMPLT | 2024-07-23 15:54 | 2024-07-23 15:55 | it won’t turn off and is going continuous | (TC 07/23/2024 03:55:29 PM, Rendon Lopez , Elizabeth) Defective activation cable. Replaced cable, tested funtions-OK
1062003 | 208405 | CM | CMPLT | CMPLT | 2024-07-23 16:40 | 2024-08-14 16:05 | CPR release handle damaged | (TC 07/23/2024 04:40:17 PM, Cortez, John) 12S requested extra bed CPR handle needs to be repaired first Requested service from Stryker (TC 08/14/2024 04:05:08 PM, Cortez, John) repaired by Aziz (stryker)
1057383 | 208406 | CM | CMPLT | CMPLT | 2024-07-23 17:06 | 2024-07-29 15:19 | MSH 525-2 CR ROOM CR 850 Internal bar scanner not working. | (TC 07/23/2024 05:10:48 PM, Dolinar, John) Internal bar scanner not working. (TC 07/23/2024 05:11:01 PM, Fan, Tak) Internal bar scanner not working. (TC 07/24/2024 03:21:55 PM, Whittington, Scott) MSH 525-2 CR ROOM CR 850 Internal bar scanner not working. RELABEL LLI SET A (TC 07/24/2024 03:21:55 PM, Whittington, Scott) MSH 525-2 CR ROOM CR 850 Internal bar scanner not working. RELABEL LLI SET A (TC 07/24/2024 03:21:55 PM, Whittington, Scott) MSH 525-2 CR ROOM CR 850 Internal bar scanner not working. RELABEL LLI SET A (TC 07/25/2024 08:21:00 AM, Whittington, Scott) MSH 525-2 CR ROOM CR 850 Internal bar scanner not working. RELABEL LLI SET A (TC 07/29/2024 03:18:29 PM, Whittington, Scott) nternal bar scanner not working. Ref# 1008565 K4187-3265 MIE 1057383 Ref# 1008565 Internal barcode reader not working. AAREN FROM CARESTREAM REPLACED INTERNAL BARCODE SCANNER | Internal bar scanner not working. Ref# 1008565 K4187-3265 MIE 1057383 Ref# 1008565 Internal barcode reader not working.
501891 | 208407 | CM | CMPLT | CMPLT | 2024-07-23 17:10 | 2024-07-31 15:48 | MSH Internal Medicine SONOSITE MTURBO Screen cracked. | (TC 07/31/2024 03:47:36 PM, Whittington, Scott) MSH Internal Medicine SONOSITE MTURBO Screen cracked. Orderded new screen. (TC 07/31/2024 03:47:36 PM, Dolinar, John) MSH Internal Medicine SONOSITE MTURBO Screen cracked. Orderded new screen. | MSH Internal Medicine SONOSITE MTURBO Screen cracked. Orderded new screen. Replaced screen.
1048472 | 208408 | CM | CMPLT | CMPLT | 2024-07-23 17:13 | 2024-07-23 17:13 | Battery error | (TC 07/24/2024 08:56:00 AM, Fan, Tak) Wrong detector screws. No part number, contacted carestream. Ordered screws. | Changed battery contacts. Performed ful detector cal.
1049905 | 208409 | CM | CMPLT | CMPLT | 2024-07-23 17:15 | 2024-07-23 17:16 | Detector losing wifi conection. |  | Replaced battery contacts. Performed full detector cal.
68848 | 208410 | CM | CMPLT | CMPLT | 2024-07-24 08:33 | 2024-07-24 08:33 | Bent plug | (TC 07/24/2024 08:33:13 AM, DeGuzman, Leila) Unbent ground pin. Device okay. Returned to service.
1050479 | 208411 | PM | CMPLT | CMPLT | 2024-07-24 09:09 | 2024-07-22 09:09 | PM-MAJOR- Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric | Performed annual PM. Tested. Functional.
52230 | 208412 | PM | CMPLT | CMPLT | 2024-07-24 11:13 | 2024-07-24 12:53 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/24/2024 12:52:37 PM, Rhuland, Courtney) Replaced: Hood Humidifier Display bezel Rechargeable 9V battery Calibrated: MAINS & VREF voltage Scale RH Functional check - OK Set 35C, 75% humidity Read 35C, 75.6% humidity Electrical safety - OK Ground resistance - 133 mOhms Current leakage - 64.3 uA
1071715 | 208413 | CM | CMPLT | CMPLT | 2024-07-24 13:05 | 2024-07-24 13:05 | Touch display was broken | (TC 07/24/2024 01:04:54 PM, Baghirzade, Nara) Tech from CAPSA came in and replaced the touch display. The motion controls are functional. Handed off the WOW to Jigar and Matt for further checks before clinical use.
| 208414 | CM | ACTIV | ISSUE | 2024-07-24 13:10 |  | Meetings | (TC 07/24/2024 01:10:11 PM, Rhuland, Courtney) PM meeting (TC 07/29/2024 12:51:29 PM, Rhuland, Courtney) Capital meeting (TC 07/31/2024 01:21:00 PM, Rhuland, Courtney) Boom PM meeting (TC 08/12/2024 02:01:00 PM, Rhuland, Courtney) Monday huddle (TC 08/13/2024 02:32:07 PM, Rhuland, Courtney) Meeting with Drager and NICU staff to test out locking brakes to minimize nitrogen leak (TC 08/16/2024 12:41:04 PM, Rhuland, Courtney) 1 on 1 (TC 08/19/2024 01:55:40 PM, Rhuland, Courtney) Monday huddle (TC 08/21/2024 01:15:14 PM, Rhuland, Courtney) Capital meeting (TC 08/22/2024 11:15:22 AM, Rhuland, Courtney) Met with Nazanin to go over boom counts for upcoming PM (TC 08/29/2024 01:18:51 PM, Rhuland, Courtney) NICU update meeting (TC 09/06/2024 11:21:01 AM, Rhuland, Courtney) Capital meeting to discuss equipment with Andrew (TC 09/09/2024 02:18:44 PM, Rhuland, Courtney) Monday huddle (TC 09/12/2024 01:49:50 PM, Rhuland, Courtney) ENT capital meeting with Andrew and Soojin (TC 09/18/2024 10:27:45 AM, Rhuland, Courtney) P3A/Capital meeting (TC 09/18/2024 12:25:32 PM, Rhuland, Courtney) Capital meeting with Andrew & dentistry (TC 09/19/2024 12:48:49 PM, Rhuland, Courtney) Capital meeting for ophthalmology clinic with Andrew & Henna (TC 09/23/2024 02:39:27 PM, Rhuland, Courtney) Monday huddle (TC 09/25/2024 12:48:44 PM, Rhuland, Courtney) Capital meeting NICU (TC 09/25/2024 01:20:04 PM, Rhuland, Courtney) Patient alarm management meeting (TC 09/26/2024 01:35:31 PM, Rhuland, Courtney) OPG capital meeting (TC 09/30/2024 01:56:24 PM, Rhuland, Courtney) Huddle & PM update (TC 10/01/2024 10:47:23 AM, Rhuland, Courtney) ENT meeting with olympus for capital (TC 10/01/2024 11:24:21 AM, Rhuland, Courtney) Walkthrough North Pod with Naz Located wall sphygs and created plan for implementation (TC 10/02/2024 01:44:27 PM, Rhuland, Courtney) ENT equipment meeting (TC 10/07/2024 02:52:52 PM, Rhuland, Courtney) monday huddle (TC 10/09/2024 11:28:10 AM, Rhuland, Courtney) Procedure light replacement meeting OPG (TC 10/17/2024 02:49:56 PM, Rhuland, Courtney) Capital meeting (TC 10/18/2024 01:12:26 PM, Rhuland, Courtney) Full team meeting (TC 10/21/2024 12:49:07 PM, Rhuland, Courtney) Philips prep meeting with NICU leadership staff (TC 10/21/2024 02:30:00 PM, Rhuland, Courtney) Colposcope meeting OPG (TC 11/04/2024 03:00:02 PM, Rhuland, Courtney) Monday huddle (TC 11/15/2024 12:07:20 PM, Rhuland, Courtney) Full team meeting (TC 11/25/2024 02:04:28 PM, Rhuland, Courtney) Monday huddle (TC 11/26/2024 11:49:45 AM, Rhuland, Courtney) Meeting with NICU and Kamran to go over alarm acknowledge changes to overview stations (TC 11/27/2024 02:50:32 PM, Rhuland, Courtney) Dialysis training (TC 11/28/2024 10:17:26 AM, Rhuland, Courtney) Capital meeting (TC 11/28/2024 03:27:29 PM, Rhuland, Courtney) ENT meeting
| 208415 | CM | CMPLT | CMPLT | 2024-07-24 13:12 | 2024-07-24 13:12 | PM update meeting | (TC 07/24/2024 01:11:57 PM, Baghirzade, Nara) Biomed PM update meeting
| 208416 | CM | CMPLT | CMPLT | 2024-07-24 13:13 | 2024-07-24 13:13 | Meeting with Naz | (TC 07/24/2024 01:12:45 PM, Baghirzade, Nara) Met with Naz to get NICU updates
1056125 | 208417 | CM | CMPLT | CMPLT | 2024-07-24 15:09 | 2024-07-24 15:09 | Error message/not warming | (TC 07/24/2024 03:09:20 PM, Canjimavadivel, Kamallesh ) Issue was port 1, removed blanket and switched port. Was able to heat up. Need to further investigate issue with port 1
1070841 | 208418 | CM | CMPLT | CMPLT | 2024-07-25 07:22 | 2024-07-24 07:24 | Phototherapy light install | (TC 07/24/2024 07:23:47 AM, Rhuland, Courtney) Installed phototherapy light Scale was 10g off - calibrated and verified scale
63438 | 208419 | CM | CMPLT | CMPLT | 2024-07-25 08:33 | 2024-07-25 10:11 | System error code 22 | (TC 07/25/2024 10:11:27 AM, DeGuzman, Leila) Ran twice, no error found. Cleaned device. Returned to service.
1071403 | 208420 | CM | CMPLT | CMPLT | 2024-07-25 10:47 | 2024-07-25 10:47 | Missing protective cover | (TC 07/25/2024 10:47:21 AM, Rhuland, Courtney) Noticed light missing cover in storage room Replaced cover with spare Functional check - OK
| 208421 | CON | CMPLT | CMPLT | 2024-07-25 11:24 | 2024-07-25 11:24 | ECG leads for echo | (TC 07/25/2024 11:24:08 AM, Fan, Tak) Part#: 2106390-001, Leadwires ECG MARQ. AHA / AMERICA, Quantity: 5
1061998 | 208422 | CM | CMPLT | CMPLT | 2024-07-25 14:18 | 2024-07-25 14:18 | broken power plug | replaced broken power cord. tested. Functional.
69360 | 208423 | FC | CMPLT | CMPLT | 2024-07-25 14:43 | 2024-07-25 14:43 | Sent down with PCU, will run functional check | (TC 07/25/2024 02:42:49 PM, Cortez, John) Check-in passed
1065910 | 208424 | FC | CMPLT | CMPLT | 2024-07-25 14:44 | 2024-07-25 14:44 | Sent down with PCU, will run functional check | (TC 07/25/2024 02:44:28 PM, Cortez, John) Check-in passed
69301 | 208425 | CM | CMPLT | CMPLT | 2024-07-25 15:11 | 2024-07-25 15:11 | Sent down with PCU, will run functional check | (TC 07/25/2024 03:10:42 PM, Cortez, John) Check-in passed
| 208426 | CM | CMPLT | CMPLT | 2024-07-25 15:45 | 2024-07-25 15:46 | Check suction regulators | (TC 07/25/2024 03:45:50 PM, DeGuzman, Leila) Checked all the suction regulator from the incoming shelves outside the shop. 16 suction regulators in total. -10 suctions are not working functionally. -6 suctions returned to service.
65869 | 208427 | PM | CMPLT | CMPLT | 2024-07-26 08:07 | 2024-07-26 08:08 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 07/25/2024 08:07:33 AM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
68064 | 208428 | CM | CMPLT | CMPLT | 2024-07-26 08:14 | 2024-07-25 08:15 | Display not working | (TC 07/25/2024 08:14:32 AM, Baghirzade, Nara) Replaced batteries for display. Assembled back together and completed functional check. Regulator is fully functional.
1061333 | 208429 | FC | CMPLT | CMPLT | 2024-07-26 08:16 | 2024-07-25 08:16 | QC checklist | (TC 07/25/2024 08:15:40 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
| 208430 | CM | CMPLT | CMPLT | 2024-07-26 08:17 | 2024-07-25 08:17 | Suction regulators functional checks | (TC 07/25/2024 08:17:18 AM, Baghirzade, Nara) Sarabeth requested to have several regulators checked as they were having issues in some of the patient rooms. Checked 7 regulators, changed batteries for 2 of them. 5/7 regulators are fully functional and no issues found with them. Returned to 11 North and updated Sarabeth.
1061207 | 208431 | CM | CMPLT | CMPLT | 2024-07-26 08:40 | 2024-07-23 08:42 | Monitor Tilting | Call re: monitor issue. Went to check. Waited for access. FOund monitor sliding back when adjusting it at an angle. Removed cover and adjusted tension. Checked functionality. Tested ok.
1063403 | 208432 | CM | CMPLT | NFF | 2024-07-26 09:22 | 2024-07-25 09:22 | "technical error" | (TC 07/25/2024 09:22:17 AM, Cortez, John) Pre-op check passed Retrieved logs, technical error 27 listed (relates to speaker" Ran panel test in service mode, passed Ran pre-op check again, passed Unable to replicate issue
1047671 | 208433 | CM | CMPLT | CMPLT | 2024-07-26 10:31 | 2024-07-30 11:35 | MSH OPG 8 PHILIPS AFFINITI TV PROBE XTAL | (TC 07/26/2024 10:31:22 AM, Whittington, Scott) MSH OPG 8 PHILIPS AFFINITI TV PROBE XTAL (TC 07/30/2024 11:27:15 AM, Whittington, Scott) MSH OPG 8 PHILIPS AFFINITI TV PROBE XTAL MAY BE SOME MEMBRANE ISSUE, BUT COMPARING TO THE 4 OTHER PROBES IN INVENTORY, NOT A GREAT DIFFERENTIATION. ADVISED LUANNE THAT THE CLINICIANS SHOULD TAKE A PICTURE OF ANY ERROR AND EMAIL TO ME | MSH OPG 8 PHILIPS AFFINITI TV PROBE XTAL REF# 0124022092
66797 | 208434 | PM | CMPLT | CMPLT | 2024-07-26 10:47 | 2024-07-26 10:49 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | PM Completed, refer to WO# 197125
69641 | 208435 | CM | CMPLT | CMPLT | 2024-07-26 10:59 | 2024-07-26 10:59 | Sent down with PCU, will run functional check | (TC 07/26/2024 10:59:02 AM, Cortez, John) Check-in passed
67959 | 208436 | CM | CMPLT | CMPLT | 2024-07-26 11:05 | 2024-07-26 11:05 | Air probe failure | (TC 07/26/2024 11:04:37 AM, Baghirzade, Nara) Ran the incubator and was consistently getting air probe failure after about 5-10 minutes. Replaced air probe. Ran the incubator again for about 30 minutes and no air probe errors present.
1055576 | 208437 | CM | CMPLT | CMPLT | 2024-07-26 11:13 | 2024-07-26 11:13 | SpO2 and temp not working | Replaced temperature probe, SpO2 probe and SpO2 cable. Installed adult cuff. Tested. Functional.
67959 | 208438 | PM | CMPLT | CMPLT | 2024-07-26 13:51 | 2024-07-26 13:53 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 07/26/2024 01:52:22 PM, Baghirzade, Nara) PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 79.6 uA Power cord resistance = 185 mohms (TC 07/26/2024 01:52:59 PM, DeGuzman, Leila) PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 79.6 uA Power cord resistance = 185 mohms
1069960 | 208439 | CM | CMPLT | NFF | 2024-07-26 13:54 | 2024-07-26 13:55 | Reported that it errors when warming milk | (TC 07/26/2024 01:53:59 PM, Baghirzade, Nara) Placed a bag with water in the warmer; first run was 5 minutes - no errors observed. Second run was 15 minutes - no errors. Third run was 15 minutes- no errors observed. Water was sufficiently warm. No fault found with the device. (TC 07/26/2024 01:55:07 PM, DeGuzman, Leila) Placed a bag with water in the warmer; first run was 5 minutes - no errors observed. Second run was 15 minutes - no errors. Third run was 15 minutes- no errors observed. Water was sufficiently warm. No fault found with the device.
| 208440 | RDS | CMPLT | CMPLT | 2024-07-26 13:56 | 2024-07-26 13:56 | Weekly Rounds: 11 North | (TC 07/26/2024 01:56:00 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. Refer to procedure for results.
1048500 | 208441 | CM | CMPLT | CMPLT | 2024-07-26 13:57 | 2024-09-25 10:46 | MSH US EPIQ FAULTY QWERTY KEYBOARD (TOP ROW OF LETTERS NOT WORKING CONSISTENTLY) ALSO REQUIRES POWER CORD AND TGC KNOB SET | (TC 07/30/2024 11:43:29 AM, Whittington, Scott) MSH US EPIQ FAULTY QWERTY KEYBOARD (TOP ROW OF LETTERS NOT WORKING CONSISTENTLY) ALSO REQUIRES POWER CORD AND TGC KNOB SET REF# 0124022133 PAARTS ORDERED AND DELIVERED (TC 07/31/2024 03:24:58 PM, Chau, Ken) MSH US EPIQ FAULTY QWERTY KEYBOARD (TOP ROW OF LETTERS NOT WORKING CONSISTENTLY) ALSO REQUIRES POWER CORD AND TGC KNOB SET MONITOR GAS STRUT NOT AVAILABLE | MSH US EPIQ FAULTY QWERTY KEYBOARD (TOP ROW OF LETTERS NOT WORKING CONSISTENTLY) ALSO REQUIRES POWER CORD AND TGC KNOB SET REF# 0124022133
| 208442 | RDS | CMPLT | CMPLT | 2024-07-26 13:57 | 2024-07-26 13:58 | Weekly Rounds: 11 South | (TC 07/26/2024 01:57:40 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 208443 | RDS | CMPLT | CMPLT | 2024-07-26 13:59 | 2024-07-26 13:59 | Weekly Rounds: NICU & RESUS | (TC 07/26/2024 01:58:48 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
32510 | 208444 | CM | CMPLT | CMPLT | 2024-07-26 14:00 | 2024-09-25 10:51 | MSH DR1 TABLE TUBE /DETECTOR ALIGNMENT | (TC 07/26/2024 10:48:00 PM, Whittington, Scott) MSH DR1 TABLE TUBE /DETECTOR ALIGNMENT | MSH DR1 TABLE TUBE /DETECTOR ALIGNMENT
| 208445 | CM | CMPLT | CMPLT | 2024-07-26 14:01 | 2024-07-26 14:01 | Meeting with Nima re: device policies | (TC 07/26/2024 02:00:56 PM, Baghirzade, Nara) Meeting with Nima re: device policies
501647 | 208446 | PM | CMPLT | CMPLT | 2024-07-26 14:15 | 2024-07-26 14:16 | PM - Annual | (TC 07/26/2024 02:15:31 PM, Green, Paul) Performed PM. Fully functional.
1054115 | 208447 | PM | CMPLT | CMPLT | 2024-07-26 15:21 | 2024-07-26 15:22 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 07/26/2024 03:21:45 PM, Rendon Lopez , Elizabeth) PM completed, all tests passed. Cleaned recorder, calibrated - OK
| 208448 | CM | CMPLT | CMPLT | 2024-07-26 17:11 | 2024-07-26 17:11 | Site coverage
| 208449 | RDS | CMPLT | CMPLT | 2024-07-29 07:12 | 2024-07-29 13:45 | Weekly check for ED
| 208450 | PA | CMPLT | CMPLT | 2024-07-29 07:46 | 2024-07-30 11:55 | Ordering sterile water for NICU altrix | (TC 07/29/2024 07:50:50 AM, Baghirzade, Nara) SC 1003279939 (TC 07/30/2024 11:54:32 AM, Baghirzade, Nara) Order # 4150281030
1061204 | 208451 | CM | CMPLT | CMPLT | 2024-07-29 08:40 | 2024-07-25 12:48 | Nitrogen Hose connection issue | CAll re: Nitrogen is leaking. Went to check. Waited for access. Spoke with staff. Checked wall connection. Tested ok. Found Y connector at the end of hose were leaking. They had two fo those. Brought it down to biomed. Dismantled both y connectors, removed old teflone tape and reapplied it. Assembled it together. Coordinatedd and Scheduled testing with the staff (required instruments from MDRD). Tested ok. Returned to OR.
1056911 | 208452 | CM | CMPLT | CMPLT | 2024-07-29 08:47 | 2024-07-24 12:52 | Screen blank | Call re: not workingl. Went to check. Unit being used. Waited for access. Spoke with staff and was informed that the screen is not displaying the image. Founf blank white screen on the monitor. Checked and verified all connections. Called JH for assistance. Was advised to check the screen button on microscope. Found they set up the monitor on a different shared screen. Issue resolved.
1055005 | 208453 | CM | CMPLT | CMPLT | 2024-07-29 08:53 | 2024-07-23 13:54 | "not wotking" | Unit was sent with tag not working. Cleaned unit. Found roll stand handle loose. FIxed it. Unit could not turn on. FOund battery was depleted. Charged unti for a while. Checked functionality. No issue found. Tested ok. Returned to service.
1058543 | 208454 | CM | CMPLT | CMPLT | 2024-07-29 09:05 | 2024-07-25 14:12 | Recording issue | Call re: unit not printing. went to check. FOund unit was not recording the alarms. Checked connections and rebooted the recorder. Rebooted system. Issue persisted. Consulted KS for assistance and worked with him to troubleshoot the issue. Found the recording device was not assigned to the appropriate central station. Rectified it. Tested ok.
1066843 | 208455 | CM | CMPLT | CMPLT | 2024-07-29 09:23 | 2024-07-23 16:23 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
1066934 | 208456 | CM | CMPLT | CMPLT | 2024-07-29 09:25 | 2024-07-23 17:24 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
68021 | 208457 | CM | CMPLT | CMPLT | 2024-07-29 09:40 | 2024-07-23 17:40 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
1066934 | 208458 | CM | CMPLT | CMPLT | 2024-07-29 09:42 | 2024-07-23 09:42 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
| 208459 | CM | CMPLT | CMPLT | 2024-07-29 10:08 | 2024-07-26 10:08 | PACU Central Station Printer setup | Worked with KC to setup laserjet printers in all central stations in PACU.
66512 | 208460 | CM | CMPLT | CMPLT | 2024-07-29 10:12 | 2024-08-01 11:47 | Recorder not feeding | (TC 07/29/2024 10:16:11 AM, Rendon Lopez , Elizabeth) Confirmed issue. Removed monitor from room and replaced with spare. (TC 07/29/2024 10:16:11 AM, Rendon Lopez , Elizabeth) Recorder has been repaired multiple times, ordering exchange recorder from Philips. (TC 07/31/2024 11:47:02 AM, Rendon Lopez , Elizabeth) Repalced recorder with exchange. Calibrated, tested all functions-OK Returned defective recorder to Philips (TC 08/01/2024 11:47:02 AM, Rendon Lopez , Elizabeth) Returned monitor to Room 18, tested connections-OK
66533 | 208461 | CM | CMPLT | CMPLT | 2024-07-29 10:18 | 2024-07-29 10:19 | Recorder missing front cover | (TC 07/29/2024 10:19:10 AM, Rendon Lopez , Elizabeth) Replaced cover that was taken out, tested printer, calibrated-OK
| 208462 | RDS | CMPLT | CMPLT | 2024-07-29 10:48 | 2024-07-29 10:48 | Weekly checks in CCU | (TC 07/29/2024 10:48:08 AM, Scero, Tomlyn) weekly checks complete. 30 J test passed
1061316 | 208463 | CON | CMPLT | CMPLT | 2024-07-29 11:09 | 2024-07-29 11:09 | Workstation laggy/cannot save. | (TC 07/26/2024 11:09:22 AM, Fan, Tak) Vpadmin volume settings were set too high for PC storage. Increased volume.
1054549 | 208464 | CM | CMPLT | NFF | 2024-07-29 13:12 | 2024-07-29 13:16 | Not printing | (TC 07/29/2024 01:15:58 PM, Rhuland, Courtney) Connected patient simulator to device Had no issues printing test sheet No fault found
1048894 | 208465 | CM | CMPLT | CMPLT | 2024-07-29 13:33 | 2024-07-29 13:34 | Physical damage | (TC 07/29/2024 01:33:31 PM, Rendon Lopez , Elizabeth) Repaired previously. Not needed, retired from service.
1047008 | 208466 | CM | CMPLT | CMPLT | 2024-07-29 13:35 | 2024-07-29 13:35 | Physical damage | (TC 07/29/2024 01:34:55 PM, Rendon Lopez , Elizabeth) Repaired previously. Not needed, retired from service.
1054197 | 208467 | CM | CMPLT | CMPLT | 2024-07-29 13:36 | 2024-08-15 14:35 | Physical damage | (TC 07/29/2024 01:36:06 PM, Rendon Lopez , Elizabeth) Confirm, kept to send out for repair (TC 08/15/2024 02:34:49 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1064894 | 208468 | FC | CMPLT | CMPLT | 2024-07-29 14:08 | 2024-07-29 14:08 | Performance verification | (TC 07/29/2024 02:07:37 PM, Canjimavadivel, Kamallesh ) Performed device self-test. all values within allowance.
1061213 | 208469 | CM | CMPLT | CMPLT | 2024-07-29 14:32 | 2024-07-29 14:32 | OR 11 Wall monitor not showing image | (TC 07/25/2024 02:32:59 PM, Huang, Jessica) received the call from OR 11 and was informed the wall monitor was not showing any image. it was in the middle of the case so had to check it on friday. (TC 07/26/2024 02:36:28 PM, Huang, Jessica) checked and did troubleshooting on the wall monitor with Rutvik. contacted Derrick for support. informed Tory once the issue got fixed.
1062121 | 208470 | PM | CMPLT | CMPLT | 2024-07-29 14:51 | 2024-07-30 10:52 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 07/30/2024 10:52:23 AM, Rendon Lopez , Elizabeth) Upgraded Software, downloaded library. PM completed, all tests passed
| 208471 | CM | CMPLT | CMPLT | 2024-07-29 14:58 | 2024-07-29 15:03 | MSH MOBILE SYSTEMS WIRELESS NO CONNECTION | (TC 07/27/2024 04:58:00 PM, Whittington, Scott) MSH MOBILE SYSTEMS WIRELESS NO CONNECTION AFTER JDMI UPGRADE, WIRELESS CONNECTIONS TO RIS & CORAL LOST ATTEMPTED TO CONNECT TO WIRELESS AND WIRED , NO SUCCESS MSH IT PRE-NAC SYSTEMS THAT USED WIRED CONNNECTION AS ONLY A BACKUP WERE NOT GRANDFATHERED INTO WHITELIST JDMI SECOND LEVEL RESOLVED CONNECTION PROBLEM WITH WIRELESS | MSH MOBILE SYSTEMS WIRELESS NO CONNECTION
| 208472 | CM | CMPLT | CMPLT | 2024-07-29 15:25 | 2024-09-25 10:51 | MSH OPG SPP ELEVATTOR LOBBY STORAGE ASSESS SYSTEMS FOR DISPOSAL | (TC 07/29/2024 03:25:33 PM, Whittington, Scott) MSH OPG SPP ELEVATTOR LOBBY STORAGE ASSESS SYSTEMS FOR DISPOSAL (TC 07/30/2024 11:24:42 AM, Whittington, Scott) MSH OPG SPP ELEVATTOR LOBBY STORAGE ASSESS SYSTEMS FOR DISPOSAL ADDITIONAL DISPOSED DEVICES FOR SCRAP FOUND ON 8TH FLR MOVED TO 3 FLR FOR RMOVAL BY HOUSEKEEPING (TC 07/30/2024 11:24:42 AM, Whittington, Scott) MSH OPG SPP ELEVATTOR LOBBY STORAGE ASSESS SYSTEMS FOR DISPOSAL ADDITIONAL DISPOSED DEVICES FOR SCRAP FOUND ON 8TH FLR MOVED TO 3 FLR FOR RMOVAL BY HOUSEKEEPING | MSH OPG SPP ELEVATTOR LOBBY STORAGE ASSESS SYSTEMS FOR DISPOSAL
1064827 | 208473 | CM | CMPLT | CMPLT | 2024-07-29 16:14 | 2024-07-29 16:14 | Front bezel for parameters is cracked/broken | (TC 07/29/2024 04:13:40 PM, Cortez, John) Replaced bezel
1061662 | 208474 | CM | CMPLT | CMPLT | 2024-07-29 16:18 | 2024-07-29 16:18 | "patient noted that the bars on the pump display were going up" | (TC 07/29/2024 04:18:19 PM, Cortez, John) Pump connected to standard mattress. Second set of LEDs shouldn't light up unless connected to truturn mattress. Spoke with Vincy. Possible that patient was using another bed with truturn before, when patient noticed problem. pump is functioning normally.
| 208475 | CM | ACTIV | ISSUE | 2024-07-29 17:24 |  | MSH MOBILE XRAY & ULTRASOUND PRE-NAC DO NOT HAVE LAN CONNECTION BACKUP | (TC 08/06/2024 02:11:03 PM, Whittington, Scott) MSH MOBILE XRAY & ULTRASOUND PRE-NAC DO NOT HAVE LAN CONNECTION BACKUP | MSH MOBILE XRAY & ULTRASOUND PRE-NAC DO NOT HAVE LAN CONNECTION BACKUP
52252 | 208476 | CM | CMPLT | CMPLT | 2024-07-30 08:28 | 2024-07-29 08:29 | Broken porthole door and hinge | (TC 07/29/2024 08:28:37 AM, Baghirzade, Nara) Replaced porthole door. Replaced porthole hinge. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 74.5%RH Incubator is fully functional.
1060030 | 208477 | CM | CMPLT | CMPLT | 2024-07-30 08:30 | 2024-09-05 07:37 | MIPS/MEPS test failing | (TC 07/29/2024 08:30:21 AM, Baghirzade, Nara) Met with Gaby as she reported that the MIPS/MEPS test is not working. First tried to do the test with a calibration syringe however the pressure doesn't hold and doesn't go above 20 on the graph. Then Gaby tried to do the test on herself and was getting the same result. Checked the patient valve, ensured that everything is connected; ensured pistons are moving properly. Possible leaks in patient valve; will have to replace the rubber seals on the pistons to confirm. (TC 07/29/2024 08:31:56 AM, Baghirzade, Nara) Practiced taking the pistons and rubbers seals out and then placing the back in. Sealing the nut with silicone. (TC 09/05/2024 07:37:11 AM, Baghirzade, Nara) Removed pistons; cleaned. Replaced pistons in the patient valve. Secured the nut on both pistons with silicone. Let it dry and then connected it back to the bodybox. Gaby ran a MIPS/MEPS cal test which passed. Also completed a biological test. MIPS/MEPS testing is functional.
1061333 | 208478 | FC | CMPLT | CMPLT | 2024-07-30 08:33 | 2024-07-29 08:33 | QC checklist | (TC 07/29/2024 08:33:17 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North
| 208479 | CM | CMPLT | CMPLT | 2024-07-30 08:36 | 2024-07-29 08:36 | Capital meeting | (TC 07/29/2024 08:36:22 AM, Baghirzade, Nara) Capital meeting
| 208480 | CM | CMPLT | CMPLT | 2024-07-30 08:37 | 2024-07-29 08:38 | Meeting with Janet | (TC 07/29/2024 08:37:30 AM, Baghirzade, Nara) Met with Janet to discuss the caregiver assignments on the MX800 monitors in NICU.
1070771 | 208481 | CM | CMPLT | CMPLT | 2024-07-30 08:40 | 2024-07-29 08:41 | Computer not turning on | (TC 07/29/2024 08:40:42 AM, Baghirzade, Nara) Computer tray moves up and down; electronically the cart is functional. Had Jigar from IT come and check the computer; computer is switching on and connected to the network. He confirmed that the computer is functional. Returned to 12 South.
1065904 | 208482 | CM | CMPLT | NFF | 2024-07-30 09:46 | 2024-07-30 09:46 | Detecting air bubbles when not present | (TC 07/30/2024 09:46:14 AM, Rhuland, Courtney) Ran air-in-line tests both wet and dry Both tests passed Functional check - OK No fault found
1069493 | 208483 | CM | CMPLT | CMPLT | 2024-07-30 10:40 | 2024-08-01 11:06 | Not turning on | (TC 07/30/2024 12:21:59 PM, Rhuland, Courtney) Worked with Samir to diagnose issue Followed voltage and resistance through circuit Battery reading 2.2V when powered by 9V adapter, but does not charge New battery ordered (TC 07/31/2024 02:34:40 PM, Rhuland, Courtney) Battery recieved - device still not turning on Battery charging but not heating After troubleshooting, found one of the wires inside heating cord was cut To be repaired tomorrow (TC 08/01/2024 11:05:37 AM, Rhuland, Courtney) Resoldered searing pen to board Installed battery Functional check - OK Returned to clinic
1055576 | 208484 | CM | CMPLT | CMPLT | 2024-07-30 12:20 | 2024-08-19 13:01 | Battery replacement PO | battery replacement PO. received exchange unit. updated SN on TMS. configured and tested. funtional
68606 | 208485 | CM | CMPLT | CMPLT | 2024-07-30 13:24 | 2024-07-30 13:24 | Replace front and rear case | (TC 07/30/2024 01:23:57 PM, DeGuzman, Leila) Replaced the front and rear case. Replaced pressure disc sensor, gasket, and disc holder. After replacing components, performed calibration, verification accuracy. Updated software, and closed Software recall work order. Performed PM, and closed PM work order
1050620 | 208486 | CM | CMPLT | CMPLT | 2024-07-30 13:33 | 2024-07-30 13:45 | Replace rear case | (TC 07/30/2024 01:44:41 PM, DeGuzman, Leila) Replaced rear case. Performed calibration and verfication accuracy. Performed PM and closed PM work order.
501791 | 208487 | ACC | CMPLT | CMPLT | 2024-07-30 14:17 | 2024-08-02 12:05 | MSH CEOU - Damaged Membrane | (TC 07/30/2024 02:17:14 PM, Chau, Ken) Ordered replacement TV probe (TC 08/02/2024 12:05:00 PM, Dolinar, John) New replacement probe SN#F088J1 (TC 08/02/2024 02:04:55 PM, Whittington, Scott) New replacement probe SN#F088J1 | New replacement probe SN#F088J1
69128 | 208488 | PM | CMPLT | CMPLT | 2024-07-31 11:58 | 2024-07-30 11:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/31/2024 11:57:49 AM, Baghirzade, Nara) Pressure sensors failed analog sensor test. Replaced both pressure sensors. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated. (TC 07/30/2024 11:57:49 AM, Baghirzade, Nara) Pressure sensors failed analog sensor test. Replaced both pressure sensors. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1067437 | 208489 | FC | CMPLT | CMPLT | 2024-07-31 11:59 | 2024-07-30 12:00 | QC checklist | (TC 07/30/2024 11:59:34 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1046541 | 208490 | CM | CMPLT | CMPLT | 2024-07-31 12:01 | 2024-07-30 12:01 | Not working | (TC 07/30/2024 12:01:12 PM, Baghirzade, Nara) Replaced batteries. Unit was not switching on. Tried a different set of batteries, but unit was staying off. Retiring the device.
| 208491 | CM | CMPLT | CMPLT | 2024-07-31 12:05 | 2024-07-30 12:05 | Supervising and teaching coop student on alaris syringes | (TC 07/30/2024 12:05:23 PM, Baghirzade, Nara) Teaching Leila (coop student) on how to replace the front case, rear case, and various parts in the alaris syringes.
69637 | 208492 | PM | CMPLT | CMPLT | 2024-07-31 12:09 | 2024-07-30 12:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 07/30/2024 12:09:03 PM, Baghirzade, Nara) Pressure sensors failed analog sensor test. Replaced both pressure sensors. Rate was calibrated. Pressure was calibrated PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
| 208493 | PA | CMPLT | CMPLT | 2024-07-31 12:13 | 2024-09-11 06:55 | Ordering parts for stock | (TC 07/31/2024 12:14:34 PM, Baghirzade, Nara) SC 1003282224 (TC 08/21/2024 07:46:38 AM, Baghirzade, Nara) Received 8/25 PO 7202439155 (TC 09/11/2024 06:54:53 AM, Baghirzade, Nara) Fully received. PO 7202439155
67890 | 208494 | CM | CMPLT | CMPLT | 2024-07-31 13:25 | 2024-10-25 09:29 | Alarming occlusion | (TC 07/31/2024 01:24:49 PM, Rendon Lopez , Elizabeth) Confirmed issue after running pump. Requested RMA to send out for repair (TC 08/14/2024 07:40:03 AM, Rendon Lopez , Elizabeth) Received complaint CN-316378, RMA SR 3063180 Sent unit out for repair (TC 10/25/2024 09:29:22 AM, Baghirzade, Nara) Received unit from repair. Attached service report. PO 7500210235
1054161 | 208495 | CM | CMPLT | CMPLT | 2024-07-31 13:43 | 2024-08-15 10:36 | weak signal | (TC 07/31/2024 01:43:55 PM, Rendon Lopez , Elizabeth) Confirmed issue. To send out for repair (TC 08/15/2024 10:35:58 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1055657 | 208496 | CM | CMPLT | CMPLT | 2024-07-31 13:50 | 2024-08-15 10:37 | Physical damage | (TC 07/31/2024 01:50:13 PM, Rendon Lopez , Elizabeth) Back membrane torn. Keep to send out for repair (TC 08/15/2024 10:37:07 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1068093 | 208497 | CM | CMPLT | CMPLT | 2024-07-31 13:54 | 2024-08-15 14:24 | Keeps stopping to work | (TC 07/31/2024 01:55:04 PM, Rendon Lopez , Elizabeth) Noisy cable, sending out for repair under warranty. (TC 08/15/2024 02:23:50 PM, Rendon Lopez , Elizabeth) Device not covered under warranty repair. Repaired under flat rate. Incoming inspection-OK
1069316 | 208498 | CM | CMPLT | CMPLT | 2024-07-31 13:59 | 2024-07-31 13:59 | Not working | (TC 07/31/2024 01:59:06 PM, Rendon Lopez , Elizabeth) Tested with different test weights, NFF.
1054418 | 208499 | CM | CMPLT | CMPLT | 2024-07-31 14:01 | 2024-08-15 10:47 | Malfunction | (TC 07/31/2024 02:01:13 PM, Rendon Lopez , Elizabeth) Confirmed TOCO Malfunction alarm. Sending out for repair under warranty. (TC 08/15/2024 10:47:15 AM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1047075 | 208500 | CM | CMPLT | CMPLT | 2024-07-31 14:14 | 2024-08-15 14:41 | Doesn't pick up FHR | (TC 07/31/2024 02:15:07 PM, Rendon Lopez , Elizabeth) Weak distorted signal. To be sent out for repair. (TC 08/15/2024 02:41:07 PM, Rendon Lopez , Elizabeth) Incoming inspection-OK
1056472 | 208501 | CM | CMPLT | CMPLT | 2024-07-31 15:02 | 2024-07-31 15:11 | "Not working" | "not working." Cleaned unit. Battery depleted. Charged for a while. Checked functionality. No issue found. Tested ok. Returned ti service.
| 208502 | CM | CMPLT | CMPLT | 2024-07-31 15:42 | 2024-09-25 10:52 | MSH MEETINGS NAZ- SONOSITE PRCHASE OR AND L&D MSH OPG SPP>HSC | (TC 07/31/2024 03:43:22 PM, Whittington, Scott) MSH MEETINGS NAZ- SONOSITE PRCHASE OR AND L&D MSH OPG SPP>HSC | MSH MEETINGS NAZ- SONOSITE PRCHASE OR AND L&D MSH OPG SPP>HSC
65938 | 208503 | CM | CMPLT | CMPLT | 2024-07-31 15:47 | 2024-07-31 15:49 | Replace rear and front case | (TC 07/30/2024 03:47:10 PM, DeGuzman, Leila) Replaced rear and front case. Replaced pressure dic gasket. (TC 07/31/2024 03:48:47 PM, DeGuzman, Leila) After replacing the cases, perforrmed calibration and accuracy. Performed PM and closed PM work order.
65977 | 208504 | CM | CMPLT | CMPLT | 2024-07-31 15:53 | 2024-07-31 15:53 | Replace rear and front case | (TC 07/31/2024 03:53:03 PM, DeGuzman, Leila) Replaced rear and front case, gasket, disc holder, and latch. Calibrated and verified accuracy. Performed PM and close PM work order.
25981 | 208505 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26148 | 208506 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26153 | 208507 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26795 | 208508 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-10 14:34 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 09/10/2024 02:34:18 PM, Rhuland, Courtney) PM completed See attached doc for details
51581 | 208509 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
62651 | 208510 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
62969 | 208511 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS
62971 | 208512 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia
62981 | 208513 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-10-02 10:50 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/02/2024 10:49:57 AM, Alvarado, Karla) 1812.1 hours Intensity light: - High: 31.1 uW/cm2/nm - Low: 13.3 uW/cm2/nm Target light is functional GND Resistance: 0.182 ohms Leakage current: 62.9 uA PM completed. (TC 10/02/2024 10:49:57 AM, Rhuland, Courtney) Supervised
63073 | 208514 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
63348 | 208515 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-10 14:38 | PM-MAJOR - Annual PM For: Electrosurgical Units | (TC 09/10/2024 02:37:42 PM, Rhuland, Courtney) PM completed See attached doc for details
63484 | 208516 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
64066 | 208517 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
64087 | 208518 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
64488 | 208519 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
64489 | 208520 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
64491 | 208521 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
65262 | 208522 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-10 11:30 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Re-installed BIOS and cockpit only image. Importes configs. Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 27 Ah and 220 min. Tested. Functional. SW.2.51 Leakage: 93 uA and 91 ohm.
65305 | 208523 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
65387 | 208524 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 12:12 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced V-battery and air filter. Run battery check. 230 minutes, 56 Ah. Tested. Functional. SW2.51.01 Leakage: 95 uA and 93 ohm.
65885 | 208525 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
65993 | 208526 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-10 08:40 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | PM completed by BD.
66049 | 208527 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-01 07:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66053 | 208528 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-01 07:39 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66054 | 208529 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-01 07:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66063 | 208530 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-01 07:33 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
66457 | 208531 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-15 11:54 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced V-battery and air filter. Run battery check. 220 minutes, 47 Ah. Did not enable ventilator functions on C500. Re-installed bios and firmware. Tested. Functional. SW2.51 Leakage: 93 uA and 90 ohm.
66467 | 208532 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-10-18 10:45 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check, 220 min, 32 Ah. Tested. Functional. SW.2.51 Leakage: 85 uA and 92 ohm.
66519 | 208533 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66520 | 208534 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-14 09:28 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/14/2024 09:27:56 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66521 | 208535 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-14 09:29 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/14/2024 09:28:54 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66669 | 208536 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66670 | 208537 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66706 | 208538 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-20 10:52 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/20/2024 10:52:24 AM, Alvarado, Karla) VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 76% Set temperature to 35C, measured 35C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 84 uA; GND Resistance = 293 mohms PM completed. (TC 09/20/2024 10:52:24 AM, Baghirzade, Nara) Supervised
66793 | 208539 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66794 | 208540 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-11-07 07:51 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 11/07/2024 07:50:54 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
66796 | 208541 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-28 10:52 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/28/2024 10:52:09 AM, DeGuzman, Leila) PM passed and completed. See procedure for details.
66798 | 208542 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66800 | 208543 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-30 09:50 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/30/2024 09:49:38 AM, DeGuzman, Leila) PM passed and completed. See procedure for details.
66801 | 208544 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-13 08:46 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/13/2024 08:46:11 AM, DeGuzman, Leila) PM passed and completed. See procedure for details. *calibrated NIBP
66808 | 208545 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-23 14:03 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/23/2024 02:02:35 PM, DeGuzman, Leila) PM passed and completed. See procedure for details.
66809 | 208546 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-02 11:57 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/02/2024 11:57:00 AM, DeGuzman, Leila) PM passed and completed. See procedure for details.
66822 | 208547 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-02 11:19 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/02/2024 11:19:17 AM, DeGuzman, Leila) PM passed and completed. See procedure for details.
66830 | 208548 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66831 | 208549 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-14 09:32 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/14/2024 09:31:46 AM, DeGuzman, Leila) PM passed and completed. Calibrated NIBP. See procedure for details.
66841 | 208550 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-01 15:21 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/01/2024 03:21:26 PM, DeGuzman, Leila) PM completed. See procedure for details.
66844 | 208551 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66845 | 208552 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-11-07 10:09 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 11/07/2024 10:08:56 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
66853 | 208553 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-09 11:41 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/09/2024 11:41:08 AM, Alvarado, Karla) PM complete. All tests passed. See procedure for details.
66856 | 208554 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66859 | 208555 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-09 13:57 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/09/2024 01:56:47 PM, Alvarado, Karla) PM completed. All tests passed. See procedure for more details.
66860 | 208556 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66863 | 208557 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-26 08:59 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/26/2024 08:58:43 AM, Alvarado, Karla) PM completed All tests passed See procedure for details
67200 | 208558 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67203 | 208559 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67205 | 208560 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-14 15:56 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/14/2024 03:55:48 PM, DeGuzman, Leila) PM passed and completed. Calibrated NIBP See procedure for details.
67206 | 208561 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-12 15:32 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 09/12/2024 03:31:33 PM, Alvarado, Karla) PM completed All tests passed See procedure for details
67374 | 208562 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-10-22 14:33 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 10/22/2024 02:33:25 PM, Rhuland, Courtney) PM completed by Spectramedx 21% - 20.9% O2 40% - 38.3% O2 60% - 60.8% O2 80% - 80.2% O2 100% - 99.1% O2
67755 | 208563 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67931 | 208564 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-28 12:28 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/28/2024 12:28:13 PM, Rhuland, Courtney) PM completed Replaced: 9V rechargeable battery Red humidifier gasket Repaired 2 porthole doors Calibrated: MAINS & VREF voltages Scale RH Functional check - OK Set 35C, 75% humidity Read 35.1C, 76.2% humidity Electrical safety - OK Ground resistance - 60 mOhms Leakage current - 83.9 uA
67932 | 208565 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-23 09:54 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/23/2024 09:53:48 AM, DeGuzman, Leila) PM completed. Replaced scale, pins were broken. VRef, Mains V, and LV were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 75%. Set temp to 35C, measured 35C. Functional tests passed. Replaced battery. Electrical safety results - Leakage current = 10.5 uA Ground resistance = 116 mohms
67965 | 208566 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-25 13:27 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/25/2024 01:26:49 PM, Alvarado, Karla) Replaced battery, corner grommet, side wall latch, and scale. During calibration the system showed error #3. According to the manual, during ADC calibration, the software detected a channel out of range. Therefore, the control board was replaced. During the functional test the RH failed. Therefore, the RH sensor was replaced. Reseated scale pins to acknowledge scale. Performed PM: VRef, Mains V, LV, and RH were calibrated. Tilt, all switches and pedals are functional. Error log was cleared. Set scale to 5kg, measured 5kg. Set temperature to 35C, measured 35C on all temperature points. RH is functional and accurate; set 75%, analyzer = 76.6%. Functional tests passed. Electrical safety results - Leakage current = 84.5 uA; GND Resistance = 184 mohms PM completed. (TC 09/25/2024 01:26:49 PM, Rhuland, Courtney) Supervised
68034 | 208567 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-05 14:35 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/05/2024 02:35:17 PM, Rhuland, Courtney) PM completed Repaired exam light Calibrated: Scale VREF & MAINS voltages Functional check - OK Set 35C, 75% humidity Read 35.1C, 76.5% humidity Electrical safety - OK Ground resistance - 49 mOhms Leakage current - 85.5 uA
68657 | 208568 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-09 11:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:17:00 PM, Patel, Rutvik) by BD
68674 | 208569 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
68936 | 208570 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 14:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68975 | 208571 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 14:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69156 | 208572 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-09 12:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:26:21 PM, Patel, Rutvik) by BD
69171 | 208573 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 15:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69443 | 208574 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69476 | 208575 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-09 12:28 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:28:25 PM, Patel, Rutvik) by BD
69631 | 208576 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-11-07 13:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/07/2024 01:59:12 PM, Baghirzade, Nara) PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
530501 | 208577 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Integrated Power Console
67359 | 208578 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
67360 | 208579 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
67362 | 208580 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
32540 | 208581 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-11 14:18 | PM - Monthly | System was down and repair was being done instead of PM.
32560 | 208582 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-04 15:29 | PM - Monthly | PM completed Mechanical PM done sc-1148868 TOS phantom a(-991.1,10.6) b(332.9,13.3) c(133.6,13.2) d(94.0,12.3) e(-105.8,10.9) f(0.9,14.2) -passed all phantom tests
32580 | 208583 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-05 21:48 | PM - Monthly | PM cancelled due to Service work at TWH Emerg CT/TGH AQ 64 CT
500230 | 208584 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-10-24 14:26 | PM - Annual | (TC 10/24/2024 02:26:17 PM, Green, Paul) Performed a PM on Cobra. Cleaen, and checked functionality, ok.
1046059 | 208585 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1046273 | 208586 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-20 13:35 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/30/2024 09:58:00 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 (TC 09/04/2024 12:58:00 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM (TC 09/13/2024 11:17:33 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300
1046274 | 208587 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-18 15:36 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/30/2024 10:00:00 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 09/04/2024 01:19:20 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM (TC 09/13/2024 10:00:00 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046275 | 208588 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-18 15:37 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/30/2024 10:01:00 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 09/04/2024 12:59:00 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 09/04/2024 01:23:00 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 09/13/2024 11:18:38 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046276 | 208589 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-20 13:35 | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | (TC 08/30/2024 10:02:00 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 08/30/2024 01:02:00 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 09/04/2024 01:24:00 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL (TC 09/13/2024 11:18:55 AM, Whittington, Scott) MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL | MSH CFRH 7flr APLIO 300 PM, SEMI-ANNUAL, ULTRASOUND GENERAL
1046286 | 208590 | PM | ACTIV | READY | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1046397 | 208591 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1046399 | 208592 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1046451 | 208593 | PM | ACTIV | READY | 2024-08-01 03:20 |  | Annual PM
1047472 | 208594 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-20 14:08 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 1st year PM after 6th year PM last year. No GS500. It has PS500. Run battery check. 27 Ah and 230 min. Tested. Functional. SW.2.51.01 Leakage: 89 uA and 93 ohm.
1047829 | 208595 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Adult
1048472 | 208596 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-12 10:49 | HARP - Semi Annual
1048472 | 208597 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-12 10:48 | PM | (TC 09/12/2024 10:48:17 AM, Fan, Tak) PM completed
1048537 | 208598 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-27 21:03 | HARP - Semi-Annual | HARP
1048537 | 208599 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-26 21:04 | PM - Semi-Annual | PM was done.
1049885 | 208600 | PM | CANCL | CANCL | 2024-08-01 03:20 | 2024-09-23 10:50 | ANNUAL PM | PM to be completed every 2 years
1050088 | 208601 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Anesthesia Units
1051358 | 208602 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-10-02 09:08 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/02/2024 09:08:29 AM, Alvarado, Karla) 4814.1 hours Intensity light: - High: 30.1 uW/cm2/nm - Low: 12.8 uW/cm2/nm Target light is functional GND Resistance: 0.232 ohms Leakage current: 53.8 uA PM completed. (TC 10/02/2024 09:08:29 AM, Rhuland, Courtney) Supervised
1053348 | 208603 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-23 12:57 | MSH BI 1272 H3 SEPT 8 TRGT HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG | (TC 08/23/2024 12:56:36 PM, Dolinar, John) MSH BI 1272 H3 SEPT 8 TRGT HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG (TC 08/23/2024 12:56:36 PM, Dolinar, John) MSH BI 1272 H3 SEPT 8 TRGT HARP SCHEDULED SEMI-ANNUAL MSH_1272_MG
1055164 | 208604 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055168 | 208605 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 10:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1055178 | 208606 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1055773 | 208607 | PM | CANCL | CANCL | 2024-08-01 03:20 | 2024-09-24 07:54 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | PM completed in Feb - updated PM schedule
1055776 | 208608 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-18 09:14 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 09/18/2024 09:11:32 AM, Alvarado, Karla) New Battery. Battery Asset: 1057735 PM completed. All tests passed. See attached check list. (TC 09/18/2024 09:13:30 AM, Rendon Lopez , Elizabeth) Supervised
1055787 | 208609 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-19 15:52 | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS | (TC 09/19/2024 03:51:43 PM, Alvarado, Karla) New Battery. Battery Asset: 1057738 PM completed. All tests passed. See attached check list. (TC 09/19/2024 03:51:43 PM, Rendon Lopez , Elizabeth) Supervised
1055799 | 208610 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1055800 | 208611 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1056665 | 208612 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR_ Annual PM for: Microscope -Surgical
1056667 | 208613 | PM | ACTIV | PV | 2024-08-01 03:20 |  | PM-MAJOR-Annual PM for : Microscope- Opthalmology
1056911 | 208614 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR_ Annual PM for: Microscope -Surgical
1057133 | 208615 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM MAJOR - ANNUAL PM for: OR Tables
1057569 | 208616 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1058057 | 208617 | PM | ACTIV | ISSUE | 2024-08-01 03:20 |  | PM-MAJOR - Annual PM For: Anesthesia Units
1058224 | 208618 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1058276 | 208619 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 10:38 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1059795 | 208620 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-13 21:23 | HARP - Semi Annual | HARP is done.
1059795 | 208621 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-12 21:24 | PM - Semi-Annual | PM done.
1059796 | 208622 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-12 21:53 | HARP - Semiannual | HARP
1059796 | 208623 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-10 21:54 | PM - Semiannual | Done.
1059810 | 208624 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-06 20:13 | HARP - Semi-Annual | HARP
1059810 | 208625 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-09-05 21:14 | PM - Semi-Annual | PM Done.
1059875 | 208626 | PM | CMPLT | CMPLT | 2024-08-01 03:20 | 2024-08-16 10:50 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD.
1061335 | 208627 | PM | CMPLT | CMPLT | 2024-08-01 03:21 | 2024-09-05 14:10 | PM-MAJOR - Annual PM For: Scales, Infant | (TC 09/05/2024 02:10:24 PM, Rendon Lopez , Elizabeth) PM completed, passed.
1063416 | 208628 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | PM MAJOR - ANNUAL PM for: OR Tables
1064324 | 208629 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | PM-MAJOR - ANNUAL PM for: Insufflators
1064692 | 208630 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064693 | 208631 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064694 | 208632 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064695 | 208633 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064696 | 208634 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064700 | 208635 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064701 | 208636 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064702 | 208637 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064703 | 208638 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064704 | 208639 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064705 | 208640 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064706 | 208641 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064713 | 208642 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064714 | 208643 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064715 | 208644 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064716 | 208645 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064717 | 208646 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064718 | 208647 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064719 | 208648 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064720 | 208649 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064721 | 208650 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064722 | 208651 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064723 | 208652 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064724 | 208653 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064725 | 208654 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064726 | 208655 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064727 | 208656 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064728 | 208657 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064729 | 208658 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064730 | 208659 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
1064733 | 208660 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064734 | 208661 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual PM
1064810 | 208662 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:28 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064811 | 208663 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:28 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064812 | 208664 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:29 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064813 | 208665 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:29 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064814 | 208666 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:29 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064815 | 208667 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:30 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064816 | 208668 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:30 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064817 | 208669 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:30 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064818 | 208670 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:31 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064819 | 208671 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:31 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064820 | 208672 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:34 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064821 | 208673 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:38 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064822 | 208674 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:39 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064823 | 208675 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:39 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064824 | 208676 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:40 | PM Major Annual: Monitors, Physiologic, Patient Transport
1064825 | 208677 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:40 | PM Major Annual: Monitors, Physiologic, Patient Transport
1064826 | 208678 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:40 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064827 | 208679 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:41 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064828 | 208680 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:41 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064829 | 208681 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:41 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064830 | 208682 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:42 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064831 | 208683 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:46 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064832 | 208684 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:47 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064833 | 208685 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:47 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064834 | 208686 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:47 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064835 | 208687 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:48 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064836 | 208688 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:48 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064837 | 208689 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:49 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064838 | 208690 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:49 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064839 | 208691 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:49 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064840 | 208692 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:50 | PM Major Annual: Monitors, Physiologic, Patient Transport
1064841 | 208693 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:50 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064842 | 208694 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:50 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064843 | 208695 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:51 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064844 | 208696 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:51 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064845 | 208697 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:53 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064846 | 208698 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:53 | PM Major Annual: Monitors, Physiologic, Patient Transport
1064847 | 208699 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:55 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064848 | 208700 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:55 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064849 | 208701 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:57 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064850 | 208702 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-08-12 09:58 | PM Major Annual: Monitors, Physiologic, Patient Transport | Device only installed last year, battery replacement not required until 3 years
1064893 | 208703 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:09 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1064894 | 208704 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:13 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1065008 | 208705 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | PM-MAJOR - Annual PM For: Aspirators, Emergency
1065011 | 208706 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:01 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1065014 | 208707 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:14 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1065015 | 208708 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:16 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1065018 | 208709 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:17 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1065020 | 208710 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:18 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1065021 | 208711 | PM | CANCL | CANCL | 2024-08-01 03:21 | 2024-09-05 11:19 | PM-MAJOR - Annual PM For: Aspirators, Emergency | PM only required after 3 years
1065639 | 208712 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | PM Major - Annual PM for: ESU
1066122 | 208713 | PM | ACTIV | ISSUE | 2024-08-01 03:21 |  | Annual Maintenance
| 208714 | CM | CMPLT | CMPLT | 2024-08-01 07:16 | 2024-07-31 07:16 | PFT/Efficess meeting | (TC 07/31/2024 07:16:07 AM, Baghirzade, Nara) PFT/efficess meeting with Steve, Adeel, and Gaby
| 208715 | CM | CMPLT | CMPLT | 2024-08-01 07:22 | 2024-10-30 13:55 | Efficess exporting data for outlook | (TC 07/31/2024 07:21:52 AM, Baghirzade, Nara) Worked on efficess to export scheduling data to transfer to outlook. Was able to export the data as a csv. file and opens in excel. Will follow up with Adeel. (TC 08/14/2024 01:44:50 PM, Baghirzade, Nara) Trying different methods of copying the file over. Need admin rights to write any file over to an encrypted key; no one in PFT has the admin password as it was set up by someone at UHN during install. At first I was able to copy the csv file over however when viewed on my laptop, the excel file is blank despite being over 200 MB in size. Now when I try to copy the file over, the computer asks for admin permissions. Computer was being used so I was not able to restart it. Will try to restart it and try again. Researching other methods of transferring this file over. Computer does not have any wireless or bluetooth capabilities. (TC 08/28/2024 11:47:03 AM, Baghirzade, Nara) After researching online on how to copy a file without admin rights, tried a few more options: -saving and opening the file in wordpad; copying the wordpad document in administrator mode: did not work. -tried running command prompts for copying the file: also did not work. -tried to change the extension of the file to other types such as word, text, and other available excel extensions. Followed up with Kamran regarding this issue. (TC 09/16/2024 07:51:04 AM, Baghirzade, Nara) Went over to PFT with Kamran. Connected to the medisoft server on the secretary computer; exported the most recent calendar events, saved file locally and then copied to the server. On a different station, accessed the server and copied the file to a USB. Created a new test calendar on outlook and exported the file successfully. Will book a time with Gaby to go over color codes and structure of calendar before completely moving over to outlook. (TC 09/24/2024 08:54:32 AM, Baghirzade, Nara) Worked on the categories, mapping, and going over the structure of the appointments on outlook calendar. Will review which sections PFT would like to keep as some are irrelevant for the outlook calendar. Set up a meeting between Steve, Gaby, Adeel, John, and myself to discuss the structure and confirm that outlook is sufficient for them. Once approved, the calendar will be live and shared amongst RTs for PFT and CPET. (TC 09/25/2024 01:33:05 PM, Baghirzade, Nara) Meeting to finalize the calendar (TC 09/25/2024 01:55:31 PM, Baghirzade, Nara) Meeting with Gaby regarding outlook calendar and details. (TC 10/01/2024 02:14:52 PM, Baghirzade, Nara) Exported latest calendar. Had issues with sharing the calendar and permissions; created a new calendar however unable to share the contents without publishing online. Tried to edit the PFT MSH calendar but i was not able to export any data into it. Followed up with Steve, he changed my permission level to owner but it did not mitigate the issue. Still unable to upload anything to the calendar. Followed up with Adeel and agreed to upload it from Steve's account. Steve's account is directly linked to the PFT MSH email and calendar. Went with Steve to PFT to export the latest efficcess calendar as techs had made appointment changes. Exported the calendar and uploaded it to the PFT MSH calendar through his account. Went through the permissions and ensured everyone was added with the right level of permissions. Went over the color categories. Created test appointments and applied category changes to test if it reflects on a different computer; takes about 5 seconds for any change to be reflected on the calendar. Gaby to circulate the calendar amongst PFT admin staff. (TC 10/01/2024 02:20:07 PM, Baghirzade, Nara) Meeting with Gaby and Steve regarding calendar structure and feedback. (TC 10/30/2024 01:54:30 PM, Baghirzade, Nara) Received positive feedback. They will carry on using Outlook calendar for their appointment scheduling.
| 208716 | CM | CMPLT | CMPLT | 2024-08-01 07:36 | 2024-07-31 07:36 | NICU nitrogen leak meeting | (TC 07/31/2024 07:36:16 AM, Baghirzade, Nara) NICU nitrogen leak meeting with draeger
| 208717 | CM | CMPLT | CMPLT | 2024-08-01 07:44 | 2024-07-31 07:45 | Power bar for 11N PSN storage area | (TC 07/31/2024 07:44:58 AM, Baghirzade, Nara) Sarabeth requested a power bar for the area where the PSNs are stored; currently only 2 outlets on the wall. Brought up the power bar, set it up, and plugged all the PSNs in.
1054979 | 208718 | CM | CMPLT | CMPLT | 2024-08-01 07:47 | 2024-07-31 07:47 | Respiratory monitoring accessories | (TC 08/01/2024 07:46:45 AM, Baghirzade, Nara) Met with Sarabeth and she mentioned they are running very low on their resp monitoring supplies. Gave Sarabeth 2 boxes each of: -RAS-45 sensors -RAS-125C sensors -LNCS Neo SpO2 sensors (TC 07/31/2024 07:46:45 AM, Baghirzade, Nara) Met with Sarabeth and she mentioned they are running very low on their resp monitoring supplies. Gave Sarabeth 2 boxes each of: -RAS-45 sensors -RAS-125C sensors -LNCS Neo SpO2 sensors
| 208719 | CM | CMPLT | CMPLT | 2024-08-01 07:51 | 2024-07-30 07:51 | Managing old/open WOs for NICU, pumps, and my areas of responsibility | (TC 07/30/2024 07:50:52 AM, Baghirzade, Nara) Managing old/open WOs for NICU, pumps, and my areas of responsibility
| 208720 | CM | CMPLT | CMPLT | 2024-08-01 07:52 | 2024-07-31 07:52 | Tried to locate the 11N PSN | (TC 07/31/2024 07:52:03 AM, Baghirzade, Nara) Sarabeth mentioned that one of their PSNs is missing. Looked through all the rooms in 11N and 11S; was unable to locate PSN 1054987
| 208721 | CM | CMPLT | CMPLT | 2024-08-01 09:56 | 2024-08-20 14:13 | Formlabs Resin Tank Order
1061207 | 208722 | CM | CMPLT | CMPLT | 2024-08-01 10:37 | 2024-08-01 15:05 | Missing Corio Matrix Issues | (TC 08/01/2024 10:37:37 AM, Reid, Emmerson ) Video matrix was not found by VMC, I rebooted the VMC, then Corio, then VMC once more and it was working. I was called again about the audio not working. The speakers were not making any sound even though the connection was fine. I went into service mode and it started working, left service mode and it stopped. Went back in one more time to service mode, it worked, and left. This caused the VMC to quickly reboot. The audio was still working but video routing stopped working. I asked if they were okay with the video settings for this case and they were. Will go back after the case to try another reboot. (TC 08/01/2024 03:04:51 PM, Reid, Emmerson ) Rebooted the Corio another time and tested video and audio. Both were working. I also helped kill a fly that was in the room.
1063787 | 208723 | CM | CMPLT | CMPLT | 2024-08-01 11:01 | 2024-09-05 11:07 | Light does not turn on | Aug 01 - call re: not working. Went to check. Found light not turning on. Brought it down to biomed Contacted vendor for RMA. Received. Created SC and requested approval. Received. Issued PO via MMC. Aug 06 - Packed, scheduled and shipped the unit. Sept 05 - Received unit back from repair. Assembled. Went to OR. Checked functionality. Cleaned unit. Tested ok.
1059828 | 208724 | PM | CMPLT | CMPLT | 2024-08-01 11:13 | 2024-08-01 12:20 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/01/2024 12:20:17 PM, DeGuzman, Leila) Drained, disinfected, rinsed and refilled with sterile water
65562 | 208725 | CM | CMPLT | CMPLT | 2024-08-01 11:18 | 2024-08-01 11:19 | Bulb replacement | (TC 08/01/2024 11:19:11 AM, Rhuland, Courtney) Replaced bulb Showed clinical staff how to change bulb
67233 | 208726 | CM | CMPLT | CMPLT | 2024-08-01 12:33 | 2024-08-07 14:16 | Display remains blank, no working | (TC 07/31/2024 12:32:38 PM, Rendon Lopez , Elizabeth) Removed warmer from L&D, confirmed issue. (TC 08/01/2024 12:32:38 PM, Rendon Lopez , Elizabeth) Testing, control board defective, to be replaced (TC 08/06/2024 02:15:48 PM, Rendon Lopez , Elizabeth) Replaced control board kit with alarm board. Testing (TC 08/07/2024 02:15:48 PM, Rendon Lopez , Elizabeth) Calibrated Scale, tested all warmer functions, all passed. Returned to L&D as spare.
| 208727 | CM | CMPLT | CMPLT | 2024-08-01 13:18 | 2024-08-01 13:18 | Helping Kamran setup the PCU for RCC | (TC 08/01/2024 01:18:26 PM, Baghirzade, Nara) Helped Kamran setup the PCU network settings for RCC and ASM on biomed laptop
1063746 | 208728 | PM | CMPLT | CMPLT | 2024-08-01 13:26 | 2024-08-01 15:29 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/01/2024 03:28:42 PM, DeGuzman, Leila) Drained, disinfected, rinsed and refilled with sterile water
1067338 | 208729 | CM | CMPLT | CMPLT | 2024-08-01 13:29 | 2024-08-07 09:00 | O2 sensor failed. Check gas connections | During pre-use check, O2 sensor test says”O2 sensor failed. Check gas connections.” Gas connections are good. Ordering O2 sensor. Received O2 sensor. PO#7202440094
| 208730 | CM | CMPLT | CMPLT | 2024-08-01 13:31 | 2024-08-01 13:31 | Worked on recall responses | (TC 08/01/2024 01:31:04 PM, Baghirzade, Nara) Worked on recall responses on onerecall and notifications forwarded from Samir for my respective areas.
1049219 | 208731 | PM | CMPLT | CMPLT | 2024-08-01 14:38 | 2024-08-01 14:38 | PM-MAJOR - Annual PM For: Central Monitor | (TC 08/01/2024 02:38:18 PM, Scero, Tomlyn) PM complete during SW patch and reboot
1049220 | 208732 | PM | CMPLT | CMPLT | 2024-08-01 14:39 | 2024-08-01 14:39 | PM-MAJOR - Annual PM For: Central Monitor | (TC 08/01/2024 02:39:28 PM, Scero, Tomlyn) PM complete during SW patch and reboot
1061443 | 208733 | CM | ACTIV | ISSUE | 2024-08-01 15:02 |  | Left Railing buttons not working. | (TC 08/01/2024 03:06:42 PM, Ferreira, Qwynn) Went to the unit to troubleshoot the bed. Power cycled the bed and tested. The function sometimes works, but it is rather spotty when it does. Some damage is noted to the side and under panel that could have loosened a cable or damaged one. contacted stryker for repair (TC 08/20/2024 02:15:13 PM, Ferreira, Qwynn) Tech came to check on bed but bed was not located. I infomred the unit to set the bed aside if they locate the bed again.
32580 | 208734 | CM | CMPLT | CMPLT | 2024-08-01 15:06 | 2024-07-29 22:06 | No radiation no scan | Ordered and replaced RTM board in Rack-1, fixed the problem.
| 208735 | CON | CMPLT | CMPLT | 2024-08-01 21:41 | 2024-09-25 10:53 | MSH SITE COVERAGE | (TC 08/01/2024 09:40:49 PM, Heron, Chris) MSH SITE COVERAGE (TC 08/01/2024 09:40:49 PM, Heron, Chris) MSH SITE COVERAGE (TC 08/13/2024 09:03:14 PM, Heron, Chris) MSH SITE COVERAGE (TC 08/14/2024 03:10:43 PM, Heron, Chris) MSH SITE COVERAGE | MSH SITE COVERAGE
| 208736 | PA | ACTIV | ISSUE | 2024-08-01 21:45 |  | MSH DR REPLACEMENT PROJECT | (TC 08/01/2024 05:45:00 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT SIEMENS YSIO XPREE SITE VISIT GRAND RIVER HOSPITAL KITCHENER (TC 08/01/2024 05:46:00 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT SIEMENS YSIO XPREE SITE VISIT GRAND RIVER HOSPITAL KITCHENER (TC 08/09/2024 12:57:39 PM, Whittington, Scott) CANON VIRTUAL PRESENTATION FOR QUEENSWAY (TC 08/12/2024 02:41:38 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT (TC 08/12/2024 02:41:38 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT (TC 08/13/2024 09:00:00 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT AUGUST 13 CAMBRIDGE MEMORIAL PHILIPS (TC 08/13/2024 09:00:30 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT AUGUST 13 CAMBRIDGE MEMORIAL PHILIPS (TC 08/14/2024 02:52:48 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT AUGUST 14 QUEENSWAY CANON (TC 08/14/2024 02:53:48 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT AUGUST 14 QUEENSWAY CANON (TC 08/14/2024 02:53:48 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT AUGUST 14 QUEENSWAY CANON (TC 08/15/2024 06:52:00 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT AUGUST 15-16 HALIFAX GE (TC 08/16/2024 01:54:01 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT AUGUST 15-16 HALIFAX GE (TC 08/21/2024 11:51:03 AM, Whittington, Scott) MSH DR REPLACEMENT PROJECT PRODUCT REVIEW AND PROJECT MEETING (TC 08/29/2024 02:52:36 PM, Whittington, Scott) MSH DR REPLACEMENT PROJECT | MSH DR REPLACEMENT PROJECT AUGUST 1 SITE VISIT AUGUST 1 GRAND RIVER HOSPITAL KITCHENER SIEMENS YSIO XPREE AUGUST 13 CAMBRIDGE MEMORIAL PHILIPS DIGITAL DIAGNOST C90 AUGUST 14 QUEENSWAY CANON OMENARA 500 AUGUST 15-16 HALIFAX GE DEFINIUM 656 HD
| 208737 | PA | ACTIV | ISSUE | 2024-08-02 07:18 |  | Ordering parts for stock | (TC 08/02/2024 07:20:32 AM, Baghirzade, Nara) SC 1003283453 (TC 08/21/2024 07:36:44 AM, Baghirzade, Nara) Received 1/5 sensors. PO 7202440667
65910 | 208738 | PM | CMPLT | CMPLT | 2024-08-02 07:46 | 2024-08-02 07:46 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | (TC 08/02/2024 07:46:16 AM, Baghirzade, Nara) Replaced rear case. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration PM passed and completed. IUI connectors and device cleaned. Error and event logs downloaded and cleared. PM schedule updated.
1059778 | 208739 | FC | CMPLT | CMPLT | 2024-08-02 07:49 | 2024-08-02 07:49 | QC checklist | (TC 08/02/2024 07:49:23 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
| 208740 | OC | CMPLT | CMPLT | 2024-08-02 07:56 | 2024-08-08 08:48 | On-call week Aug 2, 2024 - Aug 8, 2024 | (TC 08/02/2024 07:56:34 AM, Canjimavadivel, Kamallesh ) daily check report (TC 08/06/2024 08:01:00 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 08/07/2024 11:23:28 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 08/08/2024 08:47:46 AM, Canjimavadivel, Kamallesh ) Daily check report
| 208741 | CM | CMPLT | CMPLT | 2024-08-02 08:40 | 2024-08-07 11:23 | X2 batteries for biomed stock
63163 | 208742 | CM | CMPLT | CMPLT | 2024-08-02 09:04 | 2024-08-02 09:04 | Defective LSU | (TC 08/02/2024 09:03:58 AM, Baghirzade, Nara) Unit is defective - unable to enter test mode. Replaced battery and tried to enter test mode, but still unable to. Checked with Steve to swap this unit with another LSU from the RT workshop. Will retire and dispose of this device.
62267 | 208743 | PM | CMPLT | CMPLT | 2024-08-02 09:08 | 2024-08-02 09:08 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 08/02/2024 09:08:17 AM, Baghirzade, Nara) PM completed. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg (TC 08/02/2024 09:08:28 AM, DeGuzman, Leila) PM completed. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
1054976 | 208744 | CM | ACTIV | ISSUE | 2024-08-02 09:13 |  | shipped to the vendor for repair
67818 | 208745 | PM | CMPLT | CMPLT | 2024-08-02 09:48 | 2024-08-02 10:35 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/02/2024 10:34:52 AM, DeGuzman, Leila) PM passed and completed. Calibrated NIBP See procedure for details.
1056125 | 208746 | PM | CMPLT | CMPLT | 2024-08-02 11:14 | 2024-08-02 11:15 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/02/2024 11:14:31 AM, Canjimavadivel, Kamallesh ) Cleaning cycle complete
| 208747 | RDS | CMPLT | CMPLT | 2024-08-02 12:18 | 2024-08-02 08:13 | Weekly Rounds: 11 North | (TC 08/02/2024 08:12:37 AM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
1048509 | 208748 | CM | CMPLT | CMPLT | 2024-08-02 13:58 | 2024-08-02 14:12 | MSH OR SONOSITE EDGE II LOW BATTERY LIFE | (TC 08/02/2024 01:58:21 PM, Dolinar, John) MSH OR SONOSITE EDGE II LOW BATTERY LIFE BATTERY FROM DONATED LOANER SWAPPED IN (TC 08/02/2024 02:12:15 PM, Whittington, Scott) MSH OR SONOSITE EDGE II LOW BATTERY LIFE BATTERY FROM DONATED LOANER SWAPPED IN | MSH OR SONOSITE EDGE II LOW BATTERY LIFE Replaced battery with the spare Edge in MSH office.(MIE 1062082)
1063331 | 208749 | CM | CMPLT | CMPLT | 2024-08-02 14:39 | 2024-09-25 10:53 | MSH SPP C3 CHEC HSC CONNECTION | (TC 08/02/2024 02:39:09 PM, Dolinar, John) MSH SPP C3 CHEC HSC CONNECTION (TC 08/08/2024 04:17:29 PM, Whittington, Scott) MSH SPP C3 CHEC HSC CONNECTION | MSH SPP C3 CHEC HSC CONNECTION
| 208750 | RDS | CMPLT | CMPLT | 2024-08-06 08:14 | 2024-08-02 08:14 | Weekly Rounds: 11 South | (TC 08/02/2024 08:14:22 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 208751 | RDS | CMPLT | CMPLT | 2024-08-06 08:16 | 2024-08-02 08:17 | Weekly Rounds: NICU & RESUS | (TC 08/02/2024 08:15:49 AM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
| 208752 | RDS | CMPLT | CMPLT | 2024-08-06 08:23 | 2024-08-06 14:24 | Weekly check for ED
| 208753 | RDS | CMPLT | CMPLT | 2024-08-06 08:23 | 2024-08-27 13:56 | Weekly check for ICU
| 208754 | RDS | CMPLT | CMPLT | 2024-08-06 08:36 | 2024-08-06 08:37 | Weekly Checks For PACU | (TC 08/06/2024 08:37:22 AM, Reid, Emmerson ) Weekly rounds complete.
| 208755 | RDS | CMPLT | CMPLT | 2024-08-06 08:39 | 2024-08-06 08:40 | Weekly Checks For OR | (TC 08/06/2024 08:40:24 AM, Reid, Emmerson ) Rounds complete.
69099 | 208756 | PM | CMPLT | CMPLT | 2024-08-06 09:27 | 2024-08-06 09:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/06/2024 09:27:12 AM, Rhuland, Courtney) PM completed All tests passed Refer to ASM for details
1054522 | 208757 | CM | CMPLT | CMPLT | 2024-08-06 09:28 | 2024-08-06 09:28 | Intermittently turns off | (TC 08/06/2024 09:28:03 AM, Canjimavadivel, Kamallesh ) Disabled auto shutoff and adjusted standby time from 5 to 10 min. Replaced damaged LL and LA leads.
1059764 | 208758 | CM | CMPLT | CMPLT | 2024-08-06 09:42 | 2024-09-25 10:54 | MSH 15 L&D TE7 CARDIAC PROBE QUOTE REQUEST | (TC 08/06/2024 09:42:08 AM, Whittington, Scott) MSH 15 L&D TE7 CARDIAC PROBE QUOTE REQUEST | MSH 15 L&D TE7 CARDIAC PROBE QUOTE REQUEST From: Whittington, Scott Sent: August 6, 2024 9:40 AM To: Lisa Petteplace <lpetteplace@bomimed.com>; Ngo, Kim - Sinai Health System <Kim.Ngo@sinaihealth.ca> Cc: NAZ Bagherimehrbataghva, Nazanin - Sinai Health System (Nazanin.Bagherimehrbataghva@sinaihealth.ca) <Nazanin.Bagherimehrbataghva@sinaihealth.ca>; Pejcic, Sonja <Sonja.Pejcic@uhn.ca> Subject: MSH L&D TE-7 cardiac probe inquiry Re: [External] Mindray Hi These are the MSH L&D TE-7 Serial #’s CD7-8C000896 CD7-8C000897 CD7-8C000912 CD7-11004055 CD7-11004051 Thanks, Scott From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: August 6, 2024 7:32 AM To: Lisa Petteplace <lpetteplace@bomimed.com>; Ngo, Kim - Sinai Health System <Kim.Ngo@sinaihealth.ca> Subject: Re: MSH L&D TE-7 cardiac probe inquiry Re: [External] Mindray Hi Lisa I can provide a list of the TE-7 sn# from our database Thanks Scott Get Outlook for iOS ________________________________________ From: Lisa Petteplace <lpetteplace@bomimed.com> Sent: Tuesday, August 6, 2024 7:30:33 AM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Ngo, Kim - Sinai Health System <Kim.Ngo@sinaihealth.ca> Subject: Re: MSH L&D TE-7 cardiac probe inquiry Re: [External] Mindray Hi Kim There is a cardiac transducer that you could purchase. What machine do you currently have? Model and serial number would be great. That way I can assess which transducers are compatible. Regards Lisa Petteplace RDMS, CMRITO, CRGS ULTRASOUND REPRESENTATIVE 1-100 IRENE STREET WINNIPEG, MB R3T 4E1 C. 226.749.2926 T. 800.667.6276 F: 877.435.6984 www.bomimed.ca Sent from my Bell Samsung device over Canada’s largest network. ________________________________________ From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Tuesday, August 6, 2024 7:28:02 AM To: Ngo, Kim - Sinai Health System <Kim.Ngo@sinaihealth.ca> Cc: Lisa Petteplace <lpetteplace@bomimed.com>; Michelle Lee <MLee@bomimed.com> Subject: MSH L&D TE-7 cardiac probe inquiry Re: [External] Mindray CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Kim I have copied Lisa & Michelle of Bomimed/ Mindray who will be able to provide info & a quote Thanks, Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Ngo, Kim <Kim.Ngo@sinaihealth.ca> Sent: Tuesday, August 6, 2024 7:16 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] Mindray Good Morning Scott, I was wondering if you knew for the Mindray Ultrasound Machines that we have on L&D if there is a probe we could purchase to do cardiac ultrasounds? Thanks, Kim Kim Ngo, RN, BScN Clinical Co-ordinator, Labour and Delivery T 416-586-4800 x 5450 kim.ngo@sinaihealth.ca 600 University Avenue, Toronto, ON M5G 1X5
1071743 | 208759 | ACC | CMPLT | CMPLT | 2024-08-06 10:08 | 2024-08-06 10:08 | Incoming inspection | (TC 08/06/2024 10:07:55 AM, Rendon Lopez , Elizabeth) Incoming inspection and asset tagging. Delivered to AA for stocking.
1048472 | 208760 | CM | CMPLT | CMPLT | 2024-08-06 10:40 | 2024-08-06 10:40 | MSH R59 - 15th Floor - No Wifi As User Stated | (TC 08/06/2024 10:40:06 AM, Chau, Ken) User stated no wifi and worklist. Went to look at system and had full connection to network, cant find issue Spoke with users and they stated not working, wireless but when they showed the tether cable they were talking about the detector The detector wasn't selected as generator mode was selected Informed user to pick detector Performed Calibration on detector as dropped
| 208761 | RDS | CMPLT | CMPLT | 2024-08-06 11:00 | 2024-08-06 11:02 | Weekly checks in CCU | (TC 08/06/2024 11:01:46 AM, Scero, Tomlyn) Weekly checks complete. 30J tests passed.
| 208762 | PA | CMPLT | CMPLT | 2024-08-06 12:26 | 2024-08-12 12:44 | Ordering CNS-EEG-3XX Amplifier for NICU Research | (TC 08/06/2024 01:57:21 PM, Baghirzade, Nara) SC 1003284923 (TC 08/12/2024 12:44:09 PM, Baghirzade, Nara) PO 7202441378
65949 | 208763 | CM | CMPLT | CMPLT | 2024-08-06 12:36 | 2024-08-06 12:36 | Syringe calibration required when connecting to PCU | (TC 08/06/2024 12:36:13 PM, Baghirzade, Nara) Connected syringe to PCU on normal mode; immediate error message appears for syringe calibration. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. Connected syringe to PCU on normal mode; no errors messages observed.
65925 | 208764 | CM | CMPLT | CMPLT | 2024-08-06 12:38 | 2024-08-06 12:38 | Syringe calibration required when connecting to PCU | (TC 08/06/2024 12:38:05 PM, Baghirzade, Nara) Connected syringe to PCU on normal mode; immediate error message appears for syringe calibration. Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. Connected syringe to PCU on normal mode; no errors messages observed.
1065087 | 208765 | CM | CMPLT | CMPLT | 2024-08-06 12:39 | 2024-08-06 12:39 | Occlusion in line when there is no occlusion | (TC 08/06/2024 12:39:18 PM, Baghirzade, Nara) Downloaded error and event log however did not note any significant errors. Completed an analog sensor test; found the bottle-side pressure sensor to be defective. Replaced bottle-side pressure sensor and completed both rate and pressure calibrations. Both calibrations were completed successfully. Started the check-in tests and when it reached patient occlusion test, the line was occluding immediately. Tried several times however kept getting the same result. As per manual, the patient-side pressure sensor will need to be replaced. Replaced the patient-side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
1071749 | 208766 | ACC | CMPLT | CMPLT | 2024-08-06 13:24 | 2024-07-31 15:53 | Incoming inspection | (TC 07/31/2024 03:52:47 PM, Cortez, John) Unboxed, no issues with visual/functional check, Accuracy check with phantom bladder passed, ~135ml
| 208767 | CM | CMPLT | CMPLT | 2024-08-06 14:14 | 2024-08-06 14:14 | Meeting with Mykayla | (TC 08/06/2024 02:13:42 PM, Baghirzade, Nara) Meeting with new CNS (Mykayla) for 11South
1071753 | 208768 | ACC | CMPLT | CMPLT | 2024-08-06 15:34 | 2024-07-30 16:01 | Incoming inspection | (TC 07/30/2024 04:01:07 PM, Cortez, John) Unboxed, visual/functional check passed
1071754 | 208769 | ACC | CMPLT | CMPLT | 2024-08-06 15:37 | 2024-07-30 16:04 | Incoming inspection | (TC 07/30/2024 04:03:38 PM, Cortez, John) Unboxed, visual/functional check passed
1064512 | 208770 | CM | ACTIV | ISSUE | 2024-08-06 15:40 |  | ERBE Vio 3 ESU Training - Self Study - PM Parts Purchase | Aug 6 - Requested service manual from AMT Surgical. Was infomed that they cannot provide as I am not trained. Requested ER to request SM from AMT Surgical. Received. Started studying. Oct 29 - Verified all parts in PM kit. Found several are missing. Contacted vendor and requested parts info. Oct 31 - Follwed up. Was informed sales rep will visit for discussion. Met with sales rep and discussed the parts needed. Was informed that he will connect with tech supporet and provide info. Nov 01 - Info received. Created SC. Sent for approval. Received. Placed order via MMC. Nov 26 - Followed up. Awaiting response.
1071755 | 208771 | ACC | CMPLT | CMPLT | 2024-08-06 15:42 | 2024-07-30 16:04 | Incoming inspection | (TC 07/30/2024 04:04:21 PM, Cortez, John) Unboxed, visual/functional check passed
1071756 | 208772 | ACC | CMPLT | CMPLT | 2024-08-06 15:46 | 2024-07-30 16:05 | Incoming inspection | (TC 07/30/2024 04:05:16 PM, Cortez, John) Unboxed, visual/functional check passed
1071757 | 208773 | ACC | CMPLT | CMPLT | 2024-08-06 15:48 | 2024-07-31 16:06 | Incoming inspection | (TC 07/31/2024 04:06:09 PM, Cortez, John) Unboxed, visual/functional check passed
1071758 | 208774 | ACC | CMPLT | CMPLT | 2024-08-06 15:50 | 2024-07-31 16:07 | Incoming inspection | (TC 07/31/2024 04:06:50 PM, Cortez, John) Unboxed, visual/functional check passed
1071759 | 208775 | ACC | CMPLT | CMPLT | 2024-08-06 15:51 | 2024-07-31 16:08 | Incoming inspection | (TC 07/31/2024 04:08:08 PM, Cortez, John) Unboxed, visual/functional check passed
1071760 | 208776 | ACC | CMPLT | CMPLT | 2024-08-06 15:53 | 2024-07-31 16:09 | Incoming inspection | (TC 07/31/2024 04:09:26 PM, Cortez, John) Unboxed, visual/functional check passed
1071761 | 208777 | ACC | CMPLT | CMPLT | 2024-08-06 16:09 | 2024-08-02 16:12 | Incoming inspection | (TC 08/02/2024 04:11:36 PM, Cortez, John) Unboxed, visual/functional check passed Set up on dashboard, manually set report via wi-fi
| 208778 | CM | CMPLT | CMPLT | 2024-08-06 16:39 | 2024-08-06 16:39 | Site Coverage.
1051615 | 208779 | CM | CMPLT | CMPLT | 2024-08-07 10:23 | 2024-08-07 10:35 | Ground resistance high | (TC 08/07/2024 10:23:14 AM, Chukri, Samir) Contact and email Peter from Getinge . (TC 08/07/2024 10:35:08 AM, Chukri, Samir) Tested befor the email to Peter working fine.
68022 | 208780 | CM | CMPLT | CMPLT | 2024-08-07 11:15 | 2024-08-07 11:15 | O2 presure regulation failed | Tested. passed pre-use cehck 3 times. informed Rex. returned back to NICU RT
68685 | 208781 | CM | CMPLT | CMPLT | 2024-08-07 12:39 | 2024-08-07 12:40 | Network reconfiguration and functional check | (TC 08/07/2024 12:39:34 PM, Baghirzade, Nara) PCU was used to test in RCC/testing that PCU connects to RCC wifi. Kamran tested in RRC and dropped off the PCU back to MSH. Transferred current MSH network configuration. Turned on the PCU in normal mode and confirmed that SSID is mshmed and PCU connected to the alaris server successfully. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 166 mohms Leakage current: 76.8 uA
1066462 | 208782 | FC | CMPLT | CMPLT | 2024-08-07 12:43 | 2024-08-07 12:44 | QC checklist | (TC 08/07/2024 12:44:06 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
66711 | 208783 | PM | CMPLT | PMFAIL | 2024-08-07 12:47 | 2024-08-07 12:47 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/07/2024 12:47:13 PM, Baghirzade, Nara) Incubator originally was sent down to biomed for door issue; the inner panel would not hold on the hinges. Hinges were very worn out. Swapped the door with a spare one in biomed. Replaced power cord. Completed all PM steps however when running the incubator for functional check, the RH % was reading very high. Calibrated RH twice however it still kept reading high. Will have to replaced the RH sensor and run a functional check again.
66711 | 208784 | CM | CMPLT | CMPLT | 2024-08-07 12:49 | 2024-08-07 12:50 | Created by Failed PM WO # 208783 : PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/07/2024 12:50:03 PM, Baghirzade, Nara) Replaced RH sensor. PM completed. 9V rechargeable battery was replaced. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 62.5 uA Power cord resistance = 77 mohms | (TC 08/07/2024 12:47:13 PM, Baghirzade, Nara) Incubator originally was sent down to biomed for door issue; the inner panel would not hold on the hinges. Hinges were very worn out. Swapped the door with a spare one in biomed. Completed all PM steps however when running the incubator for functional check, the RH % was reading very high. Calibrated RH twice however it still kept reading high. Will have to replaced the RH sensor and run a functional check again.
| 208785 | CM | CMPLT | CMPLT | 2024-08-07 12:53 | 2024-08-07 12:53 | Meeting with Becky | (TC 08/07/2024 12:53:02 PM, Baghirzade, Nara) Met with Becky to discuss how to manage the Masimo device in PFT in regards to repairs, spares, etc.
68862 | 208786 | CM | CMPLT | CMPLT | 2024-08-07 12:55 | 2024-08-07 12:55 | Functional check | (TC 08/07/2024 12:54:46 PM, Baghirzade, Nara) Biomed tag did not have any error description. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 136 mohms Leakage current: 68.5 uA
68807 | 208787 | CM | CMPLT | CMPLT | 2024-08-07 12:57 | 2024-08-07 12:57 | User reported that it keeps beeping even when plugged in | (TC 08/07/2024 12:56:38 PM, Baghirzade, Nara) Power cord was defective; when plugged in the wall, unit operates only on battery. Replaced power cord. Downloaded error logs; did not note any significant errors. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 123 mohms Leakage current: 77.8 uA
1065938 | 208788 | CM | CMPLT | CMPLT | 2024-08-07 13:00 | 2024-08-07 13:00 | No issue/error stated on yellow biomed tag | (TC 08/07/2024 01:00:17 PM, Baghirzade, Nara) Downloaded error log; did not note any significant errors. Check-in tests passed and completed. IUI connectors and device cleaned.
| 208789 | CM | CMPLT | CMPLT | 2024-08-07 13:03 | 2024-08-07 13:03 | Checking fluke safety analyzers | (TC 08/07/2024 01:03:04 PM, Baghirzade, Nara) Samir requested to test both fluke analyzers. Testing both safety analyzers as one of them was giving a high resistance reading (more than 1 ohm). Tested both and swapped the power cords around but could not get a high reading on either. Followed up with Samir.
501070 | 208790 | CM | CMPLT | CMPLT | 2024-08-07 13:08 | 2024-08-07 13:15 | MSH CEOU ROOM 10 SYSTEM CRASHED AND WOULD NOT REBOOT . FINALLY REBOOT INTO ERROR 227 | (TC 08/07/2024 01:12:33 PM, Whittington, Scott) MSH CEOU ROOM 10 SYSTEM CRASHED AND WOULD NOT REBOOT . FINALLY REBOOT INTO ERROR 227 SEVERAL REBOOTS AND STARTS CLEAN REMOVED PT DB, WHICH HAD BEEN AT 25% | MSH CEOU ROOM 10 SYSTEM CRASHED AND WOULD NOT REBOOT . FINALLY REBOOT INTO ERROR 227
1055792 | 208791 | CM | CMPLT | CMPLT | 2024-08-07 13:46 | 2024-08-07 13:47 | Status check - Daily on call check | Status message "Plun in AC. Went to OR. Searched for the defib. Plug the AC cord. Tested ok.
1062547 | 208792 | CM | CMPLT | CMPLT | 2024-08-07 14:08 | 2024-08-07 14:08 | Bed relocation - missing from 12S | (TC 08/07/2024 02:07:58 PM, Ferreira, Qwynn) functional check - Pass Brought from 6th floor to 12S. No patient elevatros for 12S
| 208793 | CM | CMPLT | CMPLT | 2024-08-07 14:19 | 2024-08-20 14:37 | Panda warmer spare part | (TC 08/07/2024 02:34:36 PM, Rendon Lopez , Elizabeth) Ordering spare Panda warmer control board kit (TC 08/20/2024 02:36:54 PM, Rendon Lopez , Elizabeth) Received
1054976 | 208794 | CM | CMPLT | CMPLT | 2024-08-07 14:38 | 2024-08-07 14:38 | Incoming inspection and functional check | (TC 08/07/2024 02:37:49 PM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1059551 | 208795 | CM | CMPLT | CMPLT | 2024-08-07 14:38 | 2024-08-06 14:39 | Patch install | (TC 08/06/2024 02:38:38 PM, Rendon Lopez , Elizabeth) Rebooted central. Patch installation completed successfully
1062067 | 208796 | CM | CMPLT | CMPLT | 2024-08-07 14:40 | 2024-08-06 14:42 | Bed not working at all | (TC 08/06/2024 02:41:40 PM, Rendon Lopez , Elizabeth) Checked, bed Error 3009-6-5-1503 - Limited Motion Functionality. Power cycle bed, clear error logs, checked all functions, all working properly. Notified staff to swap bed if error reoccurs.
| 208797 | CM | ACTIV | ISSUE | 2024-08-07 14:49 |  | WatchBP installation | (TC 08/07/2024 02:49:06 PM, Rhuland, Courtney) Unboxed new devices Created batch asset template for new assets to be created Confirmed roll-out plan with Becky and Tomi (TC 08/09/2024 11:44:17 AM, Rhuland, Courtney) Tagged and unboxed all devices (TC 08/12/2024 11:19:11 AM, Rhuland, Courtney) Assembled 2 rolling stands and installed device (TC 08/27/2024 08:35:46 AM, Rhuland, Courtney) Issue with multiple devices throwing E3 errors Tested all devices and noted which devices were defective Working with vendor to return defective devices (TC 10/22/2024 11:15:44 AM, Rhuland, Courtney) Received replacement shipment of 4 watchBPs Unboxed and tested devices - all OK Updated asset tags, SN and warranty dates Waiting for clinic to have wall mounts installed (TC 10/28/2024 02:51:37 PM, Rhuland, Courtney) Brought over devices to 60 Murray to mount Screws needed had been discarded Found screws that will work with bracket To return this week to mount devices in rooms (TC 11/01/2024 10:32:34 AM, Rhuland, Courtney) Installed all devices on wall mounts Functional check - OK Demo'd devices to clinical staff and programmed all devices to desired settings Removed all BP Tru devices and stored in clinic managers office (TC 11/05/2024 09:47:44 AM, Rhuland, Courtney) Demo'd device for clinical staff
1055576 | 208798 | RQ | PENDG | REQST | 2024-08-07 15:26 |  | this is a test don't, respond to this WR
62703 | 208799 | FC | CMPLT | CMPLT | 2024-08-07 16:16 | 2024-08-02 16:31 | Set up for RCC location | (TC 08/02/2024 04:30:48 PM, Cortez, John) Cleaned, Replaced screen with spare, calibrated, Added spare PIM module, added leads Verified that battery health is still ok Functional check with simulator ok
| 208800 | CM | CMPLT | CMPLT | 2024-08-07 16:35 | 2024-08-30 08:57 | misc biomed - Aug 2024 | (TC 08/06/2024 04:35:04 PM, Cortez, John) RCC meeting with westpark biomed staff (TC 08/06/2024 04:38:33 PM, Cortez, John) RPCU called about "batt malfunction" alert in angio 2. Batteries already on order, still backordered. Informed Matt (temp supervisor) (TC 08/07/2024 04:40:09 PM, Cortez, John) Helped movers with RCC equipment (TC 08/07/2024 04:40:35 PM, Cortez, John) grabbed 4 PCUs for RCC from ICU (TC 08/06/2024 04:41:38 PM, Cortez, John) Tas from 12S reached out about needing bariatric bed. Found 12S bariatric bed in 6th floor storage. Informed Tas and Qwynn (TC 08/08/2024 02:41:50 PM, Cortez, John) travel to RCC (TC 08/08/2024 02:42:18 PM, Cortez, John) configured wi-fi for AED, ECG cart, delivered pumps, swapped aspirator (TC 08/09/2024 02:48:06 PM, Cortez, John) worked on philips/systems quiz (TC 08/12/2024 03:52:36 PM, Cortez, John) Installed prints to secure power cord on all vtal signs + bladder scanner (RCC) (TC 08/12/2024 03:54:24 PM, Cortez, John) Checked settings on AED (RCC) (TC 08/14/2024 04:49:07 PM, Cortez, John) Masimo 16N1 client down issue, Can't ping server Confirmed ethernet cable is ok, datadrop is correct. Couldn't find switch, should be switch2 port 44 in 16M stack Kamran reached out to Mirza, they resolved issue on their end. (TC 08/14/2024 04:52:21 PM, DeGuzman, Leila) Helped with Masimo 16N1 issue (TC 08/14/2024 04:53:19 PM, Cortez, John) systems meeting (TC 08/14/2024 04:53:55 PM, Cortez, John) Replaced batteries on philips monitors in RPCU and angio 2 room (TC 08/22/2024 04:09:58 PM, Cortez, John) Focal point alert, switch unreachable in 1st floor MCR Kamran and I went to check. Issue with outlet where are UPSes are plugged in. Power exists on other emergency power outlets however. UPSes moved over. Monitoring came back in NICU, CCU, L&D Reached out to building services, who say someone may have been working on that outlet. Will continue to monitor (TC 08/23/2024 04:19:47 PM, Cortez, John) Called for 2 WOWs on 10S. Restarting one cart resolves issue, other cart still slow. Asked ward clerk to call HELP desk (TC 08/30/2024 08:57:25 AM, Cortez, John) chs meeting
| 208801 | PA | CMPLT | CMPLT | 2024-08-08 06:47 | 2024-08-12 12:43 | Ordering sterile water for NICU Altrix | (TC 08/08/2024 06:50:54 AM, Baghirzade, Nara) SC 1003286018 (TC 08/12/2024 12:43:08 PM, Baghirzade, Nara) Received. Order #4150281511
1059618 | 208802 | CM | CMPLT | CMPLT | 2024-08-08 08:27 | 2024-08-08 08:27 | Not turning on | (TC 08/08/2024 08:26:55 AM, Reid, Emmerson ) The light was not turning on. I saw the grounding pin was broken. I replaced the power cord and it is working now.
68926 | 208803 | CM | CMPLT | CMPLT | 2024-08-08 09:39 | 2024-08-08 10:42 | Plug not working | (TC 08/08/2024 10:42:12 AM, DeGuzman, Leila) Came down issue was plug not working. Check the power cord and plug, looks fine. Battery low, charged for about an hour. Performed check-in test all passed. Ground resistance: 101 mOhm Leakage current: 7.1 uA
1053389 | 208804 | CM | CMPLT | CMPLT | 2024-08-08 09:47 | 2024-08-08 09:48 | Blank screen with beeping sound. | (TC 08/08/2024 09:48:18 AM, DeGuzman, Leila) Came down with an issue blank screen with beeping sound even with new batteries Tried with new probe and calibration plug still beeping. Retiring from service.
1071774 | 208805 | CM | CMPLT | CMPLT | 2024-08-08 10:04 | 2024-08-08 10:05 | Asset number and return to 12 South | (TC 08/08/2024 10:04:49 AM, DeGuzman, Leila) Assigned asset number and return to 12 south for the replacement of retired temp asset#: 1053389.
1071432 | 208806 | CM | CMPLT | CMPLT | 2024-08-08 10:09 | 2024-08-08 10:48 | Turning Off | Call re: not working. Went to check. Spoke with staff and was informed that the unit automatically turns off after sometime. Brought it down to biomed. Checked functionality. Ran unit for an hour. Could not replicate the issue. Tested ok. Delviered back to OR.
67891 | 208807 | CM | CMPLT | CMPLT | 2024-08-08 10:27 | 2024-08-08 15:16 | Alarms and says occlusion detected | (TC 08/07/2024 10:28:11 AM, Rendon Lopez , Elizabeth) Running with same user settings, unable to replicate issue, no errors (TC 08/08/2024 03:15:51 PM, Rendon Lopez , Elizabeth) Upgraded SW and library. Ran on user mode for hours, no fault found. Performance assurance-OK
1071775 | 208808 | ACC | CMPLT | CMPLT | 2024-08-08 10:40 | 2024-08-08 11:29 | Incoming inspection | (TC 08/08/2024 11:29:09 AM, Canjimavadivel, Kamallesh ) Assembled scale and added to database. Functional test pass
1054599 | 208809 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:36 | PIIC iX OS Patch | (TC 08/09/2024 09:35:49 AM, Rhuland, Courtney) Patch successfully completed
1059330 | 208810 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:39 | PIIC iX OS Patch | (TC 08/09/2024 09:39:11 AM, Rhuland, Courtney) Patch successfully completed
1059541 | 208811 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:33 | PIIC iX OS Patch | (TC 08/09/2024 09:33:23 AM, Rhuland, Courtney) Patch successfully completed
1059542 | 208812 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-08 10:57 | PIIC iX OS Patch | (TC 08/08/2024 10:57:20 AM, Rhuland, Courtney) Patch successfully completed
1059543 | 208813 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-08 10:59 | PIIC iX OS Patch | (TC 08/08/2024 10:58:37 AM, Rhuland, Courtney) Patch successfully completed
1059545 | 208814 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-08 11:06 | PIIC iX OS Patch | (TC 08/08/2024 11:06:06 AM, Rhuland, Courtney) Patch successfully completed
1059546 | 208815 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:37 | PIIC iX OS Patch | (TC 08/09/2024 09:36:58 AM, Rhuland, Courtney) Patch successfully completed
1059547 | 208816 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-08 10:56 | PIIC iX OS Patch | (TC 08/08/2024 10:56:04 AM, Rhuland, Courtney) Patch successfully completed
1059548 | 208817 | CM | CANCL | CANCL | 2024-08-08 10:54 | 2024-09-26 10:04 | PIIC iX OS Patch
1059550 | 208818 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:36 | PIIC iX OS Patch | (TC 08/09/2024 09:36:16 AM, Rhuland, Courtney) Patch successfully completed
1059552 | 208819 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-08 11:02 | PIIC iX OS Patch | (TC 08/08/2024 11:02:12 AM, Rhuland, Courtney) Patch successfully completed
1059553 | 208820 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-08 11:03 | PIIC iX OS Patch | (TC 08/08/2024 11:03:04 AM, Rhuland, Courtney) Patch successfully completed
1059554 | 208821 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:33 | PIIC iX OS Patch | (TC 08/09/2024 09:32:50 AM, Rhuland, Courtney) Patch successfully completed
1059555 | 208822 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-08 11:09 | PIIC iX OS Patch | (TC 08/08/2024 11:09:14 AM, Rhuland, Courtney) Patch successfully completed
1059556 | 208823 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:25 | PIIC iX OS Patch | (TC 08/09/2024 09:25:20 AM, Rhuland, Courtney) Patch successfully completed
1059557 | 208824 | CM | CMPLT | CMPLT | 2024-08-08 10:54 | 2024-08-09 09:38 | PIIC iX OS Patch | (TC 08/09/2024 09:38:17 AM, Rhuland, Courtney) Patch successfully completed
1057949 | 208825 | CM | CMPLT | CMPLT | 2024-08-08 10:57 | 2024-08-07 16:13 | "Not working" | call re: remote not working. Went to check. Waited for access. Patient on the table. Could not check the remote. Borrowed new from desk and installed it. Went to OR 12 and checked full functionality of the remote on spare OR Table. COuld not find any issue. Tested ok. Returned to service.
1055913 | 208826 | CM | CMPLT | CMPLT | 2024-08-08 11:19 | 2024-08-19 09:47 | Broken bottom mounted plate | (TC 08/08/2024 11:19:24 AM, DeGuzman, Leila) Informed John about the issue and ordering the part. Waiting for part... (TC 08/19/2024 09:47:11 AM, DeGuzman, Leila) Parts came in, installed the plate. Checked functions, all okay. Returned to service.
1061333 | 208827 | FC | CMPLT | CMPLT | 2024-08-08 11:56 | 2024-08-08 11:56 | QC checklist | (TC 08/08/2024 11:55:50 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
1054976 | 208828 | CM | CMPLT | CMPLT | 2024-08-08 11:57 | 2024-09-25 09:59 | Borrowed for PFT lab | (TC 08/08/2024 11:57:02 AM, Baghirzade, Nara) PFT's device is not functional and has to be sent to the vendor; PFT does not have a backup Masimo. Borrowed this device for PFT while their device is being repaired and will return to MBU once device is received. Transferred all accessories from PFT root to MBU root. Followed up with Jessica and Gaby. (TC 09/25/2024 09:58:36 AM, Baghirzade, Nara) Picked up the device from PFT. Completed functional testing. Returned Masimo to MBU.
1054977 | 208829 | CM | CMPLT | CMPLT | 2024-08-08 11:59 | 2024-08-08 11:59 | User did not report an issue/error | (TC 08/08/2024 11:59:11 AM, Baghirzade, Nara) Did not find any functional issue with the device. MSH config was not on the Root. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1065889 | 208830 | CM | CMPLT | CMPLT | 2024-08-08 12:01 | 2024-08-08 12:01 | Occlusion error | (TC 08/08/2024 12:01:21 PM, Baghirzade, Nara) Downloaded error and event logs; no significant issues found. Did not find any issue with the occlusion. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
69416 | 208831 | CM | CMPLT | CMPLT | 2024-08-08 12:03 | 2024-08-08 12:03 | Occlusion error | (TC 08/08/2024 12:02:53 PM, Baghirzade, Nara) Bottle-side pressure sensor is defective. Replaced bottle-side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
66053 | 208832 | CM | CMPLT | CMPLT | 2024-08-08 12:04 | 2024-08-08 12:05 | Functional check | (TC 08/08/2024 12:04:33 PM, Baghirzade, Nara) Downloaded error log; did not find any significant issues. Found the power cord to be damaged; replaced power cord. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 100 mohms Leakage current: 85.3 uA
| 208833 | RDS | CMPLT | CMPLT | 2024-08-08 12:07 | 2024-08-08 12:07 | Weekly Rounds: 11 North | (TC 08/08/2024 12:06:51 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 208834 | RDS | CMPLT | CMPLT | 2024-08-08 12:08 | 2024-08-08 12:08 | Weekly Rounds: 11 South | (TC 08/08/2024 12:08:21 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 208835 | RDS | CMPLT | CMPLT | 2024-08-08 12:10 | 2024-08-08 12:10 | Weekly Rounds: NICU & RESUS | (TC 08/08/2024 12:09:35 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
1070031 | 208836 | CM | CMPLT | CMPLT | 2024-08-08 13:09 | 2024-08-08 13:09 | humidifier alarming"no water" despite water bag full, restarting, 0 leaks. | tested 2 days. Functional.
1050071 | 208837 | CM | CMPLT | CMPLT | 2024-08-08 13:17 | 2024-08-27 08:15 | Broken housing | (TC 08/08/2024 02:19:40 PM, Canjimavadivel, Kamallesh ) Disassembled unit to assess all damage. Found multiple damaged parts. Called vendor to receive quote. Placed order for damaged parts (TC 08/09/2024 02:41:17 PM, Canjimavadivel, Kamallesh ) Received back housing. Swapped circuit boards and power supply board from damaged housing to new housing. Other items are back ordered (TC 08/26/2024 03:30:00 PM, Canjimavadivel, Kamallesh ) Replaced front housing and handle insert. (TC 08/27/2024 08:15:00 AM, Canjimavadivel, Kamallesh ) Assembled onto cart. Tested all functions.
1059764 | 208838 | CM | CMPLT | CMPLT | 2024-08-08 13:53 | 2024-09-25 10:55 | MSH 15 L&D TE7E NOT CHARGING/POWERING UP | (TC 08/08/2024 04:04:48 PM, Whittington, Scott) MSH 15 L&D TE7E NOT CHARGING/POWERING UP RECOILING POWER CORD CAP FAULTY CONNECTION UNABLE TO LOCATE MALE CORD CAP HOSPITAL GRADE ATTACHED SPARE DESKTOP POWER SUYPPLY TO SYSTEM,AVAILABLE FOR CLINICAL USE (TC 08/08/2024 04:04:48 PM, Whittington, Scott) MSH 15 L&D TE7E NOT CHARGING/POWERING UP RECOILING POWER CORD CAP FAULTY CONNECTION UNABLE TO LOCATE MALE CORD CAP HOSPITAL GRADE ATTACHED SPARE DESKTOP POWER SUYPPLY TO SYSTEM,AVAILABLE FOR CLINICAL USE | MSH 15 L&D TE7E NOT CHARGING/POWERING UP Replaced ac plug.
1063787 | 208839 | CM | CANCL | CANCL | 2024-08-08 14:09 | 2024-08-08 14:14 | Light does not turn on
1063792 | 208840 | CM | CMPLT | CMPLT | 2024-08-08 14:26 | 2024-08-08 13:20 | Light Does not work | Jul 17 - Call re: not working. Went to check. No battery present. Looked for the battery. Checked functionality. Found the knob used to turn the headlight on keeps spinning and the light can't turn on. Removed battery and brought it down to biomed. Contacted vendor for RMA and ship to address. Awaiting response. Aug 1 - Followed up. Info received. Created meditech req. Requested approval. Received. Issued PO via MMC. Scheduled, packed and shipped the unit. Aug 8 - Received unit. Checked functionality. Tested ok. Returned to OR.
67895 | 208841 | CM | CMPLT | NFF | 2024-08-08 14:28 | 2024-08-08 15:20 | Problem not specified | (TC 08/08/2024 02:28:26 PM, Rendon Lopez , Elizabeth) Running device on user mode, no alarms, no errors (TC 08/08/2024 03:19:39 PM, Rendon Lopez , Elizabeth) Ran on user mode for hours, no fault found. Performance assurance-OK
| 208842 | CM | ACTIV | ISSUE | 2024-08-08 14:30 |  | FIlac thermometer probe | (TC 08/08/2024 02:38:33 PM, Rendon Lopez , Elizabeth) Spare Biomed thermometer Filac probe
| 208843 | PA | CMPLT | CMPLT | 2024-08-08 14:42 | 2024-08-08 14:51 | TC70 Wifi card upgrade kit | (TC 08/08/2024 02:51:23 PM, Canjimavadivel, Kamallesh ) Placed order for wifi upgrade kit | Charge back when used
1046544 | 208844 | CM | CMPLT | CMPLT | 2024-08-08 14:51 | 2024-08-08 14:51 | Won't stop beeping | (TC 08/08/2024 02:51:14 PM, Rendon Lopez , Elizabeth) Replaced batteries, tested-OK. Device located in 15M
1046577 | 208845 | CM | CMPLT | CMPLT | 2024-08-08 15:06 | 2024-08-08 15:06 | Not working | (TC 08/08/2024 03:06:25 PM, Rendon Lopez , Elizabeth) Device not turning on. Replaced batteries, tried new probe, not resolved. Replaced device with spare, kept old probe, tested-Ok. Changed SN in TMS, returned to 15M
1059495 | 208846 | CM | CMPLT | CMPLT | 2024-08-08 15:08 | 2024-08-08 15:11 | Not working | (TC 08/08/2024 03:11:07 PM, Rendon Lopez , Elizabeth) Device not turning on. Replaced batteries, tried new probe, not resolved. Replaced device with spare, kept old probe, tested-Ok. Changed SN in TMS, returned to 15M
1061202 | 208847 | RQ | CMPLT | CMPLT | 2024-08-08 15:15 | 2024-08-08 15:20 | teste work order | (TC 08/08/2024 03:20:25 PM, Chukri, Samir) done test and returned | Test please inform rebaa
1069562 | 208848 | CM | CMPLT | CMPLT | 2024-08-08 16:10 | 2024-09-25 10:55 | MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 | (TC 08/08/2024 04:10:03 PM, Whittington, Scott) MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 (TC 08/09/2024 01:01:54 PM, Whittington, Scott) MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 (TC 08/09/2024 01:01:54 PM, Whittington, Scott) MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 SYSTEM SITE ID 416586VOL REF# ANDREA 1-552271163129 Probe ordered Thanks, Scott | MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 SYSTEM SITE ID 416586VOL REF# ANDREA 1-552271163129
1069363 | 208849 | CM | CMPLT | CMPLT | 2024-08-08 16:11 | 2024-09-25 10:56 | MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 | (TC 08/08/2024 04:11:13 PM, Whittington, Scott) MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4 | MSH OPG SPP 3rd Floor FAULTY PROBE C2-9-D XTAL DROP SN# 479693YP4
67631 | 208850 | PM | CMPLT | CMPLT | 2024-08-09 09:19 | 2024-08-07 09:19 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/07/2024 09:19:06 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed.
69094 | 208851 | CM | CMPLT | CMPLT | 2024-08-09 09:52 | 2024-08-09 09:54 | Channel Error | (TC 08/09/2024 09:54:28 AM, DeGuzman, Leila) Cleaned device. Downloaded error log, found error code 240.4150 - LVP bottle-side pressure sensor High. Checked analog sensor value - confirmed issue. Replaced the bottle-side pressure sensor - okay. Performed check-in tests: all passed. Returned device.
1058172 | 208852 | CM | CMPLT | CMPLT | 2024-08-09 10:02 | 2024-08-09 10:02 | Downstream Occlusion alarm when trying to run epidural | (TC 08/09/2024 10:02:05 AM, DeGuzman, Leila) Peformed functional tests: all passed No issue found.
66436 | 208853 | CM | CMPLT | CMPLT | 2024-08-09 10:03 | 2024-08-09 10:03 | Not turning on | (TC 08/09/2024 10:02:43 AM, Scero, Tomlyn) Replaced battery adaptor. Functional tests completed with an ECG simulator
63216 | 208854 | CM | CMPLT | PMFAIL | 2024-08-09 11:53 | 2024-08-09 11:53 | Grinding/Rattling noise when running | (TC 08/09/2024 11:53:02 AM, DeGuzman, Leila) Checked device, There's a cracked on the mounting hole on the back of the pump. This might damage the battery and internal components. Rotor has issue too. Need to send out.
63438 | 208855 | CM | CMPLT | PMFAIL | 2024-08-09 12:00 | 2024-08-09 12:01 | System Error | (TC 08/09/2024 12:00:35 PM, DeGuzman, Leila) Run the device for about an hour.. No system error showing.. However the rotor seems worn out and somethings off in the inside of the device.
68657 | 208856 | PM | CANCL | DUP | 2024-08-09 12:15 | 2024-08-09 12:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69307 | 208857 | PM | CMPLT | CMPLT | 2024-08-09 12:22 | 2024-08-09 12:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:21:55 PM, Patel, Rutvik) by BD
69187 | 208858 | PM | CMPLT | CMPLT | 2024-08-09 12:40 | 2024-08-09 12:40 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 08/09/2024 12:40:32 PM, Patel, Rutvik) by BD
69253 | 208859 | PM | CMPLT | CMPLT | 2024-08-09 13:03 | 2024-08-09 13:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | By BD
1057604 | 208860 | CM | CMPLT | CMPLT | 2024-08-09 13:29 | 2024-08-09 13:29 | Verification testing | (TC 08/09/2024 01:29:28 PM, Rhuland, Courtney) Tested NIBP functionality with SimCube 120/80 - 116/83 190/120 - 194/124 80/40 - 76/44 Manometer reading difference - 0.7 mmHg All within 5 mmHg Tests passed
65883 | 208861 | CM | CMPLT | CMPLT | 2024-08-09 14:18 | 2024-08-09 14:24 | Channel Error | (TC 08/09/2024 02:23:34 PM, DeGuzman, Leila) Cleaned device. Downloaded error log Downloaded event log Found error code 351.6760 - Module not calibrated. Performed Calibration and accuray verification: all passed. Performed check-in tests: all passed. Cleared error and event log PM date updated. Returned to service.
| 208862 | CM | CMPLT | CMPLT | 2024-08-09 14:28 | 2024-08-09 14:28 | Equipment check | (TC 08/09/2024 02:27:49 PM, Rhuland, Courtney) Requested to check exam lights Noted lights that were broken/missing parts Noted bulb replacement numbers for scopes Repaired manual BP hose and replaced cuff
63031 | 208863 | FC | CMPLT | CMPLT | 2024-08-09 14:40 | 2024-08-08 14:40 | Functional check | (TC 08/08/2024 02:40:01 PM, Cortez, John) Tested functionality, dropped off at RCC
69668 | 208864 | PM | CMPLT | CMPLT | 2024-08-09 14:57 | 2024-08-09 14:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | By BD
1065884 | 208865 | PM | CMPLT | CMPLT | 2024-08-09 16:58 | 2024-08-09 16:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065900 | 208866 | PM | CMPLT | CMPLT | 2024-08-09 17:00 | 2024-08-09 17:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065902 | 208867 | PM | CMPLT | CMPLT | 2024-08-09 17:03 | 2024-08-09 17:03 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
| 208868 | CM | CMPLT | CMPLT | 2024-08-09 17:08 | 2024-08-16 09:28 | Masimo Root quick connect order, biomed stock | (TC 08/09/2024 05:10:01 PM, Cortez, John) ordered x2 (TC 08/16/2024 09:28:05 AM, Cortez, John) arrived
1065908 | 208869 | PM | CMPLT | CMPLT | 2024-08-09 17:21 | 2024-08-09 17:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065912 | 208870 | PM | CMPLT | CMPLT | 2024-08-09 17:35 | 2024-08-09 17:34 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065913 | 208871 | PM | CMPLT | CMPLT | 2024-08-09 17:36 | 2024-08-09 17:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065925 | 208872 | PM | CMPLT | CMPLT | 2024-08-09 17:38 | 2024-08-09 17:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065929 | 208873 | PM | CMPLT | CMPLT | 2024-08-09 17:53 | 2024-08-09 17:53 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065937 | 208874 | PM | CMPLT | CMPLT | 2024-08-09 17:55 | 2024-08-09 17:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065945 | 208875 | PM | CMPLT | CMPLT | 2024-08-09 17:56 | 2024-08-09 17:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1065995 | 208876 | PM | CMPLT | CMPLT | 2024-08-09 17:58 | 2024-08-09 17:58 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1066023 | 208877 | PM | CMPLT | CMPLT | 2024-08-09 18:00 | 2024-08-09 18:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1066025 | 208878 | PM | CMPLT | CMPLT | 2024-08-09 18:03 | 2024-08-09 18:10 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1066031 | 208879 | PM | CMPLT | CMPLT | 2024-08-09 18:14 | 2024-08-09 18:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
1066050 | 208880 | PM | CMPLT | CMPLT | 2024-08-09 18:23 | 2024-08-09 18:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | by BD
| 208881 | CM | CMPLT | CMPLT | 2024-08-09 18:40 | 2024-08-09 18:40 | IV Pump PM/Remediation WO closure | closed all PM/remediation WO as per assigned dept.
1070801 | 208882 | CM | CMPLT | NFF | 2024-08-12 08:15 | 2024-08-12 08:16 | Air/baby mode not working | (TC 08/12/2024 08:16:06 AM, Rhuland, Courtney) Tag stated baby/air mode was not working when canopy was closed Both air mode and baby mode were functional Calibrated RH Functional check - OK
| 208883 | RDS | CMPLT | CMPLT | 2024-08-12 08:30 | 2024-08-19 14:03 | Weekly checks in CCU | (TC 08/12/2024 02:02:45 PM, Scero, Tomlyn) weekly checks complete (TC 08/19/2024 02:03:06 PM, Scero, Tomlyn) weekly checks complete
1061207 | 208884 | CM | CMPLT | CMPLT | 2024-08-12 08:46 | 2024-08-12 08:46 | Corio Matrix Error | (TC 08/12/2024 08:45:37 AM, Reid, Emmerson ) Rebooted the VMC and Corio, working now.
| 208885 | CM | ACTIV | ISSUE | 2024-08-12 08:47 |  | OR - Trumpf Spine Bridge Repair | (TC 08/12/2024 08:47:07 AM, Reid, Emmerson ) One component was missing a screw and was bent. I was able to bend it back into position. I then took one of the middle screws and moved it to the outside position. Spoke with Samir and he agreed it was okay to return this device as is while we wait for the replacement screw and bent component.
| 208886 | RDS | CMPLT | CMPLT | 2024-08-12 08:58 | 2024-08-12 09:00 | Weekly check for ED
1055933 | 208887 | CM | CMPLT | CMPLT | 2024-08-12 09:09 | 2024-08-12 09:09 | Temp cord wire exposed | (TC 08/12/2024 09:09:03 AM, DeGuzman, Leila) Replaced temperature probe. Replace BP cuff. Checked functions: all okay.
1055992 | 208888 | CM | CMPLT | CMPLT | 2024-08-12 09:12 | 2024-08-12 09:12 | Missing SpO2 sensor probe | (TC 08/12/2024 09:11:57 AM, DeGuzman, Leila) Placed SpO2 sensor probe. Checked functinality: Good
1057941 | 208889 | CM | ACTIV | ISSUE | 2024-08-12 09:21 |  | Surgical Table - Error 82C/806
1071018 | 208890 | CM | ACTIV | ISSUE | 2024-08-12 10:15 |  | SCD - Broken | (TC 08/12/2024 10:14:52 AM, Reid, Emmerson ) Emailed the vendor about getting a replacement.
1067338 | 208891 | CM | CMPLT | CMPLT | 2024-08-12 11:53 | 2024-08-12 11:53 | O2 sensor caibration failure | ordered O2 sensor from Getinge. Received and installed the sensor. Tested for 3 days. No error. Retuned back to RT. already charged to RT for sensor cost.
| 208892 | CM | CMPLT | CMPLT | 2024-08-12 12:29 | 2024-08-16 09:00 | NIBP cuffs for NICU
1059778 | 208893 | FC | CMPLT | CMPLT | 2024-08-12 12:45 | 2024-08-12 12:46 | QC checklist | (TC 08/12/2024 12:45:32 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1056441 | 208894 | CM | CMPLT | CMPLT | 2024-08-12 12:48 | 2024-09-25 10:00 | Root does not run on battery even after charging | (TC 08/12/2024 12:47:56 PM, Baghirzade, Nara) All PFT techs have complained that even after charging all day/night, the Root only runs on battery for a short period. Device is used for walks so it has to run on battery for longer periods of time. Contacted Masimo for service. Provided PFT with a borrowed Root from MBU until the original device is functional. (TC 09/25/2024 09:59:44 AM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Replaced probe well and probe. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 208895 | CM | CMPLT | CMPLT | 2024-08-12 12:50 | 2024-08-12 12:50 | OBM meeting with Natus | (TC 08/12/2024 12:50:25 PM, Baghirzade, Nara) Meeting with Natus re: NICU OBMs
69497 | 208896 | CM | CMPLT | CMPLT | 2024-08-12 12:53 | 2024-08-12 12:53 | Door does not close | (TC 08/12/2024 12:53:12 PM, Baghirzade, Nara) Spring was out of place for the door latch - replaced the spring. Door latch is functional. Check-in tests passed and completed. IUI connectors and device cleaned.
1057847 | 208897 | CM | CMPLT | CMPLT | 2024-08-12 13:49 | 2024-09-06 12:45 | shipped the device to Conmed for PM | (TC 09/06/2024 12:44:59 PM, Huang, Jessica) sent the device out for PM and put it back to the OR room (TC 09/06/2024 12:44:59 PM, Huang, Jessica) sent the device out for PM and put it back to the OR room
1070247 | 208898 | CM | CMPLT | CMPLT | 2024-08-12 14:14 | 2024-08-12 14:21 | Screen readings not showing - Very light | (TC 08/12/2024 02:21:03 PM, Rendon Lopez , Elizabeth) Replaced batteries, Tested with weigths-OK.
1069503 | 208899 | CM | CMPLT | CMPLT | 2024-08-13 07:23 | 2024-08-12 07:23 | Height rod not installed properly | (TC 08/12/2024 07:23:21 AM, Rhuland, Courtney) Height rod was off by 3cm When installed, the rod was touching the base of the scale Removed height rod and reinstalled with 3cm gap from ground Verified height with measuring tape
| 208900 | CM | CMPLT | CMPLT | 2024-08-13 08:13 | 2024-08-12 08:13 | Biomed meeting | (TC 08/12/2024 08:12:51 AM, Baghirzade, Nara) Biomed meeting
1047847 | 208901 | PM | CMPLT | CMPLT | 2024-08-13 08:19 | 2024-08-12 08:21 | PM-MAJOR-Annual PM for: Infusion Pumps, Syringe | (TC 08/12/2024 08:19:38 AM, Baghirzade, Nara) Completed calibration for: -syringe size and position -force sensor -plunger position -syringe size sensor Completed alarm/speaker tests. PM completed. See procedure document for PM results. PM schedule updated. Functional tests passed. Ran basic infusion; completed successfully. No errors occurred during infusion. Occlusion alarm is functional.
68707 | 208902 | CM | CMPLT | CMPLT | 2024-08-13 08:49 | 2024-08-13 08:49 | Not charging when plugged in | (TC 08/13/2024 08:49:29 AM, Canjimavadivel, Kamallesh ) Brought pump down from ICU. verified issue. Tested with new power cord, working ok. Replaced power cord.
1056441 | 208903 | CM | CMPLT | CMPLT | 2024-08-13 09:39 | 2024-09-25 09:58 | Sending Root for repair | (TC 08/13/2024 09:44:23 AM, Baghirzade, Nara) SC 1003288929 (TC 08/26/2024 11:30:46 AM, Baghirzade, Nara) PO 7500210510 KJV # 451083 (TC 09/25/2024 10:15:28 AM, Baghirzade, Nara) Received defective Root from Masimo. Followed up and received another Root the next day. Shipping defective Root KJV# 455964 PO 7500210510 (TC 09/25/2024 10:31:45 AM, Baghirzade, Nara) Received root replacement.
| 208904 | CM | CMPLT | CMPLT | 2024-08-13 10:33 | 2024-09-13 13:22 | Ordering M4605A batteries, biomed stock | (TC 08/13/2024 10:38:44 AM, Cortez, John) ordered x5 (TC 09/13/2024 01:22:09 PM, Cortez, John) arrived
66461 | 208905 | PM | CMPLT | CMPLT | 2024-08-13 11:38 | 2024-08-13 11:38 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3nd year PM last year. It has no GS500. Replaced V-battery and air filter. Run battery check. 220 minutes, 25 Ah. Tested. Functional. SW2.51.01 Leakage: 95 uA and 89 ohm.
| 208906 | CM | ACTIV | ISSUE | 2024-08-13 12:23 |  | Outstanding BD work orders | (TC 08/13/2024 12:23:07 PM, Baghirzade, Nara) Over 200 active and outstanding WOs after BD remediation. Going through all the WOs checking against the excel document of all completed and not completed devices from BD. (TC 08/16/2024 03:48:30 PM, Baghirzade, Nara) Finished closing all software recall and PM WOs. (TC 08/22/2024 10:49:06 AM, Baghirzade, Nara) Creating and completing PM WOs for assets that do not have them. Confirming PM status against the BD provided excel sheet. (TC 08/23/2024 01:54:11 PM, Baghirzade, Nara) Created and completed the remaining PM WOs for assets that were remediated and PM'ed but no WO was generated for the PM. Checked the completion status against the BD excel sheet. (TC 09/10/2024 02:13:41 PM, Baghirzade, Nara) Went through all active PM WOs for PCUs, LVPs, and syringes. Used BD upgrade sheet to check the remediation and pm status for the remaining devices found on TMS. Created 3 excel sheets for each modality to note any incomplete work. Will follow up with Adeel.
62658 | 208907 | CM | CMPLT | CMPLT | 2024-08-13 14:25 | 2024-08-13 14:30 | Height rod inaccurate | (TC 08/13/2024 02:30:20 PM, Rhuland, Courtney) Height rod measures 3cm off from actual height Seems height rod had been replaced and is not the proper one for this device No place to reposition rod Told clinical staff to adjust for difference when taking measurements while a replacement is looked into
1057940 | 208908 | CM | CMPLT | CMPLT | 2024-08-13 15:37 | 2024-08-13 15:38 | Not moving | Call re: not working. Went to check. Patient on the table. Waited for access. Found table not moving. Unlocked and Reboot the table. Tested ok. Returned to service.
1056665 | 208909 | CM | CMPLT | CMPLT | 2024-08-13 15:40 | 2024-08-13 15:42 | Monitors not working | Call re: Monitors not displaying microscope image. Went to check. Waited for access. Found no in on scope HD connection. Checked microscope connections. Found staff connected the cables incorrectly. Rectified it. Checked functionality. Tested ok.
1066935 | 208910 | CM | CMPLT | CMPLT | 2024-08-14 08:46 | 2024-10-04 09:08 | Parts Order - Replacement Lamp | Aug 14 - Unit is due for lamp replacement. Contacted vendor and requested info. Created SC. Requested approval. Placed order via MMC. Sept 18 - PArts received. Will schedule and replace the lamp. Oct 4 - Replaced lamp. Checked functionality. Tested ok.
1056125 | 208911 | PM | CMPLT | CMPLT | 2024-08-14 08:53 | 2024-08-14 10:26 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/14/2024 10:25:35 AM, Canjimavadivel, Kamallesh ) Cleaned hoses and nozzles Cleaning cycle complete
| 208912 | CM | CMPLT | CMPLT | 2024-08-14 08:59 | 2024-08-13 08:59 | Medfusion parts and spare parts management | (TC 08/13/2024 08:59:06 AM, Baghirzade, Nara) Completed inventory of all parts for medfusions. Did an inventory of spare parts from retired medfusion devices. Organized and labelled all parts.
| 208913 | CM | CMPLT | CMPLT | 2024-08-14 09:01 | 2024-08-13 09:01 | Meeting with Draeger and building services | (TC 08/13/2024 09:01:23 AM, Baghirzade, Nara) Meeting with Draeger and building services re: nitrogen leaks in booms
| 208914 | CM | CMPLT | CMPLT | 2024-08-14 09:03 | 2024-08-13 09:03 | Helped Courtney at 60 murray with height rod for scale in triage | (TC 08/13/2024 09:02:38 AM, Baghirzade, Nara) Helped Courtney at 60 murray with height rod for scale in triage - family medicine clinic. Height rod is off by 3 cm; let the staff know to account a 3 cm correction to their measurements.
1054111 | 208915 | PM | CMPLT | CMPLT | 2024-08-14 09:15 | 2024-08-12 09:15 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/12/2024 09:15:21 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
1048154 | 208916 | CM | CMPLT | CMPLT | 2024-08-14 09:16 | 2024-08-21 08:34 | Moving from OR to ENT Clinic | (TC 08/14/2024 09:16:04 AM, Reid, Emmerson ) I coordinated with Nara, moved the microscope from the OR to our storage room, and updated it's TMS information to reflect it's new location. (TC 08/21/2024 08:33:02 AM, Baghirzade, Nara) Met with Ana to discuss and plan the location and placement of the microscope. Worked with Courtney to move the microscope to ENT from the MCR room. Placed it in the room where Ana wanted it, and moved around other equipment as per her request.
1054112 | 208917 | PM | CMPLT | CMPLT | 2024-08-14 09:17 | 2024-08-12 09:17 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/12/2024 09:17:16 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
1054109 | 208918 | PM | CMPLT | CMPLT | 2024-08-14 09:18 | 2024-08-12 09:19 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/12/2024 09:18:56 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
1062123 | 208919 | CM | CMPLT | CMPLT | 2024-08-14 11:05 | 2024-08-14 11:05 | Pump is not delivering programmed dose | (TC 08/12/2024 11:04:47 AM, Rendon Lopez , Elizabeth) Running pump on same programmed patient, Found KVO rate lower than dose. (TC 08/14/2024 11:04:47 AM, Rendon Lopez , Elizabeth) Run pump for hours with drug setting and increased KVO, No fault found
| 208920 | CM | CMPLT | CMPLT | 2024-08-14 13:40 | 2024-08-14 13:40 | Lasers meeting | (TC 08/14/2024 01:39:47 PM, Baghirzade, Nara) Lasers meeting
67945 | 208921 | PM | CMPLT | CMPLT | 2024-08-14 13:41 | 2024-08-14 13:41 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/14/2024 01:41:12 PM, Baghirzade, Nara) Retrieved the incubator from 18th fl. storage with Leila's help. PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 80.1 uA Power cord resistance = 225 mohms
1071842 | 208922 | CM | CMPLT | CMPLT | 2024-08-14 13:59 | 2024-08-14 12:04 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071843 | 208923 | CM | CMPLT | CMPLT | 2024-08-14 14:03 | 2024-08-14 12:05 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071844 | 208924 | CM | CMPLT | CMPLT | 2024-08-14 14:05 | 2024-08-14 12:06 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071845 | 208925 | CM | CMPLT | CMPLT | 2024-08-14 14:09 | 2024-08-14 12:07 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071846 | 208926 | CM | CMPLT | CMPLT | 2024-08-14 14:12 | 2024-08-14 12:07 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071847 | 208927 | CM | CMPLT | CMPLT | 2024-08-14 14:13 | 2024-08-14 12:08 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071848 | 208928 | CM | CMPLT | CMPLT | 2024-08-14 14:14 | 2024-08-14 12:09 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071849 | 208929 | CM | CMPLT | CMPLT | 2024-08-14 14:16 | 2024-08-14 12:21 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071850 | 208930 | CM | CMPLT | CMPLT | 2024-08-14 14:17 | 2024-08-14 12:21 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071851 | 208931 | CM | CMPLT | CMPLT | 2024-08-14 14:19 | 2024-08-14 12:22 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071852 | 208932 | CM | CMPLT | CMPLT | 2024-08-14 14:21 | 2024-08-15 12:23 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071853 | 208933 | CM | CMPLT | CMPLT | 2024-08-14 14:23 | 2024-08-15 12:23 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071854 | 208934 | CM | CMPLT | CMPLT | 2024-08-14 14:24 | 2024-08-15 12:24 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071855 | 208935 | CM | CMPLT | CMPLT | 2024-08-14 14:26 | 2024-08-16 15:52 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071856 | 208936 | CM | CMPLT | CMPLT | 2024-08-14 14:28 | 2024-08-16 16:27 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071857 | 208937 | CM | CMPLT | CMPLT | 2024-08-14 14:30 | 2024-08-16 15:53 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071858 | 208938 | CM | CMPLT | CMPLT | 2024-08-14 14:33 | 2024-08-16 15:53 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071859 | 208939 | CM | CMPLT | CMPLT | 2024-08-14 14:35 | 2024-08-16 15:54 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071860 | 208940 | CM | CMPLT | CMPLT | 2024-08-14 14:36 | 2024-08-16 15:55 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1049901 | 208941 | CM | CMPLT | CMPLT | 2024-08-14 15:02 | 2024-08-14 15:02 | Battery defective
1071861 | 208942 | CM | CMPLT | CMPLT | 2024-08-14 15:13 | 2024-08-16 15:55 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071862 | 208943 | CM | CMPLT | CMPLT | 2024-08-14 15:14 | 2024-08-16 15:56 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071863 | 208944 | CM | CMPLT | CMPLT | 2024-08-14 15:16 | 2024-08-16 15:57 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071864 | 208945 | CM | CMPLT | CMPLT | 2024-08-14 15:17 | 2024-08-16 15:58 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071865 | 208946 | CM | CMPLT | CMPLT | 2024-08-14 15:18 | 2024-08-16 15:58 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071866 | 208947 | CM | CMPLT | CMPLT | 2024-08-14 15:19 | 2024-08-16 15:59 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071867 | 208948 | CM | CMPLT | CMPLT | 2024-08-14 15:19 | 2024-08-16 15:59 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071868 | 208949 | CM | CMPLT | CMPLT | 2024-08-14 15:20 | 2024-08-16 16:00 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071869 | 208950 | CM | CMPLT | CMPLT | 2024-08-14 15:20 | 2024-08-16 16:01 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071870 | 208951 | CM | CMPLT | CMPLT | 2024-08-14 15:21 | 2024-08-16 16:02 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071871 | 208952 | CM | CMPLT | CMPLT | 2024-08-14 15:22 | 2024-08-16 16:03 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071872 | 208953 | CM | CMPLT | CMPLT | 2024-08-14 15:23 | 2024-08-16 16:05 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071873 | 208954 | CM | CMPLT | CMPLT | 2024-08-14 15:25 | 2024-08-16 16:07 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071874 | 208955 | CM | CMPLT | CMPLT | 2024-08-14 15:25 | 2024-08-16 16:08 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071875 | 208956 | CM | CMPLT | CMPLT | 2024-08-14 15:27 | 2024-08-16 16:24 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071876 | 208957 | CM | CMPLT | CMPLT | 2024-08-14 15:27 | 2024-08-16 16:24 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071877 | 208958 | CM | CMPLT | CMPLT | 2024-08-14 15:29 | 2024-08-16 16:25 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071878 | 208959 | CM | CMPLT | CMPLT | 2024-08-14 15:29 | 2024-08-16 16:26 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1071879 | 208960 | CM | CMPLT | CMPLT | 2024-08-14 15:30 | 2024-08-16 16:26 | OR Asset Division - TMS Record | Divided OR assets. Recorded S/N and updated details in TMS.
1061458 | 208961 | CM | CMPLT | CMPLT | 2024-08-14 16:45 | 2024-09-12 13:28 | "controls don't work" | (TC 08/14/2024 04:45:30 PM, Cortez, John) Confirmed issue, "limited motion functionality" Power cycling only fixes motion controls on footboard, operating controls on rails cause error to return Requested service from stryker (TC 09/12/2024 01:28:26 PM, Cortez, John) Repaired by Aziz
1050672 | 208962 | CM | CMPLT | CMPLT | 2024-08-15 08:24 | 2024-08-15 08:26 | Connectivity issue | (TC 08/15/2024 08:25:59 AM, Canjimavadivel, Kamallesh ) Connected to server when tested in biomed, was able to transfer pending order.
65416 | 208963 | CM | CMPLT | CMPLT | 2024-08-15 08:29 | 2024-08-15 12:32 | Air temp probe stops working after reaching 35C | (TC 08/14/2024 08:28:50 AM, Baghirzade, Nara) Ran the incubator and set temp to 35C - RH at 75%. Air temp probe stops working after reaching 35C, immediately dips down to 19C and does not go above 22C; error "air probe failure" appears. Air probe was replaced. Ran incubator again and still getting the same issue. Will need to examine the probe harness cable (TC 08/15/2024 12:31:44 PM, Baghirzade, Nara) Replaced the cable harness. Ran another functional and did not have any air probe issues or errors. Confirmed that incubator reaches set temperature and does not drop; all temperature consistently measure the same. During functional, the RH% was reading high. Calibrated RH twice. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
1069562 | 208964 | ACC | CMPLT | CMPLT | 2024-08-15 08:42 | 2024-08-20 11:15 | MSH OPG SPP TONY CATUNTO DELIVERED AND INSTALLED REPLACEMENT PROBE SN# 496541YP4 | (TC 08/15/2024 08:41:41 AM, Whittington, Scott) AUGUST 14 MSH OPG SPP TONY CATUNTO DELIVERED AND INSTALLED REPLACEMENT PROBE SN# 496541YP4 (TC 08/19/2024 08:41:41 AM, Whittington, Scott) AUGUST 14 MSH OPG SPP TONY CATUNTO DELIVERED AND INSTALLED REPLACEMENT PROBE SN# 496541YP4 | MSH OPG SPP TONY CATUNTO DELIVERED AND INSTALLED REPLACEMENT PROBE SN# 496541YP4
1050098 | 208965 | CM | CMPLT | CMPLT | 2024-08-15 08:56 | 2024-10-17 08:11 | Screen and probe damaged | (TC 08/15/2024 08:56:20 AM, Canjimavadivel, Kamallesh ) Documented pictures and reached out to vendor for repair quote. (TC 08/16/2024 08:02:11 AM, Canjimavadivel, Kamallesh ) Received quote, requested PO for repair (TC 09/12/2024 08:40:18 AM, Canjimavadivel, Kamallesh ) Requested no charge PO to receive loaner (TC 10/17/2024 08:10:33 AM, Canjimavadivel, Kamallesh ) Received device from repair, see attached service report. Device returned to ED and removed loaner
502088 | 208966 | RQ | CMPLT | CMPLT | 2024-08-15 09:00 | 2024-08-15 10:35 | loud high pitched noise | (TC 08/15/2024 10:34:49 AM, Paz, Ramon (Tino)) Loud high pitch noise coming from the gantry. Investigated noise after patient scan. It turned out to be the UPS on a 4 year battery alarm reminder. Reset the reminder and alarm. | loud high pitched noise
| 208967 | PA | CMPLT | CMPLT | 2024-08-15 09:06 | 2024-08-21 07:35 | Ordering sterile water for pump maintenance/testing | (TC 08/15/2024 09:07:53 AM, Baghirzade, Nara) SC 1003290670 (TC 08/21/2024 07:35:25 AM, Baghirzade, Nara) Received. Order # 4150281957
66535 | 208968 | PM | CMPLT | CMPLT | 2024-08-15 09:24 | 2024-08-14 09:24 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/14/2024 09:24:25 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66518 | 208969 | PM | CMPLT | CMPLT | 2024-08-15 09:27 | 2024-08-14 09:27 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 08/14/2024 09:27:06 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
1064609 | 208970 | CM | CMPLT | CMPLT | 2024-08-15 10:10 | 2024-08-21 14:37 | Cannot detect pulse | (TC 08/15/2024 10:18:37 AM, Canjimavadivel, Kamallesh ) Isolated issue to defective probe, reached out to vendor for quote. Placed order for probe. (TC 08/21/2024 02:36:59 PM, Canjimavadivel, Kamallesh ) Replaced probe, tested ok
1054373 | 208971 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 14:05 | Audit inspection | (TC 08/15/2024 02:04:49 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1054377 | 208972 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 14:02 | Audit inspection | (TC 08/15/2024 02:01:42 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1054768 | 208973 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:59 | Audit inspection | (TC 08/15/2024 01:59:30 PM, Rhuland, Courtney) All tests passed, device is functional MISSING SPO2 FINGER SENSOR See procedure for details
1054769 | 208974 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 12:20 | Audit inspection | (TC 08/15/2024 12:20:01 PM, Rhuland, Courtney) All tests passed, device is functional See procedure for details
1054771 | 208975 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:57 | Audit inspection | (TC 08/15/2024 01:57:08 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1054774 | 208976 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:38 | Audit inspection | (TC 08/15/2024 01:38:04 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1054775 | 208977 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:32 | Audit inspection | (TC 08/15/2024 01:31:50 PM, Rhuland, Courtney) All tests passed, device is functional MISSING SPO2 FINGER SENSOR See procedure for details
1054779 | 208978 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:29 | Audit inspection | (TC 08/15/2024 01:29:20 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065157 | 208979 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:21 | Audit inspection | (TC 08/15/2024 01:21:20 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065158 | 208980 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:15 | Audit inspection | (TC 08/15/2024 01:14:57 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065159 | 208981 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 13:12 | Audit inspection | (TC 08/15/2024 01:12:28 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065160 | 208982 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 12:35 | Audit inspection | (TC 08/15/2024 12:34:51 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065161 | 208983 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 12:33 | Audit inspection | (TC 08/15/2024 12:32:52 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065162 | 208984 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 12:29 | Audit inspection | (TC 08/15/2024 12:29:14 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065163 | 208985 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 12:27 | Audit inspection | (TC 08/15/2024 12:26:46 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
1065164 | 208986 | CM | CMPLT | CMPLT | 2024-08-15 10:31 | 2024-08-15 12:25 | Audit inspection | (TC 08/15/2024 12:24:46 PM, Rhuland, Courtney) All tests passed, device is fully functional See procedure for details
65996 | 208987 | CM | CMPLT | CMPLT | 2024-08-15 10:55 | 2024-08-15 10:57 | System Error | (TC 08/15/2024 10:56:33 AM, DeGuzman, Leila) Checked the PCU Found system error code 110.6021 Service manual says replace the logic board. Nara will do the software update/ flashing. Nara did the updates, but showing the same error. In servicie manual it says replace keypad Replaced back the old logic board and save the new one. (TC 08/15/2024 10:56:33 AM, DeGuzman, Leila) Checked the PCU Found system error code 110.6021 Service manual says replace the logic board. Nara will do the software update/ flashing.
69613 | 208988 | CM | CMPLT | CMPLT | 2024-08-15 11:07 | 2024-08-15 11:07 | Channel error | (TC 08/15/2024 11:07:03 AM, DeGuzman, Leila) Downloaded error log, found error code 240.4150 Voltage too high in bottle side sensor. Checked analog sensor on ASM, confirm sensor high voltage. Replaced the bottle side sensor. Replaced the alarm light cover. Performed check-in, all tests passed. Return to service.
66840 | 208989 | PM | CMPLT | CMPLT | 2024-08-15 11:33 | 2024-08-15 14:06 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/15/2024 02:05:30 PM, DeGuzman, Leila) PM passed and completed. See procedure for details.
1061333 | 208990 | FC | CMPLT | CMPLT | 2024-08-15 12:29 | 2024-08-15 12:29 | QC checklist | (TC 08/15/2024 12:29:11 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
1070798 | 208991 | CM | CMPLT | CMPLT | 2024-08-15 12:30 | 2024-08-15 12:30 | Debris/dirt stuck under the canopy seal | (TC 08/15/2024 12:30:19 PM, Baghirzade, Nara) SA from NICU sent the omnibed down because there was debris/dirt stuck under the canopy seal and he wasn't able to get to it to clean it. Removed the seal, cleaned the seal and where the seal sits. Placed the seal back. Scale was disconnected; found the pins on the cable to be bent. Straightened all the pins and was able to plug it in. Confirmed that the scale is functional. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 76%RH Incubator is fully functional.
69138 | 208992 | CM | CMPLT | CMPLT | 2024-08-15 12:33 | 2024-08-15 12:33 | Defective pressure sensors | (TC 08/15/2024 12:33:07 PM, Baghirzade, Nara) Downloaded error log; noted multiple pressure sensor errors. Both pressure sensors failed analog test. Replaced bottle-side and patient-side pressure sensors. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
1055161 | 208993 | CM | CMPLT | CMPLT | 2024-08-15 12:34 | 2024-08-15 12:34 | Channel error | (TC 08/15/2024 12:34:10 PM, Baghirzade, Nara) Downloaded error log; noted multiple pressure sensor errors. Bottle-side pressure sensor failed analog test. Replaced bottle-side pressure sensor. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
| 208994 | CM | CMPLT | CMPLT | 2024-08-15 12:36 | 2024-08-15 12:36 | Equipment disposal in biomed | (TC 08/15/2024 12:35:35 PM, Baghirzade, Nara) Helped with equipment disposal; put all equipment to be disposed of in one place and labelled them as "for disposal". Followed up with Tabish from support services and sent him a picture of all the equipment that required disposal.
1065165 | 208995 | CM | CMPLT | CMPLT | 2024-08-15 14:39 | 2024-08-15 14:45 | Audit inspection | (TC 08/15/2024 02:44:40 PM, Rhuland, Courtney) See procedure for details
1065517 | 208996 | CM | CMPLT | CMPLT | 2024-08-15 14:39 | 2024-08-15 14:43 | Audit inspection | (TC 08/15/2024 02:43:02 PM, Rhuland, Courtney) See procedure for details
1065518 | 208997 | CM | CMPLT | CMPLT | 2024-08-15 14:39 | 2024-08-15 14:42 | Audit inspection | (TC 08/15/2024 02:42:25 PM, Rhuland, Courtney) See procedure for details
68577 | 208998 | CM | CMPLT | CMPLT | 2024-08-15 14:39 | 2024-08-15 14:41 | Audit inspection | (TC 08/15/2024 02:40:56 PM, Rhuland, Courtney) See procedure for details
1054995 | 208999 | CM | CMPLT | CMPLT | 2024-08-15 14:42 | 2024-08-15 14:42 | Error message when taking temperature | (TC 08/15/2024 02:42:09 PM, DeGuzman, Leila) Checked functionalities: all okay. Took temperature many times, no error showing. Returned to service.
1055039 | 209000 | CM | CMPLT | CMPLT | 2024-08-15 14:46 | 2024-08-15 14:46 | Battery replacement | (TC 08/15/2024 02:46:12 PM, DeGuzman, Leila) Replaced battery. Checked function, okay. Returned to service.
1066934 | 209001 | CM | CMPLT | CMPLT | 2024-08-15 15:24 | 2024-08-15 15:24 | Lamp Replacement | Coordinated, scheduled and replaced lamp. Cleaned unit. Checked functionality. Tested ok.
68021 | 209002 | CM | CMPLT | CMPLT | 2024-08-15 15:25 | 2024-08-15 15:24 | Lamp Replcement | Coordinated, scheduled and replaced lamp. Cleaned unit. Checked functionality. Tested ok.
1064745 | 209003 | CM | CMPLT | CMPLT | 2024-08-15 15:41 | 2024-08-15 15:45 | Not working | Two calls. 1) call at 10 am re: not working. Went to check. Waited for access. Was informed that it is not working with no context. Advised to bring scope set. Inserted camera head and light source. Checked functionality. No issue found. Tested ok. 2) call at 2.30 re: not working. Went to check. Waited for access. Was informed that it is not working. Found the staff was using incorrect camera head. Upon advising to use new one, they kep arguing that they used it earlier. Went to a different room and brought appropriate camera. Demonstrated functionality. Tested ok.
| 209004 | CM | CMPLT | CMPLT | 2024-08-15 17:02 | 2024-08-15 17:02 | Site coverage
1055116 | 209005 | PM | CMPLT | CMPLT | 2024-08-16 09:52 | 2024-08-16 09:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1058254 | 209006 | PM | CMPLT | CMPLT | 2024-08-16 10:25 | 2024-08-16 10:26 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1054684 | 209007 | CM | CMPLT | CMPLT | 2024-08-16 10:39 | 2024-08-16 10:39 | Not working | (TC 08/16/2024 10:39:23 AM, DeGuzman, Leila) Came down with an issue, priming doesn't work. Cleaned device Checked if theres any physical damage: Okay. Turned on the pump, tried to prime and run, showing "Pump Set Dislodge Error" Tried another set, still showing the error. Noticed the finger tab loading area is not turning when priming manually. Need to send out.
1071885 | 209008 | CM | CMPLT | CMPLT | 2024-08-16 11:42 | 2024-08-16 11:42 | Added to TMS | (TC 08/16/2024 11:41:48 AM, Rendon Lopez , Elizabeth) Found device during SW upgrade. Added to TMS, upgraded software, tested-OK
1059828 | 209009 | PM | CMPLT | CMPLT | 2024-08-16 13:38 | 2024-08-16 13:39 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/16/2024 01:39:09 PM, DeGuzman, Leila) Drained, disinfected, rinsed and refilled with sterile water.
1063746 | 209010 | PM | CMPLT | CMPLT | 2024-08-16 13:41 | 2024-08-16 13:41 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/16/2024 01:40:59 PM, DeGuzman, Leila) Drained, disinfected, rinsed and refilled with sterile water.
1065818 | 209011 | FC | CMPLT | CMPLT | 2024-08-16 14:43 | 2024-08-16 14:43 | Sent down with PCU, will run functional check | (TC 08/16/2024 02:42:52 PM, Cortez, John) functional check passed
69277 | 209012 | CM | CMPLT | CMPLT | 2024-08-16 14:47 | 2024-08-16 14:47 | Door latch broken | (TC 08/16/2024 02:46:38 PM, Cortez, John) Door latch replaced Functional check passed
69206 | 209013 | FC | CMPLT | CMPLT | 2024-08-16 14:55 | 2024-08-16 14:55 | Sent down with PCU, will run functional check | (TC 08/16/2024 02:54:41 PM, Cortez, John) functional check passed
69196 | 209014 | FC | CMPLT | CMPLT | 2024-08-16 14:56 | 2024-08-16 14:56 | Sent down with PCU, will run functional check | (TC 08/16/2024 02:56:18 PM, Cortez, John) functional check passed
69508 | 209015 | FC | CMPLT | CMPLT | 2024-08-16 14:58 | 2024-08-16 14:58 | Sent down with PCU, will run functional check | (TC 08/16/2024 02:58:08 PM, Cortez, John) functional check passed
1066039 | 209016 | FC | CMPLT | CMPLT | 2024-08-16 15:00 | 2024-08-16 15:00 | Sent down with PCU, will run functional check | (TC 08/16/2024 03:00:20 PM, Cortez, John) functional check passed
501070 | 209017 | CM | CMPLT | CMPLT | 2024-08-16 17:00 | 2024-09-03 10:06 | MSH CEOU ROOM 10 Runtime error. System down. | (TC 08/20/2024 04:49:46 PM, Whittington, Scott) MSH CEOU ROOM 10 Runtime error. System down. (TC 08/20/2024 04:50:11 PM, Dolinar, John) MSH CEOU ROOM 10 Runtime error. System down. (TC 09/03/2024 10:05:33 AM, Fan, Tak) System was repaired by Philips. Swapped system back with spare. | Runtime error. System down. Ref# 0124120235
1050672 | 209018 | CM | CMPLT | NFF | 2024-08-19 08:50 | 2024-08-19 08:50 | V5 not working | (TC 08/19/2024 08:50:27 AM, Reid, Emmerson ) Checked with sim, all leads were working. NFF
1071886 | 209019 | CM | CMPLT | CMPLT | 2024-08-19 10:36 | 2024-08-19 10:39 | TMS Update - Screen Frozen | call re: not working. Went to check. Brought it down to biomed. Found unit does not have asset tag. Found no OEM sticker on the unit. Consulted SC. Was advised to create asset with no S/N. Assigned new asset and updated TMS. Found screen frozen and turn ON/OFF button not working. Removed screws from the rear chassis and disconnected battery and rebooted the cart. Issue resolved. Checked functionalit.y Tested ok. DElviered back to 12N.
68650 | 209020 | CM | CMPLT | CMPLT | 2024-08-19 10:42 | 2024-08-19 10:42 | "error code" | (TC 08/19/2024 10:42:19 AM, Cortez, John) Checked logs, bottle-side pressure sensor errors Replaced sensor, calibrated, PM tasks completed successfully
| 209021 | CM | CMPLT | CMPLT | 2024-08-19 10:55 | 2024-08-16 10:55 | Heart function Clinic - Installation of WatchBP monitors | (TC 08/16/2024 10:55:17 AM, Scero, Tomlyn) Configured each BP monitor and installed them on the walls in the Heart function clinic rooms 1-4
68827 | 209022 | FC | CMPLT | CMPLT | 2024-08-19 11:03 | 2024-08-19 11:04 | sent down with LVP, will run functional check | (TC 08/19/2024 11:03:33 AM, Cortez, John) Check-in tasks passed, charged battery
| 209023 | CM | CMPLT | CMPLT | 2024-08-19 11:04 | 2024-08-19 11:04 | Height rod not accurate | (TC 08/19/2024 11:04:23 AM, Rhuland, Courtney) Clinic noted that their patients were measuring off When looked at, height was accurate on one half of height rod, but off on the top bar Repositioned height rod and noted the correct way to measure patients under 135cm Must measure from the bottom of bar to account for 3cm difference on bar Both sides of clinic were informed of this process and demonstrated the correct way to measure
1059741 | 209024 | CM | CMPLT | CMPLT | 2024-08-19 11:38 | 2024-08-19 11:38 | Blue light error | (TC 08/19/2024 11:38:17 AM, Rendon Lopez , Elizabeth) Confirmed problem, unable to repair. Retire from service
66028 | 209025 | CM | CMPLT | CMPLT | 2024-08-19 12:16 | 2024-08-19 14:46 | Check-in | (TC 08/19/2024 02:45:49 PM, DeGuzman, Leila) Came down with the LVP pump. Downloaded error and event log. Performed check-in, all tests passed. Cleared error and event log. Ground resistance: 95 mOhm Leakage current: 10 uA Returned to service.
69237 | 209026 | CM | CMPLT | CMPLT | 2024-08-19 12:16 | 2024-08-19 14:49 | Channel error | (TC 08/19/2024 02:48:57 PM, DeGuzman, Leila) Downloaded error and event log. Error found high voltage on bottle-side pressure sensor. Checked analog sensor in ASM: it shows patient side pressure sensor has the high voltage. Replaced the pressure sensor. Check the analog sensor, OK. Performed check-in tests: all passed. Cleared error and event log. Returned to service.
1055969 | 209027 | CM | CMPLT | CMPLT | 2024-08-19 12:24 | 2024-08-19 15:36 | Broken thumb latch | (TC 08/19/2024 03:36:29 PM, DeGuzman, Leila) Replaced the thumb latch. Checked functions: all okay. Returned to service.
68631 | 209028 | CM | CMPLT | CMPLT | 2024-08-19 14:11 | 2024-08-19 14:11 | "broken" | (TC 08/19/2024 02:11:04 PM, Cortez, John) Neutral wire on power cord open Swapped power cord 90 mohms, 73 uamps Charged battery
66525 | 209029 | CM | CMPLT | CMPLT | 2024-08-19 14:12 | 2024-09-04 14:11 | Not turning on- Power light always off | (TC 08/19/2024 02:12:24 PM, Rendon Lopez , Elizabeth) Confirmed problem, removed from OR 3 and placed spare monitor. Needs new switch board, inquire for pricing (TC 08/19/2024 02:17:59 PM, Rendon Lopez , Elizabeth) Switch board is part of small parts kit. Ordered 1 kit for repair and 1 spare. (TC 08/21/2024 11:45:49 AM, Rendon Lopez , Elizabeth) Ordering small parts kit for recorder small parts replacement (TC 08/21/2024 02:11:19 PM, Rendon Lopez , Elizabeth) Replaced switch board, tested-OK. Returned to L&D OR3
| 209030 | CM | CMPLT | CMPLT | 2024-08-19 15:10 | 2024-08-26 14:26 | Meeting | (TC 08/19/2024 03:10:10 PM, DeGuzman, Leila) Monday Huddle (TC 08/21/2024 03:18:01 PM, DeGuzman, Leila) Monthly Meeting (TC 08/26/2024 02:25:34 PM, DeGuzman, Leila) Monday Huddle
1055912 | 209031 | CM | CMPLT | CMPLT | 2024-08-19 15:40 | 2024-08-19 15:40 | Replace temp probe | (TC 08/19/2024 03:40:25 PM, DeGuzman, Leila) Came down with radical 7 that thumb latch broken. Checking all the functions, saw temp probe has exposed wire. Replaced temp probe. All functions are good. Returned to service.
32580 | 209032 | CM | CMPLT | CMPLT | 2024-08-20 07:19 | 2024-08-19 07:18 | Table noise reported | Table noise reported Service completed troubleshooting Service will contact Canon for assistance
1059272 | 209033 | CM | CMPLT | CMPLT | 2024-08-20 08:16 | 2024-09-25 09:33 | Loss of ultrasound data error | (TC 08/20/2024 08:29:38 AM, Canjimavadivel, Kamallesh ) Assessed unit and requested RMA for repair. Requesting loaner as well. (TC 08/21/2024 08:25:11 AM, Canjimavadivel, Kamallesh ) Followed up with vendor on status. Provided PO (TC 09/25/2024 09:32:57 AM, Canjimavadivel, Kamallesh ) Received device and probe back from repair. Installed back onto roll stand. Put battery on charger. See attached service report
| 209034 | CM | ACTIV | ISSUE | 2024-08-20 08:59 |  | Spare Tank regulators for RT NICU | (TC 08/20/2024 09:01:20 AM, Rendon Lopez , Elizabeth) Ordering spare tank regulators for RT NICU.
1058181 | 209035 | CM | CMPLT | CMPLT | 2024-08-20 09:17 | 2024-08-20 09:17 | Upgraded to V4.3 | (TC 08/20/2024 09:17:22 AM, Reid, Emmerson ) Upgraded firmware to V4.3 and push APS Epidural 2023 July library
1067334 | 209036 | CM | CMPLT | CMPLT | 2024-08-20 10:07 | 2024-09-10 14:26 | Not turning on | (TC 08/20/2024 10:07:10 AM, Canjimavadivel, Kamallesh ) Verified issue both with AC and batteries. Tested main console with another scale and isolated issue with console. Reached out to Medline to request quote for repair (TC 08/21/2024 08:22:37 AM, Canjimavadivel, Kamallesh ) Followed up with vendor and provided additional details (TC 08/28/2024 03:00:00 PM, Canjimavadivel, Kamallesh ) Received replacement keypad from manufacturer. Installed and tested, did not work. Contacted manufacturer for next steps (TC 09/10/2024 02:25:57 PM, Canjimavadivel, Kamallesh ) Received display head and swapped out. Checked accuracy with weights, needs calibration. Calibrated according to instructions. Updated SN in TMS
| 209037 | CM | CMPLT | CMPLT | 2024-08-20 10:18 | 2024-08-27 08:41 | NIBP cuffs for NICU
65166 | 209038 | CM | CMPLT | NFF | 2024-08-20 10:46 | 2024-08-20 10:47 | Casset not locking | (TC 08/20/2024 10:47:10 AM, Reid, Emmerson ) tested locking the casset andit worked fine. NFF
1062094 | 209039 | CM | CMPLT | CMPLT | 2024-08-20 11:20 | 2024-08-20 11:20 | failed pressure sensor calibration valve during pre-use check | Collected logs and sent them to Draeger. As per Draeger "1. Check the diaphragm on the back of the Expiratory valve, if it is not fitted correctly, it will cause this issue. 2. Check calibration valve V6 for proper operation. Replace if necessary. (PN 8415223)" replaced expiratory valve. tested. functional.
1062094 | 209040 | PM | CMPLT | CMPLT | 2024-08-20 11:24 | 2024-08-20 11:24 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. It has PS500. Run battery check. 37 Ah and 210 min. Tested. Functional. SW.2.51 Leakage: 79 uA and 84 ohm.
1053348 | 209041 | CM | CMPLT | CMPLT | 2024-08-20 12:35 | 2024-09-25 10:57 | MSH BI 1272 H3 ERRORS WHEN SENDING, CAD ERROR MSG | (TC 08/20/2024 12:35:19 PM, Whittington, Scott) MSH BI 1272 H3 ERRORS WHEN SENDING, CAD ERROR MSG SN# 3DM160700414 REF# 01580898 CARLY REMOTE ENGINEER GREG HURLEY ABLE TO CLEAR ISSUE AND SEND 8 CASES TO SECUREVIEW AND TO JDMIPACS (TC 08/19/2024 12:35:19 PM, Whittington, Scott) MSH BI 1272 H3 ERRORS WHEN SENDING, CAD ERROR MSG SN# 3DM160700414 REF# 01580898 CARLY REMOTE ENGINEER GREG HURLEY ABLE TO CLEAR ISSUE AND SEND 8 CASES TO SECUREVIEW AND TO JDMIPACS (TC 08/20/2024 04:57:41 PM, Whittington, Scott) MSH BI 1272H3 ERRORS WHEN SENDING CAD GREG HURLEY REMOTED CLEARED | MSH BI 1272 H3 ERRORS WHEN SENDING, CAD ERROR MSG
1053348 | 209042 | CM | CMPLT | CMPLT | 2024-08-20 12:39 | 2024-09-25 10:57 | MSH BI 1272 H3 TUBE FAULTS TUBE REPACEMENT | (TC 08/20/2024 04:42:17 PM, Whittington, Scott) MSH BI 1272 H3 TUBE FAULTS TUBE REPACEMENT (TC 08/23/2024 12:55:12 PM, Dolinar, John) MSH BI 1272 H3 TUBE FAULTS TUBE REPACEMENT OHYSICS | MSH BI 1272 H3 TUBE FAULTS TUBE REPACEMENT
1053347 | 209043 | CM | ACTIV | ISSUE | 2024-08-20 12:45 |  | MSH BI 1270 H1 TOMO CAL |  | MSH BI 1270 H1 TOMO CAL
1070801 | 209044 | CM | CMPLT | CMPLT | 2024-08-20 13:25 | 2024-09-17 14:13 | Air mode not working | (TC 08/20/2024 01:32:09 PM, Rhuland, Courtney) Came down for the same error so further tests conducted After running for over an hour, omnibed alarmed stating "in transition - heater off" Canopy was not moved Reached out to getinge for tech support as it is under warranty Was unable to get answers Waiting for further support (TC 08/22/2024 11:12:40 AM, Rhuland, Courtney) Benny from Getinge came on site and assessed omnibed Attempted to adjust canopy but did not resolve isssue He will order a board and sensor and return to complete the repair (TC 09/17/2024 02:12:54 PM, Rhuland, Courtney) Benny from Getinge returned and replaced defective parts Ran incubator for a few hours and alarm did not present Calibrated RH Functional check - OK Service report attached to WO
1056665 | 209045 | CM | CMPLT | CMPLT | 2024-08-20 14:14 | 2024-08-20 14:14 | Not connecting to the AV system | (TC 08/20/2024 02:14:22 PM, Reid, Emmerson ) The microscope was not connected properly, so i fixed the connection. Image was still not being displayed so I reset the microscope and it was working afterwards.
1071891 | 209046 | ACC | CMPLT | CMPLT | 2024-08-20 14:40 | 2024-08-20 15:42 | Incoming Inspection | Brought unit down from 6th floor storage. Assigned asset tag. Checked functionality. Tested ok. Delivered back to storage and awaiting on approval for clinical use.
1064355 | 209047 | CM | CMPLT | CMPLT | 2024-08-20 14:52 | 2024-08-20 14:53 | V1, V3 leads not working | (TC 08/20/2024 02:53:17 PM, Canjimavadivel, Kamallesh ) Replaced chest lead set. Tested ok with simulator
1071892 | 209048 | ACC | CMPLT | CMPLT | 2024-08-20 15:06 | 2024-08-20 15:42 | Incoming Inspection | Brought unit down from 6th floor storage. Assigned asset tag. Checked functionality. Tested ok. Delivered back to storage and awaiting on approval for clinical use.
1071893 | 209049 | ACC | CMPLT | CMPLT | 2024-08-20 15:12 | 2024-08-20 15:42 | Incoming Inspection | Brought unit down from 6th floor storage. Assigned asset tag. Checked functionality. Tested ok. Delivered back to storage and awaiting on approval for clinical use.
1071894 | 209050 | ACC | CMPLT | CMPLT | 2024-08-20 15:20 | 2024-08-20 15:41 | Incoming Inspection | Brought unit down from 6th floor storage. Assigned asset tag. Checked functionality. Tested ok. Delivered back to storage and awaiting on approval for clinical use.
1053347 | 209051 | CM | ACTIV | ISSUE | 2024-08-20 16:44 |  | MSH BI 1270 H1 REQUIRES TOMO CAL | (TC 08/20/2024 04:44:27 PM, Whittington, Scott) MSH BI 1270 H1 REQUIRES TOMO CAL postponed (TC 08/20/2024 04:45:29 PM, Dolinar, John) MSH BI 1270 H1 REQUIRES TOMO CAL postponed | MSH BI 1270 H1 REQUIRES TOMO CAL
| 209052 | RDS | CMPLT | CMPLT | 2024-08-21 08:27 | 2024-08-21 14:45 | Weekly check for ED
| 209053 | CM | CMPLT | CMPLT | 2024-08-21 13:28 | 2024-08-21 13:30 | Broken procedure bed | (TC 08/21/2024 01:29:35 PM, Rhuland, Courtney) Side pannel and foot pannel fell off Reattached both pannels Functional check - OK
1071001 | 209054 | CM | CMPLT | CMPLT | 2024-08-21 13:42 | 2024-10-03 08:35 | Torn membrane | (TC 08/21/2024 01:43:13 PM, Rhuland, Courtney) Reseated program card to clear error Inside membrane was torn Attempted to repair - not able to Reached out to customer support to get info on repair/warranty (TC 08/23/2024 12:11:38 PM, Rhuland, Courtney) Received RMA and shipping label from Medela Packaged device and dropped it off to shipping and receiving RMA #00421626 (TC 10/03/2024 08:35:05 AM, Rhuland, Courtney) Returned from vendor - all repairs completed Unboxed and verified functionality Returned to NICU
1065480 | 209055 | CM | CMPLT | CMPLT | 2024-08-21 14:20 | 2024-08-21 14:20 | Device is falling over | (TC 08/21/2024 02:20:27 PM, Reid, Emmerson ) Roll stand base had came loose and two of the four screws were missing. I found 2 screws that fit and added lock tight to all four screws and secured the base to the device. Working now.
| 209056 | CM | CMPLT | CMPLT | 2024-08-21 15:15 | 2024-08-21 15:16 | Vital Signs Functional checks | (TC 08/21/2024 03:15:36 PM, DeGuzman, Leila) Completed Welch allyn vital signs functional check, unit to be donated or disposed.
1046035 | 209057 | CM | CMPLT | CMPLT | 2024-08-22 07:08 | 2024-08-12 07:10 | MSH ER DRX EVO E5-02(ER) SYSTEM NOT RESPONDING TO MOVEMENT COMMANDS | (TC 08/11/2024 11:08:00 AM, Whittington, Scott) MSH ER DRX EVO E5-02(ER) SYSTEM NOT RESPONDING TO MOVEMENT COMMANDS REPAIRED | MSH ER DRX EVO E5-02(ER) SYSTEM NOT RESPONDING TO MOVEMENT COMMANDS
1066454 | 209058 | FC | CMPLT | CMPLT | 2024-08-22 07:26 | 2024-08-22 07:27 | QC checklist | (TC 08/21/2024 07:26:32 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
67960 | 209059 | PM | CMPLT | CMPLT | 2024-08-22 07:28 | 2024-08-21 07:29 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/21/2024 07:28:41 AM, Baghirzade, Nara) PM completed. VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 78.6 uA Power cord resistance = 248 mohms
1061447 | 209060 | CM | CMPLT | CMPLT | 2024-08-22 07:31 | 2024-08-21 07:31 | Bed exit alarm not working properly | (TC 08/21/2024 07:31:15 AM, Baghirzade, Nara) When trying to activate the bed exit alarm, was getting the error that the patient is not centered with the arm. Took a look at the scale and saw that they did not weigh the patient. Brought 160 lbs of weights; placed weights on bed and saved the weight. Then activated the bed alarm on highest sensitivity and alarm was activating once any weight was removed. Bed exit alarm is functional. Checked all motion controls and bed is functional.
1061333 | 209061 | FC | CMPLT | CMPLT | 2024-08-22 07:34 | 2024-08-21 07:35 | QC checklist | (TC 08/21/2024 07:34:36 AM, Baghirzade, Nara) QC checklist completed. Returend to 11 North.
| 209062 | CM | CMPLT | CMPLT | 2024-08-22 08:31 | 2024-08-22 08:31 | Capital meeting | (TC 08/21/2024 08:30:59 AM, Baghirzade, Nara) Capital meeting
1070464 | 209063 | CM | CMPLT | CMPLT | 2024-08-22 08:32 | 2024-08-21 08:32 | Pump is not detecting the configuration card | (TC 08/21/2024 08:31:58 AM, Baghirzade, Nara) Reseated the configuration card. Card is detected by the pump. Pump is fully functional.
67952 | 209064 | PM | CMPLT | CMPLT | 2024-08-22 09:37 | 2024-08-22 09:39 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/22/2024 09:36:51 AM, Baghirzade, Nara) Incubator was originally sent down because of "loud noises when running the incubator"; fan was improperly seated". Fixed the positioning of the fan. PM completed. 9V rechargeable battery was replaced. VRef, Mains V, and LV were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 19.1 uA Power cord resistance = 118 mohms (TC 08/22/2024 09:39:16 AM, DeGuzman, Leila) Incubator was originally sent down because of "loud noises when running the incubator"; fan was improperly seated". Fixed the positioning of the fan. PM completed. 9V rechargeable battery was replaced. VRef, Mains V, and LV were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 80%, measured on the analyzer 80% Set temperature to 36C, measured 36C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 19.1 uA Power cord resistance = 118 mohms
67971 | 209065 | CM | CMPLT | CMPLT | 2024-08-22 09:40 | 2024-08-22 09:40 | System failure 25 | (TC 08/22/2024 09:40:01 AM, Baghirzade, Nara) Fan was improperly seated. Re-seated it and turned the incubator on; no system failure alarms. RH was calibrated twice during functional check as it was reading 4-5% higher. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75.2%RH Incubator is fully functional.
1058260 | 209066 | PM | CMPLT | CMPLT | 2024-08-22 09:43 | 2024-08-22 09:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1058268 | 209067 | PM | CMPLT | CMPLT | 2024-08-22 09:57 | 2024-08-22 09:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1058280 | 209068 | PM | CMPLT | CMPLT | 2024-08-22 09:59 | 2024-08-22 09:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1058283 | 209069 | PM | CMPLT | CMPLT | 2024-08-22 10:05 | 2024-08-22 10:12 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1059880 | 209070 | PM | CMPLT | CMPLT | 2024-08-22 10:14 | 2024-08-22 10:14 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD.
1065694 | 209071 | PM | CMPLT | CMPLT | 2024-08-22 10:16 | 2024-08-22 10:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065695 | 209072 | PM | CMPLT | CMPLT | 2024-08-22 10:17 | 2024-08-22 10:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065706 | 209073 | PM | CMPLT | CMPLT | 2024-08-22 10:18 | 2024-08-22 10:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065708 | 209074 | PM | CMPLT | CMPLT | 2024-08-22 10:20 | 2024-08-22 10:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
66434 | 209075 | CM | CMPLT | CMPLT | 2024-08-22 10:20 | 2024-08-21 10:43 | Device not turning on | (TC 08/21/2024 10:42:50 AM, Scero, Tomlyn) changed the batteries and battery adaptor, didn't fix the issue. Brought it downstairs and cleaned under the battery adaptor. Unit then turned on and used on a simulator with no issues.
1065711 | 209076 | PM | CMPLT | CMPLT | 2024-08-22 10:21 | 2024-08-22 10:21 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1065725 | 209077 | PM | CMPLT | CMPLT | 2024-08-22 10:22 | 2024-08-22 10:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065744 | 209078 | PM | CMPLT | CMPLT | 2024-08-22 10:23 | 2024-08-22 10:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065746 | 209079 | PM | CMPLT | CMPLT | 2024-08-22 10:24 | 2024-08-22 10:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065759 | 209080 | PM | CMPLT | CMPLT | 2024-08-22 10:25 | 2024-08-22 10:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065763 | 209081 | PM | CMPLT | CMPLT | 2024-08-22 10:27 | 2024-08-22 10:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD
1057522 | 209082 | CM | CMPLT | CMPLT | 2024-08-22 10:44 | 2024-08-22 10:51 | SPO2 malfunction error | (TC 08/22/2024 10:50:48 AM, Scero, Tomlyn) Confirmed issue in the room. Once the unit was swapped and brought downstairs it was placed on a simulator. No issues or errors were present. After speaking with colleagues we think there was have been a presence of cleaning fluid that just needed to dry.
1061411 | 209083 | CM | ACTIV | ISSUE | 2024-08-22 10:46 |  | Communication with controller lost | (TC 08/22/2024 10:46:08 AM, Baghirzade, Nara) Unable to use the display; error "communication with main controller lost". Power-cycled the bed but did not resolve this issue. Followed up with the nurse in charge. Contacted stryker for support.
1065786 | 209084 | PM | CMPLT | CMPLT | 2024-08-22 11:41 | 2024-08-22 11:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065790 | 209085 | PM | CMPLT | CMPLT | 2024-08-22 11:42 | 2024-08-22 11:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065790 | 209086 | PM | CANCL | DUP | 2024-08-22 11:42 | 2024-09-10 07:35 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065832 | 209087 | PM | CMPLT | CMPLT | 2024-08-22 11:43 | 2024-08-22 11:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065832 | 209088 | PM | CANCL | DUP | 2024-08-22 11:43 | 2024-09-10 07:36 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065846 | 209089 | PM | CMPLT | CMPLT | 2024-08-22 11:45 | 2024-08-22 11:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065847 | 209090 | PM | CMPLT | CMPLT | 2024-08-22 11:46 | 2024-08-22 11:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065850 | 209091 | PM | CMPLT | CMPLT | 2024-08-22 11:48 | 2024-08-22 11:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065879 | 209092 | PM | CMPLT | CMPLT | 2024-08-22 11:49 | 2024-08-22 11:49 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065895 | 209093 | PM | CMPLT | CMPLT | 2024-08-22 11:50 | 2024-08-22 11:50 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065897 | 209094 | PM | CMPLT | CMPLT | 2024-08-22 11:51 | 2024-08-22 11:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065899 | 209095 | PM | CMPLT | CMPLT | 2024-08-22 11:53 | 2024-08-22 11:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065905 | 209096 | PM | CMPLT | CMPLT | 2024-08-22 11:54 | 2024-08-22 11:54 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065907 | 209097 | PM | CMPLT | CMPLT | 2024-08-22 11:55 | 2024-08-22 11:56 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1071897 | 209098 | ACC | CMPLT | CMPLT | 2024-08-22 16:43 | 2024-08-23 13:35 | Incoming inspection | (TC 08/23/2024 01:34:37 PM, Cortez, John) Unboxed, assembled, tested functionality 150 mohms, 18 uamps
1065909 | 209099 | PM | CMPLT | CMPLT | 2024-08-23 07:11 | 2024-08-23 07:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065917 | 209100 | PM | CMPLT | CMPLT | 2024-08-23 07:12 | 2024-08-23 07:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065923 | 209101 | PM | CMPLT | CMPLT | 2024-08-23 07:14 | 2024-08-23 07:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065939 | 209102 | PM | CMPLT | CMPLT | 2024-08-23 07:15 | 2024-08-23 07:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065943 | 209103 | PM | CMPLT | CMPLT | 2024-08-23 07:16 | 2024-08-23 07:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065948 | 209104 | PM | CMPLT | CMPLT | 2024-08-23 07:17 | 2024-08-23 07:17 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065949 | 209105 | PM | CMPLT | CMPLT | 2024-08-23 07:18 | 2024-08-23 07:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065954 | 209106 | PM | CMPLT | CMPLT | 2024-08-23 07:18 | 2024-08-23 07:19 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065964 | 209107 | PM | CMPLT | CMPLT | 2024-08-23 07:19 | 2024-08-23 07:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065968 | 209108 | PM | CMPLT | CMPLT | 2024-08-23 07:20 | 2024-08-23 07:20 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065975 | 209109 | PM | CMPLT | CMPLT | 2024-08-23 07:21 | 2024-08-23 07:22 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065977 | 209110 | PM | CMPLT | CMPLT | 2024-08-23 07:23 | 2024-08-23 07:23 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065983 | 209111 | PM | CMPLT | CMPLT | 2024-08-23 07:23 | 2024-08-23 07:24 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065987 | 209112 | PM | CMPLT | CMPLT | 2024-08-23 07:25 | 2024-08-23 07:25 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065990 | 209113 | PM | CMPLT | CMPLT | 2024-08-23 07:41 | 2024-08-23 07:41 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065990 | 209114 | PM | CANCL | DUP | 2024-08-23 07:41 | 2024-09-10 07:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065991 | 209115 | PM | CMPLT | CMPLT | 2024-08-23 07:42 | 2024-08-23 07:42 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065994 | 209116 | PM | CMPLT | CMPLT | 2024-08-23 07:43 | 2024-08-23 07:43 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065994 | 209117 | PM | CANCL | DUP | 2024-08-23 07:43 | 2024-09-10 07:52 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1065997 | 209118 | PM | CMPLT | CMPLT | 2024-08-23 07:44 | 2024-08-23 07:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1065998 | 209119 | PM | CMPLT | CMPLT | 2024-08-23 07:45 | 2024-08-23 07:45 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
| 209120 | CM | CMPLT | CMPLT | 2024-08-23 07:45 | 2024-08-22 13:00 | Patient monitor Games | (TC 08/22/2024 01:00:00 PM, Chukri, Samir) patient monitor games.
1066004 | 209121 | PM | CMPLT | CMPLT | 2024-08-23 07:46 | 2024-08-23 07:46 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1066026 | 209122 | PM | CMPLT | CMPLT | 2024-08-23 07:47 | 2024-08-23 07:47 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1066028 | 209123 | PM | CMPLT | CMPLT | 2024-08-23 07:48 | 2024-08-23 07:48 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1066036 | 209124 | PM | CMPLT | CMPLT | 2024-08-23 07:50 | 2024-08-23 07:51 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68589 | 209125 | PM | CMPLT | CMPLT | 2024-08-23 07:52 | 2024-08-23 07:53 | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe | Completed by BD.
68617 | 209126 | PM | CMPLT | CMPLT | 2024-08-23 07:57 | 2024-08-23 07:57 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68728 | 209127 | PM | CMPLT | CMPLT | 2024-08-23 07:58 | 2024-08-23 07:59 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68725 | 209128 | PM | CMPLT | CMPLT | 2024-08-23 08:00 | 2024-08-23 08:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68745 | 209129 | PM | CMPLT | CMPLT | 2024-08-23 08:01 | 2024-08-23 08:01 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68836 | 209130 | PM | CMPLT | CMPLT | 2024-08-23 08:02 | 2024-08-23 08:02 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68909 | 209131 | PM | CMPLT | CMPLT | 2024-08-23 08:04 | 2024-08-23 08:04 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68920 | 209132 | PM | CMPLT | CMPLT | 2024-08-23 08:06 | 2024-08-23 08:06 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
68948 | 209133 | PM | CMPLT | CMPLT | 2024-08-23 08:08 | 2024-08-23 08:08 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69023 | 209134 | PM | CMPLT | CMPLT | 2024-08-23 08:09 | 2024-08-23 08:09 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69065 | 209135 | PM | CMPLT | CMPLT | 2024-08-23 08:10 | 2024-08-23 08:11 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69115 | 209136 | PM | CMPLT | CMPLT | 2024-08-23 08:13 | 2024-08-23 08:13 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69364 | 209137 | PM | CMPLT | CMPLT | 2024-08-23 08:14 | 2024-08-23 08:14 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69486 | 209138 | PM | CMPLT | CMPLT | 2024-08-23 08:15 | 2024-08-23 08:15 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
69585 | 209139 | PM | CMPLT | CMPLT | 2024-08-23 08:16 | 2024-08-23 08:16 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | Completed by BD.
1053391 | 209140 | CM | CMPLT | CMPLT | 2024-08-23 08:59 | 2024-08-23 08:59 | Check before use and filter change | (TC 08/23/2024 08:59:23 AM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 48 mmHg Set 125 mmHg - Measured 126 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
1059778 | 209141 | FC | CMPLT | CMPLT | 2024-08-23 09:01 | 2024-08-23 09:01 | QC checklist | (TC 08/23/2024 09:00:51 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
| 209142 | PA | ACTIV | ISSUE | 2024-08-23 10:15 |  | Ordering sterile water for NICU altrix cleaning | (TC 09/10/2024 01:32:59 PM, Baghirzade, Nara) SC 1003306548
66188 | 209143 | PM | CMPLT | CMPLT | 2024-08-23 10:20 | 2024-08-23 10:21 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 08/23/2024 10:20:36 AM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
66855 | 209144 | PM | CMPLT | CMPLT | 2024-08-23 11:02 | 2024-08-23 11:25 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/23/2024 11:25:13 AM, DeGuzman, Leila) PM passed and completed. See procedure for details.
1057527 | 209145 | CM | CMPLT | CMPLT | 2024-08-23 11:18 | 2024-08-23 11:18 | Batt malfunction | (TC 08/23/2024 11:18:20 AM, Canjimavadivel, Kamallesh ) Replaced battery
1057792 | 209146 | CM | CMPLT | CMPLT | 2024-08-23 11:21 | 2024-08-23 11:21 | Swapped out arm | (TC 08/23/2024 11:20:48 AM, Canjimavadivel, Kamallesh ) Replaced mounting arm in room P31 with different style arm
1071898 | 209147 | CM | CMPLT | CMPLT | 2024-08-23 12:20 | 2024-08-23 12:20 | Inspection | (TC 08/23/2024 12:20:13 PM, Baghirzade, Nara) Inspected the device and all the contents in its box. Created asset on TMS for inventory and recall tracking only. No physical sticker on the device.
| 209148 | RDS | CMPLT | CMPLT | 2024-08-23 12:22 | 2024-08-23 12:22 | Weekly Rounds: 11 North | (TC 08/23/2024 12:22:15 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 209149 | RDS | CMPLT | CMPLT | 2024-08-23 12:25 | 2024-08-23 12:25 | Weekly Rounds: 11 South | (TC 08/23/2024 12:24:50 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 209150 | RDS | CMPLT | CMPLT | 2024-08-23 12:27 | 2024-08-23 12:27 | Weekly Rounds: NICU & RESUS | (TC 08/23/2024 12:27:17 PM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
66432 | 209151 | CM | CMPLT | CMPLT | 2024-08-23 12:52 | 2024-08-23 13:22 | Not working | (TC 08/23/2024 01:21:49 PM, Scero, Tomlyn) Came down to biomed and wouldn't turn on. I cleaned under the battery adaptor because it was filthy. After testing with a simulator there was an LL error. I replaced the leads and tested again. Unit is working fine.
| 209152 | CM | CMPLT | CMPLT | 2024-08-23 13:52 | 2024-08-23 13:53 | Helped with Stryker bed project organization | (TC 08/23/2024 01:52:48 PM, Baghirzade, Nara) Helped Mark and Matt put supplies for the PM project in the bed shop and then met with samir, becky, and stryker team to discuss some points for the PM project.
1061212 | 209153 | CM | CMPLT | CMPLT | 2024-08-23 14:00 | 2024-08-23 14:01 | Wall monitor not working | (TC 08/23/2024 02:01:05 PM, Reid, Emmerson ) Worked with Rutvik and Jessica. We changed the cables on the Corio for the recorder and wall monitor to test the signal. This worked on the recorder still but not the wall monitor. We then changed the monitor for the one in OR12 and it was still not working. We then tried to move cable 1006 over to the wall monitor on the corio and the back of the monitor, still not working. We then connected a camera directly to the monitor and this worked. We also connected a video tower to the wall monitor output and it worked. We were unsure what to do and returned the connections back to normal and the monitor started working. We will see if this holds true because nothing was changed, except the monitor. Lastly, we setup the monitor from OR11 to OR12 and it was also working.
| 209154 | CM | CMPLT | CMPLT | 2024-08-23 15:08 | 2024-08-23 15:08 | Functional Check | (TC 08/23/2024 03:08:00 PM, DeGuzman, Leila) Performed fuctional check for a welch allyn vital sign, to be retired or donated. Device still functioning.
1069388 | 209155 | CM | CMPLT | CMPLT | 2024-08-23 16:05 | 2024-08-27 16:12 | "broken* | (TC 08/23/2024 04:05:33 PM, Cortez, John) System error code B2222 on start up, discard sequence not fully completed. Completed sequence and restarted device. Device displays error B2113 when new patient therapy is started. Manual points to issue with PBP scale? Reached out to Baxter field service tech (TC 08/27/2024 04:12:27 PM, Cortez, John) Milan (Baxter) stopped by and showed how to calibrate scales. Scale may have been bumped during transport/use Error cleared
1061958 | 209156 | CM | CMPLT | CMPLT | 2024-08-23 16:19 | 2024-08-23 16:19 | "controls not working" | (TC 08/23/2024 04:18:41 PM, Cortez, John) Lock motion controls was engaged, informed nurse
1062655 | 209157 | CM | CMPLT | CMPLT | 2024-08-23 16:22 | 2024-11-05 13:55 | "right head rail bent outwards" | (TC 08/23/2024 04:22:21 PM, Cortez, John) Reached out to Stryker for quote, Mark says $917. Will ask unit if they want to go through with repair (TC 11/05/2024 01:54:56 PM, Cortez, John) WO closed, unit will not go through repair
1063000 | 209158 | PM | CMPLT | CMPLT | 2024-08-23 16:26 | 2024-08-20 16:26 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 08/20/2024 04:25:42 PM, Cortez, John) Completed cleaning cycle
| 209159 | RDS | CMPLT | CMPLT | 2024-08-26 07:36 | 2024-08-27 09:50 | Weekly check for ED
| 209160 | OC | CMPLT | CMPLT | 2024-08-26 09:00 | 2024-08-26 09:00 | On-Call Aug 16-23
| 209161 | OC | CMPLT | CMPLT | 2024-08-26 09:00 | 2024-08-26 09:01 | On-Call Aug 26-30 | (TC 08/26/2024 09:00:33 AM, Reid, Emmerson ) Completed daily checks
| 209162 | RDS | CMPLT | CMPLT | 2024-08-26 09:01 | 2024-08-26 09:05 | Weekly Checks For PACU | (TC 08/26/2024 09:05:18 AM, Reid, Emmerson ) Rounds complete
1061207 | 209163 | CM | CMPLT | CMPLT | 2024-08-26 09:08 | 2024-08-26 09:08 | Corio Matrix not working need to call tvOne | (TC 08/26/2024 09:08:05 AM, Reid, Emmerson ) Went up to fix the issue. Tried restarting the system twice and it did not resolved the problem. Maureen was able to move the second case to another room so I will get tvOne on the phone for 1pm today (TC 08/26/2024 08:28:09 AM, Reid, Emmerson ) Worked with tvOne and fully rebuilt the corio. This did not solve the issue. I tried changing the switch and this was able to fix the issue. Reached out to olympus to get a replacement switch.
1067975 | 209164 | CM | CMPLT | CMPLT | 2024-08-26 11:30 | 2024-08-26 13:57 | i18LK5 probe overheating. | SWAPPED PROBE INTO MSH 5 US RM 8 PROBLEM FOLLOWS PROBE | Swapped probes from 5th Floor MIE1067992 to see if problem reoccurs.
66705 | 209165 | CM | CMPLT | CMPLT | 2024-08-26 12:59 | 2024-08-26 12:59 | Humidity not working | (TC 08/26/2024 12:59:20 PM, Baghirzade, Nara) When trying to set the humidity, it does not go higher than 11%. Replaced the humidity sensor with a spare from another unit; ran a functional check but the humidity sensor was also defective. Replaced the humidity sensor (new). Functional tests passed and completed. RH calibrated. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
1066462 | 209166 | FC | CMPLT | CMPLT | 2024-08-26 13:01 | 2024-08-26 13:01 | QC checklist | (TC 08/26/2024 01:01:25 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1067975 | 209167 | CM | CMPLT | CMPLT | 2024-08-26 13:51 | 2024-08-28 13:17 | MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL ERROR: OVERHEAT | (TC 08/28/2024 01:16:55 PM, Whittington, Scott) MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL ERROR: OVERHEAT | MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL ERROR: OVERHEAT From: Whittington, Scott Sent: August 26, 2024 1:49 PM To: sameh.mekhaeil@medical.canon Cc: Ultrasound Support (Ultrasound.Support@uhn.ca); Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca> Subject: MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL ERROR: OVERHEAT Hi Sameh MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 SYSTEM SITE ON443937UL ERROR: OVERHEAT Swapped probe into room 8 PRISM in Main US, the probe causes same error on that system . Please provide a replacement probe. Thanks, Scott 416-402-2462
1069295 | 209168 | CM | CMPLT | CMPLT | 2024-08-26 13:53 | 2024-08-28 11:46 | MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL ERROR: OVERHEAT | (TC 08/28/2024 11:45:15 AM, Dolinar, John) MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2472163 REPLACES SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL AUG 28 2024 SHIP FAULTY BACK TO CANON | MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL ERROR: OVERHEAT From: Whittington, Scott Sent: August 26, 2024 1:49 PM To: sameh.mekhaeil@medical.canon Cc: Ultrasound Support (Ultrasound.Support@uhn.ca); Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca> Subject: MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL ERROR: OVERHEAT Hi Sameh MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2385127 SYSTEM SITE ON443937UL ERROR: OVERHEAT Swapped probe into room 8 PRISM in Main US, the probe causes same error on that system . Please provide a replacement probe. Thanks, Scott 416-402-2462
1061554 | 209169 | CM | ACTIV | ISSUE | 2024-08-26 13:54 |  | the device shutting down automatically
502044 | 209170 | RQ | CMPLT | CMPLT | 2024-08-26 13:59 | 2024-08-26 14:49 | humidity warning on ERGO | (TC 09/12/2024 02:49:12 PM, Green, Paul) another wo was open for this.
1064648 | 209171 | CM | CMPLT | CMPLT | 2024-08-26 14:02 | 2024-08-26 14:08 | Requires calibration | (TC 08/26/2024 02:07:44 PM, Baghirzade, Nara) Downloaded error log; noted error "syringe calibration failed". Full calibration and check-in: Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
1055018 | 209172 | CM | CMPLT | CMPLT | 2024-08-26 14:03 | 2024-08-26 14:03 | Battery replacement | (TC 08/26/2024 02:03:27 PM, DeGuzman, Leila) Checked device. Replaced battery.
1054977 | 209173 | CM | CMPLT | CMPLT | 2024-08-26 14:09 | 2024-08-26 14:09 | Check functionality | (TC 08/26/2024 02:09:12 PM, DeGuzman, Leila) Checked functionality, all okay. Place a new temp probe.
| 209174 | CM | CMPLT | CMPLT | 2024-08-26 14:09 | 2024-08-26 14:09 | Biomed meeting
| 209175 | CM | CMPLT | CMPLT | 2024-08-26 14:10 | 2024-08-26 14:10 | Connecting the power cords for secure connect | (TC 08/26/2024 02:10:05 PM, Baghirzade, Nara) Placed power cords on secure connects.
| 209176 | CM | CMPLT | CMPLT | 2024-08-26 14:11 | 2024-08-26 14:11 | Connecting the power cords for secure connect | (TC 08/26/2024 02:10:57 PM, Baghirzade, Nara) Connected power cords on secure connect.
69072 | 209177 | CM | CMPLT | CMPLT | 2024-08-26 14:13 | 2024-08-26 14:13 | Didn't indicate a problem | (TC 08/26/2024 02:13:16 PM, DeGuzman, Leila) Download error log Found error code 211.2000 Performed check-in, all tests passed. Calibrated rate accuracy. Device functional. Returned to service.
1053348 | 209178 | CM | ACTIV | ISSUE | 2024-08-26 14:18 |  | MSH BI 1272 H3 ERRORS AGAIN |  | MSH BI 1272 H3 ERRORS AGAIN
69444 | 209179 | CM | CMPLT | CMPLT | 2024-08-26 14:20 | 2024-08-26 14:20 | Didn't indicate a problem | (TC 08/26/2024 02:20:23 PM, DeGuzman, Leila) Checked analog sensor on ASM, both patient and bottle side sensors need to replace. Replaced both sensors, checked values: within range. Calibrated preesure, okay. Performed check-in, all tests passed. Device functional, returned to service.
| 209180 | RDS | CMPLT | CMPLT | 2024-08-26 14:37 | 2024-08-26 14:37 | Weekly checks in CCU | (TC 08/26/2024 02:37:07 PM, Scero, Tomlyn) weekly checks complete. 30J test passed.
| 209181 | CM | CMPLT | CMPLT | 2024-08-26 14:41 | 2024-08-26 08:29 | Stryker bed SW upgrade and bed PMs | (TC 08/26/2024 08:29:22 AM, Scero, Tomlyn) Stryker bed PMs and SW upgrades on 16N, 12 bedshop, and 6 storage room.
68236 | 209182 | PM | CMPLT | CMPLT | 2024-08-26 14:51 | 2024-08-26 15:10 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/26/2024 03:09:37 PM, DeGuzman, Leila) Device came down to biomed, issue was device not working. Check all functionality: All okay. PM passed and completed. See procedure for details.
| 209183 | CM | CMPLT | CMPLT | 2024-08-26 17:07 | 2024-08-26 17:07 | Site Coverage
1062282 | 209185 | PM | CMPLT | CMPLT | 2024-08-27 08:59 | 2024-09-03 09:35 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1058529 | 209187 | CM | CMPLT | CMPLT | 2024-08-27 09:53 | 2024-08-27 09:53 | Battery replacement | (TC 08/27/2024 09:52:41 AM, Canjimavadivel, Kamallesh ) Replaced both batteries
1062322 | 209188 | PM | CMPLT | CMPLT | 2024-08-27 10:14 | 2024-08-27 10:14 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062323 | 209189 | PM | CMPLT | CMPLT | 2024-08-27 10:45 | 2024-08-27 10:47 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker
1063407 | 209191 | CM | CMPLT | CMPLT | 2024-08-27 13:49 | 2024-08-27 13:49 | Exp. cassette membrance replacement | (TC 08/27/2024 01:48:43 PM, Canjimavadivel, Kamallesh ) Replaced membrance and updated count
1064212 | 209192 | CM | CMPLT | CMPLT | 2024-08-27 13:52 | 2024-08-27 13:52 | Exp. cassette membrane replacement | (TC 08/27/2024 01:52:24 PM, Canjimavadivel, Kamallesh ) Replaced membrane and updated count
1066454 | 209193 | FC | CMPLT | CMPLT | 2024-08-27 14:33 | 2024-08-27 14:34 | QC checklist | (TC 08/27/2024 02:34:03 PM, Rhuland, Courtney) QC checklist completed Returned to 11S
1061333 | 209194 | FC | CMPLT | CMPLT | 2024-08-27 14:37 | 2024-08-27 14:37 | QC checklist | (TC 08/27/2024 02:37:15 PM, Rhuland, Courtney) QC checklist completed Returned to 11N
1069389 | 209195 | CM | CMPLT | CMPLT | 2024-08-27 16:17 | 2024-08-27 16:17 | "needs service" | (TC 08/27/2024 04:17:17 PM, Cortez, John) Error B2112, effluent scale issue Similar to other asset sent down last week Milan (Baxter) came by and showed how to calibrate scale
66825 | 209196 | PM | CMPLT | CMPLT | 2024-08-28 09:12 | 2024-08-28 10:38 | PM-MAJOR - Annual PM For: Portable Patient Modules X2 | (TC 08/28/2024 10:37:44 AM, DeGuzman, Leila) PM passed and completed. See procedure for details. NIBP calibrated.
1054523 | 209197 | CM | CMPLT | CMPLT | 2024-08-28 10:06 | 2024-08-28 10:06 | V1 and V3 leads not working | (TC 08/28/2024 10:06:19 AM, Canjimavadivel, Kamallesh ) Replaced V1-V3 chest lead set. Tested ok with simulator.
| 209198 | PA | CMPLT | CMPLT | 2024-08-28 10:18 | 2024-09-04 13:11 | Zoll training sim and pads | (TC 08/28/2024 10:28:29 AM, Canjimavadivel, Kamallesh ) Placed order for training simulator equipment (TC 09/04/2024 01:10:31 PM, Canjimavadivel, Kamallesh ) Received training accessories
| 209199 | PA | CMPLT | CMPLT | 2024-08-28 10:31 | 2024-09-04 13:09 | Philips ecg leads for stock | (TC 08/28/2024 10:36:08 AM, Canjimavadivel, Kamallesh ) Placed order for ECG lead sets (TC 09/04/2024 01:09:14 PM, Canjimavadivel, Kamallesh ) Received cables and added to stock
1069295 | 209200 | ACC | CMPLT | CMPLT | 2024-08-28 11:06 | 2024-08-28 11:47 | MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2472163 REPLACES SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL | (TC 08/28/2024 11:08:34 AM, Whittington, Scott) MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2472163 REPLACES SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL AUG 28 2024 (TC 08/28/2024 11:08:43 AM, Dolinar, John) MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2472163 REPLACES SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL AUG 28 2024 | MSH US ER i800 PRISM PROBE PLI 1205BX SN# FSD2472163 REPLACES SN# FSD2385127 mie 1067975 SYSTEM SITE ON443937UL AUG 28 2024
| 209201 | CM | CMPLT | CMPLT | 2024-08-28 11:27 | 2024-08-28 11:28 | Functional Check | (TC 08/28/2024 11:27:31 AM, DeGuzman, Leila) Performed functional check on CVSM 6000 and VSM 300. To be retired or donated.
1068221 | 209202 | CM | ACTIV | ISSUE | 2024-08-28 11:50 |  | MSH US RM 6 DATABASE CORRUPTION | (TC 08/28/2024 07:49:00 AM, Whittington, Scott) MSH US RM 6 DATABASE CORRUPTION: missing images in study, mwl and pacs not responding rebuild db, (TC 08/28/2024 11:51:24 AM, Whittington, Scott) MSH US RM 6 DATABASE CORRUPTION: missing images in study, mwl and pacs not responding rebuild db, | MSH US RM 6 DATABASE CORRUPTION
1061411 | 209203 | PM | CMPLT | CMPLT | 2024-08-28 11:56 | 2024-09-03 08:57 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
68632 | 209204 | CM | CMPLT | CMPLT | 2024-08-28 13:36 | 2024-08-28 13:37 | Check-in | (TC 08/28/2024 01:36:55 PM, DeGuzman, Leila) Came down no indicated issue. Found IUI female pins were broken. Replaced IUI female. Downloaded error log and event log. Found issue for bottle-side sensor high voltage. Checked analog sensor, value was high voltage. Replaced the bottle-side sensor. Check sensor, Okay. Performed check-in tests, all passed. Return to service.
69453 | 209205 | CM | CMPLT | CMPLT | 2024-08-28 13:52 | 2024-08-28 13:53 | Check in | (TC 08/28/2024 01:52:33 PM, DeGuzman, Leila) Came down no indicated issue. Download error log and event log. No issue found. Performed check-in tests, all passed. Cleared error and event log. Return to service.
1061993 | 209206 | PM | CMPLT | CMPLT | 2024-08-28 14:13 | 2024-09-16 09:58 | PM-MAJOR-ANNUAL PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.08 ohm, 143 uA
65961 | 209207 | CM | CMPLT | CMPLT | 2024-08-28 14:40 | 2024-08-28 14:40 | Calibration required | (TC 08/28/2024 02:40:24 PM, Rhuland, Courtney) Calibration and check-in completed All tests passed Refer to ASM for details
65910 | 209208 | CM | CMPLT | CMPLT | 2024-08-28 14:44 | 2024-08-28 14:44 | Calibration required | (TC 08/28/2024 02:44:24 PM, Rhuland, Courtney) Calibration and check-in completed All tests passed Refer to ASM for details
65944 | 209209 | CM | CMPLT | CMPLT | 2024-08-28 14:46 | 2024-08-28 14:47 | Channel error | (TC 08/28/2024 02:46:53 PM, Rhuland, Courtney) Channel error 354.6770 Replaced pressure sensor Performed calibration and check-in All tests passed Refer to ASM for details
68934 | 209210 | CM | CMPLT | CMPLT | 2024-08-28 15:10 | 2024-08-28 15:10 | Check-in | (TC 08/28/2024 03:10:06 PM, DeGuzman, Leila) Came down no indicated issue. Checked device, all tests passed. Electrical safety: Okay Ground resistance: 114 mOhm Leakage current: 7.5 uA
502088 | 209211 | RQ | CMPLT | CMPLT | 2024-08-28 15:25 | 2024-09-12 14:48 | intermittent touch pad error, won't clear | (TC 09/12/2024 02:48:06 PM, Green, Paul) another wo was created for this.
32510 | 209212 | CM | ACTIV | ISSUE | 2024-08-28 15:46 |  | MSH DR1 CONING, TABLE & 4 FOOT |  | MSH DR1 CONING, TABLE & 4 FOOT Hi Mel Thanks Mel, Matt made me aware earlier today. I will check it out, and hopefully make an improvement Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Vaillancourt, Melanie <Melanie.Vaillancourt@sinaihealth.ca> Sent: Wednesday, August 28, 2024 2:49:26 PM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca> Subject: [External] Coning in DR1 Hi, good afternoon… The coning in DR1 has been off and we’ve needed to repeat/reject on larger patients because of it. Was wondering if it would be possible to have it fixed? See A#859376961 (4’ legs), 859376529 supine abdomen and 859378415 AP pelvis for reference. It is worse on 4’ legs and table work with grid from what I’ve experienced. Thanks, Mel
| 209213 | CM | CMPLT | CMPLT | 2024-08-28 17:25 | 2024-08-28 17:25 | Site coverage and removing Downtime on various US machines. |  | Site coverage and removing Downtime on various US machines.
1058181 | 209214 | CM | CMPLT | CMPLT | 2024-08-29 09:14 | 2024-08-29 09:22 | Tag - Always asks to power down after putting in settings. | (TC 08/29/2024 09:21:57 AM, DeGuzman, Leila) User error. Checked device, all tests passed. See procedure for details.
1057940 | 209215 | CM | CMPLT | CMPLT | 2024-08-29 10:05 | 2024-09-18 15:01 | Cushion Plate Broken | Aug 29 - call re: bed attachment broken. Went to check and brought it down to biomed. Contacted vendor for purchase info. Created SC and sent it for approval. Received. Issued PO via MMC. Sept 18 - Received parts. Replaced it.. Stored two cushion plates in biomed for future use. Sept 18 - Call re two OR Table accessories not working. Went to check. FOund faulty arm board and leg mount. Brought it down to biomed. Fixed Arm board and leg mount. Tested ok. Delivered back to OR.
| 209216 | PA | ACTIV | ISSUE | 2024-08-29 10:20 |  | Altrix parts for stock | (TC 08/29/2024 10:19:47 AM, Canjimavadivel, Kamallesh ) Looked up various part numbers in manual and reached out to vendor for quote (TC 09/03/2024 12:42:51 PM, Canjimavadivel, Kamallesh ) Placed order for parts (TC 09/25/2024 10:52:53 AM, Canjimavadivel, Kamallesh ) Received some items from order.
1062999 | 209217 | CM | CMPLT | CMPLT | 2024-08-29 11:18 | 2024-08-29 11:18 | "intermittently working" | (TC 08/29/2024 11:18:13 AM, Cortez, John) Amanda reported that blankets were inflating/deflating intermittently One of the hoses was leaking slightly when I was draining it, but didn't leak when I ran the warmer. Ran for an hour, different temps, switched ports. Unable to replicate issue
1054035 | 209218 | CM | CMPLT | CMPLT | 2024-08-29 13:31 | 2024-10-30 11:06 | "loss of ultrasound data" | (TC 08/29/2024 01:30:38 PM, Cortez, John) Sending original probe and console for repair (TC 09/05/2024 10:57:04 AM, Cortez, John) shipped, RMA CAS-05222-G7X4M7 (TC 10/30/2024 11:05:39 AM, Cortez, John) Returned from CHS. Tested functionality with phantom.
1049362 | 209219 | CM | CMPLT | CMPLT | 2024-08-29 14:43 | 2024-08-29 14:43 | MSH CT Injector - Side A Not Working | (TC 08/29/2024 02:42:47 PM, Chau, Ken) Side A syringe attached keeps sensing CLeaned and tested okay Informed staff
| 209220 | CM | ACTIV | ISSUE | 2024-08-29 14:56 |  | MSH IVF APLIO 300 3D IMAGING | (TC 08/29/2024 02:55:58 PM, Whittington, Scott) MSH IVF APLIO 300 3D IMAGING | MSH IVF APLIO 300 3D IMAGING
1054694 | 209221 | CM | CMPLT | CMPLT | 2024-08-29 15:49 | 2024-08-29 15:49 | Tag - Not activating | (TC 08/29/2024 03:49:21 PM, DeGuzman, Leila) Cleaned device. Power cord broken. Checked device, no fault found. Replaced power cord. Return to service.
502088 | 209222 | CM | CMPLT | CMPLT | 2024-08-29 18:24 | 2024-09-03 10:57 | HD2 gantry touch pad | (TC 08/29/2024 06:24:06 PM, Green, Paul) Troubleshot problem. HD2 inner touchpad shorted. Was able to setup workaround until part arrives. Looked up part number and ordered from Siemens. (TC 09/03/2024 03:47:32 PM, Green, Paul) Part has not arrived. Called Siemens and received waybill. Tracking number indicates it will be there later today. (TC 09/03/2024 03:49:01 PM, Green, Paul) Contacted customer and they have time to install. Travelled to site and removed tocuhpad and cleaned up cover. Installed new, cabled and covered. Tested and ok.
| 209223 | CM | CMPLT | CMPLT | 2024-08-30 07:29 | 2024-08-27 07:29 | Bed PMs with Stryker | (TC 08/27/2024 07:28:41 AM, Baghirzade, Nara) Shadowing stryker team for bed PMs.
| 209224 | CM | CMPLT | CMPLT | 2024-08-30 07:30 | 2024-08-27 07:30 | Bed PMs with styrker | (TC 08/27/2024 07:29:50 AM, Baghirzade, Nara) Shadowing stryker team for bed PMs.
1071912 | 209225 | ACC | CMPLT | CMPLT | 2024-08-30 07:55 | 2024-08-30 11:37 | Incoming inspection and installation | (TC 08/28/2024 07:55:46 AM, Baghirzade, Nara) Incoming inspection completed. Installed device with Natus rep. Functional test completed. Server connectivity is successful. (TC 08/30/2024 11:36:53 AM, Baghirzade, Nara) Completed all TMS information and asset tagging.
1071913 | 209226 | ACC | CMPLT | CMPLT | 2024-08-30 07:55 | 2024-08-30 11:38 | Incoming inspection and installation | (TC 08/28/2024 07:57:15 AM, Baghirzade, Nara) Incoming inspection completed. Installed device with Natus rep. Functional test completed. Server connectivity is successful. (TC 08/30/2024 11:37:44 AM, Baghirzade, Nara) Completed all TMS information and asset tagging.
1071914 | 209227 | ACC | CMPLT | CMPLT | 2024-08-30 07:55 | 2024-08-30 11:38 | Incoming inspection and installation | (TC 08/28/2024 07:57:45 AM, Baghirzade, Nara) Incoming inspection completed. Installed device with Natus rep. Functional test completed. Server connectivity is successful. (TC 08/30/2024 11:38:15 AM, Baghirzade, Nara) Completed all TMS information and asset tagging.
1071915 | 209228 | ACC | CMPLT | CMPLT | 2024-08-30 07:55 | 2024-08-30 11:39 | Incoming inspection and installation | (TC 08/28/2024 07:58:06 AM, Baghirzade, Nara) Incoming inspection completed. Installed device with Natus rep. Functional test completed. Server connectivity is successful. (TC 08/30/2024 11:38:44 AM, Baghirzade, Nara) Completed all TMS information and asset tagging.
| 209229 | CM | CMPLT | CMPLT | 2024-08-30 08:01 | 2024-08-29 08:01 | Bed PMs with Stryker | (TC 08/29/2024 08:01:26 AM, Baghirzade, Nara) Shadowing styker team for bed PMs.
| 209230 | CM | CMPLT | CMPLT | 2024-08-30 08:02 | 2024-08-29 08:02 | Bed PMs with Stryker | (TC 08/29/2024 08:02:16 AM, Baghirzade, Nara) Shadowing styker team for bed PMs.
52252 | 209231 | PM | CMPLT | CMPLT | 2024-08-30 11:07 | 2024-08-30 11:08 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 08/30/2024 11:07:39 AM, DeGuzman, Leila) Came down with the broken hinge on isolette door. Replaced the hinge. PM completed. 9V rechargeable battery was replaced Nov. 2023 VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. Scale was calibrated. Calibrated the scale. Scale is functional and accurate; set 5kg, measured = 5kg. Set RH to 75%, measured on the analyzer 75% Set temperature to 35C, measured 35C on all temperature points. Functional tests passed.
1071920 | 209232 | CM | CMPLT | CMPLT | 2024-08-30 11:40 | 2024-08-30 11:40 | Asset tagging and input information on TMS | (TC 08/30/2024 11:40:03 AM, Baghirzade, Nara) Completed all TMS information and asset tagging.
1071921 | 209233 | CM | CMPLT | CMPLT | 2024-08-30 11:40 | 2024-08-30 11:41 | Asset tagging and input information on TMS | (TC 08/30/2024 11:40:34 AM, Baghirzade, Nara) Completed all TMS information and asset tagging.
1071922 | 209234 | ACC | CMPLT | CMPLT | 2024-08-30 11:43 | 2024-08-30 11:43 | Incoming inspection | (TC 08/30/2024 11:43:04 AM, Baghirzade, Nara) Incoming inspection completed. Completed all TMS information and asset tagging. Functional tests passed. Dropped off the unit with Steve.
| 209235 | CM | CMPLT | CMPLT | 2024-08-30 11:44 | 2024-08-30 11:44 | NICU PICIX hardware sheet | (TC 08/30/2024 11:44:13 AM, Baghirzade, Nara) Completed the NICU PICIX Hardware list as per Kamran's request.
| 209236 | CM | CMPLT | CMPLT | 2024-08-30 11:46 | 2024-08-28 11:46 | Meeting for Bed PM WOs | (TC 08/28/2024 11:46:03 AM, Baghirzade, Nara) Meeting for Bed PM WOs
1054996 | 209237 | CM | CMPLT | CMPLT | 2024-08-30 12:51 | 2024-08-30 12:51 | Temp probe broken | (TC 08/30/2024 12:50:46 PM, DeGuzman, Leila) Temperature probe wire worn out. Replaced temperature probe. Checked functions, all tests passed. Returned to service.
1055041 | 209238 | CM | CMPLT | CMPLT | 2024-08-30 12:56 | 2024-08-30 12:56 | Battery Replacement | (TC 08/30/2024 12:55:54 PM, DeGuzman, Leila) Came down attached with the root asset #1054996 Checked the battery, its buldge or swollen. Replaced battery, and updated. Checked function, okay. Returned to service.
67454 | 209239 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
67455 | 209240 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
26231 | 209241 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
52324 | 209242 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Portable Patient Transport
52438 | 209243 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Portable Patient Transport
60165 | 209244 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Radiofrequency Generators
62928 | 209245 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-19 13:07 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 09/19/2024 01:06:30 PM, Rhuland, Courtney) Shipped out to Spectramedx for PM Returned and all tests passed Verified O2 values 21% - 20.9% 40% - 39.8% 60% - 60.9% 80% - 80.9% 100% - 99.9% Spectramedx report attached
62968 | 209246 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-31 10:28 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 10/31/2024 10:28:13 AM, Rhuland, Courtney) PM completed by Spectramedx Returned and functional check completed 21% O2 - 20.9% O2 40% O2 - 38.7% O2 60% O2 - 61.2% O2 80% O2 - 82% O2 100% O2 - 99.6% O2 Returned to RT
62982 | 209247 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-02 09:58 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/02/2024 09:57:32 AM, Alvarado, Karla) Replaced timer 0.0 hours Intensity light: - High: 34.7 uW/cm2/nm - Low: 14.9 uW/cm2/nm Target light is functional GND Resistance: 0.101 ohms Leakage current: 25.2 uA PM completed. (TC 10/02/2024 09:57:32 AM, Rhuland, Courtney) Supervised
63026 | 209248 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-05 07:34 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 09/05/2024 07:33:48 AM, Baghirzade, Nara) Vacuumed the back filter. Physical inspection ok. Irradiance: 46uw/cm2/nm Hours run: 5376.8 Electrical safety: Power cord resistance - 200 mohms Leakage current - 21.1 uA PM completed.
63048 | 209249 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-10 14:52 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 10/10/2024 02:52:25 PM, Rhuland, Courtney) PM completed by Spectramedx 21% - 20.9% 40% - 38.1% 60% - 60.8% 80% - 79.5% 100% - 98.7%
63124 | 209250 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
63133 | 209251 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
64071 | 209252 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM MAJOR Annual: Power Systems, Uninterruptible
64088 | 209253 | PM | CANCL | CANCL | 2024-09-01 03:22 | 2024-09-16 09:02 | PM MAJOR: Annual - Power Systems, Uninterruptible | Battery was replaced in 2023. Battery replacement is every 2 years.
64216 | 209254 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-21 12:47 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/17/2024 12:38:24 PM, Alvarado, Karla) NIBP- Simulated: 120/80 Results: 123/81 Simulated: 80/40 Results: 86/47 Simulated: 190/120 Results: 190/121 ECG- cable is in good condition. Simulated 60 BPM ECG Signal Resulted in 60 BPM, clear QRS complex. Simulated 90 BPM ECG Signal Resulted in 90 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 97% Simulated: 85% Results: 87% (triggered low SpO2 alarm) Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Good physical condition. Screen in good condition. All buttons working and responsive. Updated date and time. PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. Device did not work on battery. Therefore, a new battery was ordered. (TC 10/21/2024 12:46:50 PM, Rhuland, Courtney) Replaced battery Device ran on battery and battery icon appears on screen Electrical safety - OK Ground resistance - 173mOhms Leakage current - 20.3 uA
64868 | 209255 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM MAJOR - ANNUAL PM for: Warmers
64910 | 209256 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: High Flow Insufflator
64960 | 209257 | PM | ACTIV | READY | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS
64961 | 209258 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-04 15:17 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/04/2024 03:17:16 PM, Alvarado, Karla) NIBP- Simulated: 120/80 Results: 124/81 Simulated: 80/40 Results: 88/54 Simulated: 190/120 Results: 193/121 ECG- cable is in good condition. Simulated 60 BPM ECG Signal Resulted in 60 BPM, clear QRS complex. Simulated 90 BPM ECG Signal Resulted in 90 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 96% Simulated: 85% Results: 86% (triggered low SpO2 alarm) Battery is in good condition. Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Device operates when unplugged from AC, on battery. Good physical condition. Screen in good condition. All buttons working and responsive. Electrical Safety: Current leakage = 34.2 uA Power cord resistance = 239 mohms Updated date and time. PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. (TC 10/04/2024 03:17:16 PM, Rhuland, Courtney) Supervised
65153 | 209259 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65164 | 209260 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65173 | 209261 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-26 11:02 | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic | (TC 09/26/2024 11:01:45 AM, Reid, Emmerson ) The screen is not working properly, there is a line of pixels that is not working. I spoke with Wendy and Eden of PACU and included Andrew to determine if this should be repaired or retired. Based on the age of the device and cost to repair, this unit is to be retired.
65292 | 209262 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
65591 | 209263 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-04 15:14 | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS | (TC 10/04/2024 03:14:18 PM, Alvarado, Karla) Performed PM NIBP- Simulated: 120/80 Results: 123/82 Simulated: 80/40 Results: 85/51 Simulated: 190/120 Results: 191/121 ECG- cable is in good condition. Simulated 60 BPM ECG Signal Resulted in 60 BPM, clear QRS complex. Simulated 90 BPM ECG Signal Resulted in 90 BPM, clear QRS complex. Simulated arrhythmia ECG Signal Arrythmia detected on waveform clearly. ECG leads off alarm ok. SpO2- probe and cable are in good condition. Simulated: 99% Results: 99% Simulated: 95% Results: 96% Simulated: 85% Results: 86% (triggered low SpO2 alarm) Battery is in good condition. Printer checked and functional. AC indicator LED functional and lit when connecting device to AC power. Device operates when unplugged from AC, on battery. Good physical condition. Screen in good condition. All buttons working and responsive. Electrical Safety: Current leakage = 32.1 uA Power cord resistance = 280 mohms Updated date and time. PM tests passed and completed. All tests passed as per manufacturer specifications and recommendations. (TC 10/04/2024 03:14:18 PM, Rhuland, Courtney) Supervised
65592 | 209264 | PM | ACTIV | READY | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: VITAL SIGNS MONITORS
65639 | 209265 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-06 10:15 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | replaced rotary knob. Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 34 Ah and 200 min. Tested. Functional. SW.2.51 Leakage: 91 uA and 96 ohm.
65749 | 209266 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM Major Annual: Microstream CO2 extension
66024 | 209267 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-11-05 14:27 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/05/2024 02:27:00 PM, Baghirzade, Nara) Replaced front case and logic board. Replaced power cord. Applied software upgrade. Transferred network configuration and dataset files. PM passed and completed. Battery was replaced. Electrical Safety - Ground resistance: 47 mohms Leakage current: 57.7 uA Error and event logs downloaded and cleared. IUIs and device cleaned. PM schedule updated.
66460 | 209268 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-11 11:32 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 22 Ah and 230 min. Tested. Functional. SW.2.51 Leakage: 88 uA and 95 ohm.
66465 | 209269 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-17 12:41 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 31 Ah and 230 min. Tested. Functional. SW.2.51.01 Leakage: 88 uA and 99 ohm.
66474 | 209270 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-18 14:09 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. but could not finish. It needs a SW upgrade. Tested. Functional. SW.2.51 Leakage: 91 uA and 95 ohm.
66522 | 209271 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66696 | 209272 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-04 10:08 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/04/2024 10:08:06 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
66701 | 209273 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66719 | 209274 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66729 | 209275 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66732 | 209276 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66738 | 209277 | PM | CANCL | CANCL | 2024-09-01 03:22 | 2024-09-24 07:50 | PM-MAJOR - Annual PM For: Giraffe Shuttle | PM completed in Jan - updated PM schedule
66771 | 209278 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
66802 | 209279 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67236 | 209280 | PM | CANCL | CANCL | 2024-09-01 03:22 | 2024-09-24 07:49 | PM-MAJOR - Annual PM For: Infant Resuscitation System | PM completed in Jan - updated PM schedule
67766 | 209281 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67950 | 209282 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-23 14:57 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/23/2024 02:56:43 PM, Rhuland, Courtney) PM completed Repaired side door latch Calibrated: Scale Mains & VREF voltage Functional check - OK Set 35C, 75% humidity Read 35C, 75.8% humidity Electrical safety - OK Leakage current - 83 uA Ground resistance - 99 mOhms
68277 | 209283 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
68374 | 209284 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68378 | 209285 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68542 | 209286 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68554 | 209287 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68713 | 209288 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69050 | 209289 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69111 | 209290 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69190 | 209291 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69197 | 209292 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69309 | 209293 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69575 | 209294 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-11-21 10:44 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/21/2024 10:43:54 AM, Cortez, John) Updated software PM completed successfully, results logged in ASM
69657 | 209295 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69667 | 209296 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
65553 | 209297 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
500847 | 209298 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-16 07:41 | MSH APLIO XG PM - Semi Annual | (TC 10/16/2024 07:40:44 AM, Whittington, Scott) MSH APLIO XG PM - Semi Annual (TC 10/21/2024 02:51:09 PM, Whittington, Scott) MSH APLIO XG PM - Semi Annual | MSH APLIO XG PM - Semi Annual
500850 | 209299 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-16 07:41 | MSH APLIO XG PM - Semi Annual PM - Semi-Annual | (TC 10/16/2024 07:41:13 AM, Whittington, Scott) MSH APLIO XG PM - Semi Annual PM - Semi-Annual (TC 10/21/2024 02:51:39 PM, Whittington, Scott) 209299 | MSH APLIO XG PM - Semi Annual PM - Semi-Annual
32510 | 209300 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-16 07:33 | MSH DR1 PM - Semi-Annual | (TC 10/16/2024 07:33:37 AM, Whittington, Scott) MSH DR1 PM - Semi-Annual | MSH DR1 PM - Semi-Annual
32510 | 209301 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-23 20:22 | MSH DR1 HARP OCT 17 TRGTSemi Annual | (TC 10/16/2024 08:00:00 AM, Whittington, Scott) harp (TC 11/11/2024 01:25:26 PM, Whittington, Scott) MSH DR1 HARP FOLLOWUP | MSH DR1 HARP TRGT OCT 17 Semi AnnualMSH DR1 HARP TRGT OCT 17 Semi Annual
32540 | 209302 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-06 09:45 | PM - Monthly | PM completed Annual QC testing done
32560 | 209303 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-05 09:39 | PM - Monthly | PM completed Annual QC testing done
32580 | 209304 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-05 09:42 | PM - Monthly | PM completed Annual QC testing done
500800 | 209305 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-17 10:33 | MSH BIUS APLIO XG PM - Semi-Annual | (TC 09/16/2024 05:43:44 PM, Whittington, Scott) PMH BIUS APLIO XG PM - Semi-Annual (TC 09/17/2024 10:33:00 AM, Whittington, Scott) PMH BIUS APLIO XG PM - Semi-Annual (TC 09/17/2024 10:33:00 AM, Whittington, Scott) MSH BIUS APLIO XG PM - Semi-Annual (TC 09/17/2024 10:33:00 AM, Whittington, Scott) MSH BIUS APLIO XG PM - Semi-Annual | MSH BIUS APLIO XG PM - Semi-Annual
500801 | 209306 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-17 10:34 | MSH BIUS APLIO XG PM - Semi-Annual | (TC 09/17/2024 10:33:00 AM, Whittington, Scott) MSH BIUS APLIO XG PM - Semi-Annual | MSH BIUS APLIO XG PM - Semi-Annual
500802 | 209307 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-09-17 10:36 | MSH BIUS APLIO XG PM - Semi-Annual | (TC 09/16/2024 05:44:45 PM, Whittington, Scott) MSH BIUS APLIO XG PM - Semi-Annual (TC 09/16/2024 05:44:45 PM, Whittington, Scott) MSH BIUS APLIO XG PM - Semi-Annual (TC 09/17/2024 10:44:00 AM, Whittington, Scott) MSH BIUS APLIO XG PM - Semi-Annual (TC 09/17/2024 10:44:00 AM, Whittington, Scott) MSH BIUS APLIO XG PM - Semi-Annual | MSH BIUS APLIO XG PM - Semi-Annual
500857 | 209308 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-16 07:42 | MSH APLIO XG PM - Semi Annual PM - | (TC 10/16/2024 07:41:53 AM, Whittington, Scott) MSH APLIO XG PM - Semi Annual PM - (TC 10/21/2024 02:52:10 PM, Whittington, Scott) 209308 | MSH APLIO XG PM - Semi Annual PM -
501621 | 209309 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM - Annual
501888 | 209310 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-16 07:42 | MSH APLIO XG PM - Semi Annual PM - | (TC 10/16/2024 07:42:16 AM, Whittington, Scott) MSH APLIO XG PM - Semi Annual PM - | MSH APLIO XG PM - Semi Annual PM -
1046035 | 209311 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-21 14:20 | HARP - Semi Annual
1046035 | 209312 | PM | CMPLT | CMPLT | 2024-09-01 03:22 | 2024-10-21 14:19 | PM - Semi Annual | (TC 10/21/2024 02:18:57 PM, Fan, Tak) PM completed. Replaced 5 button OTC button board.
1046087 | 209313 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
1046295 | 209314 | PM | ACTIV | READY | 2024-09-01 03:22 |  | Annual PM
1046625 | 209315 | PM | ACTIV | READY | 2024-09-01 03:22 |  | Annual PM
1047362 | 209316 | PM | ACTIV | PV | 2024-09-01 03:22 |  | PM-MAJOR- Annual PM for: Testers
1048664 | 209317 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
1049677 | 209318 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
1049891 | 209319 | PM | ACTIV | ISSUE | 2024-09-01 03:22 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
1053349 | 209320 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-16 17:19 | MSH BI 1274 HARP OCT 15 TRGT SEMI-ANNUAL |  | Harp completed.
1054460 | 209321 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-09 10:31 | Lubrication cycle
1054462 | 209322 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-09 10:31 | Lubrication cycle
1054468 | 209323 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-09 10:32 | Lubrication cycle
1055801 | 209324 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1056125 | 209325 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-30 11:36 | PM-MAJOR-Annual PM For: HYPO/HYPERTHERMIA UNITS | (TC 10/29/2024 07:57:05 AM, Canjimavadivel, Kamallesh ) Replaced all flow and temp sensors. Replaced PM parts. (TC 10/30/2024 11:36:24 AM, Canjimavadivel, Kamallesh ) Completed cleaning cycle. All maintenance tests pass. Returned to storage room
1056934 | 209326 | PM | CANCL | CANCL | 2024-09-01 03:23 | 2024-09-05 08:10 | PM Major Annual: Physiologic Monitor Modules, Exhaled Carbon Dioxide | PM schedule shifted to align with rest in ED
1056943 | 209327 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056944 | 209328 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056945 | 209329 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056946 | 209330 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056947 | 209331 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056948 | 209332 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056949 | 209333 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056950 | 209334 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056951 | 209335 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056952 | 209336 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056953 | 209337 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056954 | 209338 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056955 | 209339 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056956 | 209340 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056957 | 209341 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056958 | 209342 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056959 | 209343 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056960 | 209344 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056961 | 209345 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056962 | 209346 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056963 | 209347 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056964 | 209348 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056965 | 209349 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056966 | 209350 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056967 | 209351 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056968 | 209352 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056969 | 209353 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056970 | 209354 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056971 | 209355 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056972 | 209356 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056973 | 209357 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056974 | 209358 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056975 | 209359 | PM | ACTIV | SHP | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for Bed electric
1056986 | 209360 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1056992 | 209361 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-28 13:37 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/28/2024 01:36:46 PM, Canjimavadivel, Kamallesh ) PM tests pass
1056993 | 209362 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1056994 | 209363 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1056995 | 209364 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1056996 | 209365 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1056997 | 209366 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1056998 | 209367 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1056999 | 209368 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1057000 | 209369 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1057001 | 209370 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/19/2024 02:35:34 PM, Canjimavadivel, Kamallesh ) PM tests pass
1057002 | 209371 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1057003 | 209372 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-06 11:04 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/06/2024 11:04:07 AM, Canjimavadivel, Kamallesh ) PM tests pass, battery replaced
1057005 | 209373 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-26 12:17 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/26/2024 12:17:29 PM, Canjimavadivel, Kamallesh ) Replaced battery. PM tests pass
1057007 | 209374 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 10:36 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 09/03/2024 10:35:52 AM, Canjimavadivel, Kamallesh ) Replaced battery. Verified NiBP accuracy, calibration not required
1057008 | 209375 | PM | CMPLT | PMFAIL | 2024-09-01 03:23 | 2024-11-06 13:46 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/06/2024 01:46:19 PM, Canjimavadivel, Kamallesh ) Msmt ext malfunction error when turned on
1057009 | 209376 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1057010 | 209377 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-21 14:20 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/21/2024 02:20:06 PM, Canjimavadivel, Kamallesh ) PM tests pass. Battery not installed, X3 to be kept as a spare.
1057012 | 209378 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-06 10:31 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/06/2024 10:30:55 AM, Canjimavadivel, Kamallesh ) PM tests
1057013 | 209379 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-13 09:53 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/13/2024 09:53:14 AM, Canjimavadivel, Kamallesh ) PM tests pass
1057014 | 209380 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1057015 | 209381 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-26 15:40 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/26/2024 03:39:39 PM, Canjimavadivel, Kamallesh ) Replaced battery. PM tests pass
1057016 | 209382 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-06 13:10 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/06/2024 01:09:39 PM, Canjimavadivel, Kamallesh ) PM tests pass
1057028 | 209383 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Power Systems, Uninterruptible
1057036 | 209384 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-26 14:10 | PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/26/2024 02:10:26 PM, Canjimavadivel, Kamallesh ) PM tests pass
1057083 | 209385 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-Major - Annual PM for: Wamer
1057221 | 209386 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major - Annual: Physiologic Monitor Modules, Exhaled Carbon Dioxide
1057304 | 209387 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM Major Annual: Monitors, Physiologic, Patient Transport
1057839 | 209388 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057863 | 209389 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | Annual PM - For Fluid Warmer
1058117 | 209390 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM for : Table, Operating
1058175 | 209391 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058179 | 209392 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1058606 | 209393 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-02 08:01 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 08:01:25 AM, Reid, Emmerson ) PM complete, see attached service report.
1061554 | 209394 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-02 08:07 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 08:07:16 AM, Reid, Emmerson ) PM complete, see attached service report.
1062095 | 209395 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-20 08:19 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. It has PS500. Run battery check, 210 min, 21 Ah and 25 min 2 Ah. Tested. Functional. SW.2.51 Leakage: 91 uA and 91 ohm.
1062096 | 209396 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-08 10:46 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. It has PS500. Run battery check, 210 min, 34 Ah. Tested. Functional. SW.2.51 Leakage: 85 uA and 89 ohm.
1062097 | 209397 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-09 11:00 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. It has PS500. Run battery check, 210 min, 31 Ah. Tested. Functional. SW.2.51 Leakage: 87 uA and 91 ohm.
1062098 | 209398 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-10-16 13:28 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. It has PS500. Run battery check, 220 min, 32 Ah and 25 min 2 Ah. Tested. Functional. SW.2.51 Leakage: 89 uA and 93 ohm.
1062102 | 209399 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Pressure-Vacuum Digital Meter
1062118 | 209400 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1062124 | 209401 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062125 | 209402 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062126 | 209403 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-16 09:55 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.09 ohm, 141uA
1062127 | 209404 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062128 | 209405 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 10:37 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 10:37:14 AM, Ferreira, Qwynn) Completed by Stryker
1062129 | 209406 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062130 | 209407 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-16 10:25 | PM-MAJOR-Annual PM For: Med Surg Bed
1062131 | 209408 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-16 09:55 | PM-MAJOR-Annual PM For: Med Surg Bed
1062132 | 209409 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062133 | 209410 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062134 | 209411 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062135 | 209412 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062136 | 209413 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 11:09 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:08:38 AM, Ferreira, Qwynn) Completed by Stryker
1062137 | 209414 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062138 | 209415 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062139 | 209416 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062140 | 209417 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062141 | 209418 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:46 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062142 | 209419 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062143 | 209420 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062144 | 209421 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-22 08:46 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/22/2024 08:46:11 AM, Baghirzade, Nara) Completed by Stryker.
1062145 | 209422 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062146 | 209423 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062147 | 209424 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062168 | 209425 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 11:25 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:25:17 AM, Ferreira, Qwynn) Completed by Stryker
1062169 | 209426 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-13 09:10 | PM-MAJOR-Annual PM For: Med Surg Bed | PM complete by Stryker. battery changed and SW updated. 0.07 ohm, 142 uA
1062170 | 209427 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 10:34 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 10:33:47 AM, Ferreira, Qwynn) Completed by Stryker
1062171 | 209428 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-27 08:38 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.01 ohm, 123 uA
1062172 | 209429 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062173 | 209430 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062174 | 209431 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 13:29 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 01:29:25 PM, Ferreira, Qwynn) Completed by Stryker
1062175 | 209432 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062176 | 209433 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062177 | 209434 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 11:11 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:11:29 AM, Ferreira, Qwynn) Completed by Stryker
1062178 | 209435 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-27 14:07 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by stryker.
1062179 | 209436 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-25 08:25 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Batteries changed and SW updated . 0.02 ohm, 98 uA.
1062232 | 209437 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1062233 | 209438 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1062235 | 209439 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 11:23 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:23:04 AM, Ferreira, Qwynn) Completed by Stryker
1062236 | 209440 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062237 | 209441 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:48 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062238 | 209442 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062239 | 209443 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-18 13:19 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/18/2024 01:18:39 PM, Baghirzade, Nara) PM completed by Stryker. Updated network information on asset page.
1062240 | 209444 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-25 08:15 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.04 ohm, 134 uA
1062241 | 209445 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062242 | 209446 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 11:06 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:06:21 AM, Ferreira, Qwynn) Completed by Stryker
1062243 | 209447 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-25 08:27 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Batteries changed and SW updated . 0.03 ohm, 125 uA.
1062244 | 209448 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:01 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062245 | 209449 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 10:36 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 10:35:45 AM, Ferreira, Qwynn) Completed by Stryker
1062246 | 209450 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062247 | 209451 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 08:42 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062248 | 209452 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-18 13:25 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/18/2024 01:24:43 PM, Baghirzade, Nara) PM completed by Stryker. Updated network information on asset page.
1062253 | 209453 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062254 | 209454 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062255 | 209455 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-13 09:07 | PM-MAJOR-Annual PM For: Med Surg Bed | PM complete by Stryker. battery changed and SW updated. 0.08 ohm, 128 uA
1062256 | 209456 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062257 | 209457 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 10:23 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062258 | 209458 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-18 14:11 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/18/2024 02:11:00 PM, Baghirzade, Nara) PM completed by Stryker. Updated network information on asset page.
1062259 | 209459 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-25 08:33 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Batteries changed and SW updated . 0.01 ohm, 126 uA.
1062260 | 209460 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-25 08:16 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Batteries changed and SW updated . 0.01 ohm, 136 uA.
1062261 | 209461 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:07 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062262 | 209462 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:50 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062263 | 209463 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:57 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062264 | 209464 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:42 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062265 | 209465 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-13 14:46 | PM-MAJOR-Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1062266 | 209466 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062269 | 209467 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-21 12:20 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/21/2024 12:20:13 PM, Baghirzade, Nara) PM completed by Stryker. Updated network information on asset page.
1062270 | 209468 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062271 | 209469 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-22 08:53 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/22/2024 08:53:12 AM, Baghirzade, Nara) Completed by Stryker.
1062272 | 209470 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 11:31 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:31:06 AM, Ferreira, Qwynn) Completed by Stryker
1062273 | 209471 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062274 | 209472 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-18 13:11 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/18/2024 01:11:05 PM, Baghirzade, Nara) PM completed by Stryker. Updated network information on asset page.
1062275 | 209473 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062276 | 209474 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062277 | 209475 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 11:05 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:04:51 AM, Ferreira, Qwynn) Completed by Stryker
1062278 | 209476 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062279 | 209477 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062280 | 209478 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-25 08:19 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Batteries changed and SW updated . 0.01 ohm, 123 uA.
1062281 | 209479 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062309 | 209480 | PM | ACTIV | ISSUE | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
1062312 | 209481 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062313 | 209482 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 08:18 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062314 | 209483 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:44 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062315 | 209484 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 14:04 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 02:03:22 PM, Ferreira, Qwynn) Completed by Stryker (TC 11/07/2024 02:03:55 PM, Ferreira, Qwynn) Completed by Stryker
1062316 | 209485 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-13 14:19 | PM-MAJOR-Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1062317 | 209486 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062318 | 209487 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062319 | 209488 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062320 | 209489 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-16 08:41 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.02 ohm, 128 uA
1062321 | 209490 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-16 08:44 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.01 ohm, 128 uA
1062324 | 209491 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 10:48 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 10:48:24 AM, Ferreira, Qwynn) Completed by Stryker
1062325 | 209492 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:11 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062388 | 209493 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 10:44 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062389 | 209494 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-18 13:08 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/18/2024 01:07:58 PM, Baghirzade, Nara) PM completed by Stryker. Updated network information on asset page.
1062390 | 209495 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-22 09:19 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/22/2024 09:18:33 AM, Baghirzade, Nara) Completed by Stryker.
1062391 | 209496 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062392 | 209497 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 08:48 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062393 | 209498 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 10:30 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062394 | 209499 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-03 09:31 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062395 | 209500 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-11-07 13:55 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 01:55:07 PM, Ferreira, Qwynn) Completed by Stryker
1062396 | 209501 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-25 08:28 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Batteries changed and SW updated . 0.01 ohm, 125 uA.
1062397 | 209502 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062398 | 209503 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062399 | 209504 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062400 | 209505 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062401 | 209506 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062435 | 209507 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-08-29 10:19 | PM-MAJOR-Annual PM For: Med Surg Bed
1062436 | 209508 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-09-16 10:26 | PM-MAJOR-Annual PM For: Med Surg Bed
1062437 | 209509 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062438 | 209510 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062439 | 209511 | PM | ACTIV | READY | 2024-09-01 03:23 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062440 | 209512 | PM | CMPLT | CMPLT | 2024-09-01 03:23 | 2024-08-30 09:26 | PM-MAJOR-Annual PM For: Med Surg Bed
1062441 | 209513 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062442 | 209514 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 10:16 | PM-MAJOR-Annual PM For: Med Surg Bed
1062443 | 209515 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 10:17 | PM-MAJOR-Annual PM For: Med Surg Bed
1062444 | 209516 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-11-07 11:27 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 11:27:00 AM, Ferreira, Qwynn) Completed by Stryker
1062445 | 209517 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062446 | 209518 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-03 08:37 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062447 | 209519 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 09:41 | PM-MAJOR-Annual PM For: Med Surg Bed
1062448 | 209520 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 09:28 | PM-MAJOR-Annual PM For: Med Surg Bed
1062543 | 209521 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062544 | 209522 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062545 | 209523 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062546 | 209524 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-11-07 14:17 | PM-MAJOR-Annual PM For: Med Surg Bed | (TC 11/07/2024 02:16:59 PM, Ferreira, Qwynn) Completed by Stryker
1062547 | 209525 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062548 | 209526 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-13 09:04 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.06 ohm, 71uA
1062549 | 209527 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-25 08:31 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Batteries changed and SW updated . 0.42 ohm, 62 uA.
1062550 | 209528 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 10:23 | PM-MAJOR-Annual PM For: Med Surg Bed
1062551 | 209529 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062552 | 209530 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 10:27 | PM-MAJOR-Annual PM For: Med Surg Bed
1062553 | 209531 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062554 | 209532 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 10:29 | PM-MAJOR-Annual PM For: Med Surg Bed
1062555 | 209533 | PM | ACTIV | READY | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM For: Med Surg Bed
1062557 | 209534 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-05 13:29 | PM-MAJOR-Annual PM For: Med Surg Bed | PM complete by Stryker team
1062558 | 209535 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 10:09 | PM-MAJOR-Annual PM For: Med Surg Bed
1062604 | 209536 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-09-16 10:11 | PM-MAJOR-Annual PM for Bed electric
1062605 | 209537 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062606 | 209538 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062607 | 209539 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062608 | 209540 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062609 | 209541 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062610 | 209542 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062655 | 209543 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062657 | 209544 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1062658 | 209545 | PM | ACTIV | SHP | 2024-09-01 03:24 |  | PM-MAJOR-Annual PM for Bed electric
1063193 | 209546 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-16 07:43 | MSH APLIO XG PM - Semi Annual PM - | (TC 10/16/2024 07:42:52 AM, Whittington, Scott) MSH APLIO XG PM - Semi Annual PM - (TC 10/21/2024 02:53:28 PM, Whittington, Scott) 209546 | MSH APLIO XG PM - Semi Annual PM -
1064735 | 209547 | PM | ACTIV | ISSUE | 2024-09-01 03:24 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1064736 | 209548 | PM | ACTIV | ISSUE | 2024-09-01 03:24 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
1064964 | 209549 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-10 21:21 | HARP - Semi-Annual | HARP
1064964 | 209550 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-09 19:19 | PM - Semi-Annual | PM done.
1064973 | 209551 | PM | ACTIV | ISSUE | 2024-09-01 03:24 |  | Annual PM
1065513 | 209552 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-02 07:52 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 07:51:32 AM, Reid, Emmerson ) PM complete, see attached service report.
1065514 | 209553 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-02 07:56 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 07:56:08 AM, Reid, Emmerson ) PM complete, see attached service report.
1065515 | 209554 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-02 07:59 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 07:59:23 AM, Reid, Emmerson ) PM complete, see attached service report.
1065516 | 209555 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-02 07:49 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 07:48:44 AM, Reid, Emmerson ) PM complete, see attached service report.
1066474 | 209556 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-02 08:15 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 08:15:22 AM, Reid, Emmerson ) PM complete, see attached service report.
1066475 | 209557 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-02 08:13 | PM-MAJOR - Annual PM For: Waste-Disposal Units
1066476 | 209558 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-02 08:10 | PM-MAJOR - Annual PM For: Waste-Disposal Units | (TC 10/02/2024 08:10:07 AM, Reid, Emmerson ) PM complete, see attached service report.
1066756 | 209559 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-09 10:32 | Lubrication cycle
1066757 | 209560 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-09 10:32 | Lubrication cycle
1067975 | 209561 | PM | CMPLT | CMPLT | 2024-09-01 03:24 | 2024-10-16 07:45 | MSH ER G7 PRISM PM - Semiannual | (TC 10/16/2024 07:44:36 AM, Whittington, Scott) MSH ER G7 PRISM PM - Semiannual (TC 10/21/2024 02:54:12 PM, Whittington, Scott) 209561 | MSH ER G7 PRISM PM - Semiannual
| 209562 | PM | CANCL | CANCL | 2024-09-01 03:24 | 2024-10-21 15:01 | PM-MAJOR - Annual PM For: Hemodialysis Units |  | PM Annual or operating time 6000 hours
1069389 | 209563 | PM | CANCL | CANCL | 2024-09-01 03:24 | 2024-10-21 15:04 | PM-MAJOR - Annual PM For: Hemodialysis Unit |  | Annual PM or Operting time 6000 hours
1062322 | 209564 | PM | CMPLT | CMPLT | 2024-09-03 08:09 | 2024-09-03 08:09 | Batteries Replacement Due | Completed by Stryker.
1061412 | 209565 | PM | CMPLT | CMPLT | 2024-09-03 08:12 | 2024-09-03 08:14 | Batteries Replacement Due | Completed by Stryker.
1061412 | 209566 | PM | CMPLT | CMPLT | 2024-09-03 08:16 | 2024-09-03 08:16 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062313 | 209567 | PM | CMPLT | CMPLT | 2024-09-03 08:18 | 2024-09-03 08:18 | Batteries Replacement Due | Completed by Stryker.
1061430 | 209568 | PM | CMPLT | CMPLT | 2024-09-03 08:21 | 2024-09-03 08:21 | Batteries Replacement Due | Completed by Stryker.
1061430 | 209569 | PM | CMPLT | CMPLT | 2024-09-03 08:22 | 2024-09-03 08:23 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062323 | 209570 | PM | CMPLT | CMPLT | 2024-09-03 08:25 | 2024-09-03 08:25 | Batteries Replacement Due | Completed by Stryker.
1062556 | 209571 | PM | CMPLT | CMPLT | 2024-09-03 08:27 | 2024-09-03 08:28 | Batteries Replacement Due | Completed by Stryker.
1061424 | 209572 | PM | CMPLT | CMPLT | 2024-09-03 08:29 | 2024-09-03 08:29 | Batteries Replacement Due | Completed by Stryker.
1061424 | 209573 | PM | CMPLT | CMPLT | 2024-09-03 08:31 | 2024-09-03 08:32 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062446 | 209574 | PM | CMPLT | CMPLT | 2024-09-03 08:38 | 2024-09-03 08:38 | Batteries Replacement Due | Completed by Stryker.
1062502 | 209575 | PM | CMPLT | CMPLT | 2024-09-03 08:40 | 2024-09-03 08:40 | Batteries Replacement Due | Completed by Stryker.
1062502 | 209576 | PM | CMPLT | CMPLT | 2024-09-03 08:41 | 2024-09-03 08:41 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062247 | 209577 | PM | CMPLT | CMPLT | 2024-09-03 08:43 | 2024-09-03 08:47 | Batteries Replacement Due | Completed by Stryker.
1062392 | 209578 | PM | CMPLT | CMPLT | 2024-09-03 08:49 | 2024-09-03 08:49 | Batteries Replacement Due | Completed by Stryker.
1061411 | 209579 | PM | CMPLT | CMPLT | 2024-09-03 08:57 | 2024-09-03 08:58 | Batteries Replacement Due | Completed by Stryker.
1062244 | 209580 | PM | CMPLT | CMPLT | 2024-09-03 09:02 | 2024-09-03 09:02 | Batteries Replacement Due | Completed by Stryker.
1061466 | 209581 | PM | CMPLT | CMPLT | 2024-09-03 09:04 | 2024-09-03 09:04 | Batteries Replacement Due | Completed by Stryker.
1061466 | 209582 | PM | CMPLT | CMPLT | 2024-09-03 09:05 | 2024-09-03 09:05 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
| 209583 | RDS | CMPLT | CMPLT | 2024-09-03 09:05 | 2024-09-03 10:00 | Weekly checks in CCU | (TC 09/03/2024 10:00:16 AM, Scero, Tomlyn) Weekly checks complete. 30J tests done. SpO2 cable damaged so I will order new ones.
1062261 | 209584 | PM | CMPLT | CMPLT | 2024-09-03 09:08 | 2024-09-03 09:09 | Batteries Replacement Due | Completed by Stryker.
1062325 | 209585 | PM | CMPLT | CMPLT | 2024-09-03 09:11 | 2024-09-03 09:12 | Batteries Replacement Due | Completed by Stryker.
| 209586 | CM | CMPLT | CMPLT | 2024-09-03 09:19 | 2024-10-08 09:20 | Yellow Fin Repair | Sept 03 - Found Yello Fin (S/N: A248381) was disposed by OR staff. Was requested by Reeba to take it out and send it for repair. Contacted vendor and requested RMA. Received. Created SC. Sent for approval. Recevied. Issued PO via MMC. Sept 25 - Followed up. Received repair quote. Requested approval. Received. Advised MMC to update the price and issue the PO. Awaiting delivery. Oct 07 - Received unit. Unboxed. Checked functionality. Tested ok. Delivered back to OR. Report attached.
1062066 | 209587 | PM | CMPLT | CMPLT | 2024-09-03 09:23 | 2024-09-03 09:23 | Batteries Replacement Due | Completed by Stryker.
1062066 | 209588 | PM | CMPLT | CMPLT | 2024-09-03 09:24 | 2024-09-03 09:24 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062072 | 209589 | PM | CMPLT | CMPLT | 2024-09-03 09:25 | 2024-09-03 09:26 | Batteries Replacement Due | Completed by Stryker.
1062072 | 209590 | PM | CMPLT | CMPLT | 2024-09-03 09:26 | 2024-09-03 09:26 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062090 | 209591 | PM | CMPLT | CMPLT | 2024-09-03 09:28 | 2024-09-03 09:28 | Batteries Replacement Due | Completed by Stryker.
1062090 | 209592 | PM | CMPLT | CMPLT | 2024-09-03 09:28 | 2024-09-03 09:28 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062089 | 209593 | PM | CMPLT | CMPLT | 2024-09-03 09:29 | 2024-09-03 09:29 | Batteries Replacement Due | Completed by Stryker.
1062089 | 209594 | PM | CMPLT | CMPLT | 2024-09-03 09:30 | 2024-09-03 09:30 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062394 | 209595 | PM | CMPLT | CMPLT | 2024-09-03 09:32 | 2024-09-03 09:32 | Batteries Replacement Due | Completed by Stryker.
1062282 | 209596 | PM | CMPLT | CMPLT | 2024-09-03 09:36 | 2024-09-03 09:36 | Batteries Replacement Due | Completed by Stryker.
1061447 | 209597 | PM | CMPLT | CMPLT | 2024-09-03 09:38 | 2024-09-03 09:38 | Batteries Replacement Due | Completed by Stryker.
1061447 | 209598 | PM | CMPLT | CMPLT | 2024-09-03 09:39 | 2024-09-03 09:39 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1061442 | 209599 | PM | CMPLT | CMPLT | 2024-09-03 09:41 | 2024-09-03 09:41 | Batteries Replacement Due | Completed by Stryker.
1061442 | 209600 | PM | CMPLT | CMPLT | 2024-09-03 09:41 | 2024-09-03 09:42 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062264 | 209601 | PM | CMPLT | CMPLT | 2024-09-03 09:43 | 2024-09-03 09:43 | Batteries Replacement Due | Completed by Stryker.
1062314 | 209602 | PM | CMPLT | CMPLT | 2024-09-03 09:45 | 2024-09-03 09:45 | Batteries Replacement Due | Completed by Stryker.
1062141 | 209603 | PM | CMPLT | CMPLT | 2024-09-03 09:46 | 2024-09-03 09:47 | Batteries Replacement Due | Completed by Stryker.
1062237 | 209604 | PM | CMPLT | CMPLT | 2024-09-03 09:48 | 2024-09-03 09:48 | Batteries Replacement Due | Completed by Stryker.
1062262 | 209605 | PM | CMPLT | CMPLT | 2024-09-03 09:50 | 2024-09-03 09:50 | Batteries Replacement Due | Completed by Stryker.
1062022 | 209606 | PM | CMPLT | CMPLT | 2024-09-03 09:53 | 2024-09-03 09:53 | Batteries Replacement Due | Completed by Stryker.
1062022 | 209607 | PM | CMPLT | CMPLT | 2024-09-03 09:54 | 2024-09-03 09:54 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062503 | 209608 | PM | CMPLT | CMPLT | 2024-09-03 09:55 | 2024-09-03 09:56 | Batteries Replacement Due | Completed by Stryker.
1062503 | 209609 | PM | CMPLT | CMPLT | 2024-09-03 09:56 | 2024-09-03 09:56 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1050672 | 209610 | CM | CMPLT | CMPLT | 2024-09-03 10:01 | 2024-09-03 10:01 | V3 not working | (TC 09/03/2024 10:00:51 AM, Canjimavadivel, Kamallesh ) Confirmed issue, replaced V1-V3 chest lead. Tested ok with simulator
1062263 | 209611 | PM | CMPLT | CMPLT | 2024-09-03 10:02 | 2024-09-03 10:02 | Batteries Replacement Due | Completed by Stryker.
501794 | 209612 | CM | CMPLT | CMPLT | 2024-09-03 10:04 | 2024-09-04 14:03 | Connector corner bent in, cannot connect probe. | (TC 09/03/2024 10:04:22 AM, Fan, Tak) C10-3v (SN: B37317) connector corner bent in, cannot connect probe. The case # is 0124203282. (TC 09/04/2024 02:02:51 PM, Fan, Tak) B37317 replaced by F0BL77 on Sep 4, 2024
1062257 | 209613 | PM | CMPLT | CMPLT | 2024-09-03 10:28 | 2024-09-03 10:28 | Batteries Replacement Due | Completed by Stryker.
1062393 | 209614 | PM | CMPLT | CMPLT | 2024-09-03 10:31 | 2024-09-03 10:32 | Batteries Replacement Due | Completed by Stryker.
| 209615 | CM | CMPLT | CMPLT | 2024-09-03 10:34 | 2024-09-13 13:18 | Hamilton Vent battery order | (TC 09/03/2024 10:34:06 AM, Cortez, John) ordered x2 (TC 09/13/2024 01:18:10 PM, Cortez, John) Arrived, batteries conditioned
1057007 | 209616 | CM | CMPLT | NFF | 2024-09-03 10:35 | 2024-09-03 10:35 | Not working | (TC 09/03/2024 10:34:45 AM, Canjimavadivel, Kamallesh ) Sent down with all hoses and cables. Tested functions, all passed. No fault found
1061435 | 209617 | PM | CMPLT | CMPLT | 2024-09-03 10:38 | 2024-09-03 10:38 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1061435 | 209618 | PM | CMPLT | CMPLT | 2024-09-03 10:38 | 2024-09-03 10:39 | Batteries Replacement Due | Completed by Stryker.
1061960 | 209619 | PM | CMPLT | CMPLT | 2024-09-03 10:40 | 2024-09-03 10:40 | Batteries Replacement Due | Completed by Stryker.
1061960 | 209620 | PM | CMPLT | CMPLT | 2024-09-03 10:41 | 2024-09-03 10:41 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062505 | 209621 | PM | CMPLT | CMPLT | 2024-09-03 10:42 | 2024-09-03 10:42 | Batteries Replacement Due | Completed by Stryker.
1062505 | 209622 | PM | CMPLT | CMPLT | 2024-09-03 10:43 | 2024-09-03 10:43 | PM-MAJOR-Annual PM For: Med Surg Bed | Completed by Stryker.
1062388 | 209623 | PM | CMPLT | CMPLT | 2024-09-03 10:45 | 2024-09-03 10:45 | Batteries Replacement Due | Completed by Stryker.
| 209624 | PA | CMPLT | CMPLT | 2024-09-03 10:45 | 2024-09-23 09:16 | SpO2 spare cables for stock | (TC 09/03/2024 11:19:50 AM, Scero, Tomlyn) Ordered SPO2 cables for stock (TC 09/23/2024 09:15:49 AM, Scero, Tomlyn) All items received
| 209625 | RDS | CMPLT | CMPLT | 2024-09-03 12:44 | 2024-09-03 12:46 | Daily check for ED
| 209626 | RDS | CMPLT | CMPLT | 2024-09-03 12:47 | 2024-09-03 14:32 | Daily check for ICU
1059778 | 209627 | CM | CMPLT | CMPLT | 2024-09-03 13:10 | 2024-09-03 13:10 | Front and rear case are misaligned | (TC 09/03/2024 01:10:19 PM, Baghirzade, Nara) Wound VAC rear and front cases do not align and there was a big gap. Disassembled the cases and realigned it; assembled it back. Will run QC to ensure unit is functional.
1059778 | 209628 | FC | CMPLT | CMPLT | 2024-09-03 13:11 | 2024-09-03 13:12 | QC checklist | (TC 09/03/2024 01:11:31 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1060024 | 209629 | CM | CMPLT | CMPLT | 2024-09-03 13:14 | 2024-09-03 13:14 | CMOS battery replacement | (TC 09/03/2024 01:13:57 PM, Baghirzade, Nara) Time and date were not being retained by the PC. Replaced CMOS battery. Logged in the computer and checked that boot up is normal. Updated time. Had Jessica (eeg tech) login and check also.
| 209630 | CM | CMPLT | CMPLT | 2024-09-03 13:16 | 2024-09-10 14:24 | MX40 leads | (TC 09/03/2024 01:42:59 PM, Scero, Tomlyn) ordered replacement leads for CCU (TC 09/10/2024 02:24:13 PM, Scero, Tomlyn) parts received
| 209631 | CM | ACTIV | ISSUE | 2024-09-03 13:16 |  | Closing PM WOs for Stryker beds | (TC 09/03/2024 01:16:09 PM, Baghirzade, Nara) Closing PM WOs for Stryker beds on 11 North, 11 South, and Antenatal. (TC 09/06/2024 12:49:22 PM, Baghirzade, Nara) Added MAC addressess to TMS/Assets network information tab. (TC 11/18/2024 07:46:35 AM, Baghirzade, Nara) Closed PM WOs, update network info and PM schedules.
1056936 | 209632 | CM | CMPLT | NFF | 2024-09-03 13:18 | 2024-09-03 13:18 | Won't connect | (TC 09/03/2024 01:17:36 PM, Canjimavadivel, Kamallesh ) Attached to X3 and turned on. Module was detected, tested CO2 function as well. No fault found
66475 | 209633 | CM | CMPLT | CMPLT | 2024-09-03 14:44 | 2024-09-03 14:44 | Screen is white; defective | (TC 09/03/2024 02:44:25 PM, Baghirzade, Nara) Tried to reboot and power cycle; screen remains white and is not responsive. Removed a functional display from a spare ECG cart and replaced it with the defective display. Display is functional. ECG functional checks passed. Returned ECG to Antenatal.
| 209634 | CM | CMPLT | CMPLT | 2024-09-03 15:05 | 2024-09-03 15:05 | Pcard statement. | (TC 09/03/2024 03:04:45 PM, Fan, Tak) Aug statement for tools.
1066917 | 209635 | CM | CMPLT | CMPLT | 2024-09-03 15:35 | 2024-09-27 13:26 | "right, head side rail won't lock in upright position" | (TC 09/03/2024 03:45:53 PM, Cortez, John) Put in service request with Stryker (TC 09/27/2024 01:25:37 PM, Cortez, John) Aziz (stryker) repaired bed Service report attached
1046276 | 209636 | CM | ACTIV | ISSUE | 2024-09-03 16:40 |  | MSH CFRH 7flr APLIO 300 UNABLE TO BOOT TO WINDOWS CORRUPT SOFTWARE | (TC 09/03/2024 04:40:14 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 UNABLE TO BOOT TO WINDOWS CORRUPT SOFTWARE COPY BOTH SYSTEM & DATA DISK FROM SAFE RESERVE ABLE TO BOOT UP RELOAD SETTINGS USE IP 172.30.30.20, SHOULD BE CORRECT DUE TO SEQUENTIAL ORDER (TC 09/03/2024 04:40:14 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 UNABLE TO BOOT TO WINDOWS CORRUPT SOFTWARE COPY BOTH SYSTEM & DATA DISK FROM SAFE RESERVE ABLE TO BOOT UP RELOAD SETTINGS USE IP 172.30.30.20, SHOULD BE CORRECT DUE TO SEQUENTIAL ORDER NEED TO VERIFY PT SHOW ON MWL TOMORROW (TC 09/03/2024 04:40:14 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 UNABLE TO BOOT TO WINDOWS CORRUPT SOFTWARE COPY BOTH SYSTEM & DATA DISK FROM SAFE RESERVE ABLE TO BOOT UP RELOAD SETTINGS USE IP 172.30.30.20, SHOULD BE CORRECT DUE TO SEQUENTIAL ORDER NEED TO VERIFY PT SHOW ON MWL TOMORROW (TC 09/03/2024 04:42:18 PM, Whittington, Scott) MSH CFRH 7flr APLIO 300 UNABLE TO BOOT TO WINDOWS CORRUPT SOFTWARE COPY BOTH SYSTEM & DATA DISK FROM SAFE RESERVE ABLE TO BOOT UP RELOAD SETTINGS USE IP 172.30.30.20, SHOULD BE CORRECT DUE TO SEQUENTIAL ORDER NEED TO VERIFY PT SHOW ON MWL TOMORROW (TC 09/03/2024 04:42:28 PM, Dolinar, John) MSH CFRH 7flr APLIO 300 UNABLE TO BOOT TO WINDOWS CORRUPT SOFTWARE COPY BOTH SYSTEM & DATA DISK FROM SAFE RESERVE ABLE TO BOOT UP RELOAD SETTINGS USE IP 172.30.30.20, SHOULD BE CORRECT DUE TO SEQUENTIAL ORDER NEED TO VERIFY PT SHOW ON MWL TOMORROW | MSH CFRH 7flr APLIO 300 UNABLE TO BOOT TO WINDOWS CORRUPT SOFTWARE
| 209637 | CM | ACTIV | ISSUE | 2024-09-03 17:04 |  | MSH IVF CANON PROBE 3D SN# 99117112273 BENT PIN ON CONNECTOR , POSSIBLE OTHER FAULTS | (TC 09/03/2024 05:03:49 PM, Whittington, Scott) MSH IVF CANON PROBE 3D SN# 99117112273 BENT PIN ON CONNECTOR , POSSIBLE OTHER FAULTS (TC 09/03/2024 05:05:37 PM, Dolinar, John) MSH IVF CANON PROBE 3D SN# 99117112273 BENT PIN ON CONNECTOR , POSSIBLE OTHER FAULTS (TC 09/05/2024 03:03:00 PM, Whittington, Scott) SEP MSH IVF CANON PROBE 3D SN# 99117112273 BENT PIN ON CONNECTOR , POSSIBLE OTHER FAULTS | MSH IVF CANON PROBE 3D SN# 99117112273 BENT PIN ON CONNECTOR , POSSIBLE OTHER FAULTS
1057868 | 209638 | CM | CMPLT | CMPLT | 2024-09-04 08:21 | 2024-09-17 13:52 | Ordering Reservoir Covers | (TC 09/06/2024 07:57:14 AM, Reid, Emmerson ) Got the price from the Rep and placed an order on erec (TC 09/17/2024 01:52:03 PM, Reid, Emmerson ) Covers have been recieved.
1059583 | 209639 | CM | CMPLT | CMPLT | 2024-09-04 09:33 | 2024-09-04 09:37 | Not working | (TC 09/04/2024 09:36:54 AM, Canjimavadivel, Kamallesh ) Passed calibration with tester after one attempt
1062152 | 209640 | CM | CMPLT | CMPLT | 2024-09-04 09:39 | 2024-09-04 09:39 | Not working | (TC 09/04/2024 09:38:57 AM, Canjimavadivel, Kamallesh ) Passed calibration after 2nd attempt
1054957 | 209641 | CM | CMPLT | CMPLT | 2024-09-04 09:40 | 2024-09-04 09:40 | Not reading accurately | (TC 09/04/2024 09:40:24 AM, Canjimavadivel, Kamallesh ) Passed calibration after 2nd attempt
65291 | 209642 | PM | CMPLT | CMPLT | 2024-09-04 09:45 | 2024-09-04 14:27 | PM-MAJOR - Annual PM For: WARMERS, BLOOD | (TC 09/04/2024 02:27:04 PM, Reid, Emmerson ) replaced the gasket and added the silicone to help with the seal. This stopped the leak. Completed PM, see procedure for details.
1053330 | 209643 | CM | CMPLT | CMPLT | 2024-09-04 10:28 | 2024-08-26 10:28 | Battery error | (TC 08/26/2024 10:28:29 AM, Cortez, John) Replaced battery
1062309 | 209644 | CM | CMPLT | CMPLT | 2024-09-04 10:30 | 2024-08-27 10:31 | Battery needs replacement | (TC 08/27/2024 10:30:42 AM, Cortez, John) Replaced battery
1053330 | 209645 | CM | CMPLT | CMPLT | 2024-09-04 10:40 | 2024-09-04 10:40 | Watchdog timer alarming | (TC 09/04/2024 10:39:48 AM, Canjimavadivel, Kamallesh ) Brought vent down to investigate alarm. Alarm stopped once plugged in. Battery was loose in holder. Secured battery cover and checked logs. Found multiple instances of battery removed/reinserted in past week. Returned to transport incubator and informed RT
502044 | 209646 | CM | CMPLT | CMPLT | 2024-09-04 11:15 | 2024-09-04 10:58 | Replacement of laptop | (TC 09/04/2024 11:14:42 AM, Green, Paul) Replacement laptop arrived. Powered up and backed up original system. Recorded networking and dicom nformation. (TC 09/12/2024 02:46:33 PM, Green, Paul) Restored backups on replacement laptop. Contacted Digirad to restore Worklist license. Setup network and dicom nodes. Asked clinical group for some time on system to replace. (TC 09/13/2024 02:42:28 PM, Green, Paul) From: Green, Paul <Paul.Green@uhn.ca> Sent: September 13, 2024 2:37 PM To: Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca> Cc: Green, Paul <Paul.Green@uhn.ca>; Paz, Ramon <Ramon.Paz@uhn.ca>; Mangabat, Mark Christopher <MarkChristopher.Mangabat@uhn.ca> Subject: MSH Ergo replacement laptop Barry, Ergo replacement laptop was pugged into MSH network and is located in the Symbia control room. The below ticket was opened to have the MAC address added to the MAC filter. Once this is completed, we will schedule a date to replace the laptop. Regards, Paul From: Green, Paul <Paul.Green@uhn.ca> Sent: September 13, 2024 9:49 AM To: Green, Paul <Paul.Green@uhn.ca> Subject: Ticket number 33621354 Paul Green, CET (TC 09/18/2024 04:25:18 PM, Green, Paul) Checked into ticket and still not resolved. could not access new laptop online. Checked on system and no network access as of yet. Still waiting for MAC filtering to be enabled for laptop. (TC 10/25/2024 02:28:10 PM, Green, Paul) Calibrations files were moved to new laptop, however ran into problems/corrupt calibration files on the laptop/system. Called Digirad and worked with them on this. They were able to restore old calibrations files and resolve the holes that we had on our system. Performed an Energy calibration and uniformity map. QC numbers are now with in specification.
1070842 | 209647 | CM | CMPLT | CMPLT | 2024-09-04 13:00 | 2024-09-04 13:01 | Baby probe not correlating with set temp | (TC 09/04/2024 01:01:00 PM, Rhuland, Courtney) First temp probe used would not reach temperature Replaced with new probe and reached temp, but would drop intermittently. After discussing with Jessica, found that this is normal with the radiant warmers Temperature would stay within the set ranges and would not give red alarm, only yellow warning alarm. This is to be expected Moved O2 tank and phototherapy light to righthand side of device and exam light to left as requested by Janet from NICU
1067962 | 209648 | CM | CMPLT | CMPLT | 2024-09-04 13:20 | 2024-09-04 13:20 | Bulb not working | (TC 09/04/2024 01:20:22 PM, Canjimavadivel, Kamallesh ) One of the intensity knobs was turned all the way down. Verified light was working and informed physician.
501114 | 209649 | CM | CMPLT | CMPLT | 2024-09-04 13:35 | 2024-09-05 15:25 | Lens is peeling from elements causing artifact | (TC 09/04/2024 01:34:49 PM, Fan, Tak) Confirmed issue. Ordered replacement probe. (TC 09/05/2024 03:25:00 PM, Fan, Tak) Replaced probe. F05NYB replaced by F0GBMP.
501063 | 209650 | CM | CMPLT | CMPLT | 2024-09-04 13:36 | 2024-09-09 13:06 | Operator panel unresponsive | (TC 09/04/2024 01:35:44 PM, Fan, Tak) Tested, seemed OK but techs insist issue. Ordered replacement panel. (TC 09/09/2024 01:05:41 PM, Fan, Tak) Replaced operator panel.
1062078 | 209651 | CM | CMPLT | CMPLT | 2024-09-04 14:32 | 2024-09-04 14:33 | Low light intensity | (TC 09/04/2024 02:32:42 PM, Alvarado, Karla) 1. Check the light and it was orange. 2. Change the Light Intensity Switch 3. Light Intensity Test failed 4. Replacement of LED board 5. Light Intensity Test Successful 6. Clean fans Measurement: Light Intensity Test High: 31.4 Low: 14.0 Electrical Safety Test 113 mohms 50.3 uA Hours: 2646.2 PM Completed (TC 09/04/2024 12:53:14 PM, Baghirzade, Nara) Supervised Karla
1050577 | 209652 | CM | CMPLT | CMPLT | 2024-09-04 14:36 | 2024-09-04 14:58 | "Check vent" alarm | (TC 09/04/2024 02:57:52 PM, Rendon Lopez , Elizabeth) Set up vent, run system check-OK. running pump on regular mode, no errors found. Returned to RT NICU | Vent was at sick kids while error occurred. "Life port adapter alarm did not go away, then turned off/on "check battery" alarm, but jet was always plugged in"
1061214 | 209653 | CM | CMPLT | CMPLT | 2024-09-04 14:50 | 2024-10-24 13:55 | OR 13 - Video Monitor Interference | (TC 10/02/2024 10:39:08 AM, Reid, Emmerson ) Was called again about this issue, said the scope is shutting down when the caughtery is being used. (TC 10/04/2024 09:36:16 AM, Reid, Emmerson ) Changed the OTV-S200 with the one from the Traveling Cysto Cart. Made these changes to location in TMS. After bringing in the video cart seemed to resolve the issue we are trying to change the processor to see if this is able to solve the problem. (TC 10/16/2024 06:58:00 PM, Patel, Rutvik) Coordinated and scheduled OR 13 for EMI testing with Jessica. (TC 10/21/2024 09:55:26 AM, Patel, Rutvik) Coordinated and scheduled and conducted EMI Tesing in OR 13 (TC 10/24/2024 10:37:09 AM, Patel, Rutvik) edited videos and communicated OR 13 testing results to everyone.
32540 | 209654 | CM | CMPLT | CMPLT | 2024-09-04 21:55 | 2024-09-04 21:53 | errors reported on bootup | errors reported on bootup Service completed troubleshooting Service completed system calibrations Found RTM errors/SCSI status-check condition/not ready Service will follow up on RTM errors
65416 | 209655 | CM | CMPLT | CMPLT | 2024-09-05 08:48 | 2024-09-05 08:50 | Broken Scale | (TC 09/05/2024 08:49:35 AM, Rhuland, Courtney) - Verify the scale was defected. (Damage in the plug) - Replace the scale. - Calibrate the scale. - Temperature and humidity test. - Replace both hinges from right door (view infron of the incubator).
66723 | 209656 | PM | CMPLT | CMPLT | 2024-09-05 10:07 | 2024-09-05 10:07 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/04/2024 10:06:56 AM, Rendon Lopez , Elizabeth) Cleaned and calibrated recorder. PM completed, all tests passed
65968 | 209657 | CM | CMPLT | CMPLT | 2024-09-05 10:54 | 2024-09-05 10:54 | Channel error | (TC 09/05/2024 10:54:21 AM, Rhuland, Courtney) Channel error 354.6770 - pressure sensor failure Replaced pressure sensor as it is recurring Calibrated and check in completed All tests passed Refer to ASM for details
1056420 | 209658 | CM | CMPLT | NFF | 2024-09-05 10:55 | 2024-09-05 10:55 | "can not turn on" | (TC 09/05/2024 10:54:48 AM, Cortez, John) Charged battery, didn't see any power issue all parameters tested ok
| 209659 | CM | ACTIV | ISSUE | 2024-09-05 13:17 |  | Welch Allyn CVSM Test Parts Order | Sept 5 - Identified parts for WA CVSM test kit. Contacted vendor fpr parts info. Received. Created SC> Requested approval. Placed order via MMC. Oct 29 - Some items received. Organized it. Followed up on ETA for the remaining.
1071894 | 209660 | CM | CMPLT | CMPLT | 2024-09-05 14:09 | 2024-09-05 14:09 | Changed locations | (TC 09/05/2024 02:09:10 PM, Rendon Lopez , Elizabeth) Device was moved to L&D ORs. Changed account to RT and location to L&D, relabelled device.
1054110 | 209661 | CM | CMPLT | CMPLT | 2024-09-05 14:15 | 2024-09-05 14:15 | Unable to see monitor in Fetalink | (TC 09/05/2024 02:15:22 PM, Rendon Lopez , Elizabeth) Checked monitor connections, restarted CEs several times, unable to solve problem. Notified nurse to place IT ticket for Cerner.
66720 | 209662 | CM | CMPLT | CMPLT | 2024-09-05 14:25 | 2024-09-06 11:39 | Created by Failed PM WO # 207531 : PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 09/06/2024 11:38:47 AM, Rendon Lopez , Elizabeth) Replaced recorder with repaired spare. PM completed, all tests passed. Returned to 15M - 370A. Replaced power cord. | (TC 09/05/2024 02:24:38 PM, Rendon Lopez , Elizabeth) Monitor fails recorder calibration, makes grinding noises and stucked.
32540 | 209663 | CM | CMPLT | CMPLT | 2024-09-05 21:51 | 2024-09-05 21:50 | RTM error follow up | RTM error follow up Service completed troubleshooting/testing CT CT phantom tested OK
1066462 | 209664 | FC | CMPLT | CMPLT | 2024-09-06 07:22 | 2024-09-05 07:22 | QC checklist | (TC 09/05/2024 07:22:26 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1066454 | 209665 | FC | CMPLT | CMPLT | 2024-09-06 07:23 | 2024-09-05 07:23 | QC checklist | (TC 09/05/2024 07:23:07 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1067437 | 209666 | FC | CMPLT | CMPLT | 2024-09-06 07:24 | 2024-09-05 07:24 | QC checklist | (TC 09/05/2024 07:24:17 AM, Baghirzade, Nara) Rear and front case were misaligned and there is a gap; aligned the front case to rear case. QC checklist completed. Returned to 11 South.
| 209667 | OC | CMPLT | CMPLT | 2024-09-06 08:31 | 2024-09-12 10:35 | On call Sept 6-13, 2024 | (TC 09/06/2024 08:31:03 AM, Scero, Tomlyn) Daily on call checks complete. (TC 09/09/2024 02:03:59 PM, Scero, Tomlyn) Daily on call checks complete (TC 09/10/2024 08:17:29 AM, Scero, Tomlyn) Daily on call checks complete (TC 09/11/2024 11:11:58 AM, Scero, Tomlyn) daily on call checks complete (TC 09/12/2024 10:34:47 AM, Scero, Tomlyn) daily on call checks complete (TC 09/12/2024 10:02:52 AM, Scero, Tomlyn) 12N Called at 8:47pm asking for me to come up and help them with a bariatric bed that was stuck and not moving. I explained to her that we aren't on site after hours and I could walk her through resetting the bed to see if that helps. I walked her through the steps to reset the bed and said if that didn't resolve the issue then they'd have to borrow a bed from another floor or wait until the morning when a tech arrived on site.
| 209668 | CM | CMPLT | CMPLT | 2024-09-06 09:08 | 2024-09-13 13:21 | misc biomed - Sep 2024 | (TC 09/05/2024 09:11:19 AM, Cortez, John) RCC called regarding ceiling lift, forwarded them to on-site biomed team x3333 Will reach out to RCC CNS regarding MSH biomed coverage (TC 09/13/2024 01:20:59 PM, Cortez, John) on-call daily check
65639 | 209669 | CM | CMPLT | CMPLT | 2024-09-06 10:24 | 2024-09-13 08:52 | ordering rotary knob for VN500s | received rotary knobs for VN500s. PO#7202454852
| 209670 | TEDU | CMPLT | CMPLT | 2024-09-06 10:35 | 2024-09-18 14:51 | Olympus Biomed Training Sign up | Sept 6 - Cotnacted Vendor for quote. Received. Created SC. Requested approval. Received. Issued PO via MMC. Sept 11 - Ordered VMC Switch for OR 12. Created SC. Requested approval. Received. Issued PO via MMC. Sept 18 - Received confirmation. Completed Registration.
1054469 | 209671 | CM | CMPLT | CMPLT | 2024-09-06 11:32 | 2024-09-12 13:12 | Created by Failed PM WO # 202260 : PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/12/2024 01:09:42 PM, Cortez, John) Vince (hillrom) repaired bed (TC 09/12/2024 01:12:17 PM, Cortez, John) service report attached | (TC 09/05/2024 11:29:14 AM, Cortez, John) PM fail, turn assist-left head rail , right bottom rail sensor installed nurse call cable
1059983 | 209672 | CM | CMPLT | CMPLT | 2024-09-06 11:44 | 2024-09-12 13:17 | pump running constantly on "normal", turn assist fail-left head rail | (TC 09/04/2024 11:44:05 AM, Cortez, John) Will request service (TC 09/12/2024 01:17:24 PM, Cortez, John) Vince (hillrom) repaired bed Service report attached
66713 | 209673 | CM | CMPLT | CMPLT | 2024-09-06 12:07 | 2024-09-06 12:07 | Reported that it was not switching on | (TC 09/06/2024 12:07:15 PM, Baghirzade, Nara) Did not find any issue with switching on the incubator; main switch was on. During physical inspection found one of the door latches to be very loose; when trying to tighten the screw, the screw kept rotating. Confirmed broken screw hold. Replaced latch. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 76%RH Incubator is fully functional.
| 209674 | CM | CMPLT | CMPLT | 2024-09-06 12:10 | 2024-09-06 12:10 | Working with coop student on Alaris pumps | (TC 09/06/2024 12:09:59 PM, Baghirzade, Nara) Teaching coop student Alaris pumps, syringe pumps, and PCU PMs, functions, etc.
| 209675 | CM | CMPLT | CMPLT | 2024-09-06 12:51 | 2024-09-06 12:52 | Embosser disposal | (TC 09/06/2024 12:51:36 PM, Baghirzade, Nara) As per Adeel's request, disposed of 2 embossers in the admitting dept. (TC 09/06/2024 12:52:10 PM, Alvarado, Karla) As per Adeel's request, disposed of 2 embossers in the admitting dept.
| 209676 | CM | CMPLT | CMPLT | 2024-09-06 13:20 | 2024-09-06 13:20 | Blood pressure device check | (TC 09/06/2024 01:20:22 PM, Rhuland, Courtney) Repaired manual BP bulb/cuff Replaced BP cuff on automatic BP device Functional on both devices - OK
1054466 | 209677 | CM | ACTIV | ISSUE | 2024-09-06 14:03 |  | Created by Failed PM WO # 202257 : PM-MAJOR- Factory PM for: ICU Bed Elecric |  | (TC 09/04/2024 01:56:56 PM, Cortez, John) error code, "CPR switch failed" no pendant turn assist fail-left bottom rail installed nurse call cable
1054479 | 209678 | CM | CMPLT | CMPLT | 2024-09-06 14:10 | 2024-09-13 13:09 | Created by Failed PM WO # 202266 : PM-MAJOR- Factory PM for: ICU Bed Elecric | (TC 09/13/2024 01:08:59 PM, Cortez, John) Repaired by Vince, Functionality ok, 69 deg head section, 5165 total blower hours Nurse call cable installed | (TC 09/04/2024 02:03:36 PM, Cortez, John) PM failed left siderail screen not responsive, took a while to startup, couldn't calibrate screen
1055643 | 209679 | CM | ACTIV | ISSUE | 2024-09-06 14:11 |  | Right turn assist issue | (TC 09/05/2024 02:11:14 PM, Cortez, John) turn assist fail-left bottom rail
| 209680 | CM | CMPLT | CMPLT | 2024-09-06 14:54 | 2024-10-29 09:26 | Order for Prismax scale calibration kit | (TC 09/06/2024 02:58:49 PM, Cortez, John) ordered Prismax kit (TC 10/29/2024 09:26:02 AM, Cortez, John) received
1071941 | 209681 | CM | CMPLT | CMPLT | 2024-09-09 09:03 | 2024-09-09 09:03 | Incoming Inspection - Trial | (TC 09/09/2024 09:03:22 AM, Reid, Emmerson ) Latest service records attached to WO. Device to be used for one case today only in OR16. Incoming complete.
68542 | 209682 | CM | ACTIV | ISSUE | 2024-09-09 09:17 |  | Battery failure | (TC 09/06/2024 09:17:47 AM, Rendon Lopez , Elizabeth) Charged AGM, battery indicator light green. Run device on battery power for more than 30mins with no issues. Downloaded logs, confirmed error occurred, faulty battery - Batteries were replaced during 6 year PM in 2022 - Sent logs to Draeger for suggested action (TC 09/10/2024 09:44:03 AM, Rendon Lopez , Elizabeth) Draeger advised to check battery, and battery tray connections, reset the AH in the HIT and Reset Delete Counter. Completed checks, charged, ran on battery again, no issues. (TC 09/19/2024 04:18:05 PM, Rendon Lopez , Elizabeth) Returned device to L&D as spare. Removed lamp to replaced damaged one in OR 1 #68544. Inquired part pricing.
1049786 | 209683 | CM | CMPLT | CMPLT | 2024-09-09 09:41 | 2024-09-09 09:42 | PC error blue screen | (TC 09/09/2024 09:41:51 AM, Canjimavadivel, Kamallesh ) Attempted to restart and PC did not start up. Repalced PICiX hardware woth 1049785. Ran startup and configured as MSHEMERG6S.
67895 | 209684 | CM | ACTIV | ISSUE | 2024-09-09 11:23 |  | Alarming Occlusion, no occlusion present | (TC 09/09/2024 11:23:41 AM, Rendon Lopez , Elizabeth) Confirmed issue, pump alarms once cassette is loaded. Requested RMA to return for repair.
| 209685 | CM | CMPLT | CMPLT | 2024-09-09 11:24 | 2024-09-19 14:36 | X2 batteries biomed stock | (TC 09/19/2024 02:36:25 PM, Rhuland, Courtney) Unboxed batteries and labelled with PO and price
67882 | 209686 | CM | CMPLT | CMPLT | 2024-09-09 11:35 | 2024-09-13 14:16 | Pump low battery- Doesn't charge | (TC 09/09/2024 09:23:32 AM, Rendon Lopez , Elizabeth) Replaced batteries. Upgraded Sotware and Library. Running in user mode-OK (TC 09/13/2024 02:15:42 PM, Alvarado, Karla) Performance check completed All tests passed See procedure for details
502089 | 209687 | CM | ACTIV | ISSUE | 2024-09-09 13:15 |  | Persistent error | (TC 09/09/2024 01:15:16 PM, Paz, Ramon (Tino)) SLD1 Uncommanded Motion. Need of further troubleshooting. Motor has been replaced back in dec 2022.
1071527 | 209688 | CM | CMPLT | CMPLT | 2024-09-09 13:57 | 2024-09-09 13:59 | MSH 15 IVF APLIO 500 USEED IN MAIN OR FOR EMBRYO IMPLANTATION | (TC 09/09/2024 01:57:19 PM, Whittington, Scott) MSH 15 IVF APLIO 500 USEED IN MAIN OR FOR EMBRYO IMPLANTATION PROVIDE SUPPORT IN SETTING UP SYSTEM AND STUDY MANAGEMENT TO eIVF SERVER RETURN SYSTEM OT MSH 15 OBU | MSH 15 IVF APLIO 500 USEED IN MAIN OR FOR EMBRYO IMPLANTATION
| 209689 | RDS | CMPLT | CMPLT | 2024-09-09 14:05 | 2024-09-10 10:27 | weekly checks in ccu | (TC 09/10/2024 10:27:19 AM, Scero, Tomlyn) weekly checks complete. 30J tests passed.
1069284 | 209690 | CM | CMPLT | CMPLT | 2024-09-09 14:27 | 2024-09-10 10:24 | MSH ER G7 CANON PROBE LOCK KNOB MISSING | (TC 09/10/2024 10:23:41 AM, Whittington, Scott) MSH ER G7 CANON PROBE LOCK KNOB MISSING FOUND KNOB IN LOCAL STOCK, REPLACED (TC 09/10/2024 10:23:41 AM, Whittington, Scott) MSH ER G7 CANON PROBE LOCK KNOB MISSING FOUND KNOB IN LOCAL STOCK, REPLACED (TC 09/11/2024 10:23:41 AM, Whittington, Scott) MSH ER G7 CANON PROBE LOCK KNOB MISSING FOUND KNOB IN LOCAL STOCK, REPLACED | MSH ER G7 CANON PROBE LOCK KNOB MISSING
69040 | 209691 | CM | CMPLT | CMPLT | 2024-09-10 08:53 | 2024-09-10 09:04 | Defective pressure sensors | (TC 09/10/2024 09:03:36 AM, Alvarado, Karla) Replacement of both pressure sensors (upper and lower). Performed tests in ASM: Analog Sensor, Pressure Calibration, Rate Calibration, and Check- in tests. All test passed. (TC 09/10/2024 09:22:32 AM, Baghirzade, Nara) Supervised Karla in repairs and check-in tests.
| 209692 | RDS | CMPLT | CMPLT | 2024-09-10 09:16 | 2024-09-10 09:18 | Weekly check for ED
| 209693 | RDS | CMPLT | CMPLT | 2024-09-10 09:17 | 2024-09-16 14:49 | Weekly check for ICU
67886 | 209694 | CM | CMPLT | CMPLT | 2024-09-10 09:26 | 2024-09-13 16:15 | Not working | (TC 09/09/2024 09:26:59 AM, Rendon Lopez , Elizabeth) Running on user mode, NFF. Upgraded software and library. (TC 09/13/2024 04:14:59 PM, Rendon Lopez , Elizabeth) Performance check-OK, Returned to L&D
66509 | 209695 | CM | CMPLT | CMPLT | 2024-09-10 09:29 | 2024-09-18 16:12 | Recorder not printing correct time | (TC 09/09/2024 09:29:40 AM, Rendon Lopez , Elizabeth) Recorder not priting clearly, removed from room and placed spare. (TC 09/16/2024 01:53:04 PM, Alvarado, Karla) Removed and repaired the fetal recorder. Reinstalled. Calibrated. Test passed. (TC 09/16/2024 01:56:38 PM, Rendon Lopez , Elizabeth) Supervised (TC 09/18/2024 04:12:22 PM, Rendon Lopez , Elizabeth) Returned to room 5, working properly
67894 | 209696 | CM | CMPLT | CMPLT | 2024-09-10 09:41 | 2024-09-13 15:43 | Keeps alarming occluded before pump even started | (TC 09/13/2024 03:41:38 PM, Alvarado, Karla) Performance check completed All tests passed See procedure for details (TC 09/13/2024 03:42:45 PM, Rendon Lopez , Elizabeth) Supervised
65989 | 209697 | FC | CMPLT | CMPLT | 2024-09-10 09:50 | 2024-09-10 09:51 | Functional Test | (TC 09/10/2024 09:50:48 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. All test passed. (TC 09/10/2024 09:52:17 AM, Baghirzade, Nara) Supervised Karla
1057536 | 209698 | CM | CMPLT | CMPLT | 2024-09-10 10:28 | 2024-09-10 10:35 | missing bezel | (TC 09/10/2024 10:34:32 AM, Scero, Tomlyn) replaced bezel
| 209699 | PA | CMPLT | CMPLT | 2024-09-10 14:05 | 2024-09-10 14:24 | Weights for calibration | (TC 09/10/2024 02:24:22 PM, Canjimavadivel, Kamallesh ) Placed order for weights Weights not needed, shopping cart cancelled
1059778 | 209700 | FC | CMPLT | CMPLT | 2024-09-10 14:13 | 2024-09-10 14:13 | QC checklist | (TC 09/10/2024 02:12:46 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
66013 | 209701 | CM | CMPLT | CMPLT | 2024-09-10 14:19 | 2024-09-10 14:22 | Battery does not charge | (TC 09/10/2024 02:19:13 PM, Alvarado, Karla) Tested the power cord and the issue was there. Power cord was replaced. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Alarm test failed. Turn on the monitor and the screen displayed Error code: 133.6090. According to the manual, the main speaker failed. The main speaker was replaced. Performed tests in ASM: Check- in tests. Where the GND Leakage = 88.9uA and the GND Resistance = 0.061 Ohms All test passed. Main speaker is fully functional. (TC 09/10/2024 02:22:12 PM, Baghirzade, Nara) Supervised Karla through the troubleshooting, repairs, and FC.
1053348 | 209702 | CM | ACTIV | ISSUE | 2024-09-10 15:38 |  | MSH BI 1272 H3 STEREO QC ERRORS | (TC 09/11/2024 11:34:38 AM, Whittington, Scott) MSH BI 1272 H3 STEREO QC ERRORS SN# 3DM160700414 REF# PARLEY 01596099 | MSH BI 1272 H3 STEREO QC ERRORS SN# 3DM160700414 REF# PARLEY 01596099
1054613 | 209703 | CM | CMPLT | CMPLT | 2024-09-10 16:59 | 2024-09-13 13:17 | continuously switches between AC and battery power | (TC 09/10/2024 04:59:12 PM, Cortez, John) Confirmed issue, Only seems to switch when plugged into A1-60 emergency outlet. Output voltage/freq ok Moved to normal power outlet, which seems to resolve issue. Put in building service request (TC 09/13/2024 01:17:19 PM, Cortez, John) Building services verified functionality of outlet Moved UPS back to emergency power UPS didn't switch back and forth, as it did before
1054617 | 209704 | CM | CMPLT | CMPLT | 2024-09-10 17:03 | 2024-09-13 13:14 | continuously switches between AC and battery power | (TC 09/10/2024 05:02:54 PM, Cortez, John) Confirmed issue, Only seems to switch when plugged into A1-60 emergency outlet. Output voltage/freq ok Moved to normal power outlet, which seems to resolve issue. Put in building service request (TC 09/13/2024 01:14:11 PM, Cortez, John) Building services verified functionality of outlet Moved UPS back to emergency power UPS didn't switch back and forth, as it did before
62705 | 209705 | CM | CMPLT | CMPLT | 2024-09-11 10:01 | 2024-09-11 10:02 | Main board and wifi card upgrade | (TC 09/11/2024 10:01:49 AM, Canjimavadivel, Kamallesh ) Swapped main board and wifi card and completed upgrade successfully. Restored original config and network settings
67947 | 209706 | CM | CMPLT | CMPLT | 2024-09-11 10:28 | 2024-09-11 10:44 | Humidifier leaking | (TC 09/11/2024 10:28:34 AM, Alvarado, Karla) The gasket was broken; therefore it was replaced. Replaced humidifier button Repaired porthole door Performed functional test: Temperature = 35C and Humidity = 75% All test passed. (TC 09/11/2024 10:43:42 AM, Rhuland, Courtney) Supervised Karla during repair
65964 | 209707 | CM | CMPLT | CMPLT | 2024-09-11 11:04 | 2024-09-11 11:05 | Mylar Gasket damaged | (TC 09/11/2024 11:04:38 AM, Alvarado, Karla) Replaced front case. Replaced Mylar Gasket. Performed tests in ASM: Calibration and Check- in tests. All test passed. (TC 09/11/2024 11:24:15 AM, Baghirzade, Nara) Supervised Karla in the repairs and functional checks.
1055175 | 209708 | CM | CMPLT | CMPLT | 2024-09-11 11:42 | 2024-09-11 11:44 | Latch module damaged | (TC 09/11/2024 11:42:03 AM, Alvarado, Karla) Replacement of latch. Performed tests in ASM: Check- in. All test passed. (TC 09/11/2024 11:44:10 AM, Baghirzade, Nara) Supervised
69216 | 209709 | CM | CMPLT | CMPLT | 2024-09-11 12:30 | 2024-09-11 12:30 | Defective IUI - pump is not switching on when connecting through male IUI | (TC 09/11/2024 12:30:16 PM, Baghirzade, Nara) Male IUI replaaced. During check in, rate accuracy test was failing (reading high). Rate was calibrated. Check-in tests completed. IUI connectors and device cleaned.
69452 | 209710 | CM | CMPLT | CMPLT | 2024-09-11 12:32 | 2024-09-11 12:32 | Functional check | (TC 09/11/2024 12:31:47 PM, Baghirzade, Nara) Downloaded error logs; did not note any significant errors. Check-in tests completed. IUI connectors and device cleaned.
62705 | 209711 | CM | CMPLT | CMPLT | 2024-09-11 12:37 | 2024-09-11 12:37 | Battery replacement and functional check | (TC 09/11/2024 12:37:15 PM, Baghirzade, Nara) Replaced both batteries. Ran simulation to test that all leads are functional. ECG fully functional. Returned to 11 South chemo room.
1062077 | 209712 | PM | CMPLT | CMPLT | 2024-09-11 12:54 | 2024-09-11 12:58 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 09/11/2024 12:57:36 PM, Baghirzade, Nara) Supervised (TC 09/11/2024 12:56:30 PM, Alvarado, Karla) Replaced Power Cord 1790.0 hours Intensity light: - High: 31.5 uW/cm2/nm - Low: 13.5 uW/cm2/nm Target light is functional Electrical safety results - leakage current: 48.8 uA power cord resistance: 400 mohms PM completed
32540 | 209713 | CM | CMPLT | CMPLT | 2024-09-11 14:25 | 2024-09-09 14:25 | Scanner not booting up fully | Scanner not booting up fully - comes up with an error System was rebooted a few times by the techs - didn't help Upon troubleshooting RSP1 (power supply in recon rack 1 was dead) Ordered part but Canon didn't have it in stock Found one in PMH storage - used one Replaced it and system was up and running
32540 | 209714 | CM | CMPLT | CMPLT | 2024-09-11 14:30 | 2024-09-20 20:46 | Scanner not booting up fully | Scanner not booting up fully - Checked all power supllies in the Recon unit and no issues found. Not able to collect BIR log System was shutdown using "alternate shutdown" It came back up with no issues -Going back tomorrow (Sep 12) to start the system myself to see if system is stable (TC 09/12/2024 08:36:28 AM, Debebe, Adam) System would not boot up first thing in the morning with no power to top rack. Ordred more part from Canon (TC 09/20/2024 08:46:00 PM, Khan, Muhammad) Replaced the Timer Delay Relay for top rack power supply. -System was tested multiple times before it was returned for use. (TC 09/20/2024 08:46:00 PM, Khan, Muhammad) Replaced the Timer Delay Relay for top rack power supply. -System was tested multiple times before it was returned for use.
| 209715 | CM | CMPLT | CMPLT | 2024-09-11 14:42 | 2024-09-11 14:42 | CO2 laser checks | (TC 09/11/2024 02:41:57 PM, Baghirzade, Nara) Wally from Boston Scientific was on site to check the micromanipulator connectivity to the microscope; double beams that the doctors are seeing, demo of cleaning the micromanipulator and a quick in-servicing.
| 209716 | PA | ACTIV | ISSUE | 2024-09-12 08:52 |  | Ordering parts for stock | (TC 09/12/2024 08:57:36 AM, Baghirzade, Nara) SC 1003308233 (TC 09/19/2024 07:51:01 AM, Baghirzade, Nara) Received pressure sensors. PO 7202457566
69363 | 209717 | CM | CMPLT | CMPLT | 2024-09-12 09:01 | 2024-09-12 09:02 | Door did not close | (TC 09/12/2024 09:01:25 AM, Alvarado, Karla) After doing the Binary Sensor test, the program always indicated that the door was always open. Replaced the AIL sensor. Performed tests in ASM: Binary Sensors, Analog Sensor, Rate Calibration, and Check- in. All test passed. (TC 09/12/2024 09:02:09 AM, Baghirzade, Nara) Supervised
1057869 | 209718 | CM | CMPLT | CMPLT | 2024-09-12 09:40 | 2024-09-18 15:02 | Plate Clip Broken | Sept 12 - Unit came for repair. Cleaned unit. FOund side plate clip broken. Contacted vendor for purchase info. Received. Created SC. Requested approval. Received. Placed order via MMC. Sept 18 - CLOSED THIS WO BY MISTAKE. STATUS IS STILL ACTIVE AND WAITING FOR PARTS. Sept 27 - Followed up twice. CS reached out and ifnormed the price provided was incorrect and advised to amend the PO with new pricing. Contacted MMC to have the PO updated and issued. Oct 09 - Reuqired more parts for PM. Contacted vendor for info. Created SC. Requested approval. Received. Placed order via MMC.
| 209719 | PA | CMPLT | CMPLT | 2024-09-12 10:33 | 2024-09-12 11:10 | X3 batteries | (TC 09/12/2024 11:09:58 AM, Canjimavadivel, Kamallesh ) Placed order for batteries
66458 | 209720 | PM | CMPLT | CMPLT | 2024-09-12 12:31 | 2024-09-12 12:31 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Replaced left side door. Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 17 Ah and 230 min. Tested. Functional. SW.2.51 Leakage: 89 uA and 95 ohm.
1058948 | 209721 | CM | CMPLT | CMPLT | 2024-09-12 13:14 | 2024-10-03 15:17 | Faulty Sensor | Sept 11 - Call re: not working. Went to check. Waited for access. Found error code E-98. Removed unit from the mount and replaced it with one from OR 12. Checked functionality. Tested ok. Brought faulty unit down to biomed. Checked functionality. Error code persisted. Removed cover and found bag sensor assembly corroded. Sept 12 - Removed cover and found bag sensor assembly corroded. Contacted vendor for parts info, Received. Created SC. Requested approval. Received. Placed order via MMC. Oct 03 - Parts received. REplaced sensor assembly. Checked fucntionality. Tested ok. Delivered back and mounted in OR.
| 209722 | PM | CANCL | DUP | 2024-09-12 13:43 | 2024-09-12 14:00 | PM-MAJOR_Annual PM For: Med Surg Bed | PM complete by Stryker. battery changed and SW updated
| 209723 | CM | CMPLT | CMPLT | 2024-09-12 13:53 | 2024-09-12 13:53 | functional check of culposcopes | (TC 09/12/2024 01:52:52 PM, Rhuland, Courtney) Performed functional tests on donated culposcopes ahead of use in clinics All tests passed - replaced power cord on one Asset tagged
69008 | 209724 | CM | CMPLT | CMPLT | 2024-09-12 14:14 | 2024-09-12 14:15 | Functional Test | (TC 09/12/2024 02:14:34 PM, Alvarado, Karla) Performed tests in ASM: Check- in tests completed. GND Leakage = 68.7 uA and the GND Resistance = 0.077 Ohms All test passed. (TC 09/12/2024 02:15:09 PM, Baghirzade, Nara) Supervised
1062509 | 209725 | PM | CMPLT | CMPLT | 2024-09-12 14:16 | 2024-09-12 14:20 | PM-MAJOR_Annual PM For: Med Surg Bed | PM complete by Stryker. SW upgraded to v.03.00.000.01
1062509 | 209726 | PM | CMPLT | CMPLT | 2024-09-12 14:28 | 2024-09-12 14:29 | Batteries Replacement Due | Batteries replaced by styker
| 209727 | CM | CMPLT | CMPLT | 2024-09-13 06:24 | 2024-09-12 06:24 | Review capital request for 11 South | (TC 09/12/2024 06:23:58 AM, Baghirzade, Nara) Review capital request for 11 South meeting with Andrew.
1065014 | 209728 | PM | CMPLT | CMPLT | 2024-09-13 06:25 | 2024-09-12 06:28 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 09/12/2024 06:25:55 AM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg (TC 09/12/2024 06:28:26 AM, Alvarado, Karla) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
1067437 | 209729 | FC | CMPLT | CMPLT | 2024-09-13 06:30 | 2024-09-12 06:30 | QC checklist | (TC 09/12/2024 06:29:44 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1066462 | 209730 | FC | CMPLT | CMPLT | 2024-09-13 06:31 | 2024-09-12 06:31 | QC checklist | (TC 09/12/2024 06:31:07 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1067437 | 209731 | FC | CMPLT | CMPLT | 2024-09-13 06:32 | 2024-09-12 14:32 | QC checklist | (TC 09/12/2024 02:32:00 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
68545 | 209732 | CM | CMPLT | CMPLT | 2024-09-13 09:40 | 2024-09-10 09:41 | Failing self test, keeps saying adjust APL to 20 when it is already set | (TC 09/10/2024 09:41:20 AM, Rendon Lopez , Elizabeth) AA changed breathing block, same issue. Troubleshooted device in OR, found patient circuit filter has a leak, replaced complete patient circuit. Selft test passed. Unit back and functional
52252 | 209733 | CM | CMPLT | CMPLT | 2024-09-13 09:48 | 2024-09-13 09:50 | Porthole doors broken. | (TC 09/13/2024 09:48:41 AM, Alvarado, Karla) Replaced porthole doors (SE, SW, S). Performed functional test: Temperature = 35C and Humidity = 75% All test passed. (TC 09/13/2024 09:50:00 AM, Baghirzade, Nara) Supervised
1056501 | 209734 | CM | CMPLT | CMPLT | 2024-09-13 10:12 | 2024-09-13 10:16 | Battery replacement | (TC 09/13/2024 10:15:27 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056510 | 209735 | CM | CMPLT | CMPLT | 2024-09-13 10:14 | 2024-09-13 10:14 | Battery replacement | (TC 09/13/2024 10:13:52 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055988 | 209736 | CM | CMPLT | CMPLT | 2024-09-13 10:16 | 2024-09-13 10:16 | Battery replacement | (TC 09/13/2024 10:16:07 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055989 | 209737 | CM | CMPLT | CMPLT | 2024-09-13 10:16 | 2024-09-13 10:17 | Battery replacement. | (TC 09/13/2024 10:16:46 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056504 | 209738 | CM | CMPLT | CMPLT | 2024-09-13 10:17 | 2024-09-13 10:17 | Battery replacement | (TC 09/13/2024 10:16:58 AM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055616 | 209739 | CM | CMPLT | CMPLT | 2024-09-13 10:18 | 2024-09-13 10:18 | Battery replacement. | (TC 09/13/2024 10:18:10 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
65861 | 209740 | PM | CMPLT | CMPLT | 2024-09-13 10:18 | 2024-09-13 10:18 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 09/13/2024 10:18:18 AM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
1055978 | 209741 | CM | CMPLT | CMPLT | 2024-09-13 10:19 | 2024-09-13 10:20 | Battery replacement. | (TC 09/13/2024 10:19:27 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056436 | 209742 | CM | CMPLT | CMPLT | 2024-09-13 10:20 | 2024-09-13 10:20 | Defective temperature module | (TC 09/13/2024 10:19:39 AM, Baghirzade, Nara) Replaced temperature probe and well with a spare. Functional tests passed.
1055980 | 209743 | CM | CMPLT | CMPLT | 2024-09-13 10:20 | 2024-09-13 10:21 | Battery replacement. | (TC 09/13/2024 10:20:37 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
65880 | 209744 | CM | CMPLT | NFF | 2024-09-13 10:55 | 2024-09-13 10:58 | Channel error. | (TC 09/13/2024 10:55:30 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, and Check- in tests. Download error log: did not find errors. All test passed. No fault found. (TC 09/13/2024 10:57:52 AM, Baghirzade, Nara) Supervised
1071301 | 209745 | PM | CMPLT | CMPLT | 2024-09-13 10:57 | 2024-09-13 10:59 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
| 209746 | PA | CMPLT | CMPLT | 2024-09-13 11:06 | 2024-09-23 09:04 | Clear snap table for ECG department | (TC 09/13/2024 11:24:48 AM, Scero, Tomlyn) ordered ecg tabs for Dianne in ECG (TC 09/23/2024 09:04:07 AM, Scero, Tomlyn) received and delivered to the ECG department
1071302 | 209747 | PM | CMPLT | CMPLT | 2024-09-13 11:32 | 2024-09-13 11:31 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071303 | 209748 | PM | CMPLT | CMPLT | 2024-09-13 11:34 | 2024-09-13 11:34 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071312 | 209749 | PM | CMPLT | CMPLT | 2024-09-13 11:37 | 2024-09-13 11:37 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
| 209750 | CM | CMPLT | CMPLT | 2024-09-13 11:38 | 2024-09-13 11:38 | Laser meeting | (TC 09/13/2024 11:38:20 AM, Baghirzade, Nara) Laser committee meeting
1071316 | 209751 | PM | CMPLT | CMPLT | 2024-09-13 12:01 | 2024-09-13 12:01 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071315 | 209752 | PM | CMPLT | CMPLT | 2024-09-13 12:02 | 2024-09-13 12:02 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071305 | 209753 | PM | CMPLT | CMPLT | 2024-09-13 12:33 | 2024-09-13 12:32 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071310 | 209754 | PM | CMPLT | CMPLT | 2024-09-13 12:34 | 2024-09-13 12:35 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071955 | 209755 | ACC | CMPLT | CMPLT | 2024-09-13 12:51 | 2024-09-13 14:27 | Incoming inspection | (TC 09/13/2024 12:54:50 PM, Canjimavadivel, Kamallesh ) Got boxes from receiving and assembled device with roll stand (TC 09/13/2024 02:27:12 PM, Canjimavadivel, Kamallesh ) In-service with rep
1057054 | 209756 | CM | ACTIV | ISSUE | 2024-09-13 12:56 |  | MSH Hologic Workstation #1 - Shutdown Randomly | (TC 09/12/2024 12:56:26 PM, Chau, Ken) System shutdown randomly on sonographer, looked at machine and called Hologic -They reported that this issue is a reocurring matter, they have replaced the UPS already -Spoke with Rads, the UPS did not alarm, monitor stayed on but the CPU was shutdown -Informed Hologic to order a CPU , PSU (TC 09/13/2024 12:56:26 PM, Chau, Ken) Worked with Ryan from Hologic in regards to PC shutting down -he taken picture of the PC, PSU and will be placing an order -looked into logs an noticed that the PC goes into sleep mode despite it being disabled, Ryan believes theres a setting somewhere but he doesnt have credentials, he will contact other engineers for logins and settings (TC 09/18/2024 09:32:36 AM, Chau, Ken) System is not under contract, Hologic will not replace parts System Shutdown on Radiologist again, Hologic will come in and look at SW. However, after testing system it most likely isnt SW but it will be HW - Video card System monitor was still on but no display of images, however after unplugging the monitor outputs and plugging it into other sources, the display came back online on one of the monitor being able to see the SW. Informed management and Radiologist, but nothing else we can do until system contract is resolved. (TC 09/19/2024 02:34:33 PM, Chau, Ken) Video card driver installed Barco MXRT7600 Video card physically installed AMD FirePro W7100 The issue is driver apparantly cannot be updated or changed unless you flash the GPU, however this is too complex and may cause issues. AMD Firepro is intended for all day user Barco is made for medical image processing Ordered a Barco MXRT7600 to replace (TC 10/09/2024 09:43:27 AM, Chau, Ken) Replaced and tested okay (TC 10/16/2024 10:55:45 AM, Chau, Ken) No issues for over 2 weeks Went to ask Rad and Tech, they informed it shut off yesterday Looked into device managers and observed that they have some drivers not installed -informed Hologic to provide firmwares for the missing drivers (TC 11/20/2024 02:47:15 PM, Chau, Ken) Installed new video card Barco 8700 Shipped back old PC Tested unit okay (TC 11/13/2024 02:48:55 PM, Chau, Ken) Installed new hologic workstation Station didnt come with video card, reused old video card Configured and tested unit okay
1061461 | 209757 | PM | CMPLT | CMPLT | 2024-09-13 13:00 | 2024-09-13 13:00 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071311 | 209758 | PM | CMPLT | CMPLT | 2024-09-13 13:20 | 2024-09-13 13:19 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071295 | 209759 | PM | CMPLT | CMPLT | 2024-09-13 13:21 | 2024-09-13 13:21 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1071309 | 209760 | PM | CMPLT | CMPLT | 2024-09-13 14:06 | 2024-09-13 14:06 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
1061210 | 209761 | CM | CMPLT | CMPLT | 2024-09-13 14:23 | 2024-09-13 14:23 | VMC monitor not turning on | (TC 09/13/2024 02:23:20 PM, Reid, Emmerson ) Tried changing the monitor and the power cable, neither worked. I went to reset the VMC and found the UCES was turned on. I turned it off and reset the VMC. Working now.
1061212 | 209762 | CM | CMPLT | CMPLT | 2024-09-13 14:25 | 2024-09-13 14:26 | Boom is drifting | (TC 09/13/2024 02:25:33 PM, Reid, Emmerson ) Adjusted the brakes on the boom to fix the drifting.
1061434 | 209763 | PM | CMPLT | CMPLT | 2024-09-13 14:33 | 2024-09-13 14:34 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 09/13/2024 02:33:56 PM, Cortez, John) Completed by Stryker technician
1061984 | 209764 | PM | CMPLT | CMPLT | 2024-09-13 14:35 | 2024-09-13 14:36 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
500196 | 209765 | CM | CMPLT | CMPLT | 2024-09-13 14:46 | 2024-09-18 16:23 | VG can ping Entegra, no dicom echo | (TC 09/13/2024 02:46:41 PM, Green, Paul) Troubleshot problem. Can ping Entegra, but dicom fails. Looked at settings but system was needed. Talked to GE about settings/problem. (TC 09/18/2024 04:23:26 PM, Green, Paul) Was onite and confirmed this is working now.
| 209766 | CM | CMPLT | CMPLT | 2024-09-13 14:51 | 2024-09-13 14:51 | MSH Workshop cleanup. | (TC 09/13/2024 02:50:37 PM, Green, Paul) Cleaned up workshop. Disposed of batteries down in e-waste area.
1062497 | 209767 | PM | CMPLT | CMPLT | 2024-09-13 15:08 | 2024-09-13 15:08 | PM-MAJOR_Annual PM For: Med Surg Bed | SW upgrade and PM Completed by Stryker. Checked functionality and perfomed EST. Tested ok.
| 209768 | RDS | CMPLT | CMPLT | 2024-09-16 07:55 | 2024-09-16 09:01 | Weekly check for ED
1056217 | 209769 | CM | CMPLT | CMPLT | 2024-09-16 08:04 | 2024-09-16 08:04 | Battery replacement. | (TC 09/16/2024 08:04:18 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056477 | 209770 | CM | CMPLT | CMPLT | 2024-09-16 08:24 | 2024-09-16 08:27 | Temperature probe does not work. | (TC 09/16/2024 08:23:56 AM, Alvarado, Karla) Replaced temperature probe and well with a spare. Performed tests: Temperature NIBP = set: 120/80, measured: 114/81 SpO2 = set: 95%, measured: 95% All test passed. (TC 09/16/2024 08:26:32 AM, Baghirzade, Nara) Supervised
1056220 | 209771 | CM | CMPLT | CMPLT | 2024-09-16 08:40 | 2024-09-16 08:40 | Battery replacement. | (TC 09/16/2024 08:40:29 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056211 | 209772 | CM | CMPLT | CMPLT | 2024-09-16 08:51 | 2024-09-16 08:51 | Charges cord was broken, thrown away. | (TC 09/16/2024 08:51:26 AM, Alvarado, Karla) Replaced power cord. Performed tests: Temperature NIBP = set: 120/80, measured: 115/81 SpO2 = set: 95%, measured: 95% All test passed.
1046035 | 209773 | CM | CMPLT | CMPLT | 2024-09-16 09:20 | 2024-09-16 09:20 | MSH ER Evo - Wall bucky was derailed | (TC 09/16/2024 09:20:18 AM, Chau, Ken) Placed back in the rail as it was off rail
| 209774 | RDS | CMPLT | CMPLT | 2024-09-16 09:21 | 2024-09-16 09:41 | Weekly checks in CCU | (TC 09/16/2024 09:40:43 AM, Scero, Tomlyn) weekly checks complete. 30J test passed. Room 10 bed 1062275 was stuck and needed to be reset.
1062418 | 209775 | CM | CMPLT | CMPLT | 2024-09-16 09:31 | 2024-09-16 09:31 | PM
1062070 | 209776 | CM | CMPLT | CMPLT | 2024-09-16 09:40 | 2024-09-16 09:40 | PM completed
1050633 | 209777 | CM | ACTIV | ISSUE | 2024-09-16 09:43 |  | MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED | (TC 09/16/2024 10:17:12 AM, Whittington, Scott) Good Morning Mount Sinai Hospital Toronto L&D have TE7 with Probe Voltage Error MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED DATE OF ACCEPTANCE: MARCH 18. 2019 SN# CD7-8C000912 MIE# 1050633 Thanks, Scott 416-402-2462 From: Whittington, Scott Sent: September 16, 2024 10:16 AM To: service@bomimed.com Cc: Lisa Petteplace <lpetteplace@bomimed.com>; Aman Yousif - Bomimed Mindray (ayousif@bomimed.com) <ayousif@bomimed.com>; Michelle Lee <MLee@bomimed.com>; Ultrasound Support (Ultrasound.Support@uhn.ca); Ngo, Kim - Sinai Health System <Kim.Ngo@sinaihealth.ca>; Dolan, Gillian - Sinai Health System <Gillian.Dolan@sinaihealth.ca> Subject: MSH 15 L&D TE7 system C ERROR: PROBE VOLTAGE UNDER PROTECTION SN# CD7-8C000912 SYSTEM IS DOWN Good Morning Mount Sinai Hospital Toronto L&D have a TE7 with Probe Voltage Error MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED DATE OF ACCEPTANCE: MARCH 18. 2019 SN# CD7-8C000912 MIE# 1050633 Please Advise on repair Thanks, Scott 416-402-2462 (TC 11/18/2024 11:09:57 AM, Whittington, Scott) MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED PO# 7500213484 (TC 11/19/2024 04:49:37 PM, Whittington, Scott) MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED (TC 11/19/2024 04:49:37 PM, Whittington, Scott) MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED RMA number is CRM 20412 FEDEX ON MURRAYS ACCT 357070201 TRK FROM WAYBILL 8166 5350 0884 SYSTEM PACKED IN DOUBLE BOX AND DELIVERED TO TGH DOCK FOR PICKUP | MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED DATE OF ACCEPTANCE: MARCH 18 2019 SN# CD7-8C000912 CRM 20412 /RE: MSH 15 L&D TE7C MINDRAY TE-7 ULTRASOUND sn# CD7-8C000912 REQUEST FOR RMA FOR ASSESMENT/REPAIR MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED From: Justinne Bambo <jbambo@bomimed.com> Sent: Monday, November 18, 2024 11:49:56 AM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Service <service@bomimed.com> Cc: Dolan, Gillian - Sinai Health System <Gillian.Dolan@sinaihealth.ca>; Ngo, Kim - Sinai Health System <Kim.Ngo@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca>; Lisa Petteplace <lpetteplace@bomimed.com> Subject: [External] CRM 20412 /RE: MSH 15 L&D TE7C MINDRAY TE-7 ULTRASOUND sn# CD7-8C000912 REQUEST FOR RMA FOR ASSESMENT/REPAIR MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED Hi Scott, Your RMA number is CRM 20412 FEDEX ON MURRAYS ACCT 357070201 TRK FROM WAYBILL 8166 5350 0884 Please ship your unit to this address: BOMImed 147 Hamelin St Winnipeg, MB R3T 3Z1 Please reference: CRM 20412 Please note your unit is not under warranty. Labour and parts will be charged. Thanks! Justinne Bambo Service Admin Assistant 1-100 IRENE STREET WINNIPEG, MB R3T 4E1 T. 800.667.6276 ext. 264 F: 877.435.6984 www.bomimed.ca Good Morning Mount Sinai Hospital Toronto L&D have TE7 with Probe Voltage Error MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED DATE OF ACCEPTANCE: MARCH 18. 2019 SN# CD7-8C000912 MIE# 1050633 Thanks, Scott 416-402-2462 From: Whittington, Scott Sent: September 16, 2024 10:16 AM To: service@bomimed.com Cc: Lisa Petteplace <lpetteplace@bomimed.com>; Aman Yousif - Bomimed Mindray (ayousif@bomimed.com) <ayousif@bomimed.com>; Michelle Lee <MLee@bomimed.com>; Ultrasound Support (Ultrasound.Support@uhn.ca); Ngo, Kim - Sinai Health System <Kim.Ngo@sinaihealth.ca>; Dolan, Gillian - Sinai Health System <Gillian.Dolan@sinaihealth.ca> Subject: MSH 15 L&D TE7 system C ERROR: PROBE VOLTAGE UNDER PROTECTION SN# CD7-8C000912 SYSTEM IS DOWN Good Morning Mount Sinai Hospital Toronto L&D have a TE7 with Probe Voltage Error MSH 15 L&D TE7C ERROR: TRANSMITTING VOL;TAGE IS UNDER PROTECTION, IMAGE CAN NOT BE DISPLAYED DATE OF ACCEPTANCE: MARCH 18. 2019 SN# CD7-8C000912 MIE# 1050633 Please Advise on repair Thanks, Scott 416-402-2462
1062422 | 209778 | CM | CMPLT | CMPLT | 2024-09-16 09:43 | 2024-09-16 09:43 | PM completed
1062465 | 209779 | CM | CMPLT | CMPLT | 2024-09-16 09:45 | 2024-09-16 09:45 | PM completed
1062423 | 209780 | CM | CMPLT | CMPLT | 2024-09-16 09:53 | 2024-09-16 09:53 | PM completed
1062462 | 209781 | CM | CMPLT | CMPLT | 2024-09-16 09:56 | 2024-09-16 09:56 | PM completed
1062464 | 209782 | CM | CMPLT | CMPLT | 2024-09-16 09:58 | 2024-09-16 09:58 | PM completed
1062466 | 209783 | CM | CMPLT | CMPLT | 2024-09-16 09:59 | 2024-09-16 09:59 | PM completed
1062068 | 209784 | CM | CMPLT | CMPLT | 2024-09-16 10:00 | 2024-09-16 10:00 | PM completed
1062498 | 209785 | CM | CMPLT | CMPLT | 2024-09-16 10:02 | 2024-09-16 10:02 | PM completed
1062087 | 209786 | CM | CMPLT | CMPLT | 2024-09-16 10:05 | 2024-09-16 10:05 | PM completed
1062475 | 209787 | CM | CMPLT | CMPLT | 2024-09-16 10:10 | 2024-09-16 10:10 | PM completed
1062471 | 209788 | CM | CMPLT | CMPLT | 2024-09-16 10:12 | 2024-09-16 10:12 | PM completed
1062424 | 209789 | CM | CMPLT | CMPLT | 2024-09-16 10:18 | 2024-09-16 10:18 | PM completed
1062470 | 209790 | CM | CMPLT | CMPLT | 2024-09-16 10:21 | 2024-09-16 10:21 | PM completed
1062088 | 209791 | CM | CMPLT | CMPLT | 2024-09-16 10:22 | 2024-09-16 10:22 | PM completed
1062419 | 209792 | CM | CMPLT | CMPLT | 2024-09-16 10:28 | 2024-09-16 10:28 | PM completed
1054161 | 209793 | CM | CMPLT | CMPLT | 2024-09-16 14:10 | 2024-10-21 10:56 | No Sound | (TC 09/16/2024 02:10:36 PM, Alvarado, Karla) Confirm problem weak signal. To be send out under warranty. (TC 10/21/2024 10:55:53 AM, Baghirzade, Nara) Received transducer from spectramedx. Transducer is functional. Attached service report.
1069784 | 209794 | CM | CMPLT | CMPLT | 2024-09-16 14:26 | 2024-09-16 14:27 | Black end presses into patient abdomen needs proper cover. | (TC 09/16/2024 02:26:17 PM, Alvarado, Karla) Performed test. All passed. (TC 09/16/2024 02:26:34 PM, Rendon Lopez , Elizabeth) Supervised
1054158 | 209795 | CM | CMPLT | CMPLT | 2024-09-16 14:35 | 2024-10-21 11:14 | No sound | (TC 09/16/2024 02:35:40 PM, Alvarado, Karla) Confirm problem no signal. To be send out under warranty. (TC 10/21/2024 11:14:06 AM, Baghirzade, Nara) Received unit however the vendor did not repair it as it had too much damage. Transducer will be retired and disposed of.
1053432 | 209796 | CM | CMPLT | CMPLT | 2024-09-16 14:45 | 2024-09-16 14:45 | No working | (TC 09/16/2024 02:45:01 PM, Alvarado, Karla) Noisy cable. Device repaired multiple times before. Retired from service.
1056196 | 209797 | CM | CMPLT | CMPLT | 2024-09-16 15:33 | 2024-09-16 15:33 | Black end presses into patient abdomen needs proper cover. | (TC 09/16/2024 03:33:19 PM, Alvarado, Karla) Performed test. All passed.
1054153 | 209798 | CM | CMPLT | CMPLT | 2024-09-16 15:49 | 2024-09-16 15:49 | Black end presses into patient abdomen needs proper cover. | (TC 09/16/2024 03:49:05 PM, Alvarado, Karla) Time: 0.25 Performed test. All passed.
1048003 | 209799 | CM | CMPLT | CMPLT | 2024-09-16 15:54 | 2024-09-16 15:54 | Black end presses into patient abdomen needs proper cover | (TC 09/16/2024 03:54:03 PM, Alvarado, Karla) Performed test. All passed.
1054196 | 209800 | CM | CMPLT | CMPLT | 2024-09-16 16:01 | 2024-09-16 16:02 | : Black end presses into patient abdomen needs proper cover. | (TC 09/16/2024 04:01:31 PM, Alvarado, Karla) Performed test. All passed.
| 209801 | CM | CMPLT | CMPLT | 2024-09-17 07:32 | 2024-09-16 10:32 | Broken monitor arm | (TC 09/16/2024 10:32:00 AM, Rhuland, Courtney) Arm holding MX800 monitor in room 55 did not have any tension and fell when nurse went to adjust Tried to tighten arm but would not work Replaced arm with spare Adjusted tension with monitor on arm
1054192 | 209802 | CM | CMPLT | CMPLT | 2024-09-17 07:48 | 2024-09-17 07:48 | Not detected in the monitor. | (TC 09/17/2024 07:48:02 AM, Alvarado, Karla) Performed test. All passed.
| 209803 | CM | CMPLT | CMPLT | 2024-09-17 07:48 | 2024-09-16 07:48 | Meeting with Naz re: OBMs | (TC 09/16/2024 07:48:30 AM, Baghirzade, Nara) Meeting with Naz re: OBMs
| 209804 | CM | CMPLT | CMPLT | 2024-09-17 07:49 | 2024-09-16 07:49 | Biomed huddle | (TC 09/16/2024 07:49:25 AM, Baghirzade, Nara) Biomed huddle
66470 | 209805 | CM | CMPLT | CMPLT | 2024-09-17 07:56 | 2024-09-17 07:56 | broken side door | replaced broked left side door
| 209806 | CM | ACTIV | ISSUE | 2024-09-17 07:56 |  | OBM data management | (TC 09/16/2024 07:56:10 AM, Baghirzade, Nara) Installed the OBM viewer software on biomed laptop. Was not able to connect the OBM to the software; will look through all technical documentation for instructions and will also contact Bill from Natus for instruction. Followed up with Kamran. (TC 09/17/2024 02:39:04 PM, Baghirzade, Nara) After looking through documentation from Natus, I was able to transfer patient data and view it on the OBM. A preconfigured USB is required for file transfer; preconfigured it for export/import. Allocated a second preconfigured USB for archive only. Followed up with Kamran.
66458 | 209807 | CM | CMPLT | CMPLT | 2024-09-17 07:57 | 2024-09-20 14:16 | ordering VN500 parts | ordering VN500 flow sensor cables. Received VN500 flow sensor cables. PO# 7202459585
| 209808 | CM | ACTIV | ISSUE | 2024-09-17 07:59 |  | WinSpiro software for respirology | (TC 09/16/2024 07:58:44 AM, Baghirzade, Nara) Copied over the installation file to a USB; was able to download the software using USB. Will be using the same USB when downloading the software on all resp. physician PCs. Looked through the instructions on how to use the software.
1055918 | 209809 | FC | CMPLT | CMPLT | 2024-09-17 08:02 | 2024-09-16 08:02 | Functional check | (TC 09/16/2024 08:02:13 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056450 | 209810 | CM | CMPLT | CMPLT | 2024-09-17 08:04 | 2024-09-24 12:59 | Shipping device for repair - display stuck on "software upgrade" | (TC 09/16/2024 08:04:28 AM, Baghirzade, Nara) Power cycled the Root multiple times however it boots up to the same error message. Connected Root to MICT to reload the configuration; although the write was successful, it did not clear the error message. Will initate a service request with Masimo. (TC 09/17/2024 08:48:39 AM, Baghirzade, Nara) SC 1003311116 (TC 09/19/2024 01:16:52 PM, Baghirzade, Nara) Shipped device to Masimo. PO 7500211947 KJV courier #: 455076 (TC 09/24/2024 12:59:21 PM, Baghirzade, Nara) PO 7500211947
1054985 | 209811 | CM | CMPLT | CMPLT | 2024-09-17 08:08 | 2024-09-24 13:00 | Shipping device for repair - power entry module loose and broken | (TC 09/16/2024 08:07:40 AM, Baghirzade, Nara) Power entry module is loose and not in place; disassembled the root from the back cover and found that the metal bracket that holds the power entry module in place is broken. Will initiate a service request with Masimo. (TC 09/17/2024 08:52:24 AM, Baghirzade, Nara) SC 1003311117 (TC 09/19/2024 01:18:00 PM, Baghirzade, Nara) Shipped device to Masimo. PO 7500211947 KJV courier #: 455075 (TC 09/24/2024 01:00:15 PM, Baghirzade, Nara) PO 7500211947
1050675 | 209812 | CM | CMPLT | CMPLT | 2024-09-17 08:43 | 2024-09-17 08:43 | Not detected in the monitor. | (TC 09/17/2024 08:42:53 AM, Alvarado, Karla) Confirm problem no signal.
1071966 | 209813 | CM | CMPLT | CMPLT | 2024-09-17 08:53 | 2024-09-17 08:53 | Battery does not work. | (TC 09/17/2024 08:52:51 AM, Alvarado, Karla) Replaced battery from retired transducer. Performed test. Added Asset to the system. All passed.
67655 | 209814 | CM | CMPLT | CMPLT | 2024-09-17 09:19 | 2024-09-17 09:19 | Battery does not work | (TC 09/17/2024 09:18:43 AM, Alvarado, Karla) Replaced battery from retired transducer. Performed test. All passed.
67708 | 209815 | CM | CMPLT | CMPLT | 2024-09-17 09:39 | 2024-09-17 09:39 | Battery does not work. | (TC 09/17/2024 09:39:10 AM, Alvarado, Karla) Replaced battery from retired transducer. Blue light; therefore, the device does not work. Board issue.
1063310 | 209816 | CM | CMPLT | CMPLT | 2024-09-17 09:43 | 2024-09-17 09:43 | checking the laser and supporting the case | (TC 09/05/2024 09:43:10 AM, Huang, Jessica) Irina required to check the laser before the case supported the case in OR 3
1063310 | 209817 | CM | CMPLT | CMPLT | 2024-09-17 09:45 | 2024-09-17 09:45 | assisting Wally to check the device
1071827 | 209818 | CM | CMPLT | CMPLT | 2024-09-17 09:51 | 2024-09-17 09:51 | new incoming inspection
1071828 | 209819 | CM | CMPLT | CMPLT | 2024-09-17 09:52 | 2024-09-17 09:52 | new incoming inspection
1071829 | 209820 | CM | CMPLT | CMPLT | 2024-09-17 09:53 | 2024-09-17 09:53 | new incoming inspection
67711 | 209821 | CM | CMPLT | CMPLT | 2024-09-17 09:53 | 2024-09-17 09:53 | Battery does not work. | (TC 09/17/2024 09:52:56 AM, Alvarado, Karla) Replaced battery from retired transducer. Blue light; therefore, the device does not work. Board issue.
1071888 | 209822 | CM | CMPLT | CMPLT | 2024-09-17 09:54 | 2024-09-17 09:54 | new incoming inspection
1071889 | 209823 | CM | CMPLT | CMPLT | 2024-09-17 10:02 | 2024-09-17 10:02 | new incoming inspection
1071890 | 209824 | CM | CMPLT | CMPLT | 2024-09-17 10:03 | 2024-09-17 10:03 | new incoming inspection
1070862 | 209825 | CM | CMPLT | CMPLT | 2024-09-17 10:04 | 2024-09-17 10:04 | new incoming inspection
1046429 | 209826 | CM | CMPLT | CMPLT | 2024-09-17 10:08 | 2024-09-17 10:08 | Battery does not work. | (TC 09/17/2024 10:08:14 AM, Alvarado, Karla) Replaced battery from retired transducer. Performed test. All passed.
1057537 | 209827 | CM | CMPLT | CMPLT | 2024-09-17 10:23 | 2024-09-20 12:18 | BP pump not inflating | (TC 09/17/2024 02:11:48 PM, Scero, Tomlyn) Ordered NIBP pump (TC 09/20/2024 12:17:30 PM, Scero, Tomlyn) BP pump received and replaced. Tested the BP using a simulator. All tests passed.
1055658 | 209828 | CM | CMPLT | CMPLT | 2024-09-17 10:23 | 2024-10-21 11:13 | TOCO Malfunction | (TC 09/16/2024 10:23:52 AM, Rendon Lopez , Elizabeth) TOCO Malfunction alarm. Sending out under warranty repair (TC 10/21/2024 11:12:12 AM, Baghirzade, Nara) Received transducer. Attached service report. (TC 10/21/2024 11:13:28 AM, Baghirzade, Nara) PO 7500211946
1059161 | 209829 | CM | CMPLT | CMPLT | 2024-09-17 10:24 | 2024-09-20 12:12 | BP not inflating after high systolic BP >300mmHg | (TC 09/17/2024 02:02:42 PM, Scero, Tomlyn) pump ordered from ereq (TC 09/20/2024 12:12:15 PM, Scero, Tomlyn) BP pump received. Replaced the pump and tested the BP on the simulator. BP passed with no issues.
1061209 | 209830 | CM | CMPLT | CMPLT | 2024-09-17 10:26 | 2024-09-17 10:26 | OR lights not connecting | (TC 09/17/2024 10:26:30 AM, Reid, Emmerson ) OR lights were not connecting to the WiFi. They nurse had tried to troubleshoot and turned the wifi off on the lights. After I turned it back on they started working. They also reported a smell in the room, I only noticed the the smoke evac filter was on low life so I sugested they change it. (TC 09/17/2024 10:26:30 AM, Reid, Emmerson ) OR lights were not connecting to the WiFi. They nurse had tried to troubleshoot and turned the wifi off on the lights. After I turned it back on they started working. They also reported a smell in the room, I only noticed the the smoke evac filter was on low life so I sugested they change it.
1061217 | 209831 | CM | CMPLT | CMPLT | 2024-09-17 10:29 | 2024-09-20 09:14 | Audio Recording not working | (TC 09/17/2024 10:28:39 AM, Reid, Emmerson ) Clinincal staff want to make a recording with Audio on Sept 20, 2024. I teted this is several room and have not been able to get the audio in the recording. I also spoke with Derrick from Olympus and he is going to get back to me with some troubleshooting. (TC 09/20/2024 09:14:10 AM, Reid, Emmerson ) Went and setup the audio in OR6 and the camera. Had to restart the widi router for the camera to work. Also went up a second time to help them make adjustments. Working now
1071967 | 209832 | CM | CMPLT | CMPLT | 2024-09-17 10:38 | 2024-09-17 10:38 | Battery does not work. | (TC 09/17/2024 10:38:08 AM, Alvarado, Karla) Replaced battery from retired transducer. Performed test. All passed.
| 209833 | CM | CMPLT | CMPLT | 2024-09-17 10:45 | 2024-09-23 09:13 | X3 bezels for stock | (TC 09/17/2024 02:01:32 PM, Scero, Tomlyn) ordered parts from ereq (TC 09/23/2024 09:13:04 AM, Scero, Tomlyn) received
1068220 | 209834 | CM | ACTIV | ISSUE | 2024-09-17 11:03 |  | MSH 5 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | (TC 09/17/2024 11:03:14 AM, Whittington, Scott) MSH 5 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | MSH 5 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 SAMEH
1071750 | 209835 | CM | CMPLT | CMPLT | 2024-09-17 11:06 | 2024-09-17 11:06 | new incoming inspection
1071751 | 209836 | CM | CMPLT | CMPLT | 2024-09-17 11:14 | 2024-09-17 11:14 | new incoming inspection
1071752 | 209837 | CM | CMPLT | CMPLT | 2024-09-17 11:15 | 2024-09-17 11:15 | new incoming inspection
1071927 | 209838 | CM | CMPLT | CMPLT | 2024-09-17 11:19 | 2024-09-17 11:19 | new incoming inspection
1071928 | 209839 | CM | CMPLT | CMPLT | 2024-09-17 11:31 | 2024-09-17 11:31 | new incoming inspection
1071929 | 209840 | CM | CMPLT | CMPLT | 2024-09-17 11:34 | 2024-09-17 11:34 | new incoming inspection
1071930 | 209841 | CM | CMPLT | CMPLT | 2024-09-17 11:35 | 2024-09-17 11:35 | new incoming inspection
1063123 | 209842 | CM | CMPLT | CMPLT | 2024-09-17 11:36 | 2024-09-17 11:36 | not working | (TC 09/03/2024 11:35:48 AM, Huang, Jessica) replaced batteries and tested it, working fine
1047211 | 209843 | CM | CMPLT | CMPLT | 2024-09-17 11:37 | 2024-09-17 11:37 | not working | (TC 08/20/2024 11:37:38 AM, Huang, Jessica) replaced batteries, tested it, working fine
1069567 | 209844 | CM | CMPLT | CMPLT | 2024-09-17 11:39 | 2024-09-17 11:39 | replacing the footpedal | (TC 08/01/2024 11:39:57 AM, Huang, Jessica) replaced the bipolar foot pedal
| 209845 | CM | CMPLT | CMPLT | 2024-09-17 12:30 | 2024-09-17 12:30 | Broken exam light | (TC 09/17/2024 12:30:02 PM, Rhuland, Courtney) Loose screw holding neck of exam light onto base Tightened screw to secure lamp Functional check - OK
66461 | 209846 | CM | CMPLT | CMPLT | 2024-09-17 12:39 | 2024-09-17 12:39 | Nebulizer device check failed | replaced and adjusted both O2 and air pressure regulators. Passed device check.
1064493 | 209847 | CM | ACTIV | ISSUE | 2024-09-17 13:29 |  | warranty exchange
1053391 | 209848 | CM | CMPLT | CMPLT | 2024-09-17 14:29 | 2024-09-17 14:31 | Before and after patient use check | (TC 09/17/2024 02:30:34 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
1051383 | 209849 | CM | CMPLT | CMPLT | 2024-09-17 14:31 | 2024-09-26 10:45 | Keeps turning off. Over heating | (TC 09/17/2024 02:31:36 PM, Alvarado, Karla) Hours: 5267.2 GND Resistance: 0.7 Ohms Leaved the equipment running for 30 min. The fan did not work. Test the voltage to prove that it is a fan issue. Sent an email to the company to get a new fan kit. (TC 09/26/2024 10:44:49 AM, Alvarado, Karla) Replaced power cord. Replaced fan. Replaced enclosure. Leaved the equipment running for 30 min. Tests results: - Light Intensity small pad: 80.75µW•cm-2•nm-1 - Hours: 5268.7 - Electrical safety results - Leakage current = 18.7 uA; GND Resistance = 243 mohms All tests passed. (TC 09/26/2024 10:44:49 AM, Rhuland, Courtney) Supervised (TC 10/21/2024 10:24:21 AM, Baghirzade, Nara) Received full order. SC 1003311756 PO 7202462437
1066118 | 209850 | CM | CMPLT | CMPLT | 2024-09-17 14:32 | 2024-09-17 14:32 | Broken cable | (TC 09/17/2024 02:31:58 PM, Baghirzade, Nara) Replaced cable. Tested doppler; functional.
1071181 | 209851 | CM | CMPLT | CMPLT | 2024-09-17 14:33 | 2024-09-17 14:33 | Functional check; error not stated by user | (TC 09/17/2024 02:33:29 PM, Baghirzade, Nara) Downloaded error log; did not note any errors. Rate failed during first check-in (reading high). Rate was calibrated. Check-in tests completed. IUI connectors and device cleaned.
1071187 | 209852 | CM | CMPLT | CMPLT | 2024-09-17 14:35 | 2024-09-17 14:35 | Functional check; error not stated by user | (TC 09/17/2024 02:35:23 PM, Baghirzade, Nara) Downloaded error log; did not note any errors. Rate failed during first check-in (reading high). Rate was calibrated. Check-in tests completed. IUI connectors and device cleaned.
1071186 | 209853 | CM | CMPLT | CMPLT | 2024-09-17 14:37 | 2024-09-17 14:37 | Functional check; error not stated by user | (TC 09/17/2024 02:36:45 PM, Baghirzade, Nara) Downloaded error log; did not note any errors. Check-in tests completed. IUI connectors and device cleaned.
1071180 | 209854 | CM | CMPLT | CMPLT | 2024-09-17 14:37 | 2024-09-17 14:37 | Functional check; error not stated by user | (TC 09/17/2024 02:37:19 PM, Baghirzade, Nara) Downloaded error log; did not note any errors. Check-in tests completed. IUI connectors and device cleaned.
1071130 | 209855 | CM | CMPLT | CMPLT | 2024-09-17 14:38 | 2024-09-17 14:38 | Functional check; error not stated by user | (TC 09/17/2024 02:38:06 PM, Baghirzade, Nara) Downloaded error log; did not note any errors. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 98 mohms Leakage current: 67.3 uA
| 209856 | CM | ACTIV | ISSUE | 2024-09-17 14:38 |  | HL-90 O-rings | HL-90 PM o-rings
1050106 | 209857 | CM | CMPLT | CMPLT | 2024-09-17 14:42 | 2024-09-17 14:42 | Frame was off the hinge | (TC 09/17/2024 02:42:03 PM, Baghirzade, Nara) Tightened screw and placed the cover back on; frame moves easily and does not come off.
1068221 | 209858 | CM | ACTIV | ISSUE | 2024-09-17 18:31 |  | MSH 5 rm 6 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | (TC 09/17/2024 06:34:14 PM, Whittington, Scott) MSH 5 rm 6 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | MSH 5 rm 6 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7
1067992 | 209859 | CM | ACTIV | ISSUE | 2024-09-17 18:35 |  | MSH 5 rm 8 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | (TC 09/17/2024 08:24:00 PM, Whittington, Scott) MSH 5 rm 8 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | MSH 5 rm 8 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7
1070800 | 209860 | CM | CMPLT | CMPLT | 2024-09-18 08:14 | 2024-09-18 08:14 | Phototherapy light install | (TC 09/18/2024 08:14:14 AM, Rhuland, Courtney) Switched O2 tank and exam light Installed phototherapy light to right side
1071973 | 209861 | ACC | CMPLT | CMPLT | 2024-09-18 08:52 | 2024-09-18 07:56 | Incoming Inspection | (TC 09/18/2024 07:56:21 AM, Reid, Emmerson ) Setup the device with gas tanks and brought up to OR - Pod 3 storage room. I also informed Reeba and Andrew of this.
68799 | 209862 | CM | CMPLT | CMPLT | 2024-09-18 09:59 | 2024-09-18 09:59 | Broken. Not working. No turning on. | (TC 09/18/2024 09:59:22 AM, Alvarado, Karla) Replaced power cord. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage = 82.9uA and the GND Resistance = 0.078 Ohms All test passed.
| 209863 | CM | CMPLT | CMPLT | 2024-09-18 10:29 | 2024-09-18 10:29 | Exam light replacement OPG | (TC 09/18/2024 10:29:10 AM, Rhuland, Courtney) Clinic had spare light Replaced broken light with new light Functional check - OK
1065793 | 209864 | CM | CMPLT | CMPLT | 2024-09-18 10:44 | 2024-09-18 10:45 | Broken. Not working. No turning on. | (TC 09/18/2024 10:44:52 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Binary Sensors, Analog Sensor and Check- in. All test passed.
1054308 | 209865 | CM | CMPLT | CMPLT | 2024-09-18 11:22 | 2024-09-18 11:23 | O2 sat need to be replaced. | (TC 09/18/2024 11:22:56 AM, Alvarado, Karla) Replaced SpO2 sensor. Performed tests: Temperature NIBP SpO2 = set: 95%, measured: 95% All test passed.
1071975 | 209866 | CM | CMPLT | CMPLT | 2024-09-18 11:52 | 2024-09-18 11:53 | Repair | Unit not recorded in TMS. Recorded unit and assigned asset tag. Unit came with tag "Not working". Found screen frozen. Rebooted unit. FOund battery almost depleted. Charged unit for a while. Checked functionality. Tested ok. Returned to 12N.
1071976 | 209867 | ACC | CMPLT | CMPLT | 2024-09-18 12:08 | 2024-09-18 12:10 | Incoming inspection | (TC 09/18/2024 12:07:52 PM, Alvarado, Karla) Incoming inspection. Tested. Functional.
1071978 | 209868 | CM | CMPLT | CMPLT | 2024-09-18 13:06 | 2024-09-18 13:08 | Incoming inspection | (TC 09/18/2024 01:07:38 PM, Alvarado, Karla) Incoming inspection. Tested. Functional.
1071979 | 209869 | ACC | CMPLT | CMPLT | 2024-09-18 13:19 | 2024-09-18 13:21 | Incoming inspection | (TC 09/18/2024 01:20:32 PM, Alvarado, Karla) Incoming inspection. Tested. Functional.
1071980 | 209870 | ACC | CMPLT | CMPLT | 2024-09-18 13:30 | 2024-09-18 13:30 | Incoming inspection | (TC 09/18/2024 01:30:06 PM, Alvarado, Karla) Incoming inspection. Tested. Functional.
1071981 | 209871 | ACC | CMPLT | CMPLT | 2024-09-18 13:32 | 2024-09-18 13:32 | Incoming Inspection | (TC 09/18/2024 01:32:03 PM, Reid, Emmerson ) Asset created, and device is tagged. Incoming complete.
1061450 | 209872 | PM | CMPLT | CMPLT | 2024-09-18 14:06 | 2024-09-18 14:08 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker
1061450 | 209873 | PM | CMPLT | CMPLT | 2024-09-18 14:11 | 2024-09-18 14:11 | Batteries Replacement Due | Battery replaced by Stryker
1061455 | 209874 | PM | CMPLT | CMPLT | 2024-09-18 14:18 | 2024-09-18 14:19 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker
1061455 | 209875 | PM | CMPLT | CMPLT | 2024-09-18 14:20 | 2024-09-18 14:20 | Batteries Replacement Due | Battery replaced by Styker
501071 | 209876 | CM | CMPLT | CMPLT | 2024-09-18 14:49 | 2024-09-30 13:17 | MSH RM#12 - CEOU - System No Screen Images | (TC 09/18/2024 02:48:45 PM, Chau, Ken) Top monitor black, bottom monitor bright white Restart machine, both monitor doesn't display Restarted again, system and monitor came back -BEP has been ordered as issue comes and goes, FSE to replace BEP (TC 09/26/2024 12:02:41 PM, Chau, Ken) Replaced UMB System boots on and runs now Waiting for Philips FSE to provide code to load SW (TC 09/30/2024 01:17:06 PM, Whittington, Scott) MSH RM#12 - CEOU - System No Screen Images SYSTEM BOOTED UP WITH ERROR 5, TRANSDUCER NOT SELECTED. NO BUTTONS ON TCS SELECTED TRANSDUCER BUTTON ON MAIN OP PNL TCS CAME UP NORMALLY, PROBE SELECTED, AND DISPLAYING PROPERLY
1071984 | 209877 | CM | CMPLT | CMPLT | 2024-09-18 15:22 | 2024-09-18 15:22 | Incoming Inspection | (TC 09/18/2024 03:22:22 PM, Ferreira, Qwynn) Tested for Gabe, working and ready for use
1062496 | 209878 | PM | CMPLT | CMPLT | 2024-09-19 08:46 | 2024-09-19 08:47 | PM-MAJOR_Annual PM For: Med Surg Bed | PM complete by Stryker. SW upgraded
1062496 | 209879 | PM | CMPLT | CMPLT | 2024-09-19 08:49 | 2024-09-19 08:49 | Batteries Replacement Due | Batteries replaced by Stryker
| 209880 | CM | CMPLT | CMPLT | 2024-09-19 09:00 | 2024-10-04 14:58 | 99% IPA Barrel
1059093 | 209881 | CM | CMPLT | CMPLT | 2024-09-19 10:18 | 2024-09-19 10:18 | Not working | (TC 09/19/2024 10:18:06 AM, Alvarado, Karla) Performed calibration through the calibration plug to verify 37C. Test passed.
| 209882 | PA | CANCL | CANCL | 2024-09-19 10:19 | 2024-09-19 10:21 | Ordering parts for stock - syringe pressure sensors
1060030 | 209883 | PA | CMPLT | CMPLT | 2024-09-19 10:47 | 2024-09-26 10:37 | Ordering patient valve | (TC 09/19/2024 10:49:35 AM, Baghirzade, Nara) SC 1003313153 (TC 09/26/2024 10:36:33 AM, Baghirzade, Nara) PO 7500212302
1047226 | 209884 | CM | CMPLT | CMPLT | 2024-09-19 10:52 | 2024-09-19 10:53 | Not working | (TC 09/19/2024 10:52:20 AM, Alvarado, Karla) Confirm issue. Replace thermometer with spare one. Changed SN in TMS. Test passed.
| 209885 | PA | CMPLT | CMPLT | 2024-09-19 11:14 | 2024-09-25 11:42 | PIC + cables for BIS | (TC 09/19/2024 11:18:07 AM, Scero, Tomlyn) Cables ordered on ereq (TC 09/25/2024 11:41:48 AM, Scero, Tomlyn) items received and delivered to the OR
| 209886 | CM | ACTIV | ISSUE | 2024-09-19 12:31 |  | MSH BIUS OBTAIN SURPLUS PROBES FROM WCH | (TC 09/19/2024 12:31:27 PM, Whittington, Scott) MSH BIUS OBTAIN SURPLUS PROBES FROM WCH | MSH BIUS OBTAIN SURPLUS PROBES FROM WCH
| 209887 | CM | CMPLT | CMPLT | 2024-09-19 13:24 | 2024-09-26 08:13 | NIBP cuffs for NICU
1053144 | 209888 | FC | CMPLT | NFF | 2024-09-19 13:43 | 2024-09-24 14:40 | 1053144 - Preventive Maintenance | Testing PM scheule functionality in TMS.
1059605 | 209889 | FC | ACTIV | ISSUE | 2024-09-19 13:43 |  | 1059605 - Preventive Maintenance
1050257 | 209890 | CM | CMPLT | CMPLT | 2024-09-19 13:46 | 2024-09-19 13:49 | States “channel error” | (TC 09/19/2024 01:46:09 PM, Alvarado, Karla) Downloaded error log, did not find any errors. Performed analog sensor test and upper sensor failed. Replaced upper sensor. Tried to complete a check-in but pressure sensor kept failing. Flashed pump. Tried to calibrate the pressure sensors; found the lower sensor failed. Replaced lower sensor. Completed pressure sensor calibration. Check-in tests passed. (TC 09/19/2024 01:49:16 PM, Baghirzade, Nara) Supervised
1053506 | 209891 | ACC | CMPLT | CMPLT | 2024-09-19 13:52 | 2024-09-19 13:52 | MSH 12 BIUS PROVIDE PLT 1204 PROBE, FORMERLY AT SPORTS CLINIC | (TC 09/19/2024 01:52:20 PM, Whittington, Scott) MSH 12 BIUS PROVIDE PLT 1204 PROBE, FORMERLY AT SPORTS CLINIC | MSH 12 BIUS PROVIDE PLT 1204 PROBE, FORMERLY AT SPORTS CLINIC
1049905 | 209892 | CM | CMPLT | CMPLT | 2024-09-19 14:31 | 2024-10-18 17:32 | MSH R65 COLLIMATOR LED LIGHT SOURCE DIMINISHED BRIGHTNESS | (TC 09/19/2024 02:32:11 PM, Whittington, Scott) MSH R65 COLLIMATOR LED LIGHT SOURCE DIMINISHED BRIGHTNESS (TC 10/22/2024 08:38:43 PM, Whittington, Scott) OCT 18 MSH R65 COLLIMATOR LED LIGHT SOURCE DIMINISHED BRIGHTNESS COLLIMAOR REPLACED, DAP SWAPPED | MSH R65 COLLIMATOR LED LIGHT SOURCE DIMINISHED BRIGHTNESS K5804-4514 Replaced collimator Ref# 1054717
1056477 | 209893 | CM | CMPLT | CMPLT | 2024-09-19 16:10 | 2024-09-19 16:13 | BP not working | (TC 09/19/2024 04:12:31 PM, Alvarado, Karla) Replaced blood pressure tube. Performed tests: Temperature, NIBP, and SpO2. All test passed.
68544 | 209894 | CM | CMPLT | CMPLT | 2024-09-19 16:19 | 2024-09-19 16:20 | Lamp damaged | (TC 09/19/2024 04:19:54 PM, Rendon Lopez , Elizabeth) Replaced lamp from spare AGM #68544.
1064423 | 209895 | PM | CMPLT | CMPLT | 2024-09-20 07:24 | 2024-09-19 15:30 | PM-ANNUAL MAJOR | (TC 09/19/2024 03:30:00 PM, Alvarado, Karla) 851.0 hours Intensity light: - High: 32.2 uW/cm2/nm - Low: 13.8 uW/cm2/nm Target light is functional GND Resistance: 0.120 ohms Leakage current: 50.8 uA All test passed. PM completed.
1061333 | 209896 | FC | CMPLT | CMPLT | 2024-09-20 07:44 | 2024-09-19 07:44 | QC Checklist | (TC 09/19/2024 07:44:05 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
69292 | 209897 | CM | CMPLT | CMPLT | 2024-09-20 07:45 | 2024-09-19 07:45 | Defective bottle-side pressure sensor | (TC 09/19/2024 07:45:26 AM, Baghirzade, Nara) Downloaded error log; did not note any errors. Ran analog test and found the bottle-side pressure sensor to be defective (stays at 4.9V). Replaced pressure sensor. Rate was calibrated. Check-in tests completed. IUI connectors and device cleaned.
1062767 | 209898 | CM | CMPLT | CMPLT | 2024-09-20 07:46 | 2024-09-20 07:46 | Not working. | (TC 09/20/2024 07:46:26 AM, Alvarado, Karla) Replaced batteries. Test passed.
67952 | 209899 | CM | CMPLT | CMPLT | 2024-09-20 07:49 | 2024-09-19 07:49 | Fan improperly seated | (TC 09/19/2024 07:48:38 AM, Baghirzade, Nara) Re-seated fan. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 76%RH Incubator is fully functional.
65892 | 209900 | CM | CMPLT | CMPLT | 2024-09-20 07:50 | 2024-09-19 07:51 | Broken front case and pressure sensor replacement | (TC 09/19/2024 07:50:58 AM, Baghirzade, Nara) Replaced: -front case -pressure sensor -mylar gasket Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
63215 | 209901 | CM | CMPLT | CMPLT | 2024-09-20 07:55 | 2024-09-19 07:55 | Missing loading door | (TC 09/19/2024 07:55:00 AM, Baghirzade, Nara) Placed door (took off from pump that is being sent for repairs). Ran a feed for several hours; no issues or errors found. Pump is functional.
62705 | 209902 | CM | CMPLT | CMPLT | 2024-09-20 10:47 | 2024-09-20 10:54 | User reported that there is no ECG reading | (TC 09/20/2024 10:53:40 AM, Baghirzade, Nara) Connected all leads to simulator; functional tests passed. All leads are functional. Printed the reading; printer is functional. ECG device is functional.
1055803 | 209903 | CM | CMPLT | CMPLT | 2024-09-20 13:32 | 2024-09-20 13:32 | Clinician saying the cardiac monitor wires weren't working | (TC 09/20/2024 01:31:49 PM, Scero, Tomlyn) Replaced missing defib pads and performed 30J test with no issues. Connected my simulator and tested the philips patient leads. No issues were found with the leads, all leads showing proper waveforms. Notified Nana that the defib is ok to use. Also noted to her not to have electrodes on the leads when not in use.
1067975 | 209904 | CM | ACTIV | ISSUE | 2024-09-20 13:40 |  | MSH 3 ER rm G7 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | (TC 09/20/2024 01:39:37 PM, Whittington, Scott) MSH 3 ER rm G7 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | MSH 3 ER rm G7 i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 SAMEH
1067979 | 209905 | CM | ACTIV | ISSUE | 2024-09-20 13:44 |  | MSH 20 SPORTS CLINIC i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | (TC 09/20/2024 01:49:49 PM, Whittington, Scott) MSH 20 SPORTS CLINIC i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 | MSH 20 SPORTS CLINIC i800 PRISM SOFTWARE UPGRADE VERSION 6.5> VERSION 7 SAMEH
66474 | 209906 | CM | CMPLT | CMPLT | 2024-09-20 14:13 | 2024-09-19 14:12 | It needs SW upgrade to 2.51.01 | It has new LI-On battery PS500. Draeger rep forgot to upgrade the SW. Did not passed battery check. upgraded SW to 2.51.01. passed battery check. 230 min and 43 Ah.
502453 | 209907 | CM | ACTIV | ISSUE | 2024-09-20 14:52 |  | MSH OPG SPP RM C2 VOLUSON HAS LATERAL BLANK ACROSS MIDDLE OF BEAM IMAGE | (TC 09/20/2024 02:53:34 PM, Whittington, Scott) MSH OPG SPP RM C2 VOLUSON HAS LATERAL BLANK ACROSS MIDDLE OF BEAM IMAGE ADJUSTED VIRTUAL TGC CONTROLS, NO AFFECT REBOOTED, IMAGE RETURNED TO NORMAL | MSH OPG SPP RM C2 VOLUSON HAS LATERAL BLANK ACROSS MIDDLE OF BEAM IMAGE
1064418 | 209908 | CM | CMPLT | CMPLT | 2024-09-20 15:53 | 2024-10-21 08:21 | Poor doors with no seals, loose latch, and temperature drop during use. | (TC 10/21/2024 08:20:32 AM, Ferreira, Qwynn) Replaced the seals with 3d printed ones using formlabs biomed 50a Tightened the latch on the port door Tightened latch on the corner main door Ran the device to test temperature. Temperature was stable at 33degrees. Problem was no seals to keep heat in causing the temperature to drop quickly. Noticed LED charging light not on when charging. Notified manufacturer to find out issue.
| 209909 | RDS | CMPLT | CMPLT | 2024-09-23 07:14 | 2024-09-20 07:14 | Weekly Rounds: 11 North | (TC 09/20/2024 07:13:57 AM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 209910 | RDS | CMPLT | CMPLT | 2024-09-23 07:48 | 2024-09-20 07:48 | Weekly Rounds: 11 South | (TC 09/20/2024 07:48:20 AM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 209911 | RDS | CMPLT | CMPLT | 2024-09-23 07:50 | 2024-09-20 07:51 | Weekly Rounds: NICU & RESUS | (TC 09/20/2024 07:50:40 AM, Baghirzade, Nara) Weekly rounds in NICU & RESUS completed. See procedure for results.
32560 | 209912 | CM | CMPLT | CMPLT | 2024-09-23 08:15 | 2024-09-20 19:00 | UndermA (Fixed) | (TC 09/18/2024 08:15:38 AM, Khan, Muhammad) Under mA (Fixed) -Tube has high slice count (1175411) -Unable to do IF -Fails on exposure test under different settings -mA signal can't be seen on signal show -HV shows some spike -Oredered tube and HV cable -Replaced Tube first but error didn't clear -Replaced HV cable and it was still the same (TC 09/19/2024 08:15:38 AM, Khan, Muhammad) -Further troubleshooting was done -Ordered and replaced HV unit -Replaced HV unit -This time system passese on IF and exposure test on Small Focus but fails on Large Focus -When checking for continuity on the high voltage cable the resistance between L focus and resistance b/n Small Focus was the same (11 ohms) Which can mean the cable was shorted. -Replaced HV cable with on that we had on site and that resolved the issue. (TC 09/20/2024 08:15:38 AM, Khan, Muhammad) -Ordred and replaced HV cable -System passed IF cal -Tube alignment was done -Seasoning goes only half way and fails - twice with XC Overcurrent error. Canon support suggested skipping seasoning as it was over heating. -Calibrations were done and system returned for use. (TC 09/23/2024 03:27:36 PM, Debebe, Adam) Assisted Canon in shipping back all parts.
| 209913 | RDS | CMPLT | CMPLT | 2024-09-23 08:45 | 2024-09-24 09:42 | Weekly check for ED
| 209914 | RDS | CMPLT | CMPLT | 2024-09-23 08:46 | 2024-10-08 14:57 | Weekly check for ICU
1061217 | 209915 | CM | CMPLT | NFF | 2024-09-23 09:02 | 2024-09-23 09:02 | PC not working | (TC 09/23/2024 09:02:24 AM, Reid, Emmerson ) Went to look at it but they solved the issue before I came up.
63724 | 209916 | CM | CMPLT | NFF | 2024-09-23 09:03 | 2024-09-23 09:04 | Not Reading | (TC 09/23/2024 09:03:33 AM, Reid, Emmerson ) Tested with a simulator and it is reading on all leads. NFF
1062999 | 209917 | CM | CMPLT | NFF | 2024-09-23 09:14 | 2024-09-23 09:14 | Not working | (TC 09/23/2024 09:14:15 AM, Canjimavadivel, Kamallesh ) Turned on device and ran treatment on max temp setting. Was able to reach temp and all ports working. No issues or alarms present.
| 209918 | RDS | CMPLT | CMPLT | 2024-09-23 10:05 | 2024-09-25 11:38 | Weekly checks in CCU | (TC 09/25/2024 11:38:20 AM, Scero, Tomlyn) weekly checks complete. 30J tests passed.
67961 | 209919 | PM | CMPLT | CMPLT | 2024-09-23 10:20 | 2024-09-23 10:20 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 09/23/2024 10:20:23 AM, Alvarado, Karla) Incubator was sent down to biomed; the system showed error #25. According to the manual, fan requires inspection. Checked the fan and found that it was not properly seated. Re-seated the fan. Performed PM, where VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. RH was calibrated. RH is functional and accurate; set 75%, analyzer = 75.6%. Set scale to 5kg, measured 5kg. Set temperature to 35C, measured 35C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 72.5 uA; GND Resistance = 171 mohms PM completed. (TC 09/23/2024 10:20:23 AM, Baghirzade, Nara) Supervised
1047853 | 209920 | CM | CMPLT | CMPLT | 2024-09-23 11:08 | 2024-10-31 13:25 | Defective clutch/coarse travel | (TC 09/23/2024 12:06:03 PM, Alvarado, Karla) Disassembled the device. Replaced clutch. Put grease on the leadscrew and clutch. Assembled the device. Continuous force sensor error when the device is turned on. Replaced plunger head assembly. However, the force sensor keep showing error. Tested a new data transmission cable and confirmed the data cable is defective. The cable will be replaced. (TC 09/23/2024 08:52:21 AM, Baghirzade, Nara) Disassembled device. Data transmission cable was replaced however its too long on the plungerhead side. Turned the device on to ensure the error was cleared; no errors on start-up. Will disassemble the device to reseat the cable. (TC 09/24/2024 02:29:58 PM, Baghirzade, Nara) Plunger cable/data transmission cable was re-seated. Device was assembled by Karla. When trying to calibrate, the device prompts that values are out of range. When moving the plunger back and forth, noted that it was rigid and loud; will have to diasssemble and inspect. (TC 10/03/2024 03:00:00 PM, Alvarado, Karla) Disassembled device. Re-seated the leadscrew. (TC 10/04/2024 01:00:00 PM, Alvarado, Karla) Re-seated syringe barrel clamp and syringe plunger holders. Assembled the device. (TC 10/31/2024 01:25:14 PM, Rhuland, Courtney) Assembled plungerhead and tested device Clicking noise was present and coarse travel error present Opened device and found leadscrew nut to be loose, causing bushing to be out of place Reseated bushing and tightened nut Reseated C-clip at wormgear to hold in place Reassembled device - performed occlusion test and test passed, no errors PM to be completed in PM WO
65616 | 209921 | CM | CMPLT | CMPLT | 2024-09-23 11:18 | 2024-10-29 13:54 | Lead Issue | Sept 23 - Request from physician to take a look at the unit. Went to check. Found one of the ecg/eeg lead missing. Replaced it. Was also requested to order a couple of full sets to be kept as spare. Contacted vendor. Requested info. Received. Created SC. Requested approval. Received. Placed order via MMC. Oc 29 - Cables received. Organized it in inventory.
66470 | 209922 | PM | CMPLT | CMPLT | 2024-09-23 11:39 | 2024-09-23 11:40 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 18 Ah and 230 min. Tested. Functional. SW.2.51 Leakage: 99 uA and 95 ohm.
500732 | 209923 | CM | CMPLT | CMPLT | 2024-09-23 13:55 | 2024-09-25 13:28 | MSH-20-470 APLIO XG SETUP IN SPORTS CLINIC AFTER FAILURE OF PRISM RX BD | (TC 09/23/2024 01:54:39 PM, Whittington, Scott) MSH-20-470 APLIO XG SETUP IN SPORTS CLINIC AFTER FAILURE OF PRISM RX BD (TC 09/25/2024 08:52:01 AM, Whittington, Scott) MSH-20-470 APLIO XG SETUP IN SPORTS CLINIC AFTER FAILURE OF PRISM RX BD PM WHEN DELIVERED TO MSH 20 (TC 09/25/2024 08:52:01 AM, Whittington, Scott) MSH-20-470 APLIO XG SETUP IN SPORTS CLINIC AFTER FAILURE OF PRISM RX BD PM WHEN DELIVERED TO MSH 20 | MSH-20-470 APLIO XG SETUP IN SPORTS CLINIC AFTER FAILURE OF PRISM RX BD
1067979 | 209924 | CM | CMPLT | CMPLT | 2024-09-23 14:14 | 2024-09-25 11:16 | MSH 20 SPORTS CLINIC PRISM I800 NOT SCANNING SAMEH DIAGNOSED ERROR AS RX RC BD FAILURE | (TC 09/23/2024 02:17:44 PM, Whittington, Scott) MSH 20 SPORTS CLINIC PRISM I800 NOT SCANNING SAMEH DIAGNOSED ERROR AS RX RC BD FAILURE (TC 09/24/2024 12:38:44 PM, Whittington, Scott) MSH 20 SPORTS i800 PRISM NOT SCANNING RXRC BD FAIL SAMEH WILL BRING BD (TC 09/24/2024 12:38:44 PM, Whittington, Scott) MSH 20 SPORTS i800 PRISM NOT SCANNING RXRC BD FAIL SAMEH WILL BRING BD (TC 09/24/2024 12:38:44 PM, Whittington, Scott) MSH 20 SPORTS i800 PRISM NOT SCANNING RXRC BD FAIL SAMEH WILL BRING BD (TC 09/24/2024 10:38:00 PM, Whittington, Scott) MSH 20 SPORTS i800 PRISM NOT SCANNING RXRC BD FAIL SAMEH WILL BRING BD SEPT 24 BOARD REPLACED OK UPGRADE WAS NOT INVOLVED IN FAULT, JUST COINCIDENCE From: sameh.mekhaeil@medical.canon <sameh.mekhaeil@medical.canon> Sent: September 24, 2024 4:19 PM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Giannetti, Gabriella - Sinai Health System <Gabriella.Giannetti@sinaihealth.ca> Cc: Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca>; Kumaran.Thayumanavan@medical.canon; Ryan.Francis@medical.canon; Ashley.Sestito@medical.canon; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca>; Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] Software upgrade-V7.0 Hello everyone, The problem has been resolved. The error was due to RX RC pwb malfunction. Regards, Sameh Sent from my iPhone On Sep 23, 2024, at 12:20?PM, SAMEH MEKHAEIL/CMSCA <sameh.mekhaeil@medical.canon> wrote: Hi Scott, I am off today. However, based on the error message you sent me, you need to replace the RX-RC PWB, part# PM30-44224*F 2 RX-RC-ASSY-V5 The problem has nothing to do with the upgrade that was done on Friday. We only upgraded the system software; no hardware upgrade was included in this upgrade. Regards, Sameh From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: September 23, 2024 11:35 AM To: SAMEH MEKHAEIL/CMSCA <sameh.mekhaeil@medical.canon>; Giannetti, Gabriella - Sinai Health System <Gabriella.Giannetti@sinaihealth.ca> Cc: Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca>; KUMARAN THAYUMANAVAN/CMSCA <Kumaran.Thayumanavan@medical.canon>; RYAN FRANCIS/CMSCA <Ryan.Francis@medical.canon>; ASHLEY SESTITO/CMSCA <Ashley.Sestito@medical.canon>; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca>; Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] Software upgrade-V7.0 Hi Gabby I will check it out shortly Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Giannetti, Gabriella <Gabriella.Giannetti@sinaihealth.ca> Sent: Monday, September 23, 2024 10:17:58 AM To: sameh.mekhaeil@medical.canon <sameh.mekhaeil@medical.canon>; Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca>; Kumaran.Thayumanavan@medical.canon <Kumaran.Thayumanavan@medical.canon>; Ryan.Francis@medical.canon <Ryan.Francis@medical.canon>; Ashley.Sestito@medical.canon <Ashley.Sestito@medical.canon>; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca>; Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: RE: [External] Software upgrade-V7.0 Hi all, From what I know the upgrade occurred on Friday to our ultrasound machine in the Dovigi clinic however this morning one of our physicians has told me the US machine is not working. Is someone available to come check on this today at some point? Please let me know, thanks so much! Gabby Giannetti, RN, MBA (she/her) Manager, Dovigi Orthopaedic Sports Medicine Clinic Adjunct Lecturer, Lawrence S. Bloomberg Faculty of Nursing, University of Toronto T 416-586-4800 ext. 3843 C 437-522-4914 gabriella.giannetti@sinaihealth.ca sinaihealth.ca 600 University Avenue, Toronto, ON M5G 1X5 (TC 09/25/2024 10:38:00 AM, Whittington, Scott) MSH 20 SPORTS i800 PRISM NOT SCANNING RXRC BD FAIL SAMEH WILL BRING BD SEPT 24 BOARD REPLACED OK UPGRADE WAS NOT INVOLVED IN FAULT, JUST COINCIDENCE From: sameh.mekhaeil@medical.canon <sameh.mekhaeil@medical.canon> Sent: September 24, 2024 4:19 PM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Giannetti, Gabriella - Sinai Health System <Gabriella.Giannetti@sinaihealth.ca> Cc: Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca>; Kumaran.Thayumanavan@medical.canon; Ryan.Francis@medical.canon; Ashley.Sestito@medical.canon; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca>; Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] Software upgrade-V7.0 Hello everyone, The problem has been resolved. The error was due to RX RC pwb malfunction. Regards, Sameh Sent from my iPhone On Sep 23, 2024, at 12:20?PM, SAMEH MEKHAEIL/CMSCA <sameh.mekhaeil@medical.canon> wrote: Hi Scott, I am off today. However, based on the error message you sent me, you need to replace the RX-RC PWB, part# PM30-44224*F 2 RX-RC-ASSY-V5 The problem has nothing to do with the upgrade that was done on Friday. We only upgraded the system software; no hardware upgrade was included in this upgrade. Regards, Sameh From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: September 23, 2024 11:35 AM To: SAMEH MEKHAEIL/CMSCA <sameh.mekhaeil@medical.canon>; Giannetti, Gabriella - Sinai Health System <Gabriella.Giannetti@sinaihealth.ca> Cc: Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca>; KUMARAN THAYUMANAVAN/CMSCA <Kumaran.Thayumanavan@medical.canon>; RYAN FRANCIS/CMSCA <Ryan.Francis@medical.canon>; ASHLEY SESTITO/CMSCA <Ashley.Sestito@medical.canon>; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca>; Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: Re: [External] Software upgrade-V7.0 Hi Gabby I will check it out shortly Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Giannetti, Gabriella <Gabriella.Giannetti@sinaihealth.ca> Sent: Monday, September 23, 2024 10:17:58 AM To: sameh.mekhaeil@medical.canon <sameh.mekhaeil@medical.canon>; Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Calabrese, Angela - Sinai Health System <Angela.Calabrese@sinaihealth.ca>; Kumaran.Thayumanavan@medical.canon <Kumaran.Thayumanavan@medical.canon>; Ryan.Francis@medical.canon <Ryan.Francis@medical.canon>; Ashley.Sestito@medical.canon <Ashley.Sestito@medical.canon>; Ivo, Barry - Sinai Health System <Barry.Ivo@sinaihealth.ca>; Colley, Krista - Sinai Health System <Krista.Colley@sinaihealth.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Subject: RE: [External] Software upgrade-V7.0 Hi all, From what I know the upgrade occurred on Friday to our ultrasound machine in the Dovigi clinic however this morning one of our physicians has told me the US machine is not working. Is someone available to come check on this today at some point? Please let me know, thanks so much! Gabby Giannetti, RN, MBA (she/her) Manager, Dovigi Orthopaedic Sports Medicine Clinic Adjunct Lecturer, Lawrence S. Bloomberg Faculty of Nursing, University of Toronto T 416-586-4800 ext. 3843 C 437-522-4914 gabriella.giannetti@sinaihealth.ca sinaihealth.ca 600 University Avenue, Toronto, ON M5G 1X5 | MSH 20 SPORTS CLINIC PRISM I800 NOT SCANNING SAMEH DIAGNOSED ERROR AS RX RC BD FAILURE
1054462 | 209925 | CM | CMPLT | CMPLT | 2024-09-23 14:32 | 2024-09-24 09:03 | Door latch stuck | (TC 09/23/2024 09:02:59 AM, Canjimavadivel, Kamallesh ) Removed latch and applied lubricant to free up movement. Cleaned spring and tested latch. Needed to open up device to reattach latch. (TC 09/24/2024 09:02:59 AM, Canjimavadivel, Kamallesh ) Finished reassmbling and ran device to test all functions and alarms.
1070763 | 209926 | CM | CMPLT | CMPLT | 2024-09-23 14:42 | 2024-09-23 14:43 | Repositioning phototherapy light | (TC 09/23/2024 02:42:41 PM, Rhuland, Courtney) Located device Moved phototherapy light and O2 tank to right side and exam light to left side Requested by Janet in NICU Functional - OK
1070798 | 209927 | CM | CMPLT | CMPLT | 2024-09-23 14:42 | 2024-09-23 14:44 | Repositioning phototherapy light | (TC 09/23/2024 02:44:11 PM, Rhuland, Courtney) Located device Moved phototherapy light and O2 tank to right side and exam light to left side Requested by Janet in NICU Functional - OK
1050604 | 209928 | CM | CMPLT | CMPLT | 2024-09-24 07:43 | 2024-09-23 15:00 | Scale broken / not working. | (TC 09/23/2024 03:00:00 PM, Alvarado, Karla) Replaced batteries. Tested 25lb, 40lb, and 65 lb. All tests passed.
1056453 | 209929 | CM | CMPLT | CMPLT | 2024-09-24 08:09 | 2024-09-24 08:09 | O2 sat broken. | (TC 09/24/2024 08:08:53 AM, Alvarado, Karla) Replaced temperature prob. Replaced SpO2 sensor. Performed tests: Temperature, NIBP, and SpO2 All test passed.
1055981 | 209930 | CM | CMPLT | CMPLT | 2024-09-24 08:21 | 2024-09-24 08:22 | Battery replacement | (TC 09/24/2024 08:21:45 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055922 | 209931 | CM | CMPLT | CMPLT | 2024-09-24 08:47 | 2024-09-24 08:47 | O2 sat broken. | (TC 09/24/2024 08:47:23 AM, Alvarado, Karla) Replaced power cord Replaced SpO2 sensor. Performed tests: Temperature NIBP = set: 120/80, measured: 113/83 SpO2 = set: 95%, measured: 95% All test passed.
1056442 | 209932 | CM | CMPLT | CMPLT | 2024-09-24 08:50 | 2024-09-24 08:50 | Power cord broken. | (TC 09/24/2024 08:50:13 AM, Alvarado, Karla) Replaced power cord Performed tests: Temperature NIBP = set: 120/80, measured: 115/84 SpO2 = set: 95%, measured: 95% All test passed.
1072003 | 209933 | ACC | CMPLT | CMPLT | 2024-09-24 11:16 | 2024-09-24 11:40 | Incoming inspection | (TC 09/24/2024 11:40:23 AM, Alvarado, Karla) Incoming inspection completed. Radical was configured. Functional test passed.
1053348 | 209934 | CM | ACTIV | ISSUE | 2024-09-24 11:24 |  | MSH BI 1272 H3 FILTER WHEEL ARTIFACT ON AG & CU CEM ON QC | (TC 09/24/2024 12:14:44 PM, Whittington, Scott) MSH BI 1272 H3 FILTER WHEEL ARTIFACT ON AG & CU CEM ON QC SN# 3DM160700414 REF# TIM 01605530 ROSE FOUND ARTIFACTS DURING QC WITH SO YOUJG, MARKDED FOR SERVICE AFFECTED IMAGES RYAN HAD SUGGESTED CLEANING PHANTOM AND RE DOING QC, IF ARTIFACT ON ALL 4 FILTERS, APPEARS TO BE ONLY 2 FILTERS AFFECTED (TC 10/01/2024 10:59:32 AM, Whittington, Scott) MSH BI 1272 H3 FILTER WHEEL ARTIFACT ON AG & CU CEM ON QC SN# 3DM160700414 REF# TIM 01605530 | MSH BI 1272 H3 FILTER WHEEL ARTIFACT ON AG & CU CEM ON QC SN# 3DM160700414 REF# TIM 01605530
| 209935 | PM | CMPLT | CMPLT | 2024-09-24 12:13 | 2024-09-24 12:13 | Test asset PM Experiment | PM Schedule experiment
1067975 | 209936 | CM | CMPLT | CMPLT | 2024-09-24 12:30 | 2024-09-24 12:30 | MSH 3 ER rm G7 i800 PRISM RANDON ANNOTATION | (TC 09/24/2024 12:29:47 PM, Whittington, Scott) MSH 3 ER rm G7 i800 PRISM RANDON ANNOTATION REPLACED QWERTY KYBD FROM TWH STOCK (TC 09/24/2024 12:29:47 PM, Whittington, Scott) MSH 3 ER rm G7 i800 PRISM RANDON ANNOTATION REPLACED QWERTY KYBD FROM TWH STOCK (TC 09/24/2024 01:55:14 PM, Whittington, Scott) MSH 3 ER rm G7 i800 PRISM RANDON ANNOTATION REPLACED QWERTY KYBD FROM TWH STOCK (TC 09/24/2024 07:55:00 AM, Whittington, Scott) MSH 3 ER rm G7 i800 PRISM RANDON ANNOTATION REPLACED QWERTY KYBD FROM TWH STOCK (TC 09/24/2024 07:55:00 AM, Whittington, Scott) MSH 3 ER rm G7 i800 PRISM RANDON ANNOTATION REPLACED QWERTY KYBD FROM TWH STOCK | MSH 3 ER rm G7 i800 PRISM RANDON ANNOTATION
1049905 | 209937 | CM | CMPLT | CMPLT | 2024-09-24 12:34 | 2024-09-25 11:11 | MSH R65 MOBILE STUCK IN PT ELEVATOR CAR # 7 | (TC 09/24/2024 01:57:35 PM, Whittington, Scott) MSH R65 MOBILE STUCK IN PT ELEVATOR CAR # 7 | MSH R65 MOBILE STUCK IN PT ELEVATOR CAR # 7
1068221 | 209938 | CM | CMPLT | CMPLT | 2024-09-24 12:36 | 2024-09-26 12:36 | MSH 5 US PRISM ANNOTATION AUTO DELETED FROM ABDO STUDIES | (TC 09/24/2024 12:36:11 PM, Whittington, Scott) From: Angaran, Crystal <Crystal.Angaran@sinaihealth.ca> Sent: September 24, 2024 11:17 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] RE: Rm6 Canon machine on 5th floor ultrasound Hey Scott, I may have figured out what the problem on Rm6 was. There is an Auto Delete button that was selected on the abdominal setting…so that may have been why the annotation was deleted after unfreezing the image. We’ll keep an eye on this machine to see if it happens again. Crystal From: Angaran, Crystal Sent: September 24, 2024 10:11 AM To: 'scott.whittington@uhn.on.ca' <scott.whittington@uhn.on.ca> Subject: Rm6 Canon machine on 5th floor ultrasound Hi Scott, With regards to the Rm6 machine on the 5th floor, when you scan on the abdominal setting, every time you unfreeze the image the annotation disappears. However, when the tech was scanning on the Leg Doppler setting, the annotation did not disappear and remained on the screen when unfreezing the image. Crystal (TC 09/24/2024 12:36:11 PM, Whittington, Scott) From: Angaran, Crystal <Crystal.Angaran@sinaihealth.ca> Sent: September 24, 2024 11:17 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] RE: Rm6 Canon machine on 5th floor ultrasound Hey Scott, I may have figured out what the problem on Rm6 was. There is an Auto Delete button that was selected on the abdominal setting…so that may have been why the annotation was deleted after unfreezing the image. We’ll keep an eye on this machine to see if it happens again. Crystal From: Angaran, Crystal Sent: September 24, 2024 10:11 AM To: 'scott.whittington@uhn.on.ca' <scott.whittington@uhn.on.ca> Subject: Rm6 Canon machine on 5th floor ultrasound Hi Scott, With regards to the Rm6 machine on the 5th floor, when you scan on the abdominal setting, every time you unfreeze the image the annotation disappears. However, when the tech was scanning on the Leg Doppler setting, the annotation did not disappear and remained on the screen when unfreezing the image. Crystal (TC 09/26/2024 12:36:11 PM, Whittington, Scott) From: Angaran, Crystal <Crystal.Angaran@sinaihealth.ca> Sent: September 24, 2024 11:17 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Subject: [External] RE: Rm6 Canon machine on 5th floor ultrasound Hey Scott, I may have figured out what the problem on Rm6 was. There is an Auto Delete button that was selected on the abdominal setting…so that may have been why the annotation was deleted after unfreezing the image. We’ll keep an eye on this machine to see if it happens again. Crystal From: Angaran, Crystal Sent: September 24, 2024 10:11 AM To: 'scott.whittington@uhn.on.ca' <scott.whittington@uhn.on.ca> Subject: Rm6 Canon machine on 5th floor ultrasound Hi Scott, With regards to the Rm6 machine on the 5th floor, when you scan on the abdominal setting, every time you unfreeze the image the annotation disappears. However, when the tech was scanning on the Leg Doppler setting, the annotation did not disappear and remained on the screen when unfreezing the image. Crystal | MSH 5 US PRISM ANNOTATION AUTO DELETED FROM ABDO STUDIES
64267 | 209939 | CM | CMPLT | CMPLT | 2024-09-24 12:52 | 2024-09-24 12:53 | Project - J13 option (RS232 board) upgrade in monitor | (TC 09/24/2024 12:52:44 PM, Rhuland, Courtney) Monitor in room did not have board Requested by PIONIRS study to install Patient monitored on X2 while monitor was removed. Disassembled the rear of the monitor; inserted the RS232 board. Put monitor back together. Switched the monitor on, went in to service mode and changed the driver to DtOut1 for the RS232. RS232 SN# DE48609377 Confirmed with study that data was being transferred.
1056443 | 209940 | CM | CMPLT | CMPLT | 2024-09-24 13:18 | 2024-09-24 13:18 | Reading is off and not accurate. | (TC 09/24/2024 01:17:54 PM, Alvarado, Karla) Performed tests: Temperature NIBP = set: 120/80, measured: 116/84 SpO2 = set: 95%, measured: 95% All test passed.
| 209941 | CM | CMPLT | CMPLT | 2024-09-24 13:34 | 2024-09-24 13:38 | Thermometer troubleshooting | (TC 09/24/2024 01:38:20 PM, Rhuland, Courtney) Multiple thermometers stuck with OFF on screen and temperature reading in small font on top. Could not find any information in manual regarding this setting Attempted troubleshooting but could not turn off this function Called Cardinal health and was told this is a common issue and they do not have a fix for it other than sending device in for repair. Preparing to replace fleet on next capital so devices will not be sent out - to be retired.
| 209942 | PA | CMPLT | CMPLT | 2024-09-24 14:06 | 2024-09-24 15:22 | Powerstrip for stock | (TC 09/24/2024 03:22:30 PM, Canjimavadivel, Kamallesh ) Placed order for powerstrips
1066899 | 209943 | CM | ACTIV | ISSUE | 2024-09-24 15:34 |  | Periodical alarms and malfunctions about every 2 or 3 hours | (TC 09/24/2024 03:34:05 PM, Alvarado, Karla) Pump was running for 3 different periods of time (8 hours for 3 days). The third day, the pump showed error #23. According to the manual, “Motor Failsafe Time Out Error. Possible damaged gearbox encoder or flex cable.” Cardinal health service was contacted to send it for repair. (TC 09/25/2024 08:43:51 AM, Canjimavadivel, Kamallesh ) Requested no chage PO
1055977 | 209944 | CM | ACTIV | ISSUE | 2024-09-25 11:03 |  | Broken pins - shipping for repair | (TC 09/25/2024 12:58:45 PM, Baghirzade, Nara) 1st pin on the radical keeps getting stuck; 2nd pin is broken off. Created a service request with Masimo. (TC 10/02/2024 07:49:34 AM, Baghirzade, Nara) SC 1003321401 (TC 10/02/2024 10:58:16 AM, Baghirzade, Nara) RMA 46444 KJV # 457177 PO 7500212682 (TC 10/31/2024 12:02:05 PM, Baghirzade, Nara) Received radical 7 from repair. The packing slip indicates the wrong PO#; reached out to Fan and Masimo to fix this. Configured the radical 7.
65264 | 209945 | PM | CMPLT | CMPLT | 2024-09-25 11:11 | 2024-09-25 11:11 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check. 16 Ah and 220 min. Tested. Functional. SW.2.51 Leakage: 89 uA and 93 ohm.
66474 | 209946 | CM | CMPLT | CMPLT | 2024-09-25 11:13 | 2024-09-24 11:13 | unit not passing battery check | collected logs. created help ticket. as per Draeger tech support upgraded unit SW from 2.51 to 2.51.01 but Weight option on the Start/standby screen is missing. Tried to put it manually and I copied config from another 2.51.01 unit too. But both did not work. Contacted Drager again. started new patient on startup screen as per David. It worked. Returned unit back to service.
1055585 | 209947 | CM | CMPLT | NFF | 2024-09-25 12:16 | 2024-09-25 12:16 | Machine is not working. | (TC 09/25/2024 12:16:18 PM, Alvarado, Karla) Performed tests: Temperature NIBP= set: 120/80, measured: 117/84 SpO2 = set: 95%, measured: 95% All test passed. No Fault Found
| 209948 | CM | CMPLT | CMPLT | 2024-09-25 13:34 | 2024-09-25 13:34 | NICU FY-Capital meeting | (TC 09/25/2024 01:34:03 PM, Baghirzade, Nara) NICU FY-Capital meeting
| 209949 | CM | CMPLT | CMPLT | 2024-09-25 13:35 | 2024-09-25 13:35 | NICU silence alarm button from overview stations issue | (TC 09/25/2024 01:34:59 PM, Baghirzade, Nara) NICU silence alarm button from overview stations issue meeting
1054985 | 209950 | CM | CMPLT | CMPLT | 2024-09-25 13:54 | 2024-09-25 13:54 | Root configuration and functional testing | (TC 09/25/2024 01:53:58 PM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Replaced temperature probe and well. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1054987 | 209951 | CM | CMPLT | CMPLT | 2024-09-25 13:57 | 2024-09-25 10:55 | Reported that SpO2 not reading properly | (TC 09/25/2024 10:55:02 AM, Baghirzade, Nara) Used SpO2 simulator to check for accuracy: set - 90% measured - 90% set - 95% measured - 95% set - 100% measured - 100% Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 209952 | CM | CMPLT | CMPLT | 2024-09-25 14:52 | 2024-09-25 14:53 | Camera in black and white | (TC 09/25/2024 02:51:36 PM, Rhuland, Courtney) Very old video unit displayed on monitor in black and white Attempted to change settings on both monitor and adapter, but no change Video system outdated - to be replaced under capital in the coming year
1054963 | 209953 | CM | CMPLT | NFF | 2024-09-25 14:59 | 2024-09-25 15:00 | BP cuff not functional. | (TC 09/25/2024 02:59:36 PM, Alvarado, Karla) Performed tests: Temperature NIBP = set: 120/80, measured: 113/84 SpO2 = set: 95%, measured: 95% All test passed. NFF
| 209954 | OC | CMPLT | CMPLT | 2024-09-25 16:48 | 2024-09-25 16:48 | On-call - Sep 13-19 2024
1056055 | 209955 | CM | CMPLT | CMPLT | 2024-09-25 20:07 | 2024-09-25 20:06 | Padestal Control Console not working. | Repaired the control console, re-installl the software handed the system in good working condition.
1072008 | 209956 | ACC | CMPLT | CMPLT | 2024-09-26 07:20 | 2024-09-30 14:41 | Incoming inspection | (TC 09/30/2024 02:40:42 PM, Cortez, John) Unboxed, completed check-in successfully
1072009 | 209957 | ACC | CMPLT | CMPLT | 2024-09-26 07:29 | 2024-09-30 12:52 | Incoming inspection | (TC 09/30/2024 12:52:23 PM, Cortez, John) Unboxed, completed check-in successfully
1072010 | 209958 | ACC | CMPLT | CMPLT | 2024-09-26 07:32 | 2024-09-30 14:40 | Incoming inspection | (TC 09/30/2024 02:39:53 PM, Cortez, John) Unboxed, completed check-in successfully
1072011 | 209959 | ACC | CMPLT | CMPLT | 2024-09-26 07:38 | 2024-09-30 14:39 | Incoming inspection | (TC 09/30/2024 02:39:12 PM, Cortez, John) Unboxed, completed check-in successfully
1072012 | 209960 | ACC | CMPLT | CMPLT | 2024-09-26 07:44 | 2024-09-30 14:38 | Incoming inspection | (TC 09/30/2024 02:37:36 PM, Cortez, John) Unboxed, completed check-in successfully
1072013 | 209961 | ACC | CMPLT | CMPLT | 2024-09-26 07:45 | 2024-09-30 14:36 | Incoming inspection | (TC 09/30/2024 02:35:38 PM, Cortez, John) Unboxed, completed check-in successfully
1072014 | 209962 | ACC | CMPLT | CMPLT | 2024-09-26 07:47 | 2024-09-30 14:34 | Incoming inspection | (TC 09/30/2024 02:34:01 PM, Cortez, John) Unboxed, completed check-in successfully
1072015 | 209963 | ACC | CMPLT | CMPLT | 2024-09-26 07:49 | 2024-09-30 14:33 | Incoming inspection | (TC 09/30/2024 02:33:16 PM, Cortez, John) Unboxed, completed check-in successfully
1072016 | 209964 | ACC | CMPLT | CMPLT | 2024-09-26 07:53 | 2024-09-30 14:30 | Incoming inspection | (TC 09/30/2024 02:29:34 PM, Cortez, John) Unboxed, completed check-in successfully
1072017 | 209965 | ACC | CMPLT | CMPLT | 2024-09-26 07:56 | 2024-09-30 14:26 | Incoming inspection | (TC 09/30/2024 02:25:56 PM, Cortez, John) Unboxed, completed check-in successfully
1072018 | 209966 | ACC | CMPLT | CMPLT | 2024-09-26 07:57 | 2024-09-30 14:24 | Incoming inspection | (TC 09/30/2024 02:24:29 PM, Cortez, John) Unboxed, completed check-in successfully
1061898 | 209967 | PM | CMPLT | CMPLT | 2024-09-26 08:01 | 2024-09-26 08:01 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.07 ohm, 137 uA
1072019 | 209968 | ACC | CMPLT | CMPLT | 2024-09-26 08:04 | 2024-09-30 14:21 | Incoming inspection | (TC 09/30/2024 02:20:58 PM, Cortez, John) Unboxed, completed check-in successfully
1061416 | 209969 | PM | CMPLT | CMPLT | 2024-09-26 08:05 | 2024-09-26 08:05 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. batteries changed and SW updated . 0.07 ohm, 126 uA
1072020 | 209970 | ACC | CMPLT | CMPLT | 2024-09-26 08:10 | 2024-09-30 14:19 | Incoming inspection | (TC 09/30/2024 02:18:53 PM, Cortez, John) Unboxed, completed check-in successfully
1072021 | 209971 | ACC | CMPLT | CMPLT | 2024-09-26 08:13 | 2024-09-30 14:12 | Incoming inspection | (TC 09/30/2024 02:11:57 PM, Cortez, John) Unboxed, completed check-in successfully
1072022 | 209972 | ACC | CMPLT | CMPLT | 2024-09-26 08:14 | 2024-09-30 13:07 | Incoming inspection | (TC 09/30/2024 01:07:15 PM, Cortez, John) Unboxed, completed check-in successfully
1072023 | 209973 | ACC | CMPLT | CMPLT | 2024-09-26 08:16 | 2024-09-30 13:04 | Incoming inspection | (TC 09/30/2024 01:04:24 PM, Cortez, John) Unboxed, completed check-in successfully
| 209974 | PA | CMPLT | CMPLT | 2024-09-26 10:19 | 2024-10-21 14:03 | Ordering mattress hose assembly | (TC 09/26/2024 10:20:38 AM, Baghirzade, Nara) SC 1003317707 (TC 10/21/2024 02:03:13 PM, Baghirzade, Nara) PO 7500212464
1056450 | 209975 | CM | CMPLT | CMPLT | 2024-09-26 10:57 | 2024-09-26 10:57 | Incoming inspection and configuration load | (TC 09/26/2024 10:57:08 AM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1051133 | 209976 | CM | CMPLT | CMPLT | 2024-09-26 11:13 | 2024-09-26 11:14 | P01 error | (TC 09/26/2024 11:13:54 AM, Rhuland, Courtney) Replaced probe with spare Replaced batteries Functional check - OK
67744 | 209977 | CM | CMPLT | CMPLT | 2024-09-26 11:42 | 2024-09-26 11:42 | Battery replacement | (TC 09/26/2024 11:41:40 AM, Alvarado, Karla) Replaced battery. (TC 09/26/2024 11:41:40 AM, Rhuland, Courtney) Supervised
1056439 | 209978 | FC | CMPLT | CMPLT | 2024-09-26 12:40 | 2024-09-26 12:40 | Functional testing | (TC 09/26/2024 12:39:43 PM, Baghirzade, Nara) Confirmed with Becky that this device can go back to 11 North. Retrieved from MCR room. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055915 | 209979 | CM | CMPLT | CMPLT | 2024-09-26 12:41 | 2024-09-26 12:41 | Functional testing | (TC 09/26/2024 12:41:01 PM, Baghirzade, Nara) Confirmed with Becky that this device can go back to 11 North. Retrieved from MCR room. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1057120 | 209980 | CM | CMPLT | CMPLT | 2024-09-26 12:42 | 2024-09-26 12:50 | Repurposed for 11 South | (TC 09/26/2024 12:49:51 PM, Baghirzade, Nara) Confirmed with Becky that this device can be repurposed for 11 South. Retrieved from MCR room. Updated location and account. Printed new asset tag. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional. (TC 09/26/2024 12:49:51 PM, Baghirzade, Nara) Confirmed with Becky that this device can be repurposed for 11 South. Retrieved from MCR room. Updated location and account. Printed new asset tag. No patient cable and SpO2 sensor found; provided a LNCS patient cable and SpO2 sensor. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1061320 | 209981 | CM | CMPLT | CMPLT | 2024-09-26 12:55 | 2024-09-26 12:55 | Repurposed for 11 South | (TC 09/26/2024 12:55:06 PM, Baghirzade, Nara) Confirmed with Becky that this device can be repurposed for 11 South. Retrieved from MCR room. Updated location and account. Printed new asset tag. Attached to Root 1057120
1056435 | 209982 | CM | CMPLT | CMPLT | 2024-09-26 12:56 | 2024-09-26 12:56 | Repurposed for 11 South | (TC 09/26/2024 12:56:14 PM, Baghirzade, Nara) Confirmed with Becky that this device can be repurposed for 11 South. Retrieved from MCR room. Updated location and account. Printed new asset tag. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1056493 | 209983 | CM | CMPLT | CMPLT | 2024-09-26 12:57 | 2024-09-26 12:57 | Repurposed for 11 South | (TC 09/26/2024 12:56:45 PM, Baghirzade, Nara) Confirmed with Becky that this device can be repurposed for 11 South. Retrieved from MCR room. Updated location and account. Printed new asset tag. Attached to Root 1056435
1061320 | 209984 | CM | CMPLT | CMPLT | 2024-09-26 12:58 | 2024-09-26 12:58 | Battery replacement | (TC 09/26/2024 12:57:54 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055965 | 209985 | CM | CMPLT | CMPLT | 2024-09-26 12:58 | 2024-09-26 12:58 | Battery replacement | (TC 09/26/2024 12:58:24 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1057614 | 209986 | CM | CMPLT | CMPLT | 2024-09-26 12:59 | 2024-09-26 12:59 | Battery replacement | (TC 09/26/2024 12:58:52 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1057615 | 209987 | CM | CMPLT | CMPLT | 2024-09-26 13:00 | 2024-09-26 13:01 | Repurposed as a biomed spare device | (TC 09/26/2024 01:00:33 PM, Baghirzade, Nara) Confirmed with Becky that this device can be repurposed for biomed spare. Retrieved from MCR room. Updated location and account. Printed new asset tag. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1055906 | 209988 | CM | CMPLT | CMPLT | 2024-09-26 13:01 | 2024-09-26 13:01 | Repurposed as a biomed spare device | (TC 09/26/2024 01:01:28 PM, Baghirzade, Nara) Confirmed with Becky that this device can be repurposed for biomed spare. Retrieved from MCR room. Updated location and account. Printed new asset tag. Previously Root had a PatientSafetyNet configuration; changed configuration to MSH VSM. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1066454 | 209989 | FC | CMPLT | CMPLT | 2024-09-26 13:03 | 2024-09-26 13:03 | QC checklist | (TC 09/26/2024 01:03:03 PM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
1065409 | 209990 | CM | ACTIV | ISSUE | 2024-09-26 15:00 |  | IVF Visit - Parts Order | Sept 26 - IVF manager reached out to explore and see if their vital sign machine can be equiped with capnography. Went to check and did a walk through. Found one WA monitor with capnography capability. Was inrmed that they would like to start utilizing it. Checked device record. Found CO2 was never calibrated. Advised manager to awaiti unitl I am able to calibraqte the monitor. Came back. Reached out to vendor and requested parts info. Awaiting response. Oct 7/8 - Followed up 3 time within past two weeks. Was infoemd to reach out to Air liquid America. Cantacted them. Was informed to reach out to Vital aire canada. Contacted them and followed up twice. Received parts info. Created SC. Requested approval. Received. Placed order via MMC.
32510 | 209991 | CM | ACTIV | ISSUE | 2024-09-26 16:37 |  | MSH DR1 TUBE HEAD SWITCHES HAVE NO EFFECT | (TC 09/26/2024 04:37:34 PM, Whittington, Scott) MSH DR1 TUBE HEAD SWITCHES HAVE NO EFFECT (TC 09/25/2024 11:37:00 PM, Whittington, Scott) MSH DR1 TUBE HEAD SWITCHES HAVE NO EFFECT | MSH DR1 TUBE HEAD SWITCHES HAVE NO EFFECT
1057384 | 209992 | CM | ACTIV | ISSUE | 2024-09-26 16:41 |  | MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED | (TC 10/11/2024 04:34:04 PM, Whittington, Scott) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED SEAN YOUNG INSTALLED PC-HP5700 E7400 PN# SP9G6512-3 MIRZA WHITELISTED MAC, UNABLE TO MAKE CONNECTION ON EITHER JACK MIRZA UNABLE TO PING (TC 10/11/2024 04:34:51 PM, Dolinar, John) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED SEAN YOUNG INSTALLED PC-HP5700 E7400 PN# SP9G6512-3 MIRZA WHITELISTED MAC, UNABLE TO MAKE CONNECTION ON EITHER JACK MIRZA UNABLE TO PING (TC 09/27/2024 07:42:00 PM, Whittington, Scott) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED (TC 10/17/2024 01:34:00 PM, Whittington, Scott) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED SEAN YOUNG INSTALLED PC-HP5700 E7400 PN# SP9G6512-3 MIRZA WHITELISTED MAC, UNABLE TO MAKE CONNECTION ON EITHER JACK MIRZA UNABLE TO PING (TC 10/17/2024 01:34:00 PM, Dolinar, John) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED SEAN YOUNG INSTALLED PC-HP5700 E7400 PN# SP9G6512-3 MIRZA WHITELISTED MAC, UNABLE TO MAKE CONNECTION ON EITHER JACK MIRZA UNABLE TO PING (TC 10/17/2024 01:34:00 PM, Dolinar, John) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED SEAN YOUNG INSTALLED PC-HP5700 E7400 PN# SP9G6512-3 MIRZA WHITELISTED MAC, UNABLE TO MAKE CONNECTION ON EITHER JACK MIRZA UNABLE TO PING (TC 09/27/2024 07:42:00 PM, Whittington, Scott) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED (TC 09/27/2024 07:42:00 PM, Whittington, Scott) MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED | MSH CR 2B NO NETWORK CONNECTION, PT STUDY STRANDED SEAN YOUNG INSTALLED PC-HP5700 E7400 PN# SP9G6512-3
1063808 | 209993 | CM | ACTIV | ISSUE | 2024-09-26 16:44 |  | MSH 15 SPP VOLUSON VIDEO OUTPUT CABLE | (TC 09/26/2024 04:44:03 PM, Whittington, Scott) MSH 15 SPP VOLUSON VIDEO OUTPUT CABLE REQUIRES LONG SVGA CABLE | MSH 15 SPP VOLUSON VIDEO OUTPUT CABLE
32809 | 209994 | CM | ACTIV | ISSUE | 2024-09-26 16:47 |  | MSH DR1 CAR3ESTREAM LLI CUTOFF OF IMAGE AT FEET | (TC 09/27/2024 03:48:24 PM, Whittington, Scott) MSH DR1 CAR3ESTREAM LLI CUTOFF OF IMAGE AT FEET (TC 10/08/2024 03:30:43 PM, Whittington, Scott) MSH DR1 CAR3ESTREAM LLI CUTOFF OF IMAGE AT FEET CONFIRMED BY EXPOSURE WITH REVOLUTION MOBILE THAT ARTIFACT IS CREATED IN DR1, LIKELY UPPER CONES NOT OPENING COMPLETELY (TC 09/27/2024 05:33:00 PM, Whittington, Scott) MSH DR1 CAR3ESTREAM LLI CUTOFF OF IMAGE AT FEET | MSH DR1 CAR3ESTREAM LLI CUTOFF OF IMAGE AT FEET
1048472 | 209995 | CM | CMPLT | CMPLT | 2024-09-27 07:49 | 2024-09-27 07:51 | MSH R59 QUYEN REPORTED LOWER BATTERY CHARGE SLOT NOT WORKING | (TC 09/27/2024 07:51:39 AM, Whittington, Scott) From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: September 24, 2024 9:48 AM To: Ultrasound Support <Ultrasound.Support@uhn.ca>; Ho, Quyen - Sinai Health System <Quyen.Ho@sinaihealth.ca> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca> Subject: Re: [External] R59 issue Hi Quyen I just placed a small detector battery in the lower charge slot of L&D Resuc R59. The battery which is new and had been in reserve in our shop , is charging correctly. Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Whittington, Scott <Scott.Whittington@uhn.ca> Sent: Tuesday, September 24, 2024 8:00:45 AM To: Ultrasound Support <Ultrasound.Support@uhn.ca>; Ho, Quyen - Sinai Health System <Quyen.Ho@sinaihealth.ca> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca> Subject: Re: [External] R59 issue Hi Quyen I will be there shortly on 15 Thanks Scott Get Outlook for iOS ________________________________________ From: Ho, Quyen <Quyen.Ho@sinaihealth.ca> Sent: Tuesday, September 24, 2024 7:59:36 AM To: Whittington, Scott <Scott.Whittington@uhn.ca>; Ultrasound Support <Ultrasound.Support@uhn.ca> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca> Subject: RE: [External] R59 issue Yes, I checked the pins and they are not bent. For some reasons, it is not charging, the battery which was in there is dead and I tried to charge another battery and the message box came up on the collimator box saying that the bottom charger cannot charge. Quyen. From: Whittington, Scott [mailto:Scott.Whittington@uhn.ca] Sent: September 23, 2024 1:29 PM To: Ultrasound Support <Ultrasound.Support@uhn.ca>; Ho, Quyen <Quyen.Ho@sinaihealth.ca> Cc: Schwietz, Matthew <Matthew.Schwietz@sinaihealth.ca> Subject: Re: [External] R59 issue CAUTION: This email originated from outside of the organization. Do not click links or open attachments unless you recognize the sender and know the content is safe. Hi Quyen Lower charging slot appears to be ok. No broken pins Thanks Scott 416-402-2462 Get Outlook for iOS ________________________________________ From: Ho, Quyen <Quyen.Ho@sinaihealth.ca> Sent: Monday, September 23, 2024 9:15:45 AM To: Ultrasound Support <Ultrasound.Support@uhn.ca> Cc: Schwietz, Matthew - Sinai Health System <Matthew.Schwietz@sinaihealth.ca> Subject: [External] R59 issue Hi Scott, For R59 on 15th floor, the bottom charger is not working. Can someone please take a look at it? Thank you, Quyen. | MSH R59 QUYEN REPORTED LOWER BATTERY CHARGE SLOT NOT WORKING
1058532 | 209996 | CM | CMPLT | NFF | 2024-09-27 08:22 | 2024-09-27 08:23 | Not calibrating properly | (TC 09/27/2024 08:22:35 AM, Reid, Emmerson ) Tested with the phamton and it was working and reading consistently. Self test also passed. I am sending it back up with 3 SpO2 sets as well that I got working.
1061423 | 209997 | PM | CMPLT | CMPLT | 2024-09-27 09:48 | 2024-09-27 09:48 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.09 ohm, 141 uA
1061429 | 209998 | PM | CMPLT | CMPLT | 2024-09-27 09:49 | 2024-09-27 09:49 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.08 ohm, 131 uA
1061432 | 209999 | PM | CMPLT | CMPLT | 2024-09-27 09:51 | 2024-09-27 09:51 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.07 ohm, 140 uA
1061448 | 210000 | PM | CMPLT | CMPLT | 2024-09-27 09:52 | 2024-09-27 09:52 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.07 ohm, 129 uA
1061457 | 210001 | PM | CMPLT | CMPLT | 2024-09-27 09:53 | 2024-09-27 09:54 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.09 ohm, 142 uA
1061464 | 210002 | PM | CMPLT | CMPLT | 2024-09-27 09:55 | 2024-09-27 09:55 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.10 ohm, 138 uA
1061468 | 210003 | PM | CMPLT | CMPLT | 2024-09-27 09:56 | 2024-09-27 09:57 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.08 ohm, 130 uA
1061470 | 210004 | PM | CMPLT | CMPLT | 2024-09-27 09:57 | 2024-09-27 09:57 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.08 ohm, 143 uA
1061962 | 210005 | PM | CMPLT | CMPLT | 2024-09-27 09:58 | 2024-09-27 09:59 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.08 ohm, 143 uA
1061964 | 210006 | PM | CMPLT | CMPLT | 2024-09-27 10:00 | 2024-09-27 10:00 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.08 ohm, 130 uA
1061414 | 210007 | PM | CMPLT | CMPLT | 2024-09-27 10:46 | 2024-09-27 10:46 | PM-MAJOR-ANNUAL PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.003 ohm, 128 uA
1061415 | 210008 | PM | CMPLT | CMPLT | 2024-09-27 10:48 | 2024-09-27 10:48 | PM-MAJOR-ANNUAL PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.007 ohm, 134 uA
1061441 | 210009 | PM | CMPLT | CMPLT | 2024-09-27 10:50 | 2024-09-27 10:50 | PM-MAJOR-ANNUAL PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.07 ohm, 136 uA
1061965 | 210010 | PM | CMPLT | CMPLT | 2024-09-27 10:51 | 2024-09-27 10:52 | PM-MAJOR-ANNUAL PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.009 ohm, 136 uA
1062029 | 210011 | PM | CMPLT | CMPLT | 2024-09-27 10:53 | 2024-09-27 10:53 | PM-MAJOR-ANNUAL PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.02 ohm, 123 uA
1062259 | 210012 | PM | CMPLT | CMPLT | 2024-09-27 11:13 | 2024-09-27 11:14 | Batteries Replacement Due | Battery replaced by stryker
1062469 | 210013 | PM | CMPLT | CMPLT | 2024-09-27 11:33 | 2024-09-27 11:33 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.01 ohm, 123 uA
1061417 | 210014 | PM | CMPLT | CMPLT | 2024-09-27 11:34 | 2024-09-27 11:34 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.016 ohm, 127 uA
1061444 | 210015 | PM | CMPLT | CMPLT | 2024-09-27 11:35 | 2024-09-27 11:36 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.03 ohm, 74 uA
1061460 | 210016 | PM | CMPLT | CMPLT | 2024-09-27 11:37 | 2024-09-27 11:37 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.03 ohm, 119 uA
1061991 | 210017 | PM | CMPLT | CMPLT | 2024-09-27 11:38 | 2024-09-27 11:38 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.06 ohm, 130 uA
1062028 | 210018 | PM | CMPLT | CMPLT | 2024-09-27 11:40 | 2024-09-27 11:40 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.01 ohm, 125 uA
1061988 | 210019 | PM | CMPLT | CMPLT | 2024-09-27 11:41 | 2024-09-27 11:41 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.039 ohm, 119 uA
1061987 | 210020 | PM | CMPLT | CMPLT | 2024-09-27 11:42 | 2024-09-27 11:42 | PM-MAJOR-Annual PM For: Med Surg Bed | PM completed by Stryker. Battery changed and SW updated. 0.02 ohm, 102 uA
66713 | 210021 | CM | CMPLT | CMPLT | 2024-09-27 12:27 | 2024-09-27 12:27 | Scale not working. | (TC 09/27/2024 12:26:40 PM, Alvarado, Karla) Replaced scale wire harness. Replaced power cord. Performed functional test: Temperature = 35C and Humidity = 72.2% Electrical safety results - Leakage current = 86.5 uA; GND Resistance = 104 mohms All test passed. (TC 09/27/2024 12:27:05 PM, Rhuland, Courtney) Supervised
1062504 | 210022 | PM | CMPLT | CMPLT | 2024-09-27 12:43 | 2024-09-27 12:49 | PM-MAJOR_Annual PM For: Med Surg Bed | PM completed by Stryker.
1071886 | 210023 | CM | CMPLT | CMPLT | 2024-09-27 13:29 | 2024-09-27 13:34 | Repair | "not working" Brought it to biomed. FOund cart not turning ON. FOund power cord connector loose and not attached. Connected it and secured it. Charged batteries for a while. Checked functionalit.y Tested ok. Returned to 12N.
1065725 | 210024 | CM | CMPLT | NFF | 2024-09-27 14:02 | 2024-09-27 14:02 | Alarming battery going die. | (TC 09/27/2024 02:01:48 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Download error log: did not find errors. Electrical safety results - Leakage current = 88.5 uA; GND Resistance = 92 mohms All test passed. No Fault Found.
65899 | 210025 | CM | CMPLT | NFF | 2024-09-27 14:03 | 2024-09-27 14:03 | Alarming battery going die. Pumps not working – turning on – “Battery low” despite plugged in. | (TC 09/27/2024 02:03:12 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in. Download error log: did not find errors. All test passed. No Fault Found.
66035 | 210026 | CM | CMPLT | NFF | 2024-09-27 14:04 | 2024-09-27 14:04 | Alarming battery going die. | (TC 09/27/2024 02:04:20 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Download error log: did not find errors. Electrical safety results - Leakage current = 70.5 uA; GND Resistance = 200 mohms All test passed. No Fault Found.
1062178 | 210027 | PM | CMPLT | CMPLT | 2024-09-27 14:47 | 2024-09-27 14:47 | Batteries Replacement Due | battery replaced by stryker
1065556 | 210028 | CM | CMPLT | CMPLT | 2024-09-27 14:57 | 2024-09-27 14:57 | Temperature verification | (TC 09/27/2024 02:57:23 PM, Alvarado, Karla) Temperature tested with a probe cover: Low: set: 36.0, measured: 36.1 Medium: set: 38.5, measured: 38.4 High: set: 41, measured: 40.9 All tests passed.
1065557 | 210029 | CM | CMPLT | CMPLT | 2024-09-27 14:58 | 2024-09-27 14:58 | Temperature verification | (TC 09/27/2024 02:58:19 PM, Alvarado, Karla) Temperature tested without a probe cover: Low: set: 36, measured: 36.1 Medium: set: 38.5, measured: 38.5 High: set: 41, measured: 41 All tests passed.
502231 | 210030 | CM | CMPLT | CMPLT | 2024-09-27 17:51 | 2024-09-27 17:51 | Heat sensor ON problem needs to be fixed, and Thelas problem has to be looked. | Fixed.
| 210031 | CM | ACTIV | ISSUE | 2024-09-28 11:51 |  | MSH 5 US PRISM SSD B/U AFTER UPGRADE |  | MSH 5 US PRISM SSD B/U AFTER UPGRADE
63872 | 210032 | CM | CMPLT | CMPLT | 2024-09-30 08:36 | 2024-09-30 08:36 | Missing door (mini) | (TC 09/30/2024 08:36:01 AM, Alvarado, Karla) Incubator was sent down to biomed; missing door (mini). Replaced flip door. Performed PM, where VRef, Mains V, LV, and RH were calibrated. Tilt functions, all switches and pedals are functional. Error log was cleared. RH is functional and accurate; set 75%, analyzer = 77.8%. Set scale to 5kg, measured 5kg. Set temperature to 35C, measured 35C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 80.4 uA; GND Resistance = 127 mohms PM completed.
1071121 | 210033 | CM | CMPLT | NFF | 2024-09-30 08:55 | 2024-09-30 08:55 | Displaying “channel error”. Code: 240.4150 | (TC 09/30/2024 08:54:57 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Download error log: did not find errors. Electrical safety results - Leakage current = 66.8 uA; GND Resistance = 86 mohms All test passed. No Fault Found (NFF)
1047041 | 210034 | CM | CMPLT | CMPLT | 2024-09-30 09:13 | 2024-09-30 09:13 | Pump is running fluids faster than rate set on pump. | (TC 09/30/2024 09:13:08 AM, Alvarado, Karla) Downloaded error log, sensor broken. Performed analog sensor test and upper sensor failed. Replaced upper sensor. Performed tests in ASM: Analog Sensor and Check- in. All test passed.
| 210035 | RDS | CMPLT | CMPLT | 2024-09-30 09:34 | 2024-09-30 13:51 | Weekly checks in CCU | (TC 09/30/2024 01:51:06 PM, Scero, Tomlyn) weekly checks complete. 30 J tests passed.
| 210036 | CM | CMPLT | CMPLT | 2024-09-30 09:58 | 2024-10-07 09:26 | CONMED Training Sign Up | Sept 30 - Contacted vendor for Quote. Received. Requested clarification on items covered and itenerary. Received. Was provided with Registration link. Filled out form. Awaiting for confirmation. Oct 03 - Followed up. Confirmation received. Created SC. Sent for approval. Received. Issued PO vial MMC. Oct 07 - Followed up. Confirmation received from vendor.
1071201 | 210037 | CM | CMPLT | NFF | 2024-09-30 10:03 | 2024-09-30 10:04 | Pump is running fluids faster than rate set on pump. | (TC 09/30/2024 10:03:36 AM, Alvarado, Karla) Performed tests in ASM: Analog Sensor and Check- in. Downloaded error log, did not find any errors. All test passed.
1072034 | 210038 | ACC | CMPLT | CMPLT | 2024-09-30 10:24 | 2024-09-30 10:25 | Incoming Inspection | (TC 09/30/2024 10:24:58 AM, Reid, Emmerson ) Unboxed and labelled. Incoming Complete.
65166 | 210039 | CM | CMPLT | CMPLT | 2024-09-30 11:12 | 2024-09-30 11:12 | Casset was not attaching | (TC 09/30/2024 11:12:20 AM, Reid, Emmerson ) Tried connecting a casset and it worked. NFF
66433 | 210040 | CM | CMPLT | CMPLT | 2024-09-30 13:53 | 2024-09-30 14:21 | LL not working | (TC 09/30/2024 02:21:24 PM, Scero, Tomlyn) Replaced patient leads when I confirmed that the LL lead wasn't reading and giving an error. Tested with a simulator and working fine. Returned to CCU
1057302 | 210041 | CM | ACTIV | ISSUE | 2024-09-30 13:57 |  | MSH OPG SPP D7 GE VOLUSON WILL NOT TXFR STUDIES TO HSC INTERMITTENT | (TC 09/30/2024 01:57:32 PM, Whittington, Scott) MSH OPG SPP D7 GE VOLUSON WILL NOT TXFR STUDIES TO HSC INTERMITTENT (TC 10/01/2024 11:28:12 AM, Whittington, Scott) MSH OPG SPP D7 GE VOLUSON WILL NOT TXFR STUDIES TO HSC INTERMITTENT | MSH OPG SPP D7 GE VOLUSON WILL NOT TXFR STUDIES TO HSC INTERMITTENT
1072035 | 210042 | ACC | CMPLT | CMPLT | 2024-09-30 13:59 | 2024-09-30 14:25 | Incoming Inspection | New unit. Unboxed and Assembled unit. Assigned asset tag and recorded in TMS. Set up charger and Charged battery. Checked functionality. Tested ok.
65639 | 210043 | CM | CMPLT | CMPLT | 2024-09-30 14:03 | 2024-09-30 14:03 | nebulizer test failed during device check | nebulizer test failed during device check. replaced both Air and O2 pressure regulators. adjusted and tested. Functional. Passed device check 3 times.
69032 | 210044 | CM | CMPLT | CMPLT | 2024-09-30 15:07 | 2024-09-30 15:07 | Machine is showing error, need to be fixed. | (TC 09/30/2024 03:06:56 PM, Alvarado, Karla) Downloaded error log, sensor broken. Performed analog sensor test and upper pressure sensor failed. Replaced upper pressure sensor. Performed tests in ASM: Analog Sensor and Check- in. All test passed.
26111 | 210045 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26119 | 210046 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26131 | 210047 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
26230 | 210048 | PM | CANCL | CANCL | 2024-10-01 03:23 | 2024-10-10 15:06 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
27024 | 210049 | PM | CANCL | CANCL | 2024-10-01 03:23 | 2024-10-22 14:31 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
52109 | 210050 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
52428 | 210051 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Lasers, Nd:YAG, Ophthalmic
62485 | 210052 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | ANNUAL PM FOR: Electrosurgical Unit
62640 | 210053 | PM | CMPLT | CMPLT | 2024-10-01 03:23 | 2024-10-31 11:56 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 10/31/2024 11:56:04 AM, Baghirzade, Nara) Calibrated touch. Calibrated recorder. All functional tests passed. PM completed.
63019 | 210054 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR-Annual PM for:MR-Diagnostics Incubator
63301 | 210055 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
63835 | 210056 | PM | CMPLT | CMPLT | 2024-10-01 03:23 | 2024-10-21 13:21 | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile | (TC 09/20/2024 01:21:25 PM, Huang, Jessica) received the device from mother and baby PM completed
64076 | 210057 | PM | PENDG | REQST | 2024-10-01 03:23 |  | Replace battery
64077 | 210058 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM Major Annual - Power Systems, Uninterruptible
64086 | 210059 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM MAJOR - Power Systems, Uninterruptible
64089 | 210060 | PM | PENDG | REQST | 2024-10-01 03:23 |  | Replace battery
65126 | 210061 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65131 | 210062 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65155 | 210063 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65163 | 210064 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
65572 | 210065 | PM | CMPLT | CMPLT | 2024-10-01 03:23 | 2024-11-04 07:19 | PM-MAJOR - Annual PM For: YAG laser, gyne | (TC 11/04/2024 07:19:25 AM, Baghirzade, Nara) PM completed by Dornier. Attached service report. Facilitated the retrieval and drop off of the laser from the OR. (TC 11/04/2024 12:45:00 PM, Patel, Rutvik) Facilitated the retrieval and drop off of the laser from the OR. (TC 11/05/2024 03:53:23 PM, Patel, Rutvik) While transporting unit for PM yesterday, the casters on the cart came off and unit was dropped on the floor raising concenrs regarding the stability of the cart carrying the unit. Communicated to OR staff. Was advised to explore other options for the new cart. In the interim, brought the old cart down to biomed and fixed casters to cart legs using teflon tape and glue. (TC 11/06/2024 02:43:00 PM, Patel, Rutvik) Researched and explored new cart options for the LASER. (TC 11/07/2024 10:44:44 AM, Patel, Rutvik) Researched and explored new cart options for the LASER. Reached out to vendors to get pricing. Compiled all info and sent to OR staff for review.
65876 | 210066 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Syringe
66391 | 210067 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Air-oxygen Blender
66463 | 210068 | PM | CMPLT | CMPLT | 2024-10-01 03:23 | 2024-11-25 10:20 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check, 220 min, 47 Ah. Tested. Functional. SW.2.51 Leakage: 81 uA and 93 ohm.
66515 | 210069 | PM | CMPLT | CMPLT | 2024-10-01 03:23 | 2024-11-27 08:48 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 11/27/2024 08:47:31 AM, Baghirzade, Nara) PM completed. Recorder cleaned and calibrated. Touch calibrated. Functional testing passed; NIBP, SpO2, printer, and 4 transducer ports are functional.
66697 | 210070 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66698 | 210071 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66699 | 210072 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66702 | 210073 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66712 | 210074 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
66721 | 210075 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66724 | 210076 | PM | ACTIV | ISSUE | 2024-10-01 03:23 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66728 | 210077 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66757 | 210078 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66874 | 210079 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
67385 | 210080 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners
67764 | 210081 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67770 | 210082 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | ANNUAL PM - Catarac
67948 | 210083 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
68367 | 210084 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68547 | 210085 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68553 | 210086 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Anesthesia Units
68893 | 210087 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69298 | 210088 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69434 | 210089 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-21 15:18 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/21/2024 03:17:32 PM, Cortez, John) PM completed successfully, results logged in ASM Updated software
69539 | 210090 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
531100 | 210091 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X2
32540 | 210092 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-02 10:14 | PM - Monthly | PM completed Mechanical PM done TOS phantom a(-996.1,9.3) b(333.8,14.1) c(133.9,13.8) d(95.3,12.6) e(-104.2,12.8) f(0.8,14.1) -passed all phantom tests
32560 | 210093 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-26 09:38 | PM - Monthly | PM completed Mechanical PM done a (-996.7,9.4) b (333.7,14.2) c (133.6,13.2) d (95.3,13.8) e(-104.4,12.4) f(0.81,14.3) -passed all phantom tests
32580 | 210094 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-26 09:44 | PM - Monthly | PM completed Mechanical PM done Tos phantom a(-996.1,9.3) b(333.2,14.1) c(133.8,13.6) d(95.1,12.9) e(-104.8,12.1) f(0.8,14.6) -passed all phantom tests
501648 | 210095 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM - Annual
502088 | 210096 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-03 15:49 | PM - Semi-Annual | (TC 10/03/2024 03:48:43 PM, Green, Paul) PM3 performed as per procedure.
502089 | 210097 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM - Semi-Annual | (TC 10/31/2024 01:57:06 PM, Paz, Ramon (Tino)) Performed PM3 on the system a per OEM procedures. Noticed some tube oil on the floor. Attached PM report.
502231 | 210098 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-25 17:44 | PM - Semi-Annual | A complete PM waS DONE.
502231 | 210099 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | HARP - Semi Annual
1046453 | 210100 | PM | ACTIV | READY | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1048154 | 210101 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | ANNUAL PM - Microscope
1049878 | 210102 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-04 07:22 | PM-MAJOR - Annual PM For: YAG laser, gyne | (TC 11/04/2024 07:21:42 AM, Baghirzade, Nara) Facilitated the coordination of retrieval and drop off the laser in the OR.
1050238 | 210103 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
1051767 | 210104 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM MAJOR - ANNUAL PM for: Electrosurgical Unit
1053347 | 210105 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-11 17:32 | MSH BI 1270 HARP SEMI-ANNUAL | (TC 10/11/2024 01:12:07 PM, Dolinar, John) MSH BI 1270 HARP SEMI-ANNUAL & POST TUBE REPLACEMENT | MSH BI 1270 HARP SEMI-ANNUAL & POST TUBE REPLACEMENT
1053347 | 210106 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | MSH BI 1270 -PM SEMI-ANNUAL
1053379 | 210107 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR- Annual PM For : Ventilators, Intensive Care, Neonatal/Pediatric
1054108 | 210108 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
1054113 | 210109 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
1054114 | 210110 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
1054433 | 210111 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-26 12:44 | PM-MAJOR - Annual PM For: Aspirators, Emergency | (TC 11/26/2024 12:44:02 PM, Baghirzade, Nara) PM completed. Battery replaced. No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
1055788 | 210112 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1055789 | 210113 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1055790 | 210114 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1055792 | 210115 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1056055 | 210116 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | HARP - Semi Annual
1056055 | 210117 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-11 17:49 | PM - Semi Annual | pm WAS DONE.
1057506 | 210118 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-12 15:42 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/12/2024 03:42:23 PM, Canjimavadivel, Kamallesh ) Upgraded to latest rev N software and uploaded config. PM tests pass
1057513 | 210119 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057515 | 210120 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-25 10:52 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/25/2024 10:51:45 AM, Canjimavadivel, Kamallesh ) PM tests pass
1057516 | 210121 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-06 13:31 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/06/2024 01:31:27 PM, Canjimavadivel, Kamallesh ) PM tests pass
1057517 | 210122 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057518 | 210123 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-25 10:27 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/25/2024 10:26:31 AM, Canjimavadivel, Kamallesh ) Replaced battery. Upgraded to latest version of rev N. Replaced bezel. PM tests pass
1057519 | 210124 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-20 13:29 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/20/2024 01:28:42 PM, Canjimavadivel, Kamallesh ) PM tests pass. Updated software to rev N.00.06
1057520 | 210125 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057521 | 210126 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057523 | 210127 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-13 09:30 | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057524 | 210128 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-26 11:48 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/26/2024 11:47:41 AM, Canjimavadivel, Kamallesh ) Upgraded to latest version of rev N. Replaced battery. PM tests pass
1057525 | 210129 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057526 | 210130 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057527 | 210131 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057529 | 210132 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057530 | 210133 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-25 11:33 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/25/2024 11:33:16 AM, Canjimavadivel, Kamallesh ) PM tests pass
1057531 | 210134 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-11-20 15:11 | PM-MAJOR - Annual PM For: Portable Patient Modules X3 | (TC 11/20/2024 03:10:40 PM, Canjimavadivel, Kamallesh ) PM tests pass. Upgraded to rev N.00.06
1057533 | 210135 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Portable Patient Modules X3
1057571 | 210136 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057572 | 210137 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057574 | 210138 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057602 | 210139 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR ANNUAL PM for: Ventilators, Intensive Care, Adult
1057840 | 210140 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057844 | 210141 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-Major - ANNUAL PM for: Elecrosurgical Unit
1058196 | 210142 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1059541 | 210143 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059542 | 210144 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059543 | 210145 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059545 | 210146 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059546 | 210147 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059547 | 210148 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059548 | 210149 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059553 | 210150 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059554 | 210151 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059556 | 210152 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1059557 | 210153 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Central Monitor
1060518 | 210154 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | HARP - Semi Annual
1060518 | 210155 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-04 17:51 | PM - Semi Annual | PM was done.
1062099 | 210156 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-03 12:38 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 3rd year PM after 2nd year PM last year. No GS500. It has PS500. Run battery check, 200 min, 42 Ah. Tested. Functional. SW.2.51 Leakage: 90 uA and 97 ohm.
1062660 | 210157 | PM | CMPLT | CMPLT | 2024-10-01 03:24 | 2024-10-16 12:57 | PM-MAJOR - ANNUAL PM for: Insufflators | (TC 07/03/2024 01:07:05 PM, Huang, Jessica) sent the device for PM and replaced the conector installed it in the OR room
1064495 | 210158 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1064583 | 210159 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: DEFIBRILLATOR/MONITORS
1064666 | 210160 | PM | ACTIV | READY | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1065114 | 210161 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | Annual PM
1065115 | 210162 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | Annual PM
1065167 | 210163 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065168 | 210164 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065169 | 210165 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065170 | 210166 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065171 | 210167 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065173 | 210168 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065174 | 210169 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065176 | 210170 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Chairs, Examination/Treatment
1065226 | 210171 | PM | ACTIV | READY | 2024-10-01 03:24 |  | PM-MAJOR - Annual PM For: Fluid managment system. |  | Company HOLOGIC will need to do PM ( 1 800 442 9892 , Connor Blair ) .Check with Reeba for any contract.
1070151 | 210172 | PM | ACTIV | ISSUE | 2024-10-01 03:24 |  | Quaterly PM for: 3D Pritner
67915 | 210173 | CM | CMPLT | CMPLT | 2024-10-01 08:18 | 2024-10-01 08:18 | not heating. | (TC 10/01/2024 08:18:28 AM, Alvarado, Karla) Replaced corner grommet. Performed functional test. Tilt functions, all switches and pedals are functional. RH is functional and accurate; set 75%, analyzer = 75.3%. Set scale to 5kg, measured 5kg. Set temperature to 35C, measured 34.8C on all temperature points. Functional tests passed.
1055670 | 210174 | CM | CMPLT | CMPLT | 2024-10-01 08:26 | 2024-10-07 11:27 | Incident report #81526 investigation | (TC 10/01/2024 01:06:20 PM, Canjimavadivel, Kamallesh ) Turned on and verified battery health, within good health. Tested with biomed test blade, was not able to recreate issue. Looked into existing recalls related to device. Found correction action for Core 15 that is similar to problem experienced.. Reached out to vendor for further comment on if correction action is applicable to Core 10 devices as well. (TC 10/02/2024 10:26:30 AM, Canjimavadivel, Kamallesh ) Followed up with vendor, sales rep coming on site Friday to provide recommendation (TC 10/04/2024 11:26:44 AM, Canjimavadivel, Kamallesh ) Vendor came in to assess and provide recommendation (TC 10/07/2024 11:26:44 AM, Canjimavadivel, Kamallesh ) Completed investigation report and sent to Samir to attach to incident. Overall issue is suspected with blade. Emailed ED staff to update
1049104 | 210175 | CM | CMPLT | CMPLT | 2024-10-01 09:02 | 2024-10-01 09:03 | Needs Calibration | (TC 10/01/2024 09:03:18 AM, Reid, Emmerson ) Wendy reached out to me to have a look at the scale. She said it was needing calibration, and the height measuring pole was loose. I used locktight and secured the pole in place. I also checked the weight in 50lb increments up to 200lbs, At 200lbs it read 199.8lbs. No need for calibration. I brought it back up, leveled the scale, and let the unit know it was working.
66460 | 210176 | CM | CMPLT | CMPLT | 2024-10-01 10:43 | 2024-11-07 08:58 | ordering neo flow PCB | unit is giving neo flow error. ordering neo flow PCB as per Draeger technical service. Received PCB. PO#7202466404. Installed it and calibrated and tested the unit. Functional. Returned back to service. Informed RT.
1050480 | 210177 | CM | CMPLT | CMPLT | 2024-10-01 10:52 | 2024-10-01 10:52 | expired exp. caset. membrane | replaced expiratory casette membrane. reset timer. Tested. Functional.
| 210178 | CM | CMPLT | CMPLT | 2024-10-01 10:53 | 2024-10-01 10:53 | Check for MIB boards NICU | (TC 10/01/2024 10:53:16 AM, Rhuland, Courtney) Checked all rooms to verify that all MX800 monitors had MIB (RS232) boards installed for PIONIRS study Verified number of rooms with completed work orders All rooms had board and spare in biomed as well.
1065647 | 210179 | CM | CMPLT | CMPLT | 2024-10-01 11:01 | 2024-10-01 11:01 | Temp probe physically damaged | (TC 10/01/2024 11:00:59 AM, Cortez, John) Replaced probe
1049905 | 210180 | CM | CMPLT | CMPLT | 2024-10-01 11:21 | 2024-10-07 13:22 | MSH R65 GORAN REPORTED THAT COLLIMATOR BRAKE DID NOT HOLD WHEN ANGLE LEFT/RIGHT | (TC 10/01/2024 02:18:24 PM, Whittington, Scott) MSH R65 GORAN REPORTED THAT COLLIMATOR BRAKE DID NOT HOLD WHEN ANGLE LEFT/RIGHT INTERMITTENT,OBSERVED WHEN TUBE ARM EXTENDED OUT AND PUSHED BACK IN | MSH R65 GORAN REPORTED THAT COLLIMATOR BRAKE DID NOT HOLD WHEN ANGLE LEFT/RIGHT INTERMITTENT,OBSERVED WHEN TUBE ARM EXTENDED OUT AND PUSHED BACK IN
68014 | 210181 | CM | CMPLT | CMPLT | 2024-10-01 12:42 | 2024-10-01 12:42 | Not working | (TC 10/01/2024 12:41:50 PM, Alvarado, Karla) Replaced key board.
32560 | 210182 | CM | CMPLT | CMPLT | 2024-10-01 13:33 | 2024-09-30 13:32 | RDD errors reported | RDD errors reported Service replace RDD jbod unit CT phantom tested OK
| 210183 | CM | CMPLT | CMPLT | 2024-10-01 13:51 | 2024-10-01 13:52 | Sphygmometer installation North Pod | (TC 10/01/2024 01:51:33 PM, Rhuland, Courtney) Installed sphygmometers onto headwall in all procedure/exam rooms (TC 10/01/2024 01:52:00 PM, Alvarado, Karla) Installed sphygmometers onto headwall in all procedure/exam rooms
66723 | 210184 | CM | CMPLT | CMPLT | 2024-10-01 14:09 | 2024-10-01 14:09 | Reported not printing | (TC 10/01/2024 02:08:37 PM, Baghirzade, Nara) Room was very dark, unable to troubleshoot in the room. Removed the fetal monitor from the room. Found that the paper was incorrectly loaded. Cleaned printer head. Re-seated the paper. Printing is functional. Tested the transducer and SpO2; device is fully functional.
| 210185 | CM | CMPLT | CMPLT | 2024-10-01 14:10 | 2024-10-01 14:14 | US and Toco transducer checks | (TC 10/01/2024 02:10:29 PM, Baghirzade, Nara) Tested toco and US transducers; found 5 defective transducers. Will ship them out. Returned all functional transducers to Antenatal.
1056974 | 210186 | CM | ACTIV | ISSUE | 2024-10-01 14:23 |  | Unable to lower head, error 304 | (TC 10/01/2024 02:22:59 PM, Rhuland, Courtney) Bed controls for rail were locked Unlocked and able to control all controls from side rail Error code due to faulty foot switch Bed fully funtional Looking into quote for switch replacement
1068092 | 210187 | CM | CMPLT | CMPLT | 2024-10-01 14:23 | 2024-10-01 14:24 | Reported not working | (TC 10/01/2024 02:23:33 PM, Baghirzade, Nara) Tested transducer; no issue found. No physical damage. Fully functional.
69246 | 210188 | CM | CMPLT | CMPLT | 2024-10-01 14:38 | 2024-10-01 14:53 | Not working. Message “safety clamp is open”. | (TC 10/01/2024 02:38:07 PM, Alvarado, Karla) Downloaded error log, door sensor. Performed binary sensor test and door sensor failed. Replaced door latch. Performed tests in ASM: Check- in. All test passed. (TC 10/01/2024 02:52:49 PM, Alvarado, Karla) Downloaded error log, door sensor. Performed binary sensor test and door sensor failed. Replaced door latch. Performed tests in ASM: Rate calibration (twice) and Check- in. All test passed.
| 210189 | CM | CMPLT | CMPLT | 2024-10-01 14:39 | 2024-10-01 14:39 | test | test (TC 10/01/2024 02:39:15 PM, Chukri, Samir) test
530499 | 210190 | PM | CMPLT | CMPLT | 2024-10-01 14:44 | 2024-10-22 11:24 | PM Major - Delivery Module SEVO | (TC 10/01/2024 02:44:04 PM, Scero, Tomlyn) Pm parts replaced. Device tests to follow. (TC 10/22/2024 11:24:15 AM, Scero, Tomlyn) 6 year PM tests complete and passed. Device tested on AGM 68546 during PM.
67895 | 210191 | CM | ACTIV | ISSUE | 2024-10-02 07:26 |  | Shipping for repair | (TC 10/02/2024 07:40:42 AM, Baghirzade, Nara) SC 1003321374 (TC 10/25/2024 09:27:24 AM, Baghirzade, Nara) KJV # 461112
1064566 | 210192 | CM | CMPLT | CMPLT | 2024-10-02 08:04 | 2024-10-02 08:04 | Pad not working. | (TC 10/02/2024 08:03:58 AM, Alvarado, Karla) Pad was tested and came down. Retired pad.
1064566 | 210193 | PM | CMPLT | CMPLT | 2024-10-02 08:11 | 2024-10-02 08:12 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/02/2024 08:11:39 AM, Alvarado, Karla) Tests results: - Light Intensity small pad: 83 µW•cm-2•nm-1 - Hours: 1318.7 hours - Electrical safety results - Leakage current = 10.5 uA; GND Resistance = 488 mohms All tests passed. PM completed.
| 210194 | MEP | ACTIV | ISSUE | 2024-10-02 08:27 |  | MSH OBUS RFP OCTOBER 2024 | (TC 10/02/2024 08:28:20 AM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 OBTAINED ACCESS TO BONFIRE FOUND THAT THERE WERE NO MIF FORMS REQUESTED, NOR MDS2 FORMS REQUESTED ADVISED JANICE MOUNT MMC BUYER, TO SEND TO ALL VENDORS A REQUEST TO FILL OUT A MIF, MSH IT NAC FORM, MSH IT VENDOR FORM, AND TO PROVIDE MDS2 FOR THEIR PROPOSALS. AE TITLES TO BE USED: MSH_OBUS_MND MSH_OBUS_CAN MSH_OBUS_SAM MSH_OBUS_GE1 MSH_OBUS_GE2 MSH_OBUS_SMN MSH_OBUS_PHP (TC 10/03/2024 01:02:17 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 (TC 10/04/2024 01:03:08 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 SAMSUNG HERA W10 DELIVERED AND SETUP (TC 10/09/2024 03:33:45 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 SAMSUNG HERA W10 DELIVERED AND SETUP SAMSUNG SETUPP AND PROGRAMMED INVOGP SPP D1 (TC 10/11/2024 01:17:27 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 GE 20 ACC (TC 10/11/2024 01:17:38 PM, Dolinar, John) MSH OBUS RFP OCTOBER 2024 GE 20 ACC (TC 10/15/2024 12:11:00 PM, Whittington, Scott) MSH OB US RFP DEMO OCTOBER 2024 GE 22 ACC (TC 10/18/2024 03:00:21 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 PHILIPS EPIQ (TC 10/21/2024 03:01:10 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 PHILIPS EPIQ (TC 10/24/2024 04:23:19 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 BOMIMED MINDRAY RESONA i9 DELIVERED. COMMENCE SETUP (TC 10/24/2024 04:24:58 PM, Dolinar, John) MSH OBUS RFP OCTOBER 2024 BOMIMED MINDRAY RESONA i9 DELIVERED. COMMENCE SETUP (TC 10/24/2024 04:24:58 PM, Dolinar, John) MSH OBUS RFP OCTOBER 2024 BOMIMED MINDRAY RESONA i9 DELIVERED. COMMENCE SETUP (TC 10/25/2024 04:58:00 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 MINDRAY i9 ACCEPTANCE AND SETUP (TC 10/28/2024 04:07:29 PM, Whittington, Scott) MSH OBUS RFP PHP CEOU TEST IMAGES (TC 10/28/2024 04:07:29 PM, Whittington, Scott) MSH OBUS RFP PHP CEOU TEST IMAGES (TC 10/30/2024 05:49:45 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024mindray 19 acc ACC TO CEOU (TC 10/30/2024 05:51:01 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 SIEMENS SEQUOIA DELIVERY &SETUP (TC 10/31/2024 05:52:04 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 SIEMENS SEQUOIA DELIVERY & SETUP (TC 11/04/2024 10:19:11 AM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 SIEMENS ON-SITE BRIAN: MENKE: RICHARD NEWMAN (TC 11/05/2024 03:18:31 PM, Dolinar, John) MSH OBUS RFP OCTOBER 2024 SIEMENS ON-SITE BRIAN: MENKE: RICHARD NEWMAN (TC 11/08/2024 04:02:20 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 GE VOLUSON ACCEPTANCE (TC 11/08/2024 04:05:13 PM, Chau, Ken) MSH OBUS RFP OCTOBER 2024 GE VOLUSON ACCEPTANCE (TC 11/08/2024 04:08:27 PM, Fan, Tak) MSH OBUS RFP OCTOBER 2024 GE VOLUSON ACCEPTANCE (TC 11/11/2024 01:18:40 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 GE 22 ACC (TC 11/11/2024 01:19:19 PM, Dolinar, John) MSH OBUS RFP OCTOBER 2024 GE 22 ACC (TC 11/11/2024 01:21:16 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 samsung ACC (TC 11/12/2024 03:03:57 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 samsung ACC (TC 11/15/2024 02:07:00 PM, Whittington, Scott) MSH OBS RFP OCTOBER 2024 CANON ACCEPTANCE GE & SAMSUNG DEPARTURE (TC 11/19/2024 04:51:40 PM, Whittington, Scott) MSH OBUS RFP OCTOBER 2024 CANON 3D PROBE DELIVERY | MSH OBUS RFP OCTOBER 2024 AE TITLES TO BE USED: MSH_OBUS_MND MSH_OBUS_CAN MSH_OBUS_SAM MSH_OBUS_GE1 MSH_OBUS_GE2 MSH_OBUS_SMN MSH_OBUS_PHP
| 210195 | RDS | CMPLT | CMPLT | 2024-10-02 09:55 | 2024-10-02 10:25 | Weekly check for ED
67242 | 210196 | ACC | CMPLT | CMPLT | 2024-10-02 10:22 | 2024-10-02 10:23 | Incoming inspection | (TC 10/02/2024 10:22:32 AM, Baghirzade, Nara) Received replacement pump. Incoming inspection completed. Updated SN on TMS. Attached power cord and mount. Functional tests passed. Attached performance certification.
1054693 | 210197 | ACC | CMPLT | CMPLT | 2024-10-02 10:23 | 2024-10-02 10:27 | Incoming inspection | (TC 10/02/2024 10:27:24 AM, Baghirzade, Nara) Received replacement pump. Incoming inspection completed. Updated SN on TMS. Attached power cord and mount. Functional tests passed. (TC 10/02/2024 10:27:24 AM, Baghirzade, Nara) Received replacement pump. Incoming inspection completed. Updated SN on TMS. Attached power cord and mount. Functional tests passed. Attached performance certification.
1054696 | 210198 | ACC | CMPLT | CMPLT | 2024-10-02 10:31 | 2024-10-02 10:31 | Incoming inspection | (TC 10/02/2024 10:31:17 AM, Baghirzade, Nara) Received replacement pump. Incoming inspection completed. Updated SN on TMS. Attached power cord and mount. Functional tests passed. Attached performance certification.
1054669 | 210199 | ACC | CMPLT | CMPLT | 2024-10-02 10:48 | 2024-10-02 10:48 | Incoming inspection | (TC 10/02/2024 10:48:18 AM, Baghirzade, Nara) Received replacement pump. Incoming inspection completed. Updated SN on TMS. Attached power cord and mount. Functional tests passed. Attached performance certification.
63437 | 210200 | CM | CMPLT | CMPLT | 2024-10-02 10:50 | 2024-10-02 10:50 | Broken mount | (TC 10/02/2024 10:49:36 AM, Baghirzade, Nara) Replaced mount. Functional tests passed.
1056503 | 210201 | CM | CMPLT | CMPLT | 2024-10-02 12:12 | 2024-10-02 12:13 | Battery replacement | (TC 10/02/2024 12:12:30 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056445 | 210202 | FC | CMPLT | CMPLT | 2024-10-02 12:13 | 2024-10-02 12:13 | Functional testing | (TC 10/02/2024 12:13:03 PM, Baghirzade, Nara) Replaced SpO2 finger sensor. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1069644 | 210203 | PM | CMPLT | CMPLT | 2024-10-02 12:41 | 2024-10-02 12:41 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/02/2024 12:41:19 PM, Alvarado, Karla) Tests results: - Light Intensity small pad: 76.9 µW•cm-2•nm-1 - Hours: 353.5 hours - Electrical safety results - Leakage current = 10.8 uA; GND Resistance = 168 mohms All tests passed. PM completed.
1069776 | 210204 | CM | CMPLT | CMPLT | 2024-10-02 13:30 | 2024-10-02 13:30 | Not working | (TC 10/02/2024 01:30:26 PM, Canjimavadivel, Kamallesh ) Probe tip had what appeared to be earwax on it. Cleaned tip and verified calibration. Passed after 2 attempts
| 210205 | CM | CMPLT | CMPLT | 2024-10-02 14:24 | 2024-10-02 09:22 | Laser safety meeting | TC 10/02/2024 09:21:56 AM, Baghirzade, Nara) Laser safety meeting with Nicholas Brouse, Adeel, and Jessica
1049047 | 210206 | CM | ACTIV | ISSUE | 2024-10-02 14:36 |  | warranty exchange
502239 | 210207 | CM | CMPLT | CMPLT | 2024-10-02 14:43 | 2024-10-04 12:59 | MSH ER POSUC SONOSITE XPORTE NO POWER | (TC 10/02/2024 02:43:31 PM, Whittington, Scott) MSH ER POSUC SONOSITE XPORTE NO POWER (TC 10/03/2024 02:43:56 PM, Whittington, Scott) MSH ER POCUS XPORTE NO POWER REPLACED POWER CORD, ALLOWED TO CHARGE UP RETURNED TO ER | MSH ER POSUC SONOSITE XPORTE NO POWER
32520 | 210208 | CM | CMPLT | CMPLT | 2024-10-02 14:53 | 2024-10-07 13:21 | MSH DR2 COLLIMATOR LOSING POWER/CONNECTION WHEN SWIVELLED | (TC 10/02/2024 02:53:03 PM, Whittington, Scott) MSH DR2 COLLIMATOR LOSING POWER/CONNECTION WHEN SWIVELLED (TC 10/07/2024 01:21:25 PM, Whittington, Scott) MSH DR2 COLLIMATOR LOSING POWER/CONNECTION WHEN SWIVELLED REPAIR CABLING | MSH DR2 COLLIMATOR LOSING POWER/CONNECTION WHEN SWIVELLED
1072049 | 210209 | ACC | CMPLT | CMPLT | 2024-10-02 15:04 | 2024-10-02 15:04 | Incoming Inspection | (TC 10/02/2024 03:04:22 PM, Reid, Emmerson ) Tagged and setup device, incoming complete.
1072050 | 210210 | ACC | ACTIV | ISSUE | 2024-10-02 15:12 |  | Incoming Inspection
1072051 | 210211 | ACC | ACTIV | ISSUE | 2024-10-02 15:26 |  | Incoming Inspection
1056580 | 210212 | CM | CMPLT | CMPLT | 2024-10-03 07:57 | 2024-10-03 07:58 | Battery replacement | (TC 10/03/2024 07:57:58 AM, Rhuland, Courtney) Replaced dead battery in AED3 Sent data over wifi to dashboard
1049922 | 210213 | CM | CMPLT | CMPLT | 2024-10-03 08:00 | 2024-10-03 08:03 | Configuration changes | (TC 10/03/2024 08:00:29 AM, Rhuland, Courtney) Nurse requested touch screen calibration and have only sys/dia pressure be displayed Power cycled and put into service mode Changed config so only sys/dia pressure is displayed Calibrated touch screen
1071913 | 210214 | CM | CMPLT | CMPLT | 2024-10-03 08:18 | 2024-10-21 14:23 | No tracing and no cursors in the acquisition window | (TC 10/02/2024 08:18:36 AM, Baghirzade, Nara) During patient test, the doctor was able to stimulate however there was no tracing. The device had to be rebooted several times to get the tracing to work. Cursor also disappeared and doctor was not able to move between traces. Swapped the base unit with the old base unit, ran the HW programmer. Tested the amplifier and stimulus with Vinod; EMG is functional. Will troubleshoot with Natus. (TC 10/21/2024 02:23:29 PM, Baghirzade, Nara) Received replacement base unit. Swapped base units. Ran HW programmer and diagnostics successfully. Vinod tested NCS and EMG; no issues found. Will send defective base unit back to Natus.
62698 | 210215 | CM | CMPLT | CMPLT | 2024-10-03 08:47 | 2024-10-02 08:48 | multiple leads need replacement | (TC 10/02/2024 08:48:21 AM, Cortez, John) ECG tech requested 5x replacement leads, and keyboard cover
| 210216 | CM | CMPLT | CMPLT | 2024-10-03 09:07 | 2024-11-06 14:26 | Order for Prismax PM kits, PM tools | (TC 10/01/2024 10:35:35 AM, Cortez, John) Reached out to Baxter for part pricing (TC 10/03/2024 10:37:02 AM, Cortez, John) Placed order (TC 10/29/2024 09:43:21 AM, Cortez, John) Tools arrived, PM kits still incoming (TC 10/31/2024 01:37:55 PM, Cortez, John) 3x pm kits arrived (TC 11/04/2024 02:59:49 PM, Cortez, John) 3x PM kits arrived (TC 11/06/2024 02:25:59 PM, Cortez, John) x2 kits arrived
1061455 | 210217 | CM | ACTIV | ISSUE | 2024-10-03 09:09 |  | Limited motion functionality error; head portion does not move | (TC 10/03/2024 09:08:50 AM, Baghirzade, Nara) Limited motion functionality error; head portion of the bed is not moving. Power cycled the bed however it did not resolve the issue. All other motion controls are working except for the head portion. Created service request with Stryker.
1062022 | 210218 | CM | CMPLT | CMPLT | 2024-10-03 09:13 | 2024-10-03 09:13 | Damaged AUX power cord | (TC 10/03/2024 09:13:18 AM, Baghirzade, Nara) Replaced AUX power cord. Main power cord in good condition. Confirmed that there is power delivered to the aux power outlets. Checked all motion controls including both side panels and patient panels. Bed is functional.
66731 | 210219 | FC | CMPLT | CMPLT | 2024-10-03 09:17 | 2024-10-03 09:20 | Power cord and functional tests | (TC 10/03/2024 09:19:33 AM, Baghirzade, Nara) Replaced power cord. Functional testing passed; NIBP, SpO2, printer, and 4 transducer ports are functional.
62640 | 210220 | CM | CMPLT | CMPLT | 2024-10-03 09:24 | 2024-10-03 12:50 | Screen is glitching | (TC 10/03/2024 12:49:31 PM, Baghirzade, Nara) Screen was glitching and was intermittently turning on and off. Disassembled display and re-seated the ribbon cable. Display is no longer glitching, tested the display in different positions and no issues with it. Functional tests passed.
1063746 | 210221 | PM | CMPLT | CMPLT | 2024-10-03 10:02 | 2024-10-03 10:02 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/03/2024 10:01:58 AM, Alvarado, Karla) Drained, disinfected, rinsed and refilled with sterile water (TC 10/03/2024 10:01:58 AM, Rhuland, Courtney) Supervised
1049050 | 210222 | CM | ACTIV | ISSUE | 2024-10-03 10:47 |  | warranty exchange
1054118 | 210223 | CM | CMPLT | CMPLT | 2024-10-03 11:14 | 2024-10-24 08:18 | Sending for repair | (TC 10/03/2024 11:30:33 AM, Baghirzade, Nara) SC 1003322719 (TC 10/03/2024 12:51:19 PM, Baghirzade, Nara) Sent with KJV. KJV# 457452 (TC 10/24/2024 08:17:31 AM, Baghirzade, Nara) PO 7500212772 (TC 10/24/2024 08:17:51 AM, Baghirzade, Nara) Received unit from spectramedx. Unit is too damaged and not repaired. Will be retired. Attached service report.
1047986 | 210224 | CM | CMPLT | CMPLT | 2024-10-03 11:16 | 2024-10-24 08:21 | Sending for repair | (TC 10/03/2024 11:31:22 AM, Baghirzade, Nara) SC 1003322719 (TC 10/03/2024 12:52:09 PM, Baghirzade, Nara) Sent with KJV. KJV# 457452 (TC 10/22/2024 08:39:57 AM, Baghirzade, Nara) Updating cost after repair assessment. (TC 10/24/2024 08:20:35 AM, Baghirzade, Nara) PO 7500212772 Received transducer back from spectramedx. Transducer is functional. Attached service report.
1072052 | 210225 | CM | CMPLT | CMPLT | 2024-10-03 11:17 | 2024-10-24 08:26 | Sending for repair | (TC 10/03/2024 11:32:00 AM, Baghirzade, Nara) SC 1003322719 (TC 10/03/2024 12:52:28 PM, Baghirzade, Nara) Sent with KJV. KJV# 457452 (TC 10/22/2024 08:40:41 AM, Baghirzade, Nara) Updating cost after repair assessment. (TC 10/24/2024 08:26:16 AM, Baghirzade, Nara) PO 7500212772 Received transducer from spectramedx. Transducer is functional. Attached service report.
1072053 | 210226 | CM | CMPLT | CMPLT | 2024-10-03 11:17 | 2024-10-24 08:30 | Sending for repair | (TC 10/03/2024 11:32:37 AM, Baghirzade, Nara) SC 1003322719 (TC 10/03/2024 12:52:53 PM, Baghirzade, Nara) Sent with KJV. KJV# 457452 (TC 10/22/2024 08:40:54 AM, Baghirzade, Nara) Updating cost after repair assessment. (TC 10/24/2024 08:29:32 AM, Baghirzade, Nara) PO 7500212772 Received transducer from spectramedx. Transducer is functional. Attached service report.
1061333 | 210227 | FC | CMPLT | CMPLT | 2024-10-03 13:38 | 2024-10-03 13:38 | QC checklist | (TC 10/03/2024 01:38:24 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
62704 | 210228 | CM | CMPLT | CMPLT | 2024-10-03 13:45 | 2024-10-03 13:45 | damaged keyboard cover | replaced keyboard cover
64258 | 210229 | CM | CMPLT | CMPLT | 2024-10-03 14:43 | 2024-10-03 14:44 | !!Check equipment alarm | (TC 10/03/2024 02:43:32 PM, Rhuland, Courtney) X2 was from a transport so equipment label was not compatible Had nurses swap out X2 with NICU label Alarm cleared - brought transport X2 down to biomed
1064507 | 210230 | CM | CMPLT | CMPLT | 2024-10-03 15:36 | 2024-10-03 15:36 | Software update | (TC 10/03/2024 03:36:23 PM, Cortez, John) Software update to 1.10
1064509 | 210231 | CM | CMPLT | CMPLT | 2024-10-03 15:38 | 2024-10-03 15:38 | Software update | (TC 10/03/2024 03:38:05 PM, Cortez, John) Software update to 1.10
1054321 | 210232 | CM | CMPLT | CMPLT | 2024-10-04 07:16 | 2024-10-08 15:27 | "not lighting up" | (TC 10/01/2024 07:15:42 AM, Cortez, John) Both heads ok. Seems to be issue with handles. One confirmed defective, the other worked intermittently during testing Reached out to Amico. They will send warranty replacement (TC 10/08/2024 03:27:05 PM, Cortez, John) Diagnostic station arrived, Replaced both handles, kept diagnostic station Shipped defective handles back
1064693 | 210233 | CM | CMPLT | CMPLT | 2024-10-04 09:39 | 2024-10-04 09:39 | Moved to Traveling Cysto Cart | (TC 10/04/2024 09:39:19 AM, Reid, Emmerson ) Was possibly having issues, moved from Boom to cart to see.
1064750 | 210234 | CM | CMPLT | CMPLT | 2024-10-04 09:42 | 2024-10-04 09:43 | Moved from Cysto Cart to OR 13 | (TC 10/04/2024 09:42:38 AM, Reid, Emmerson ) See WO 1061214
1061213 | 210235 | CM | ACTIV | ISSUE | 2024-10-04 10:00 | 2024-10-09 13:24 | Ordering DVI boom cover | (TC 10/04/2024 10:03:31 AM, Reid, Emmerson ) Placed an order for the part (TC 10/09/2024 01:23:20 PM, Patel, Rutvik) Installed switch. Rebooted VMC. Checked functionality. Tested ok.
65921 | 210236 | CM | CMPLT | CMPLT | 2024-10-04 10:17 | 2024-10-04 10:17 | Received from SickKids | (TC 10/04/2024 10:17:03 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. All tests passed.
1062779 | 210237 | CM | ACTIV | ISSUE | 2024-10-04 10:30 |  | Ordering ECG leads | (TC 10/04/2024 10:37:47 AM, Baghirzade, Nara) SC 1003323499
1072074 | 210238 | CM | CMPLT | CMPLT | 2024-10-04 11:54 | 2024-10-04 11:54 | Incoming inspection. | (TC 10/04/2024 11:54:02 AM, Alvarado, Karla) Moved unit from 6th floor storage room to 4th floor. Tested weigh. Tests passed. Assigned asset tag.
32560 | 210239 | CM | CMPLT | CMPLT | 2024-10-04 12:38 | 2024-10-01 12:38 | Unusual noise coming from gantry during rotation | Unusual noise coming from gantry during rotation Noise is more pronounced at higher rotational speed Troubleshooting done and found a loose scre touching the top chasis Tightened the screw and checked all other screws for tightness Returned system for use.
1047469 | 210240 | PM | CMPLT | CMPLT | 2024-10-04 12:42 | 2024-10-04 12:42 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 1st year PM after 6th year PM last year. No GS500. It has PS500. Run battery check, 220 min, 68 Ah. Tested. Functional. SW.2.51 Leakage: 87 uA and 93 ohm.
1058998 | 210241 | CM | ACTIV | ISSUE | 2024-10-04 12:59 |  | no heat
| 210242 | CM | CMPLT | CMPLT | 2024-10-04 14:51 | 2024-10-21 09:52 | Newark Parts order | Contacted for part# and pricing. Created SC. Sent for approval. Placed order via MMC. Parts received. Sorted and organized in shelves.
| 210243 | OC | CMPLT | CMPLT | 2024-10-07 08:08 | 2024-10-07 08:08 | replaced the wall regulator in 14 south | (TC 10/05/2024 08:08:11 AM, Huang, Jessica) received the call from 14 south on October 5 at 3:38 pm regarding the wall regulator. The clincial team was not able to replace it and the patient could not be moved to the different room. arrived at the hospital around 5 pm and replaced the regulator. Tested it with the nurse. Left the hospital 5:30 pm. (TC 10/05/2024 08:08:11 AM, Huang, Jessica) received the call from 14 south on October 5 at 3:38 pm regarding the wall regulator. The clincial team was not able to replace it and the patient could not be moved to the different room. arrived at the hospital around 5 pm and replaced the regulator. Tested it with the nurse. Left the hospital 5:30 pm.
1054522 | 210244 | CM | CMPLT | CMPLT | 2024-10-07 09:34 | 2024-10-07 09:34 | V4 lead not working | (TC 10/07/2024 09:33:40 AM, Canjimavadivel, Kamallesh ) Replaced V4-V6 chest lead set. Also resoldered multiple pins on PIM that were loose.
1062603 | 210245 | CM | ACTIV | ISSUE | 2024-10-07 10:12 |  | warranty replacement
1057279 | 210246 | CM | CMPLT | CMPLT | 2024-10-07 10:16 | 2024-10-23 10:11 | error message and needs to be sent for repair | (TC 10/23/2024 10:11:12 AM, Baghirzade, Nara) PO 7500212905
| 210247 | CM | CMPLT | CMPLT | 2024-10-07 10:40 | 2024-10-09 14:18 | Inventory Check | (TC 10/07/2024 10:40:01 AM, Rhuland, Courtney) Checked rooms on 3rd and 8th floor for equipment and collected SN for devices that were not tagged. (TC 10/09/2024 02:18:11 PM, Rhuland, Courtney) Asset tagged new exam tables on 8th floor
1054976 | 210248 | CM | CMPLT | CMPLT | 2024-10-07 10:40 | 2024-10-07 10:42 | Missing plug. | (TC 10/07/2024 10:40:30 AM, Alvarado, Karla) Replaced power cord. Performed tests: Temperature NIBP = set: 120/80, measured: 117/84 SpO2 = set: 95%, measured: 95% All test passed. (TC 10/07/2024 10:42:24 AM, Huang, Jessica) Supervised
1060002 | 210249 | CM | CMPLT | CMPLT | 2024-10-07 10:49 | 2024-10-07 10:49 | LL lead damaged | (TC 10/07/2024 10:49:04 AM, Canjimavadivel, Kamallesh ) Replaced LL lead and V1-V3 chest lead. Tested ok with simulator
| 210250 | PA | CMPLT | CMPLT | 2024-10-07 10:59 | 2024-10-17 10:33 | TC70 leads for stock | (TC 10/07/2024 11:08:07 AM, Canjimavadivel, Kamallesh ) Placed order for leads (TC 10/17/2024 10:33:02 AM, Canjimavadivel, Kamallesh ) Received some leads, added to stock
66008 | 210251 | CM | CMPLT | CMPLT | 2024-10-07 11:23 | 2024-10-07 11:24 | Plug broken. | (TC 10/07/2024 11:23:48 AM, Alvarado, Karla) Replaced power cord. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage = 93.0uA and the GND Resistance = 0.101 Ohms All tests passed.
65942 | 210252 | CM | CMPLT | CMPLT | 2024-10-07 11:42 | 2024-10-08 08:17 | It is not keeping consistent pressure | (TC 10/07/2024 10:00:00 AM, Alvarado, Karla) Performed tests in ASM: Download error log and Check- in tests. Download error log: pressure sensor failure. Replaced pressure sensor and gasket. (TC 10/08/2024 08:17:06 AM, Alvarado, Karla) Performed tests in ASM: Calibration and Check- in tests. All tests passed.
1072085 | 210253 | CM | ACTIV | ISSUE | 2024-10-07 13:00 |  | shipped the monitor for repair
32520 | 210254 | CM | ACTIV | ISSUE | 2024-10-07 13:24 |  | MSH DR2 WOULD NOT EXPOSE; FD PROTOCOL MPRS TEST FAIL; OFFSET MPRSTEST FAIL | (TC 10/07/2024 01:24:26 PM, Whittington, Scott) MSH DR2 WOULD NOT EXPOSE; FD PROTOCOL MPRS TEST FAIL; OFFSET MPRSTEST FAIL (TC 10/11/2024 07:33:14 AM, Whittington, Scott) MSH DR2 WOULD NOT EXPOSE; FD PROTOCOL MPRS TEST FAIL; OFFSET MPRSTEST FAIL (TC 10/11/2024 07:33:14 AM, Whittington, Scott) MSH DR2 WOULD NOT EXPOSE; FD PROTOCOL MPRS TEST FAIL; OFFSET MPRSTEST FAIL (TC 10/11/2024 07:33:14 AM, Whittington, Scott) MSH DR2 WOULD NOT EXPOSE; FD PROTOCOL MPRS TEST FAIL; OFFSET MPRSTEST FAIL | MSH DR2 WOULD NOT EXPOSE; FD PROTOCOL MPRS TEST FAIL; OFFSET MPRSTEST FAIL
1057384 | 210255 | CM | ACTIV | ISSUE | 2024-10-07 14:07 |  | MSH CR850 BLUE SCREEN K4187-3354 | (TC 10/08/2024 03:28:10 PM, Whittington, Scott) MSH CR850 BLUE SCREEN K4187-3354 REF# MAURICE 1048574 SEAN ORDERING HP5700 PC (TC 10/11/2024 01:20:02 PM, Whittington, Scott) MSH CR850 BLUE SCREEN K4187-3354 REF# MAURICE 1048574 SEAN ON SITE WITH HP5700 PC | MSH CR850 BLUE SCREEN K4187-3354 REF# MAURICE 1048574
1066025 | 210256 | CM | CMPLT | CMPLT | 2024-10-07 14:24 | 2024-10-07 14:24 | one programmed it, it went to error/warning message. | (TC 10/07/2024 02:24:22 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in. Download error log: keyboard failure. Keyboard replaced. Performed tests in ASM: Check- in. Tests passed.
| 210257 | CM | CMPLT | CMPLT | 2024-10-07 14:26 | 2024-10-07 14:26 | Replaced power bar on code blue crash cart | (TC 10/07/2024 02:25:59 PM, Canjimavadivel, Kamallesh ) Ground pin on power bar was damaged. Ordered replacement power bar and swapped out.
| 210258 | RDS | CMPLT | CMPLT | 2024-10-07 14:31 | 2024-10-08 12:54 | Weekly check for ED
1062180 | 210259 | PM | CMPLT | CMPLT | 2024-10-07 14:58 | 2024-10-07 14:58 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 10/07/2024 02:58:11 PM, Alvarado, Karla) 1207.0 hours Intensity light: - High: 32.5 uW/cm2/nm - Low: 14.1 uW/cm2/nm Target light is functional GND Resistance: 0.180 ohms Leakage current: 48.2 uA All test passed. PM completed.
502089 | 210260 | CM | ACTIV | ISSUE | 2024-10-07 15:18 |  | error | (TC 10/07/2024 03:18:08 PM, Green, Paul) Found a persistant error occurred. Was able to clear it and home the system.
1053347 | 210261 | CM | CMPLT | CMPLT | 2024-10-08 08:21 | 2024-10-17 13:57 | MSH BI 1270 H1 TUBE ARC RYAN REPLACING TUBE AND GENERATOR DRAWER | (TC 10/08/2024 03:18:38 PM, Whittington, Scott) MSH BI 1270 H1 TUBE ARC RYAN REPLACING TUBE AND GENERATOR DRAWER (TC 10/08/2024 03:18:47 PM, Chau, Ken) MSH BI 1270 H1 TUBE ARC RYAN REPLACING TUBE AND GENERATOR DRAWER (TC 10/09/2024 09:40:41 AM, Chau, Ken) Replaced workstation video card System operating without issues Informed Radiologist and left my number to call me if system still shuts down Old video card was not Barco, over-heat temperature Radiologist informed monitor was off again yesterday and today Helped Noor with tube alignment Noor to finish AEC and reassemble system by end of today (TC 10/17/2024 01:55:26 PM, Whittington, Scott) MSH BI 1270 H1 TUBE ARC RYAN REPLACING TUBE AND GENERATOR DRAWER VERIFYING NOORS PIRHANA DOSIMETER READING WITH ACCUGOLD RADCAL (TC 10/17/2024 01:56:50 PM, Dolinar, John) MSH BI 1270 H1 TUBE ARC RYAN REPLACING TUBE AND GENERATOR DRAWER VERIFYING NOORS PIRHANA DOSIMETER READING WITH ACCUGOLD RADCAL | MSH BI 1270 H1 TUBE ARC RYAN REPLACING TUBE AND GENERATOR DRAWER Tube replaced.
| 210262 | CON | ACTIV | ISSUE | 2024-10-08 08:24 |  | MSH BI TMIST TRIAD | (TC 10/08/2024 03:20:09 PM, Whittington, Scott) MSH BI TMIST TRIAD TRIAD BY IROC ACR | MSH BI TMIST TRIAD From: Kostyrka, Andriy <Andriy.Kostyrka@sinaihealth.ca> Sent: October 8, 2024 8:14 AM To: Whittington, Scott <Scott.Whittington@uhn.ca> Cc: Sousa, Rosemary - Sinai Health System <Rosemary.Sousa@sinaihealth.ca> Subject: [External] RE: Triad Hi Scott, I believe you were originally setting up Triad on MSH PCs. After software update users not able to send images anymore. Can you help us to resolve the issue. Sincerely, Andriy Kostyrka Desktop / Network Analyst T 416-586-4800 x 2040 From: Sousa, Rosemary Sent: October 03, 2024 2:49 PM To: Kostyrka, Andriy <Andriy.Kostyrka@sinaihealth.ca> Subject: Triad Importance: High Hi Andrew, Triad is looking to obtain some technical support from the Sinai IT team. If you can be available – I can send you a zoom meeting invitation for you to join in. Let me know.. Thanks, Rosemary Sousa TMIST Research Assistant I Phone: 416-596-4200 x 4406 Email: rosemary.sousa@sinaihealth.ca 600 University Avenue, Toronto, ON M5G 1X5 A patient care, teaching and research centre affiliated with University of Toronto.
67950 | 210263 | CM | CMPLT | CMPLT | 2024-10-08 09:32 | 2024-10-08 09:39 | latch broken. | (TC 10/08/2024 09:33:19 AM, Alvarado, Karla) Replaced side door latch-south and inside wall latch. Replaced power cord. Performance functional test. RH set 75%, analyzer = 74.5%. Set scale to 5kg, measured 5kg. Set temperature to 35C, measured 35C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 86.8 uA; GND Resistance = 88 mohms (TC 10/08/2024 09:38:38 AM, Chukri, Samir) Supervised
1055908 | 210264 | CM | CMPLT | NFF | 2024-10-08 09:41 | 2024-10-08 09:44 | No accurate reading | (TC 10/08/2024 09:43:45 AM, Alvarado, Karla) Performed tests: Temperature NIBP= set: 120/80, measured: 116/84 SpO2= set: 95%, measured: 95% All test passed. No Fault Found (NFF)
1054987 | 210265 | CM | ACTIV | ISSUE | 2024-10-08 09:47 |  | Not working properly. | (TC 10/08/2024 09:46:46 AM, Alvarado, Karla) Performed tests: Temperature NIBP = set: 120/80, measured: 116/81 SpO2 = set: 95%, measured: 95% All test passed.
| 210266 | CM | ACTIV | ISSUE | 2024-10-08 09:50 |  | Yello Fin (S/N: 241401) repair | Oct 08 - call re: not working. Went to check. Tested unit in OR 12. Found adjustment handle stiff and unit does not move freely. Unit will be sent to vendor for repair. Contacted AMT surgical for RMA. Awaiting response. Oct 09 - Followed up. RMA received. Created SC. Requested approval. Received. Issued PO via MMC. Packed unit. Scheduled KJV Pick up and shipped the unit. Oct 29 - Followed up. Repair estimate recevied. Sent it for approval Received. Advised MMC to update the PO and send it to vendor. Nov 15 - Followed up. Was informed AMT has parts back ordered from OEM and are awaiting for the parts. Will ship the unit once it is repaired.
| 210267 | CM | ACTIV | ISSUE | 2024-10-08 10:43 |  | OR Table Stirrup Repair | informed Reeba the item needed to be sent out for repair and confirmed with Baxter it could only be done in the states attached all the emails repair qupte would be provide from Baxter Trump US once the item has beed inspected
1058526 | 210268 | CM | ACTIV | ISSUE | 2024-10-08 11:47 |  | Battery replacement | (TC 10/08/2024 12:55:15 PM, Canjimavadivel, Kamallesh ) Swapped out battery 1, need to order relacement for battery 2
1058167 | 210269 | CM | ACTIV | ISSUE | 2024-10-08 11:48 |  | Battery replacement | (TC 10/08/2024 12:54:24 PM, Canjimavadivel, Kamallesh ) Swapped out battery 1, need to order relacement for battery 2
1054286 | 210270 | CM | CMPLT | CMPLT | 2024-10-08 13:06 | 2024-10-30 15:14 | Not turning on | (TC 10/09/2024 03:08:45 PM, Canjimavadivel, Kamallesh ) Probe damaged, replaced. Device still not turning on after prolonged standby. Attempted to clean out probe well and sensor area. (TC 10/30/2024 03:14:02 PM, Canjimavadivel, Kamallesh ) Installed new probe and well, turning on normally
| 210271 | PA | CMPLT | CMPLT | 2024-10-08 13:13 | 2024-10-17 10:09 | Temp probes for stock | (TC 10/08/2024 02:16:01 PM, Canjimavadivel, Kamallesh ) Placed order for temp probes (TC 10/17/2024 10:09:02 AM, Canjimavadivel, Kamallesh ) Probes received
1051709 | 210272 | CM | CMPLT | CMPLT | 2024-10-08 13:40 | 2024-10-08 13:41 | Not working | (TC 10/08/2024 01:40:34 PM, Alvarado, Karla) Replaced batteries and screen.
1051734 | 210273 | CM | CMPLT | CMPLT | 2024-10-08 15:07 | 2024-10-08 15:07 | No Suction - Low pressure | (TC 10/08/2024 03:07:19 PM, Ferreira, Qwynn) Worked with Samir, the pin in the centre of the knob controlling the suction pressure was bent out of position. Placed the pin back into position and tested the suction - Working and tested 5- 60cmhg as they typically use 10-15cmhg Returned to L&D
67874 | 210274 | CM | CMPLT | NFF | 2024-10-08 15:11 | 2024-10-08 15:12 | It does not load cassettes properly | (TC 10/08/2024 03:11:36 PM, Alvarado, Karla) Performed Functional Test No fault found.
67886 | 210275 | CM | CMPLT | NFF | 2024-10-08 15:13 | 2024-10-08 15:13 | Epidural occluded. | (TC 10/08/2024 03:12:52 PM, Alvarado, Karla) Performed Functional Test No fault found.
67879 | 210276 | CM | CMPLT | NFF | 2024-10-08 15:14 | 2024-10-09 08:06 | Cassette won’t connect correctly. Displayed message. | (TC 10/08/2024 03:13:52 PM, Alvarado, Karla) Replaced batteries. “Turn pump with latch facing down” message showed up while it is priming. Performed Functional Test. (TC 10/09/2024 08:05:43 AM, Alvarado, Karla) Performed Functional Test. Test passed. No fault found.
1064830 | 210277 | CM | CMPLT | CMPLT | 2024-10-08 15:49 | 2024-10-17 11:04 | ART parameter not working | (TC 10/17/2024 11:03:40 AM, Canjimavadivel, Kamallesh ) Found IBP labels were both ART, changed one to different label and was able to capture waveform.
1067437 | 210278 | FC | CMPLT | CMPLT | 2024-10-09 09:02 | 2024-10-09 09:02 | QC checklist | (TC 10/09/2024 09:02:23 AM, Rhuland, Courtney) QC checklist completed
1057350 | 210279 | CM | ACTIV | ISSUE | 2024-10-09 09:24 |  | not reading temperature.
1064842 | 210280 | CM | CMPLT | CMPLT | 2024-10-09 09:39 | 2024-10-09 09:39 | Missing battery | (TC 10/09/2024 09:39:05 AM, Canjimavadivel, Kamallesh ) Brought up and installed replacement battery
1057350 | 210281 | CM | CMPLT | CMPLT | 2024-10-09 09:46 | 2024-10-09 09:47 | not reading temperature | (TC 10/09/2024 09:46:38 AM, Alvarado, Karla) Performance calibration with tester and passed.
1062336 | 210282 | CM | CMPLT | NFF | 2024-10-09 09:51 | 2024-10-09 09:51 | not working. | (TC 10/09/2024 09:51:25 AM, Alvarado, Karla) Test passed. No Fault Found (NFF)
| 210283 | PA | CMPLT | CMPLT | 2024-10-09 10:58 | 2024-10-24 11:40 | Ordering SpO2 Cables for PACU | (TC 10/09/2024 11:00:18 AM, Reid, Emmerson ) Communicated with unit and put in the order (TC 10/24/2024 11:39:36 AM, Reid, Emmerson ) Recieved and brought to PACU
66437 | 210284 | CM | CMPLT | CMPLT | 2024-10-09 13:45 | 2024-10-09 13:45 | V lead not working | (TC 10/09/2024 01:45:15 PM, Scero, Tomlyn) Replaced ECG 5 lead grabber and tested ECG function iwth a simulator. Replaced battery adaptor as it was very warn.
502366 | 210285 | CM | CMPLT | CMPLT | 2024-10-09 14:14 | 2024-10-11 09:46 | Leftspeaker broken | (TC 10/09/2024 02:13:52 PM, Fan, Tak) Removed old speaker, ordered replacement. (TC 10/11/2024 09:46:00 AM, Fan, Tak) Replaced speaker.
1049905 | 210286 | CM | CMPLT | CMPLT | 2024-10-09 14:31 | 2024-10-09 14:31 | Arm rocker switch broken | (TC 10/09/2024 02:30:50 PM, Fan, Tak) Replaced switch. Ordered parts.
501070 | 210287 | CM | ACTIV | ISSUE | 2024-10-09 15:42 |  | MSH OPG CEOU ROOM 10 IU22 DOESNT HAVE COMPLETE MEASUREMENT PACKAGE | (TC 10/09/2024 03:42:21 PM, Whittington, Scott) MSH OPG CEOU ROOM 10 IU22 DOESNT HAVE COMPLETE MEASUREMENT PACKAGE RESTORED SETTINGS FROM ROOM 11 IU22 | MSH OPG CEOU ROOM 10 IU22 DOESNT HAVE COMPLETE MEASUREMENT PACKAGE
| 210288 | CON | ACTIV | ISSUE | 2024-10-09 21:48 |  | MSH IVF MEETING ON PURCHASE OF US SYSTEMS | (TC 10/09/2024 09:48:05 PM, Whittington, Scott) MSH IVF MEETING ON PURCHASE OF US SYSTEMS | MSH IVF MEETING ON PURCHASE OF US SYSTEMS
1072040 | 210289 | CM | CMPLT | CMPLT | 2024-10-10 09:22 | 2024-10-10 09:22 | new incoming inspection
1072041 | 210290 | CM | CMPLT | CMPLT | 2024-10-10 09:22 | 2024-10-10 09:22 | new incoming inspection
1072042 | 210291 | CM | CMPLT | CMPLT | 2024-10-10 09:23 | 2024-10-10 09:23 | new incoming inspection
1072043 | 210292 | CM | CMPLT | CMPLT | 2024-10-10 09:23 | 2024-10-10 09:23 | new incoming inspection
1072044 | 210293 | CM | CMPLT | CMPLT | 2024-10-10 09:24 | 2024-10-10 09:24 | new incoming inspection
1072045 | 210294 | CM | CMPLT | CMPLT | 2024-10-10 09:25 | 2024-10-10 09:25 | new incoming inspection
1072046 | 210295 | CM | CMPLT | CMPLT | 2024-10-10 09:25 | 2024-10-10 09:25 | new incoming inspection
1072047 | 210296 | CM | CMPLT | CMPLT | 2024-10-10 09:26 | 2024-10-10 09:26 | new incoming inspection
1072048 | 210297 | CM | CMPLT | CMPLT | 2024-10-10 09:26 | 2024-10-10 09:26 | new incoming inspection
67959 | 210298 | CM | CMPLT | CMPLT | 2024-10-10 09:35 | 2024-10-10 09:35 | Broken Panel | (TC 10/10/2024 09:35:26 AM, Alvarado, Karla) Replaced porthole. Add scale plate. Set RH to 75%, analyzer = 75.0%. Set scale to 5kg, measured 5kg. Set temperature to 35C, measured 35C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 78.7 uA; GND Resistance = 207 mohms
66706 | 210299 | CM | CMPLT | CMPLT | 2024-10-10 10:03 | 2024-10-10 10:04 | Needs O2 tank holder | (TC 10/10/2024 10:02:47 AM, Alvarado, Karla) Installed O2 tank holder. (TC 10/10/2024 10:03:45 AM, Rhuland, Courtney) supervised
| 210300 | CM | ACTIV | ISSUE | 2024-10-10 10:30 |  | NICU ascom testing | (TC 10/10/2024 10:30:34 AM, Rhuland, Courtney) Simulated various alarms to test how ascom phones alerted on REV 4 (TC 10/22/2024 02:51:29 PM, Rhuland, Courtney) Issue with rooms 21 and 22 not sending alarms to ASCOM phones tested alarms multiple times simulating tachycardia and bradycardia but no alerts were being sent. Issue ongoing. (TC 10/23/2024 02:02:38 PM, Rhuland, Courtney) Issue ongoing in NICU Found 27 rooms that are having the issue Issue lies with how Philips is sending out the location ID Rooms ending in ^MSH are not sending to ASCOM Investigated further with Nima, including ICU that is also potentially affected. (TC 10/24/2024 02:06:45 PM, Rhuland, Courtney) Tested secondary alerting work around for temporarily resolving issue with ASCOM phones not alerting (TC 10/30/2024 01:59:26 PM, Rhuland, Courtney) Worked with Philips and IT to admit and readmit patients to fix the issue with secondary alerting
1072111 | 210301 | CM | CMPLT | CMPLT | 2024-10-10 10:31 | 2024-10-10 11:28 | Incoming inspection | (TC 10/10/2024 11:27:57 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged. Electrical safety pass
1072112 | 210302 | CM | CMPLT | CMPLT | 2024-10-10 10:35 | 2024-10-10 11:27 | Incoming inspection | (TC 10/10/2024 11:27:06 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged. Electrical safety pass
1072113 | 210303 | CM | CMPLT | CMPLT | 2024-10-10 10:36 | 2024-10-10 11:28 | Incoming inspection | (TC 10/10/2024 11:28:27 AM, Canjimavadivel, Kamallesh ) Added to database and asset tagged. Electrical safety pass
63875 | 210304 | CM | CMPLT | CMPLT | 2024-10-10 11:23 | 2024-10-10 11:23 | Broken mattress tray support | (TC 10/10/2024 11:22:56 AM, Rhuland, Courtney) Replaced mattress tray support with spare Functional check - OK
1072114 | 210305 | CM | CMPLT | CMPLT | 2024-10-10 11:46 | 2024-10-10 13:20 | 6 year PM parts replaced | (TC 10/10/2024 01:19:55 PM, Scero, Tomlyn) Finished replacing all 6 year pm parts. Self test complete and no issues or leaks.
1059764 | 210306 | CM | CMPLT | CMPLT | 2024-10-10 15:06 | 2024-10-10 15:06 | MSH 15 L&D TE7E CORD CAP RIPPED FROM POWER CORD WHEN BEING UNPLUGGED; FOUND THE CORD CAP IN ELECTRICAL SHOP | (TC 10/10/2024 03:06:03 PM, Whittington, Scott) MSH 15 L&D TE7E CORD CAP RIPPED FROM POWER CORD WHEN BEING UNPLUGGED; FOUND THE CORD CAP IN ELECTRICAL SHOP REATTACHED OK | MSH 15 L&D TE7E CORD CAP RIPPED FROM POWER CORD WHEN BEING UNPLUGGED; FOUND THE CORD CAP IN ELECTRICAL SHOP
1053348 | 210307 | CM | ACTIV | ISSUE | 2024-10-10 15:08 |  | MSH BI 1272 H3 STEREO BIOPSY CAL FAILED | (TC 10/10/2024 03:07:59 PM, Whittington, Scott) MSH BI 1272 H3 STEREO BIOPSY CAL FAILED | MSH BI 1272 H3 STEREO BIOPSY CAL FAILED REF# 01618409
| 210308 | CM | ACTIV | ISSUE | 2024-10-10 15:29 |  | misc biomed - Oct - Dec 2024 | (TC 10/10/2024 03:29:21 PM, Cortez, John) Helped install secondary monitor in PFT front desk (TC 10/04/2024 03:58:15 PM, Cortez, John) ICU Philips rev 4 nurse call testing (TC 10/11/2024 11:42:15 AM, Cortez, John) CCU recorder reported not working. Power cord was loose on power supply. Restarted central and verified functionality (TC 10/15/2024 04:23:04 PM, Cortez, John) Issue with 10S WOW #8. Not powering on. Fuse blew. Replaced (TC 10/17/2024 04:28:32 PM, Cortez, John) Amanda reported another X3 with inaccurate IBP readings. 8/13 X3 modules checked to be inaccurate with patient simulator Turns out Cal. factor setting is related. Running Cal. press option with simulator sets Cal. factor back to 200. Fixes inaccuracy. However we need to clarify purpose of these settings/options with Philips clinical rep. Technical support, unsure of correct usage of these settings. (TC 10/28/2024 03:11:18 PM, Cortez, John) Assisted off hours with syringe pump/IAware issue (TC 11/08/2024 11:26:07 AM, Cortez, John) Checked equipment in new 16N surge space before Nov 12 go-live 8x beds 2x vital signs 1x scale 1x aed 1x aspirator (TC 11/15/2024 02:52:01 PM, Cortez, John) Radiology called about their central monitor in angio control room. Display not showing vitals from angio 2 Muhammad from JDMI resolved similar issue a couple of years ago Reached out to Scott and Muhammad
1061681 | 210309 | CM | ACTIV | ISSUE | 2024-10-10 15:31 |  | Pump not running | (TC 10/08/2024 03:31:16 PM, Cortez, John) Confirmed issue, swapped with spare (TC 10/10/2024 03:31:46 PM, Cortez, John) Reached out to Stryker. 3 year warranty lapsed. Possibly the compressor assembly, or PCBA
1071407 | 210310 | CM | CMPLT | CMPLT | 2024-10-10 15:41 | 2024-10-10 15:41 | System failure 25 | (TC 10/10/2024 03:41:10 PM, Rhuland, Courtney) Reseated fan Calibrated RH Functional check - OK
1064850 | 210311 | CM | CMPLT | CMPLT | 2024-10-10 15:52 | 2024-10-17 16:34 | IBP discrepancy with edwards trial monitor | (TC 10/04/2024 03:55:27 PM, Cortez, John) Confirmed issie with X3 module, reading ~10 mmhg off with simulator (TC 10/17/2024 04:34:08 PM, Cortez, John) Ran Cal. Press with simulator Fixes inaccuracy, at least when testing with patient simulator Need to verify corrrect usage of this setting/option with Philips, and relay info to ICU staff
1064508 | 210312 | CM | CMPLT | CMPLT | 2024-10-10 16:42 | 2024-10-10 16:42 | Software update to 1.10 | (TC 10/10/2024 04:42:17 PM, Cortez, John) Update completed
65416 | 210313 | CM | CMPLT | CMPLT | 2024-10-11 07:56 | 2024-10-11 07:56 | Broken Door | (TC 10/11/2024 07:56:04 AM, Alvarado, Karla) Replaced wall hinge snap post (NE door).
65492 | 210314 | CM | CMPLT | CMPLT | 2024-10-11 08:45 | 2024-10-18 16:22 | "0 working" | (TC 10/10/2024 08:45:34 AM, Cortez, John) Just reads 35 deg C , Not a probe issue (TC 10/18/2024 04:21:59 PM, Cortez, John) swapped unit with spare. set this asset to inactive
60022 | 210315 | CM | CMPLT | CMPLT | 2024-10-11 09:41 | 2024-10-11 09:42 | Burning smell reported. | (TC 10/11/2024 09:41:44 AM, Alvarado, Karla) Run the equipment for 40 min in both modes (open and close). Replaced porthole hinge cover (SE door). Set RH to 75%, analyzer = 73.4%. Calibrated scale. Set scale to 5kg, measured 5kg. Set temperature to 35C, measured 35C on all temperature points. Functional tests passed. Electrical safety results - Leakage current = 77.8 uA; GND Resistance = 312 mohms (TC 10/11/2024 09:41:44 AM, Arslan, Olcay) Supervised
1055074 | 210316 | CM | CMPLT | CMPLT | 2024-10-11 10:30 | 2024-10-11 10:30 | Repurposed for 10S | (TC 10/11/2024 10:29:41 AM, Cortez, John) Moved to 10S
1066460 | 210317 | CM | CMPLT | CMPLT | 2024-10-11 11:00 | 2024-10-11 11:00 | QC checklist | (TC 10/11/2024 11:00:25 AM, Alvarado, Karla) QC checklist completed. Returned to 11 south. (TC 10/11/2024 11:00:25 AM, Arslan, Olcay) Supervised
1057733 | 210318 | CM | CMPLT | CMPLT | 2024-10-11 11:25 | 2024-10-16 11:22 | Battery Fail | (TC 10/11/2024 11:29:22 AM, Chukri, Samir) Ordered two batteries one for stock in Biomed workshop. Qwynn switched the battery from another incabator ( battery replaced Jul 2024) . Michelle needs this transport incubator and cann't wait the new battery. (TC 10/16/2024 11:21:09 AM, Chukri, Samir) Received two batteries installed one and keep one on stock. (TC 10/16/2024 11:22:05 AM, Ferreira, Qwynn) Installed new battery
1051809 | 210319 | CM | CMPLT | CMPLT | 2024-10-11 11:33 | 2024-10-11 11:33 | Malfunctioning. Readings are too high. | (TC 10/11/2024 11:33:07 AM, Alvarado, Karla) Calibrated. Replaced probe. Test passed
63203 | 210320 | CM | ACTIV | ISSUE | 2024-10-11 13:42 |  | Machine does not recognize the set in. | (TC 10/11/2024 01:42:29 PM, Alvarado, Karla) The 4th selection button is not working properly. Door is not keep it closed. Machine recognize the set in. (No fault found) Kept the device running for around 3 hours and feeding completed.
1056895 | 210321 | CM | ACTIV | ISSUE | 2024-10-11 13:43 |  | replaced the power adapter for the scale OR | (TC 10/11/2024 01:25:20 PM, Huang, Jessica) received the call from OR 1 and replaced the power adapter on the ScaleOR. Contacted Olympus for the quote and ordered spare power adaters.
1066462 | 210322 | CM | CMPLT | CMPLT | 2024-10-11 13:45 | 2024-10-28 13:22 | QC checklist | (TC 10/11/2024 01:44:50 PM, Alvarado, Karla) Power charger is not working. Holder broken. (TC 10/28/2024 01:21:53 PM, Baghirzade, Nara) Replaced solution hanger. Replaced power supply. QC checklist completed. Returned to 11 South.
1059763 | 210323 | CM | CMPLT | CMPLT | 2024-10-11 16:23 | 2024-10-17 13:46 | MSH 15 L&D MSH_LD_TE7_D NOT CHARGING, NO POWER | (TC 10/11/2024 04:23:07 PM, Whittington, Scott) MSH 15 L&D MSH_LD_TE7_D NOT CHARGING, NO POWER REPLACEMENT CORD CAP CONNECTIONS HAD BEEN LOOSENED BY STRESS ON POWER CORD DURING UNPLUGGING (TC 10/11/2024 04:25:24 PM, Dolinar, John) MSH 15 L&D MSH_LD_TE7_D NOT CHARGING, NO POWER REPLACEMENT CORD CAP CONNECTIONS HAD BEEN LOOSENED BY STRESS ON POWER CORD DURING UNPLUGGING REATTACH CORD CAP OK (TC 10/17/2024 01:23:00 PM, Whittington, Scott) MSH 15 L&D MSH_LD_TE7_D NOT CHARGING, NO POWER REPLACEMENT CORD CAP CONNECTIONS HAD BEEN LOOSENED BY STRESS ON POWER CORD DURING UNPLUGGING (TC 10/17/2024 01:25:00 PM, Dolinar, John) MSH 15 L&D MSH_LD_TE7_D NOT CHARGING, NO POWER REPLACEMENT CORD CAP CONNECTIONS HAD BEEN LOOSENED BY STRESS ON POWER CORD DURING UNPLUGGING REATTACH CORD CAP OK | MSH 15 L&D MSH_LD_TE7_D NOT CHARGING, NO POWER
| 210324 | OC | CMPLT | CMPLT | 2024-10-15 08:11 | 2024-10-22 19:30 | On-call Oct 11-24, 2024 | (TC 10/12/2024 10:15:00 AM, Reid, Emmerson ) Was called by 12N because they could not plug the LSU (asset# 1055770) from the crash cart into the wall. This was due to a bent power plug. I was able to bend the plug back with pliers and return the cable to the device. I also noticed that the power bar on crash cart BCLS # 13 is missing the ground pin and needs to be replaced. (TC 10/15/2024 09:30:00 PM, Reid, Emmerson ) Was called by the ICU and 18:30. They said that 4 of their CRRT are down and they have no more that are working. I was told that "someone would die tonight if one was not working". I came in and contacted John and Samir to advise. John let me know that there was another in biomed that was working so I brought it up and set it up with the nurse who tried the other devices (Olga). In the first setup she selected to use the syringe when she didn't mean to so we needed to start over. I set it up again with another nurse and we were able to get it working. The TL (Julie) let me know that they may need a second CRRT for another patient so I grabbed another device from the storage room and setup the priming with Julie. This worked so I let them know the second device was ready and I left. (TC 10/22/2024 07:30:00 PM, Reid, Emmerson ) Nima riched out for me to check on the NICU beds 21 & 22 because they were not getting ASCOM alerts. I remoted in and checked that the names, demographic, and vital data was coming through. Worked with Nima and Kamran to determine this was not a Philips issue.
1059778 | 210325 | CM | CMPLT | CMPLT | 2024-10-15 09:12 | 2024-10-15 09:12 | QC checklist | (TC 10/15/2024 09:12:07 AM, Rhuland, Courtney) QC checklist completed Returned to 11S
1061333 | 210326 | CM | CMPLT | CMPLT | 2024-10-15 09:12 | 2024-10-15 09:16 | QC checklist | (TC 10/15/2024 09:15:57 AM, Rhuland, Courtney) QC checklist completed Returned to 11N
1066454 | 210327 | CM | CMPLT | CMPLT | 2024-10-15 09:12 | 2024-10-15 09:17 | QC checklist | (TC 10/15/2024 09:16:43 AM, Rhuland, Courtney) QC checklist completed Returned to 14S
1054521 | 210328 | CM | CMPLT | CMPLT | 2024-10-15 09:25 | 2024-10-15 09:29 | Lead 5 not working | (TC 10/15/2024 09:28:40 AM, Reid, Emmerson ) Replaced the keyboard cover, RA lead and V4-V6 leads. Tested and working now.
| 210329 | RDS | CMPLT | CMPLT | 2024-10-15 10:34 | 2024-10-15 10:34 | Weekly checks in CCU | (TC 10/15/2024 10:34:24 AM, Scero, Tomlyn) weekly checks complete. 30J test passed
66434 | 210330 | CM | CMPLT | CMPLT | 2024-10-15 10:44 | 2024-10-15 11:16 | V lead not working | (TC 10/15/2024 11:15:33 AM, Scero, Tomlyn) replaced ECG 5 lead grabbers. Tested with patient simulator. All waveforms showing properly.
1072130 | 210331 | ACC | CMPLT | CMPLT | 2024-10-15 11:20 | 2024-10-15 11:58 | Incoming inspection | (TC 10/15/2024 11:58:14 AM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK
1072131 | 210332 | ACC | CMPLT | CMPLT | 2024-10-15 11:20 | 2024-10-15 11:59 | Incoming inspection | (TC 10/15/2024 11:59:01 AM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK
1072132 | 210333 | ACC | CMPLT | CMPLT | 2024-10-15 11:20 | 2024-10-15 12:00 | Incoming inspection | (TC 10/15/2024 11:59:32 AM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK
1072133 | 210334 | ACC | CMPLT | CMPLT | 2024-10-15 11:20 | 2024-10-15 12:03 | Incoming inspection | (TC 10/15/2024 12:02:39 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK
1072134 | 210335 | ACC | CMPLT | CMPLT | 2024-10-15 11:20 | 2024-10-15 12:03 | Incoming inspection | (TC 10/15/2024 12:03:00 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK
1072135 | 210336 | ACC | CMPLT | CMPLT | 2024-10-15 11:20 | 2024-10-15 12:03 | Incoming inspection | (TC 10/15/2024 12:03:22 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK
1072136 | 210337 | ACC | CMPLT | CMPLT | 2024-10-15 11:20 | 2024-10-15 12:04 | Incoming inspection | (TC 10/15/2024 12:03:54 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK
66459 | 210338 | PM | CMPLT | CMPLT | 2024-10-15 12:55 | 2024-10-15 12:55 | PM-MAJOR - Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric, High-Frequency | Performed 4th year PM after 3rd year PM last year. No GS500. It has PS500. Run battery check, 220 min, 43 Ah. Tested. Functional. SW.2.51 Leakage: 86 uA and 90 ohm.
| 210339 | CM | ACTIV | ISSUE | 2024-10-15 13:28 |  | Broken camera head
1064422 | 210340 | CM | CMPLT | CMPLT | 2024-10-15 13:46 | 2024-10-15 13:48 | No turning on / black screen. | (TC 10/15/2024 01:48:23 PM, Alvarado, Karla) Replaced keyboard. Replaced keyboard cover Performed functional test. Test passed. (TC 10/15/2024 01:48:23 PM, Dolinar, John) Supervised
1071973 | 210341 | CM | ACTIV | ISSUE | 2024-10-16 08:32 |  | INC Investigation - #83321 - Hose Malfunction | Oct 15 - Call re: Hose issue. Went to check and spoke with staff. Found a bubble formed on the hose that connects from gas purifier to cryo unit. Disconnected everything and released pressure. Removed machine from service and arranged for the spare one. Took images and send to vendor and requested replacement. Further investigation to be done. Oct 24 - Request from Reeba to file SAFER. Gethered sime details and Filed SAFER. Oct 25 - Upon inspecting, It was found that the damaged hose (that came with new unit) did not have any regulatory markings when we received it from the vendor. Contacted vendor (stat medical) again and requested an explanation. Was informed that they have reached out to manufacturer (cooper surgical) and are awaiting their response. Requested stat medical to put us directly in touch with the OEM. Awaiting respose. Oct 28 - Found another SAFER filed by the OR nursing team lead. Closed my safer. Started working on Investigation report. Oct 29 - Received questions from OEM. Answered and sent our concens to them and requested clarification. Oct 31 - Followed up. Awaiting response. Nov 05 - Followed up again with both cooper surgical and stat medical. Awaiting response. Nov 18-20 - Followed up. Was informed that both stat medical and cooper surgical are investigating on their end and are awaiting to hear back from their product surveillance and Health canada. Requested stat mediacal to send technician to inspect the hose and certify the unit. Awaiting response. Initiated medical device problem with health canada. Could not submit the form after filling out all details. Attempted multiple times and in different browsers. Issue still persisted. Reached out to health canada support for assistace.
1050605 | 210342 | CM | CMPLT | CMPLT | 2024-10-16 09:32 | 2024-10-16 09:32 | Scale does not turn on. | (TC 10/16/2024 09:32:26 AM, Alvarado, Karla) Replaced batteries. Tested 10lb, 20, 25lb, and 50lb. All tests passed.
1055585 | 210343 | CM | CMPLT | NFF | 2024-10-16 10:03 | 2024-10-16 10:03 | it needs to be calibrated. | (TC 10/16/2024 10:03:24 AM, Alvarado, Karla) Performed tests: Temperature NIBP= set: 120/80, measured: 116/83, 118/85, 115/84. SpO2 = set: 95% and 40 b/m, measured: 95% and 40 b/m. All test passed. No Fault Found
1055593 | 210344 | CM | CMPLT | CMPLT | 2024-10-16 10:31 | 2024-10-16 10:31 | BP function is not working. | (TC 10/16/2024 10:31:07 AM, Alvarado, Karla) Replaced blood pressure tube. Performed tests: Temperature NIBP= set: 120/80, measured: 117/84. SpO2 = set: 85% and 80 b/m, measured: 83% and 80 b/m. All test passed.
1064386 | 210345 | CM | CMPLT | CMPLT | 2024-10-16 11:44 | 2024-10-16 11:44 | Not working. | (TC 10/16/2024 11:43:56 AM, Alvarado, Karla) Replaced batteries. Test passed.
1054522 | 210346 | CM | CMPLT | CMPLT | 2024-10-16 11:55 | 2024-10-16 11:55 | the ECG machine was making a sound. | (TC 10/16/2024 11:55:04 AM, Alvarado, Karla) The equipment was send to the department making a noise; the reason found it was low battery. While performing the functional test, the V4-V6 leads were making a lot of noise; therefore the lead chest set (V4-V6 standard length) were replaced. The equipment was left running and charging for almost 24 hours without showing any issue. Performed functional test. Test passed. Before sending the equipment back, it was noticed that it could not turn off.
1063746 | 210347 | PM | CMPLT | CMPLT | 2024-10-16 12:15 | 2024-10-16 12:15 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/16/2024 12:14:55 PM, Alvarado, Karla) Drained, disinfected, rinsed and refilled with sterile water.
1071527 | 210348 | CM | ACTIV | ISSUE | 2024-10-16 13:01 |  | MSH 15 IVF FRAGILE PT MOCK PROCEDURE | (TC 10/16/2024 01:01:32 PM, Whittington, Scott) MSH 15 IVF FRAGILE PT MOCK PROCEDURE | MSH 15 IVF FRAGILE PT MOCK PROCEDURE
1050476 | 210349 | CM | CMPLT | CMPLT | 2024-10-16 13:32 | 2024-10-16 13:32 | expiratory cassette membrane expired | replaced expiratory cassette membrane. reset counter. tested. passed device check.
66016 | 210350 | CM | CMPLT | NFF | 2024-10-16 13:52 | 2024-10-16 13:52 | Battery does not charge. | (TC 10/16/2024 01:52:27 PM, Alvarado, Karla) Left the equipment charge. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage = 90.1 uA and the GND Resistance = 0.125 Ohms Download error log: did not find errors. All test passed.
1059828 | 210351 | PM | CMPLT | CMPLT | 2024-10-16 14:36 | 2024-10-16 14:36 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/16/2024 02:36:10 PM, Alvarado, Karla) Drained, disinfected, rinsed and refilled with sterile water.
| 210352 | CON | ACTIV | ISSUE | 2024-10-16 15:11 |  | Toronto Western Aneurysms
1046196 | 210353 | CM | CMPLT | CMPLT | 2024-10-16 15:27 | 2024-10-23 15:11 | Strange noise coming from gantry | Strange noise coming from gantry. Several tests done between patients. Gets worse when rotating and during exposure. Unable to determine with the covers closed. Booked the room for Friday at 10 AM Martin will be on site as well. Oct 23-x-ray tube replaced/calibrated Service completed annual CT QC
1054351 | 210354 | CM | CMPLT | CMPLT | 2024-10-16 15:35 | 2024-10-16 15:35 | Screen reads “SEC”. | (TC 10/16/2024 03:35:17 PM, Alvarado, Karla) Team leader asked to remove the security mode Removed security mode. Replaced batteries.
1054356 | 210355 | CM | CMPLT | CMPLT | 2024-10-16 15:36 | 2024-10-16 15:37 | Screen reads “SEC”. | (TC 10/16/2024 03:36:31 PM, Alvarado, Karla) Team leader asked to remove the security mode Removed security mode.
1054287 | 210356 | CM | CMPLT | CMPLT | 2024-10-16 15:42 | 2024-10-16 15:42 | Screen reads “SEC”. | (TC 10/16/2024 03:42:09 PM, Alvarado, Karla) Team leader asked to remove the security mode Removed security mode. Replaced batteries.
1055071 | 210357 | CM | CMPLT | CMPLT | 2024-10-16 15:43 | 2024-10-16 15:43 | Screen reads “SEC”. | (TC 10/16/2024 03:43:22 PM, Alvarado, Karla) Team leader asked to remove the security mode Removed security mode.
1065522 | 210358 | CM | ACTIV | ISSUE | 2024-10-17 08:28 |  | failed caliscan test
1055594 | 210359 | CM | CMPLT | CMPLT | 2024-10-17 08:52 | 2024-10-17 08:52 | not letting RN to do BP. It is asking for password. | (TC 10/17/2024 08:51:59 AM, Alvarado, Karla) Replaced probe well. Performed tests: Temperature NIBP = set: 120/80, measured: 117/84. SpO2 = set: 85% and 80 b/m, measured: 83% and 80 b/m. All test passed.
| 210360 | CM | CMPLT | CMPLT | 2024-10-17 08:58 | 2024-10-21 12:38 | Batteries for biomed stock
1055944 | 210361 | CM | CMPLT | CMPLT | 2024-10-17 09:13 | 2024-10-22 15:03 | needs a new battery | shipping Masimo Root for repair. Received new unit. sn.2000172783. PO# 7500213404
63801 | 210362 | CM | CMPLT | NFF | 2024-10-17 10:34 | 2024-10-17 10:35 | not turning on. | (TC 10/17/2024 10:34:27 AM, Alvarado, Karla) Performed functional test. Test passed. No Fault Found. (TC 10/17/2024 10:35:10 AM, Arslan, Olcay) Supervised
| 210363 | CM | CMPLT | CMPLT | 2024-10-17 13:43 | 2024-10-29 14:33 | Chair not working | (TC 10/17/2024 02:30:12 PM, Rhuland, Courtney) Chair in ENT clinic not working Fuses were blown at the power cord Replaced fuses and fuses blew again immediately Called MTI and was told it could be bridge rectifier Tested bridge rectifier with multimeter - found to be faulty Placed order for bridge rectifier (TC 10/29/2024 02:33:12 PM, Rhuland, Courtney) Recieved bridge rectifier Replaced bridge rectifier in chair Replaced both fuses at power cord Reassembled chair Chair is fully functional
1054351 | 210364 | CM | CMPLT | NFF | 2024-10-17 14:56 | 2024-10-17 14:58 | Screen reads “SEC”. | (TC 10/17/2024 02:57:34 PM, Alvarado, Karla) No problem found.
1070570 | 210365 | CM | CMPLT | CMPLT | 2024-10-18 07:29 | 2024-10-18 07:29 | Loose rolling base | (TC 10/18/2024 07:28:41 AM, Rhuland, Courtney) Removed pump from stand Tightened screw from base and botom of stand Reattached pump Functional check - OK
| 210366 | CM | CMPLT | CMPLT | 2024-10-18 08:10 | 2024-10-29 13:52 | Formlabs Order
1047006 | 210367 | CM | CMPLT | NFF | 2024-10-18 08:43 | 2024-10-18 08:43 | Alarm keeps going off | (TC 10/18/2024 08:43:01 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Download error log: did not find errors. All test passed. No fault find
68627 | 210368 | CM | CMPLT | CMPLT | 2024-10-18 08:44 | 2024-10-18 08:44 | Alarm keeps going off | (TC 10/18/2024 08:44:01 AM, Alvarado, Karla) PCU was not charging. Therefore, the power cord was replaced. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage =72.4 uA and the GND Resistance = 0.160 Ohms All test passed. Main speaker is fully functional.
1055731 | 210369 | PM | CMPLT | CMPLT | 2024-10-18 09:29 | 2024-10-18 10:07 | PM-MAJOR - ANNUAL FOR: MICROSTREAM EXTENSION | (TC 10/18/2024 10:06:57 AM, Canjimavadivel, Kamallesh ) PM complete, see procedure
1055721 | 210370 | PM | CMPLT | CMPLT | 2024-10-18 10:12 | 2024-10-18 10:23 | PM-MAJOR - ANNUAL FOR: MICROSTREAM EXTENSION | (TC 10/18/2024 10:22:40 AM, Canjimavadivel, Kamallesh ) PM complete, see procedure
68775 | 210371 | CM | CMPLT | NFF | 2024-10-18 10:49 | 2024-10-18 10:49 | it does not work | (TC 10/18/2024 10:49:18 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage =78 uA and the GND Resistance = 0.230 Ohms All test passed. No fault find
1065921 | 210372 | CM | CMPLT | CMPLT | 2024-10-18 10:51 | 2024-10-18 10:51 | it does not work | (TC 10/18/2024 10:51:05 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Download error log: did not find errors. All test passed. No fault find
68894 | 210373 | CM | CMPLT | CMPLT | 2024-10-18 11:32 | 2024-10-18 11:32 | it does not work | (TC 10/18/2024 11:32:12 AM, Alvarado, Karla) While measuring the GND resistance, the value was too high, almost 400mOhms. Therefore, the power cord was replaced for security. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage =68uA and the GND Resistance = 231mOhms All test passed.
66507 | 210374 | CM | CMPLT | CMPLT | 2024-10-18 15:20 | 2024-10-18 15:20 | Recorder not printing | (TC 10/18/2024 03:20:24 PM, Alvarado, Karla) Removed and repaired the fetal recorder. Reinstalled. Calibrated. Test passed.
1069393 | 210375 | CM | CMPLT | CMPLT | 2024-10-18 16:08 | 2024-10-18 16:13 | Technical error, scale issue | (TC 10/18/2024 04:08:13 PM, Cortez, John) Users reported difficulty in using device. Called after hours Checked device next day, confirmed issue with scale. calibration off Tools/kit still on order. Reached out to Baxter for assistance Scale was calibrated by David (field support) Patient simulation test passed
1069388 | 210376 | CM | CMPLT | CMPLT | 2024-10-18 16:15 | 2024-10-18 16:15 | Technical error, scale issue | (TC 10/18/2024 04:14:44 PM, Cortez, John) Users reported difficulty in using device. Called after hours Checked device next day, confirmed issue with scale. calibration off Tools/kit still on order. Reached out to Baxter for assistance Scale was calibrated by David (field support) Patient simulation test passed
1069394 | 210377 | CM | CMPLT | NFF | 2024-10-18 16:17 | 2024-10-15 16:18 | ""purple" line giving error, won't prime" | (TC 10/15/2024 04:17:35 PM, Cortez, John) Was able to prime and run patient simulation for an hour. NFF
67097 | 210378 | CM | CMPLT | CMPLT | 2024-10-18 16:20 | 2024-10-15 16:20 | "missing 02 monitor" | (TC 10/15/2024 04:19:50 PM, Cortez, John) Added spo2 sensor
500868 | 210379 | CM | ACTIV | ISSUE | 2024-10-20 23:42 |  | MSH US RM 4 APLIO XG UNABLE TO DFRG, REQ HDD MNTCE | (TC 10/22/2024 08:33:21 PM, Whittington, Scott) MSH US RM 4 APLIO XG UNABLE TO DFRG, REQ HDD MNTCE | MSH US RM 4 APLIO XG UNABLE TO DFRG, REQ HDD MNTCE
1053391 | 210380 | CM | CMPLT | CMPLT | 2024-10-21 09:23 | 2024-10-21 09:23 | Check before and after patient use | (TC 10/21/2024 09:23:19 AM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 49 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 201 mmHg Set 125 mmHg - Measured 124 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device, it was off by a few minutes. Device is functional. Returned to 11S.
1054987 | 210381 | FC | CMPLT | NFF | 2024-10-21 09:27 | 2024-10-21 09:28 | Reported not working | (TC 10/21/2024 09:28:08 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1069391 | 210382 | CM | CMPLT | CMPLT | 2024-10-21 10:02 | 2024-10-21 10:52 | Air detected while in use | (TC 10/21/2024 10:51:55 AM, Canjimavadivel, Kamallesh ) Connected to test set and ran simulation, was not able to reproduce error. Reached out to unit for additional info, nurse worked overnight and not available for comment. Device returned to service
63216 | 210383 | CM | ACTIV | SHP | 2024-10-21 10:30 |  | Sending for repair/exchange under consignment | (TC 10/21/2024 10:41:13 AM, Baghirzade, Nara) SC 1003333418 (TC 10/23/2024 08:52:59 AM, Baghirzade, Nara) RMA: WO-00104696 KJV#: 460537 PO 7500213548
63438 | 210384 | CM | ACTIV | SHP | 2024-10-21 10:31 |  | Sending for repair/exchange under consignment | (TC 10/21/2024 10:40:18 AM, Baghirzade, Nara) SC 1003333418 (TC 10/23/2024 08:55:46 AM, Baghirzade, Nara) RMA: WO-00104697 KJV#: 460538 PO 7500213548
1054684 | 210385 | CM | ACTIV | SHP | 2024-10-21 10:32 |  | Sending for repair/exchange under consignment | (TC 10/21/2024 10:39:34 AM, Baghirzade, Nara) SC 1003333418 (TC 10/23/2024 08:51:02 AM, Baghirzade, Nara) RMA: WO-00104698 KJV#: 460536 PO 7500213548
| 210386 | RDS | CMPLT | CMPLT | 2024-10-21 11:15 | 2024-10-21 13:32 | Weekly checks in CCU | (TC 10/21/2024 01:31:58 PM, Scero, Tomlyn) Checks in CCU complete and 30J tests passed.
66504 | 210387 | CM | CMPLT | CMPLT | 2024-10-21 11:58 | 2024-10-21 11:58 | Recorder not printing | (TC 10/21/2024 11:58:07 AM, Alvarado, Karla) Removed monitor from room, confirmed problem. Removed and repaired the fetal recorder. Ribbon cable and rubber band were replaced. Reinstalled. Calibrated. Test passed.
1066475 | 210388 | CM | CMPLT | CMPLT | 2024-10-21 12:57 | 2024-10-21 12:57 | not working | (TC 10/16/2024 12:57:21 PM, Huang, Jessica) Beena from PSA contacted regarding the device was in OR 12. Checked the device and it was functional. Infromed Beena and put it back to the service.
1057047 | 210389 | CM | CMPLT | CMPLT | 2024-10-21 12:59 | 2024-10-21 12:59 | replaced the temperature probe
1063121 | 210390 | CM | CMPLT | CMPLT | 2024-10-21 13:00 | 2024-10-21 13:00 | cleaned the sensor and passed calibration
1063123 | 210391 | CM | CMPLT | CMPLT | 2024-10-21 13:02 | 2024-10-21 13:02 | cleaned the sensor and passed calibration
1064422 | 210392 | CM | CMPLT | CMPLT | 2024-10-21 13:14 | 2024-10-21 13:14 | Not taking readings | (TC 10/21/2024 01:13:45 PM, Canjimavadivel, Kamallesh ) Found faulty patient data cable. Replaced with new cable, tested ok.
| 210393 | PA | CMPLT | CMPLT | 2024-10-21 13:18 | 2024-10-30 14:21 | TC70 cables for stock | (TC 10/21/2024 01:36:51 PM, Canjimavadivel, Kamallesh ) Placed order for stock cables (TC 10/30/2024 02:21:25 PM, Canjimavadivel, Kamallesh ) Received cables, added to stock
1057607 | 210394 | CM | CMPLT | CMPLT | 2024-10-21 13:23 | 2024-10-21 13:23 | replaced the BP hose and the power cord
1056895 | 210395 | CM | CMPLT | CMPLT | 2024-10-21 13:28 | 2024-10-21 13:28 | reset the PM notification message | (TC 09/30/2024 01:27:52 PM, Huang, Jessica) received request from Reeba in regard to check the device. Dr. Monteiro saw the PM notification message while he was using the device. Contacted Steve to reset the message and informed the OR team
| 210396 | ACC | CMPLT | CMPLT | 2024-10-21 14:07 | 2024-10-21 14:07 | Philips Epiq demo | (TC 10/18/2024 02:07:04 PM, Fan, Tak) Demo for SPP ultrasound RFP
63203 | 210397 | CM | CMPLT | CMPLT | 2024-10-21 14:22 | 2024-10-21 14:22 | Door not closing properly | (TC 10/21/2024 02:22:05 PM, Baghirzade, Nara) Disassembled the case and reseated the loading area; door closes properly. Ran a feed at various rates; no issues or errors found. Pump is fully functional.
| 210398 | CM | CMPLT | CMPLT | 2024-10-21 14:27 | 2024-10-21 14:27 | File recovery | (TC 10/21/2024 02:27:18 PM, Baghirzade, Nara) Gaby accidently deleted all CPET physician interpretations for year 2022, 2023, and 2024. The files did not go in the recycle bin. Looked for any backups or any other path that the files went to however was not able to find it. Finally found the option to see "previous versions of a file" and was able to recover all physician interpretations.
32580 | 210399 | CM | CMPLT | CMPLT | 2024-10-21 20:55 | 2024-10-21 20:53 | abnormal x-ray condition reported/ no x-ray | abnormal x-ray condition reported/ no x-ray Service completed troubleshooting Service completed IF cals Service tested BIR logs Found Recon unit failures reported CT rebooted/tested OK
1051433 | 210400 | CM | CMPLT | CMPLT | 2024-10-22 07:40 | 2024-10-31 11:51 | Sending for repair | (TC 10/22/2024 07:57:50 AM, Baghirzade, Nara) SC 1003334175 (TC 10/29/2024 10:33:02 AM, Baghirzade, Nara) Updated repair cost after assessment. (TC 10/31/2024 07:58:42 AM, Baghirzade, Nara) PO 7500213595 (TC 10/31/2024 11:50:33 AM, Baghirzade, Nara) Received unit back from repair. Functional test passed. Attached service report. Dropped off at L&D.
1050752 | 210401 | CM | CMPLT | CMPLT | 2024-10-22 07:41 | 2024-10-31 11:52 | Sending for repair | (TC 10/22/2024 07:57:15 AM, Baghirzade, Nara) SC 1003334175 (TC 10/29/2024 10:33:37 AM, Baghirzade, Nara) Updated repair cost after assessment. (TC 10/31/2024 07:59:22 AM, Baghirzade, Nara) PO 7500213595 (TC 10/31/2024 11:52:03 AM, Baghirzade, Nara) Received unit back from repair. Functional test passed. Attached service report. Dropped off at L&D.
1060134 | 210402 | CM | CMPLT | CMPLT | 2024-10-22 07:42 | 2024-10-31 11:53 | Sending for repair | (TC 10/22/2024 07:56:45 AM, Baghirzade, Nara) SC 1003334175 (TC 10/29/2024 10:34:06 AM, Baghirzade, Nara) Updated repair cost after assessment. (TC 10/31/2024 07:59:53 AM, Baghirzade, Nara) PO 7500213595 (TC 10/31/2024 11:52:53 AM, Baghirzade, Nara) Received unit back from repair. Functional test passed. Attached service report. Dropped off at L&D.
1057452 | 210403 | CM | CMPLT | CMPLT | 2024-10-22 07:43 | 2024-10-31 11:54 | Sending for repair | (TC 10/22/2024 07:56:12 AM, Baghirzade, Nara) SC 1003334175 (TC 10/31/2024 08:00:10 AM, Baghirzade, Nara) PO 7500213595 (TC 10/31/2024 11:53:39 AM, Baghirzade, Nara) Received unit back from repair. Functional test passed. Attached service report. Dropped off at L&D.
1048572 | 210404 | CM | ACTIV | SHP | 2024-10-22 08:36 |  | Measuring PEEP incorrectly; requesting service | (TC 10/22/2024 08:35:44 AM, Baghirzade, Nara) Reached out to Bailey from Bunnell for service. (TC 10/23/2024 12:29:38 PM, Baghirzade, Nara) RA#S2024-2037 SC 1003335806 (TC 10/24/2024 09:33:07 AM, Baghirzade, Nara) KJV# 460811
65943 | 210405 | CM | CMPLT | CMPLT | 2024-10-22 08:48 | 2024-10-22 08:49 | it is not connecting. | (TC 10/18/2024 08:48:01 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Plunger position accuracy failed. Knob and barrel clamp have to be relocated. (TC 10/22/2024 08:48:39 AM, Alvarado, Karla) Replaced Knob. Relocated barrel clamp. Performed Calibration and Check- in test in ASM. All test passed.
68793 | 210406 | CM | CMPLT | NFF | 2024-10-22 09:07 | 2024-10-22 09:07 | it does not turn on. | (TC 10/22/2024 09:07:23 AM, Alvarado, Karla) Equipment was low battery. Keep it charging. Performed check-In tests in ASM. . Where the GND Leakage = 69 uA and the GND Resistance = 0.140 Ohms All tests passed. No Fault Find
68385 | 210407 | CM | CMPLT | CMPLT | 2024-10-22 11:13 | 2024-10-24 09:08 | DVIA module not working, according to note | (TC 10/24/2024 09:07:43 AM, Scero, Tomlyn) Diva module was tested on AGM and working fine. All gas delivery tests passed.
1062542 | 210408 | CM | ACTIV | SHP | 2024-10-22 12:56 |  | OBM freezes; sending for repair | (TC 10/23/2024 09:30:20 AM, Baghirzade, Nara) SC 1003335446 (TC 10/29/2024 11:12:56 AM, Baghirzade, Nara) RMA# 10485814 KJV# 461599 PO 7500213680
| 210409 | CM | CMPLT | NFF | 2024-10-22 14:04 | 2024-10-22 14:04 | CAPSA med cart, unable to access med bins (Asset#M31865) | (TC 10/22/2024 02:03:57 PM, Canjimavadivel, Kamallesh ) Request tag stated no power, unable to access med bins. Plugged in cart and battery level was at 50% and charging. Entered PIN and was able to open bins. Returned to 10S
| 210410 | CM | CMPLT | CMPLT | 2024-10-22 14:07 | 2024-10-22 14:07 | Fetal monitor recorder repair | (TC 10/22/2024 02:07:29 PM, Baghirzade, Nara) Worked on repairing the recorder; cleaned and greased the gears. Cleaned the roller. Placed the recorder back in the fetal monitor. Calibration failed the first time. Cleaned the roller and the top again; calibration was successful. Recorder is fully functional. Kept recorder as a spare.
1062542 | 210411 | CM | CMPLT | CMPLT | 2024-10-22 14:12 | 2024-10-25 13:43 | Archiving patient data | (TC 10/22/2024 02:11:36 PM, Baghirzade, Nara) Archiving patient data from OBM to a USB drive before sending the OBM for repair. (TC 10/23/2024 11:14:33 AM, Baghirzade, Nara) Archiving patient data (TC 10/24/2024 10:21:57 AM, Baghirzade, Nara) Continuing to archive patient data (TC 10/24/2024 01:52:50 PM, Baghirzade, Nara) Completed data archiving however one of the patient files corrupted the USB drive. Will place an order for a new USB drive and archive. (TC 10/25/2024 01:22:16 PM, Baghirzade, Nara) Received new USB stick. Archived all patient data. (TC 10/25/2024 01:42:48 PM, Baghirzade, Nara) Followed up with Dr. Deshpande and Dr. Diambomba re: shipping the device and archiving all patient data.
1055944 | 210412 | CM | CMPLT | CMPLT | 2024-10-22 15:09 | 2024-10-22 15:09 | Battery not holding charge | Shipped unit for exchange to Masimo. Received new unit. sn.2000172783. PO# 7500213404. Installed Sinai config, temp probe and adult cuff. Tested. Functional.
| 210413 | RDS | CMPLT | CMPLT | 2024-10-22 15:10 | 2024-10-22 15:18 | Weekly check in ED
| 210414 | RDS | CMPLT | CMPLT | 2024-10-22 15:10 | 2024-10-22 15:13 | Weekly check in ICU
| 210415 | CM | CMPLT | CMPLT | 2024-10-22 17:06 | 2024-10-22 17:06 | Site coverage
502055 | 210416 | CM | ACTIV | ISSUE | 2024-10-22 20:45 |  | MSH 15 SPP PROBE TV FAULT? | (TC 10/22/2024 08:44:53 PM, Whittington, Scott) MSH 15 SPP PROBE TV FAULT? (TC 11/06/2024 11:15:28 AM, Whittington, Scott) MSH 15 SPP PROBE TV FAULT? REF# | MSH 15 SPP PROBE TV FAULT? ROOM C 3535 OLD PROBE SOAK ROOM SN# B0R3MX REF# 0104658689
501585 | 210417 | CM | ACTIV | ISSUE | 2024-10-22 20:48 |  | MSH OPG SPP C HALLWAY END IU22 OVER HEAT | (TC 10/22/2024 08:48:24 PM, Whittington, Scott) MSH OPG SPP C HALLWAY END IU22 OVER HEAT (TC 11/06/2024 02:29:35 PM, Whittington, Scott) MSH OPG SPP C HALLWAY END IU22 OVER HEAT | MSH OPG SPP C HALLWAY END IU22 OVER HEAT ROOM C 3535 OLD PROBE SOAK ROOM
1055898 | 210418 | CM | ACTIV | ISSUE | 2024-10-23 08:29 |  | Missing plug and base broken.
1061333 | 210419 | CM | CMPLT | CMPLT | 2024-10-23 09:30 | 2024-10-23 09:31 | QC checklist | (TC 10/23/2024 09:30:32 AM, Alvarado, Karla) QC checklist completed. Returned to 11N
1055898 | 210420 | CM | CMPLT | CMPLT | 2024-10-23 09:37 | 2024-10-23 09:37 | Missing plug and base broken. | (TC 10/23/2024 09:36:54 AM, Alvarado, Karla) Replaced the rolling stand and mount. Performed tests: Temperature NIBP = set: 120/80, measured: 118/84 SpO2 = set: 95%, measured: 95% All test passed.
1057279 | 210421 | ACC | CMPLT | CMPLT | 2024-10-23 10:19 | 2024-10-23 10:43 | Received replacement Root | (TC 10/23/2024 10:42:40 AM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
1066462 | 210422 | PA | CMPLT | CMPLT | 2024-10-23 10:48 | 2024-10-28 13:13 | Ordering parts for repair | (TC 10/23/2024 10:59:04 AM, Baghirzade, Nara) SC 1003335609 (TC 10/28/2024 01:13:20 PM, Baghirzade, Nara) PO 7202476734 | Defective power supply and solution hanger
65894 | 210423 | CM | CMPLT | CMPLT | 2024-10-23 10:51 | 2024-10-23 10:52 | Channel error | (TC 10/23/2024 10:52:21 AM, Rhuland, Courtney) Channel error 351.6740 present Unable to replicate error Ran multiple test infusions and performed check-in All tests passed
65877 | 210424 | CM | CMPLT | CMPLT | 2024-10-23 11:07 | 2024-10-23 12:07 | Channel error | (TC 10/23/2024 12:06:40 PM, Rhuland, Courtney) Channel error 351.6740 present in error logs Unable to replicate error Ran check-in and multiple test infusions All tests passed
| 210425 | PA | CMPLT | CMPLT | 2024-10-23 11:08 | 2024-10-23 11:17 | Ordering temperature probe and well parts for stock | (TC 10/23/2024 11:09:10 AM, Baghirzade, Nara) SC 1003335663 (TC 10/23/2024 11:17:16 AM, Baghirzade, Nara) Someone else had already ordered parts fort stock. Cancelled shopping cart.
1055948 | 210426 | CM | CMPLT | CMPLT | 2024-10-23 11:09 | 2024-10-23 11:10 | BP inflation malfunction. | (TC 10/23/2024 11:09:34 AM, Alvarado, Karla) Replaced temperature probe. Replaced blood pressure tube. Performed tests: Temperature NIBP= set: 120/80, measured: 117/84 SpO2 = set: 95%, measured: 95% All test passed.
1056494 | 210427 | CM | CMPLT | CMPLT | 2024-10-23 11:39 | 2024-10-23 11:39 | Battery replacement. | (TC 10/23/2024 11:39:18 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056007 | 210428 | CM | CMPLT | CMPLT | 2024-10-23 11:40 | 2024-10-23 11:40 | Battery replacement. | (TC 10/23/2024 11:40:19 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056531 | 210429 | CM | CMPLT | CMPLT | 2024-10-23 11:41 | 2024-10-23 11:41 | Battery replacement. | (TC 10/23/2024 11:41:09 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
66494 | 210430 | CM | CMPLT | NFF | 2024-10-23 13:27 | 2024-10-23 13:27 | Reported that its not working | (TC 10/23/2024 01:27:14 PM, Baghirzade, Nara) Did not find any issue with the pump. Tested another unit and it is functioning the same. Was able to build up max pressure and release it without issue.
1054981 | 210431 | CM | CMPLT | CMPLT | 2024-10-23 13:57 | 2024-11-05 09:27 | Does not connect to Radical/sending for repair | (TC 10/23/2024 01:56:56 PM, Baghirzade, Nara) Root does not detect Radical when connected; power cycled and tried with other Radical devices but the Root does not detect the Radical devices. Requested service from Masimo. (TC 10/24/2024 10:19:26 AM, Baghirzade, Nara) SC 1003336546 (TC 11/05/2024 09:26:38 AM, Baghirzade, Nara) PO 7500213824
69291 | 210432 | CM | ACTIV | ISSUE | 2024-10-23 14:18 |  | it is not working. | (TC 10/23/2024 02:18:01 PM, Alvarado, Karla) Downloaded error log: sensor broken. Performed analog sensor test and pressure calibration but the calibration failed. Replaced upper sensor. Performed pressure calibration, rate calibration (twice), and check-in test. All tests passed
66431 | 210433 | CM | CMPLT | CMPLT | 2024-10-23 14:46 | 2024-10-28 14:07 | not working | (TC 10/25/2024 09:47:49 AM, Scero, Tomlyn) Device dropping it's WiFi signal at times and losing connection to the network. Kamran suggested upgrading the SW. Upgraded SW from B.06.59 to B.06.91 (TC 10/28/2024 09:49:03 AM, Scero, Tomlyn) admitting "biomed" patient to the tele pack to I could monitor a patient on the simulator for the day and check that the signal wasn't dropped. (TC 10/28/2024 02:07:22 PM, Scero, Tomlyn) Device tested all day and log were reviewed. No data was missing and signal seemed to stay consistant all day. Battery adaptor was replaced as well.
32510 | 210434 | CM | ACTIV | ISSUE | 2024-10-23 20:21 |  | MSH DR1 GENERATOR ERRORS , POWER , POSSIBLE D160 BD FAULT, NO POWER OUT TO PRECISION 5V SUPPLY , ETC | (TC 10/23/2024 08:20:52 PM, Whittington, Scott) MSH DR1 GENERATOR ERRORS , POWER , POSSIBLE D160 BD FAULT, NO POWER OUT TO PRECISION 5V SUPPLY , ETC (TC 10/24/2024 04:44:23 PM, Whittington, Scott) MSH DR1 GENERATOR ERRORS , POWER , POSSIBLE D160 BD FAULT, NO POWER OUT TO PRECISION 5V SUPPLY , ETC (TC 10/24/2024 04:44:36 PM, Whittington, Scott) MSH DR1 GENERATOR ERRORS , POWER , POSSIBLE D160 BD FAULT, NO POWER OUT TO PRECISION 5V SUPPLY , ETC REPLACED D160, SNT (FAILED), MONITOR (FAILED) | MSH DR1 GENERATOR ERRORS , POWER , POSSIBLE D160 BD FAULT, NO POWER OUT TO PRECISION 5V SUPPLY , ETC
1055039 | 210435 | CM | CMPLT | CMPLT | 2024-10-24 08:10 | 2024-10-24 08:10 | BP not inflating | (TC 10/24/2024 08:10:16 AM, Ferreira, Qwynn) Cuff was worn and tubing was cracked swapped with different hose and cuff and was able to inflate. Old cuff had a hole near the connector leaking some air out Tested and inflated as normal Sent back
1064418 | 210436 | CM | ACTIV | ISSUE | 2024-10-24 08:24 |  | Main door broken by right latch | (TC 10/23/2024 08:29:00 AM, Ferreira, Qwynn) Met with Michael from AME to discus the break and other issues related to the MRI. Took photos of the broken corner as well as the none working LED. May need to purchase a new hood for the device
| 210437 | PA | CMPLT | CMPLT | 2024-10-24 08:44 | 2024-10-28 13:57 | Ordering loctite for OR table clips | (TC 10/24/2024 08:47:20 AM, Baghirzade, Nara) SC 1003336329 (TC 10/28/2024 01:57:04 PM, Baghirzade, Nara) PO 7202477826
1056426 | 210438 | CM | CMPLT | CMPLT | 2024-10-24 08:46 | 2024-10-24 08:46 | It is not working. | (TC 10/24/2024 08:46:18 AM, Alvarado, Karla) Added a power cord. Performed tests: Temperature NIBP= set: 120/80, measured: 115/83 SpO2 = set: 95%, measured: 95% All test passed.
66816 | 210439 | CM | CMPLT | NFF | 2024-10-24 08:47 | 2024-10-24 08:47 | BP malfunction. | (TC 10/24/2024 08:47:26 AM, Alvarado, Karla) Performed NIBP tests: set: Adult 120/80, measured: 117/79, 118/77. set: Low 80/40, measured: 81/40. set: Neo 70/40, measured: 71/44 All test passed. No Fault Found (NFF).
1056473 | 210440 | CM | CMPLT | CMPLT | 2024-10-24 08:51 | 2024-10-29 08:33 | battery not hold charge | shipping Masimo Root for exchange (TC 10/29/2024 08:32:51 AM, Baghirzade, Nara) RMA# 47110 KJV# 461542 PO 7500213823
1050476 | 210441 | CM | CMPLT | CMPLT | 2024-10-24 08:56 | 2024-11-18 11:16 | ordering Philips dongles for Servo-Ns | we received 10 x Philips dongles. PO#7202477829. delivered to Rex.
1060149 | 210442 | CM | CMPLT | CMPLT | 2024-10-24 10:00 | 2024-10-25 09:15 | patient data cable not holding | (TC 10/25/2024 09:15:24 AM, Scero, Tomlyn) Replaced keyboard cover that was ripped. LL not working, replaced LL lead. Patient data cable wasn't clicking into the PIM properly. We don't have any new cables in stock currently so I gave them an old cable that works for now until they arrive.
1056436 | 210443 | CM | CMPLT | NFF | 2024-10-24 11:04 | 2024-10-24 11:04 | When measuring BP, it displays "time cut off" | (TC 10/24/2024 11:04:04 AM, Baghirzade, Nara) When booting up the Root, noticed that the Radical was not detected. Re-seated Radical and then measured the BP. Did not find the error that the staff reported. Shut down both Radical and Root; booted up again and was able to measure BP without error. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 210444 | PA | CMPLT | CMPLT | 2024-10-24 11:30 | 2024-11-21 09:41 | Ordering SpO2 Cables for PACU | (TC 11/21/2024 09:40:52 AM, Reid, Emmerson ) Received the order
| 210445 | CM | CMPLT | CMPLT | 2024-10-24 11:54 | 2024-10-24 11:55 | Philips L&D Impact Meeting | (TC 10/24/2024 11:54:37 AM, Baghirzade, Nara) Philips L&D Impact Meeting
502044 | 210446 | PM | CMPLT | CMPLT | 2024-10-24 11:59 | 2024-10-24 11:59 | PM - Semi-Annual | (TC 10/24/2024 11:59:07 AM, Green, Paul) Replaced laptop. Performed PM on system. Topped up alcohol based coolant. Removed and regreased up/down achme screw and linear bearings. Cleaned system and ran uniformity corrections.
| 210447 | PA | CMPLT | CMPLT | 2024-10-24 14:56 | 2024-10-24 15:01 | Probe well for SureTemp | (TC 10/24/2024 03:01:00 PM, Canjimavadivel, Kamallesh ) Called vendor and received quote, placed order for probe wells
22692 | 210448 | PM | CMPLT | CMPLT | 2024-10-25 07:46 | 2024-11-08 12:26 | PM-MAJOR - Annual PM For: Oxygen-Air Proportioners | (TC 11/08/2024 12:26:16 PM, Rhuland, Courtney) PM completed by Spectramedx Returned and functional check completed 21% O2 - 20.9% O2 40% O2 - 39.2% O2 60% O2 - 60.6% O2 80% O2 - 80.2% O2 100% O2 - 99.9% O2 Returned to RT
68377 | 210449 | CM | CMPLT | CMPLT | 2024-10-25 07:59 | 2024-10-24 09:42 | PATO sensor failed | (TC 10/24/2024 09:42:00 AM, Scero, Tomlyn) Confirmed PATO sensor defective. No O2 reading and PGA cannot be calibrated. Replaced PATO sensor and self test passed. Confirmed ok during gas tests.
68799 | 210450 | CM | CMPLT | NFF | 2024-10-25 08:17 | 2024-10-25 08:17 | Channel error – lad beeping | (TC 10/25/2024 08:17:06 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage = 70.4uA and the GND Resistance = 120 mOhms Download error log: did not find errors. All test passed.
| 210451 | OC | CMPLT | CMPLT | 2024-10-25 08:24 | 2024-10-31 08:00 | On-call week Oct 25-31 | (TC 10/25/2024 08:24:01 AM, Canjimavadivel, Kamallesh ) On-call daily check report (TC 10/25/2024 07:26:00 PM, Canjimavadivel, Kamallesh ) Call from pharmacy Friday Oct 25, 7:26 PM. They had issue with their incubator and were directed from building services to contact us. I verified incubator was not for neonatal use and was used for lab samples. Informed them this equipment is not our responsibility and that they should reach out to a vendor or manufacturer for support. (TC 10/27/2024 01:43:00 AM, Canjimavadivel, Kamallesh ) Call from 14N Sunday Oct 27 at 1:43 AM. There was an issue with a stryker bed where the controls only worked on one rail and not the other or footboard. Asked if spare bed was available and nurse said there was one available. Asked them to use spare bed and to leave note on bed and set aside as biomed would further assess on Monday. (TC 10/27/2024 01:43:00 AM, Canjimavadivel, Kamallesh ) Call from 14N Sunday Oct 27 at 1:43 AM. There was an issue with a stryker bed where the controls only worked on one rail and not the other or footboard. Asked if spare bed was available and nurse said there was one available. Asked them to use spare bed and to leave note on bed and set aside as biomed would further assess on Monday. Did not need to come on site (TC 10/28/2024 09:23:20 AM, Canjimavadivel, Kamallesh ) On-call daily report and filled out logbook for weekend calls (TC 10/29/2024 08:47:06 AM, Canjimavadivel, Kamallesh ) On-call daily report (TC 10/30/2024 08:00:00 AM, Canjimavadivel, Kamallesh ) On-call daily report (TC 10/31/2024 08:00:00 AM, Canjimavadivel, Kamallesh ) On-call daily report
| 210452 | PA | CMPLT | CMPLT | 2024-10-25 08:33 | 2024-11-14 10:15 | Nitronox replacement parts | (TC 10/25/2024 08:36:47 AM, Baghirzade, Nara) SC 1003337415 (TC 10/29/2024 11:20:59 AM, Baghirzade, Nara) PO 7500213849 1 item on backorder: Nitronox demand valve (TC 11/14/2024 10:14:46 AM, Baghirzade, Nara) Received full order. PO 7500213849
| 210453 | CM | ACTIV | ISSUE | 2024-10-25 10:51 |  | MSH SPP E2- Setup for Mindray i9 Demo | Setup PACs, WL, SR, Wifi
32520 | 210454 | CM | ACTIV | ISSUE | 2024-10-25 11:18 |  | MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT | (TC 10/25/2024 11:17:44 AM, Whittington, Scott) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT PLAN TO REPALCE DETECTOR ARM (TC 10/28/2024 03:56:25 PM, Whittington, Scott) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT (TC 10/28/2024 03:56:36 PM, Whittington, Scott) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT (TC 10/28/2024 03:57:10 PM, Heron, Chris) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT (TC 10/28/2024 03:57:59 PM, Dolinar, John) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT (TC 10/29/2024 04:53:47 PM, Whittington, Scott) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT (TC 10/29/2024 04:54:10 PM, Heron, Chris) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT (TC 10/29/2024 04:54:45 PM, Dolinar, John) MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT | MSH DR2 DETECTOR BETA MOVEMENT FAULTY; GEARBOX OIL HAS LEAKED OUT
62151 | 210455 | CM | CMPLT | CMPLT | 2024-10-25 13:46 | 2024-10-25 13:46 | Moved to 6th fl. storage | (TC 10/25/2024 01:46:12 PM, Baghirzade, Nara) Followed up with Naz as we are waiting for a storage solution for the EEG data. Confirmed with Naz to move the OBM to 6th fl. storage area. Placed sign on it stating "do not use, do not move from storage area".
62152 | 210456 | CM | CMPLT | CMPLT | 2024-10-25 13:50 | 2024-10-25 13:50 | Moved to 6th fl. storage | (TC 10/25/2024 01:50:22 PM, Baghirzade, Nara) Followed up with Naz as we are waiting for a storage solution for the EEG data. Confirmed with Naz to move the OBM to 6th fl. storage area. Placed sign on it stating "do not use, do not move from storage area".
| 210457 | RDS | CMPLT | CMPLT | 2024-10-25 13:52 | 2024-10-25 13:52 | Weekly Rounds: 11 North | (TC 10/25/2024 01:52:27 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 210458 | RDS | CMPLT | CMPLT | 2024-10-25 13:54 | 2024-10-25 13:54 | Weekly Rounds: 11 South | (TC 10/25/2024 01:54:06 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
69067 | 210459 | CM | CMPLT | CMPLT | 2024-10-25 13:58 | 2024-10-25 13:58 | it is not working | (TC 10/25/2024 01:58:15 PM, Alvarado, Karla) Download error log: Bottle pressure sensor broken. Replaced module latch. Replaced front case. Replaced upper pressure sensor. While the pressure calibration was being performed, the low pressure sensor failed. Therefore, low pressure sensor was also replaced. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. All test passed.
68962 | 210460 | CM | CMPLT | NFF | 2024-10-25 14:27 | 2024-10-25 14:27 | it is not working. | (TC 10/25/2024 02:27:17 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage = 87.7 uA and the GND Resistance = 137 mOhms Download error log: did not find errors. All test passed.
1048158 | 210461 | CM | CMPLT | CMPLT | 2024-10-28 08:25 | 2024-10-28 08:25 | it is not working. | (TC 10/28/2024 08:25:16 AM, Alvarado, Karla) Replaced battery.
| 210462 | RDS | CMPLT | CMPLT | 2024-10-28 09:04 | 2024-10-28 09:05 | Weekly checks in CCU | (TC 10/28/2024 09:04:33 AM, Scero, Tomlyn) Weekly checks in CCU complete. 30J tests passed. Showed Karla around the CCU as well
66433 | 210463 | CM | CMPLT | CMPLT | 2024-10-28 09:47 | 2024-10-28 14:09 | Unit does not turn on | (TC 10/28/2024 02:08:37 PM, Scero, Tomlyn) Battery adaptor was replaced and new batteries installed. Ran on a simulator and didn't find any issues with the device.
| 210464 | RDS | CMPLT | CMPLT | 2024-10-28 10:26 | 2024-10-28 10:29 | Weekly check for ED
| 210465 | RDS | CMPLT | CMPLT | 2024-10-28 10:27 | 2024-10-28 14:55 | Weekly check for ICU
1069392 | 210466 | CM | CMPLT | CMPLT | 2024-10-28 10:30 | 2024-10-28 10:31 | Requires maintenance message on screen | (TC 10/28/2024 10:30:34 AM, Canjimavadivel, Kamallesh ) Was able to override message. Changed PM schedule so message will not be present. PM will be done next month
1055691 | 210467 | PA | CMPLT | CMPLT | 2024-10-28 10:36 | 2024-11-18 12:53 | Broken; requesting replacement unit under consignment | (TC 10/28/2024 10:41:27 AM, Baghirzade, Nara) SC 1003338635 (TC 10/30/2024 01:29:18 PM, Baghirzade, Nara) PO 7500213927 (TC 10/30/2024 01:29:35 PM, Baghirzade, Nara) Attached test certificate. Aleks from Stryker will drop off the roll stand for the mistral warmer. (TC 11/18/2024 12:53:15 PM, Baghirzade, Nara) Stand was dropped off by Aleks. Mounted unit to stand and sent to main OR.
| 210468 | PA | CMPLT | CMPLT | 2024-10-28 10:52 | 2024-11-04 07:40 | Ordering PCU backup speakers for stock | (TC 10/28/2024 11:04:29 AM, Baghirzade, Nara) SC 1003338658 (TC 11/04/2024 07:40:12 AM, Baghirzade, Nara) PO 7202478769
1054769 | 210469 | CM | CMPLT | NFF | 2024-10-28 11:11 | 2024-10-28 11:11 | Blood pressure not reading | (TC 10/28/2024 11:10:56 AM, Rhuland, Courtney) Tested blood pressure functionality with both cuffs that were sent with device No issues after taking multiple readings Functional check - OK
32510 | 210470 | CM | ACTIV | ISSUE | 2024-10-28 11:30 |  | MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT | (TC 10/28/2024 11:30:33 AM, Whittington, Scott) MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT (TC 10/28/2024 11:30:33 AM, Whittington, Scott) MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT (TC 10/28/2024 03:55:27 PM, Heron, Chris) MSH DR1 TABLE ERROR (TC 11/01/2024 12:33:37 PM, Whittington, Scott) MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT (TC 10/29/2024 08:33:00 PM, Whittington, Scott) MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT (TC 10/29/2024 08:33:00 PM, Whittington, Scott) MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT (TC 11/01/2024 08:33:00 AM, Whittington, Scott) MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT APPEARSR TO BE A SITUATION WITH THE EMERGENCY STOP WIRING OR CIRCUITRY ABLE TO MOVE TABLE VIA ARISTOS SERVICE SW' BUT EMERGENCY STOP ACTIVATED ERROR APPEARS IN CLINICAL MODE. SWITCHES ARE BOTH IN OPERATE POSITION (TC 11/01/2024 08:33:00 AM, Whittington, Scott) MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT APPEARSR TO BE A SITUATION WITH THE EMERGENCY STOP WIRING OR CIRCUITRY ABLE TO MOVE TABLE VIA ARISTOS SERVICE SW' BUT EMERGENCY STOP ACTIVATED ERROR APPEARS IN CLINICAL MODE. SWITCHES ARE BOTH IN OPERATE POSITION | MSH DR1 TABLE ERROR 30168 APID 22 TARGET & POT DISTANCE TOO GREAT
66045 | 210471 | CM | CMPLT | CMPLT | 2024-10-28 12:00 | 2024-10-28 12:00 | System Error | (TC 10/28/2024 12:00:14 PM, Alvarado, Karla) Turn on the monitor and the screen displayed Error code: 133.6080. According to the manual, the backup speaker failed. The backup speaker was replaced. Performed tests in ASM: Check- in tests. Where the GND Leakage = 82.8 uA and the GND Resistance = 107 mOhms All test passed. Backup speaker is fully functional.
68646 | 210472 | CM | CMPLT | NFF | 2024-10-28 12:02 | 2024-10-28 12:02 | System Error | (TC 10/28/2024 12:02:04 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Download error log: did not find errors. All test passed. No fault find
67890 | 210473 | ACC | CMPLT | CMPLT | 2024-10-28 13:09 | 2024-10-28 13:09 | Incoming and Functional check | (TC 10/28/2024 01:08:52 PM, Baghirzade, Nara) Received the unit from vendor repair. Uploaded the L&D Epidural library. Printed and placed asset tag. Completed functional checks. See procedure for results.
1061333 | 210474 | CM | CMPLT | CMPLT | 2024-10-28 13:14 | 2024-10-28 13:14 | QC checklist | (TC 10/28/2024 01:13:56 PM, Alvarado, Karla) QC checklist completed. Returned to 11N
65951 | 210475 | CM | CMPLT | CMPLT | 2024-10-28 13:23 | 2024-10-28 13:23 | Reported that its not pumping enough pressure | (TC 10/28/2024 01:23:27 PM, Baghirzade, Nara) Front and rear case were not aligning properly; reassembled the front/rear case. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional. No fault found.
65981 | 210476 | CM | CMPLT | CMPLT | 2024-10-28 13:25 | 2024-10-28 13:26 | Channel error 351.6770 | (TC 10/28/2024 01:25:43 PM, Baghirzade, Nara) Pressure sensor disc was previously re-seated and is found to be defective. On physical inspection, found that the front case was broken. Replaced: -front case -mylar gasket -pressure disc flag -pressure sensor disc Completed full calibration of: -barrel clamp -plunger force -plunger position -pressure disc calibration Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional
1066890 | 210477 | CM | CMPLT | CMPLT | 2024-10-28 13:30 | 2024-10-28 13:30 | Occlusion alarms | (TC 10/28/2024 01:30:04 PM, Baghirzade, Nara) User reported that there are occlusion alarms even with no occlusion present. Ran the pump all day on Oct 24th and 25th at varying rates; did not find any issues with the pump. Ran the pump Oct. 28th; did not find any issues. No fault found. Pump is functional.
| 210478 | CM | CMPLT | CMPLT | 2024-10-28 14:09 | 2024-10-28 14:09 | Stryker Lifenet meeting | (TC 10/28/2024 02:09:18 PM, Baghirzade, Nara) Stryker Lifenet meeting
1059828 | 210479 | PM | CMPLT | CMPLT | 2024-10-28 14:49 | 2024-10-28 14:50 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 10/28/2024 02:49:36 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1061217 | 210480 | CM | CMPLT | CMPLT | 2024-10-28 15:40 | 2024-10-28 15:45 | USB Export issue | Call re: image transfer issue. Went to check. Waited for access. Spoke with surgeon and was requested to transfer patient images into USB. USB export not enabled. Advised to arrange for another USB. Still same issue. Contacetd olympus. Was informed to format USB into NTFS format. Formatted USB. Waited for them to finish procedure. Completed export. Dropped off USB to the surgeon. Tested ok.
1046396 | 210481 | CM | CMPLT | CMPLT | 2024-10-28 16:03 | 2024-10-28 16:07 | Not working | Found unit was retired in TMS but was being used in OR 16. Call re: not working. Went to check. Waited for access. FOund unit was out of sync and filter was loose. Fixef it. Synched unit with ESU. Checked fuctionality. Tested ok.
1063331 | 210482 | CM | ACTIV | ISSUE | 2024-10-28 16:04 |  | MSH OPG 3 SPP C3 PROBE NOT ACTIVATING | (TC 10/28/2024 04:05:57 PM, Whittington, Scott) MSH OPG 3 SPP C3 PROBE NOT ACTIVATING | MSH OPG 3 SPP C3 PROBE NOT ACTIVATING
1046381 | 210483 | CM | CMPLT | CMPLT | 2024-10-28 16:08 | 2024-10-28 16:08 | Error EC 204 |  | Error EC 204 Reset the transducer connection. Tested OK
| 210484 | CM | CMPLT | CMPLT | 2024-10-29 09:16 | 2024-11-14 13:23 | OR table remote clips | (TC 10/29/2024 09:16:08 AM, Baghirzade, Nara) Tightened all 3 screws on remote for OR table in OR1. (TC 10/29/2024 10:36:37 AM, Baghirzade, Nara) Clip and screws in OR4 is missing entirely; requested quote from vendor. OR5 clip screws also missing. (TC 11/06/2024 10:48:34 AM, Baghirzade, Nara) Multiple remotes had screw holds that were broken off and found inside the remotes. Disassembled the remotes and fixed the screw holds to the remote. Replaced the clips. OR 2, 3, and 4 completed. (TC 11/14/2024 01:22:56 PM, Baghirzade, Nara) Replaced OR5 clip and any other clips that were broken or loose.
| 210485 | CM | CMPLT | CMPLT | 2024-10-29 10:17 | 2024-10-29 10:18 | Stethoscope repair | (TC 10/29/2024 10:17:36 AM, Rhuland, Courtney) Medical education (60 Murray) reached out for repair of electronic stethoscope Picked up device and re-soldered new battery cable to board Functional check - OK unable to repair 3 others missing ear pieces
| 210486 | CM | ACTIV | ISSUE | 2024-10-29 10:20 |  | Boom PM work | (TC 10/29/2024 10:19:37 AM, Rhuland, Courtney) Removed all equipment from SIMS lab to prepare for Drager to complete boom PM in room (TC 11/15/2024 09:52:22 AM, Rhuland, Courtney) Removed equipment from room 11 Replaced equipment in procedure room (TC 11/26/2024 11:49:18 AM, Rhuland, Courtney) Put equipment back into room 11
1054432 | 210487 | FC | CMPLT | CMPLT | 2024-10-29 10:36 | 2024-10-29 10:36 | Functional testing | (TC 10/29/2024 10:35:58 AM, Baghirzade, Nara) No physical damage or issues. Test 1 - <100 mmHg Test 2 - >300 mmHg Test 3 - >500 mmHg Test 4 - >450 mmHg
1055599 | 210488 | CM | CMPLT | CMPLT | 2024-10-29 10:37 | 2024-10-29 10:38 | Not charging | (TC 10/29/2024 10:37:49 AM, Canjimavadivel, Kamallesh ) Radical was needed to transport patient. Was locked into Root and did not have a battery. Installed battery and removed locking tab.
1058058 | 210489 | CM | CMPLT | CMPLT | 2024-10-29 12:22 | 2024-11-04 14:38 | C700 display not working properly | (TC 10/29/2024 03:02:30 PM, Scero, Tomlyn) Emailed Draeger after looking through the service connect documents. Took pictures and sent them to tech support. They agreed with me and said it looks like a backlight issue and try replacing the LC display. removed the display from the the AGM for further inspection (TC 11/04/2024 02:38:25 PM, Scero, Tomlyn) LC display was received and replaced. After I reassembled the monitor I put it back on the device to confirm. Monitor is working perfectly. Device put into service mode to calibrate the touch screen. Pm can now be continued.
1055591 | 210490 | CM | CMPLT | CMPLT | 2024-10-29 12:47 | 2024-10-29 12:47 | Missing BP connection. | (TC 10/29/2024 12:47:15 PM, Alvarado, Karla) Replaced blood pressure tube. Performed tests: Temperature NIBP= set: 120/80, measured: 115/82 SpO2 = set: 95%, measured: 95% All test passed.
1063259 | 210491 | PM | CMPLT | CMPLT | 2024-10-29 13:03 | 2024-11-06 12:44 | PM-MAJOR - Annual PM For: Phototherapy Units, Visible Light, Hyperbilirubinemia | (TC 11/06/2024 12:44:26 PM, Alvarado, Karla) Replaced case. Tests results: - Light Intensity small pad: 77.6 µW•cm-2•nm-1 - Hours: 1978.4 hours - Electrical safety results - Leakage current = 10.4 uA; GND Resistance = 268 mohms All tests passed. PM completed.
| 210492 | PA | ACTIV | ISSUE | 2024-10-29 13:39 |  | Ordering respiratory monitoring accessories for stock | (TC 10/29/2024 01:53:29 PM, Baghirzade, Nara) SC 1003340228 (TC 11/04/2024 07:24:33 AM, Baghirzade, Nara) Received replacement tapes and neo sensors only. PO 7202480184 (TC 11/07/2024 10:34:45 AM, Baghirzade, Nara) Received SpO2 sensors (QTY 6). PO 7202479880 (TC 11/11/2024 07:33:23 AM, Baghirzade, Nara) Received 4 RAS-45 sensors. PO 7202480184
| 210493 | CM | CMPLT | CMPLT | 2024-10-29 14:00 | 2024-10-29 14:00 | Laser safety meeting | (TC 10/29/2024 02:00:17 PM, Baghirzade, Nara) Laser safety meeting
| 210494 | CM | CMPLT | CMPLT | 2024-10-29 14:34 | 2024-10-29 14:34 | Checking equipment for PMs | (TC 10/29/2024 02:34:07 PM, Baghirzade, Nara) Checking equipment that requires PMs on antenatal unit
| 210495 | CM | CMPLT | CMPLT | 2024-10-29 14:57 | 2024-11-01 13:24 | Draeger Zeus repair parts | (TC 10/29/2024 02:59:12 PM, Scero, Tomlyn) called Draeger support to ask for product price updates and then ordered parts. (TC 11/01/2024 01:24:18 PM, Scero, Tomlyn) Parts were received.
| 210496 | PA | CMPLT | CMPLT | 2024-10-30 08:15 | 2024-11-04 07:23 | Ordering clips for OR table remotes | (TC 10/30/2024 08:18:25 AM, Baghirzade, Nara) SC 1003340636 (TC 11/04/2024 07:22:43 AM, Baghirzade, Nara) PO 7202480093
| 210497 | PA | CMPLT | CMPLT | 2024-10-30 08:40 | 2024-11-04 07:41 | Ordering LVP pressure sensors for stock | (TC 10/30/2024 08:43:35 AM, Baghirzade, Nara) SC 1003340682 (TC 11/04/2024 07:40:58 AM, Baghirzade, Nara) PO 7202480163
63027 | 210498 | PA | CMPLT | CMPLT | 2024-10-30 10:16 | 2024-11-05 13:31 | Ordering items for repair | (TC 10/30/2024 10:22:16 AM, Baghirzade, Nara) SC 1003340832 (TC 11/05/2024 01:30:40 PM, Baghirzade, Nara) PO 7202482146
1054987 | 210499 | CM | CMPLT | CMPLT | 2024-10-30 10:51 | 2024-10-30 10:51 | Respiratory monitoring accessories | (TC 10/30/2024 10:51:22 AM, Baghirzade, Nara) Dropped off: 1 box of RAS-45 1 box of RAS-125C 1 box of LNCS Neo SpO2
1056453 | 210500 | CM | CMPLT | CMPLT | 2024-10-30 10:55 | 2024-10-30 10:55 | Respiratory monitoring accessories | (TC 10/30/2024 10:54:44 AM, Baghirzade, Nara) Dropped off: 1 box of RAS-45 1 box of RAS-125C 1 box of LNCS Neo SpO2 2 boxes of LNCS DCI SpO2
| 210501 | CM | CMPLT | CMPLT | 2024-10-30 10:58 | 2024-10-30 10:58 | Water bath analyzer voltage verification | (TC 10/30/2024 10:58:01 AM, Baghirzade, Nara) Blood bank received a new water bath analyzer; Melanie requested to check the voltage and amperage for their audit. Measured voltage and amperage with fluke electrical safety analyzer and filled the one part of the form for this. Plugged the device in emergency power outlet.
1059778 | 210502 | FC | CMPLT | CMPLT | 2024-10-30 13:37 | 2024-10-30 13:37 | QC checklist | (TC 10/30/2024 01:36:48 PM, Baghirzade, Nara) Noticed that the case was not properly aligning; aligned the case so its properly sealed. QC checklist completed. Returned to 11 South.
1055925 | 210503 | CM | CMPLT | NFF | 2024-10-30 13:43 | 2024-10-30 13:43 | Reported not working | (TC 10/30/2024 01:43:11 PM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional. No fault found.
65952 | 210504 | CM | CMPLT | CMPLT | 2024-10-30 13:44 | 2024-10-30 13:44 | Missing barcode sticker | (TC 10/30/2024 01:43:51 PM, Baghirzade, Nara) Printed and reattached barcode sticker. Check-in tests passed and completed. IUI connectors and device cleaned. Device is functional.
69326 | 210505 | CM | CMPLT | CMPLT | 2024-10-30 13:45 | 2024-10-30 13:45 | Channel error when connecting to PCU | (TC 10/30/2024 01:45:14 PM, Baghirzade, Nara) Downloaded error log; found bottle-side pressure sensor failure errors. When doing the analog test, found that the patient-side pressure sensor is also defective. Repalced both pressure sensors. During check-in test, noticed that the pump was making very loud noises. Diassembled and found that the motor was not sitting properly, causing the gears to grind. Re-seated the motor and added a little bit of grease. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
67965 | 210506 | CM | CMPLT | CMPLT | 2024-10-30 14:07 | 2024-10-30 14:07 | Scale not recognized | (TC 10/30/2024 02:07:29 PM, Rhuland, Courtney) Sclae does not show up when plugged in Tried changing scale and did not show up Popped cable out from front pannel and tested continuity Reseated pins and tightened connections and all pins had continuity Tested scale connection multiple times and scale continued to be recognized Functional check - OK
1061211 | 210507 | CM | ACTIV | ISSUE | 2024-10-30 14:49 |  | pinpoint camera head not working | (TC 10/30/2024 02:52:01 PM, Reid, Emmerson ) SC 1003341355
| 210508 | CM | CMPLT | CMPLT | 2024-10-30 15:55 | 2024-10-30 15:56 | Suction tree Repair | Broekn. Went to check and brought it down to biomed. Found suction regulator cracked. Disposed it. Found tubing and filter contaminated with liquid residue. Removed everything and disposed it. Repalced new tubing and canister, Replaced suction regulator and tightned it. Checked functionality. Tested ok. Returned to OR.
1048500 | 210509 | CM | CMPLT | CMPLT | 2024-10-30 17:28 | 2024-11-06 11:13 | MSH 5 US EPIQ WILL NOT AWAKEN AFTER PUTTING INTO SLEEP MODE BETWEEN PT IN NICU, SYSEM HAS TO REBOOT, AND WIRELESS HAS TO BE RESLECTED | (TC 10/30/2024 05:32:18 PM, Whittington, Scott) MSH 5 US EPIQ WILL NOT AWAKEN AFTER PUTTING INTO SLEEP MODE BETWEEN PT IN NICU, SYSEM HAS TO REBOOT, AND WIRELESS HAS TO BE RESLECTED CHECKED BATTERY, IT IS OLDER JULY 2018 BATTERY HAS FULL CHAHARGE DOWNLOAD LOGS (TC 11/06/2024 11:12:46 AM, Whittington, Scott) MSH 5 US EPIQ WILL NOT AWAKEN AFTER PUTTING INTO SLEEP MODE BETWEEN PT IN NICU, SYSEM HAS TO REBOOT, AND WIRELESS HAS TO BE RESLECTED NO RECURRENCE OF ISSUE | MSH 5 US EPIQ WILL NOT AWAKEN AFTER PUTTING INTO SLEEP MODE BETWEEN PT IN NICU, SYSEM HAS TO REBOOT, AND WIRELESS HAS TO BE RESLECTED
1049905 | 210510 | CM | ACTIV | ISSUE | 2024-10-30 18:14 |  | MSH R65 PC FAULT | (TC 10/04/2024 08:14:00 PM, Whittington, Scott) MSH R65 PC FAULT | MSH R65 PC FAULT
1054521 | 210511 | CM | CMPLT | CMPLT | 2024-10-31 08:38 | 2024-10-31 09:20 | Not working | (TC 10/31/2024 09:19:33 AM, Canjimavadivel, Kamallesh ) Replaced V1-V3 lead set and RL lead. Tested okay with simulator
1062010 | 210512 | PM | ACTIV | READY | 2024-10-31 10:01 |  | PM-MAJOR_Annual PM For: Med Surg Bed
1062020 | 210513 | PM | ACTIV | READY | 2024-10-31 10:04 |  | Batteries Replacement Due
1064356 | 210514 | CM | CMPLT | CMPLT | 2024-10-31 10:08 | 2024-10-31 10:08 | Not working | (TC 10/31/2024 10:08:10 AM, Canjimavadivel, Kamallesh ) Replaced V1-V3 chest leads set
66242 | 210515 | CM | ACTIV | ISSUE | 2024-10-31 11:12 |  | Welch Allyn Service Tool Installation | Oct 31 - Installed Welch Allyn Service Tool after completing the WA CVSM Training. Received Gold license activation code but could not activate it due to error. Went through instructions but issue persisted. Contacted WA tech support. Waited on hold for half an hour. Was advised to install partner connect first. Awaiting instruction from tech support.
1059796 | 210516 | CM | CMPLT | CMPLT | 2024-10-31 11:14 | 2024-10-31 11:14 | No image in the reference monitor. | Fixed.
1047859 | 210517 | CM | CMPLT | CMPLT | 2024-10-31 14:26 | 2024-11-01 14:14 | Plunger does not move | (TC 10/31/2024 02:26:20 PM, Baghirzade, Nara) Tried to run a functional, plunger does not move. Power cycled but did not help resolve the issue. Disassembled device. Checked the data transmission cable; all good, no tears in the wire. Replaced the motor. Assembled the device however one of the cables is in the way when trying to move the plunger. Will disassemble and re-seat the cable. (TC 11/01/2024 10:23:59 AM, Baghirzade, Nara) Disassembled the device. Re-seated the data transmission (plunger) cable; checked that it was in good condition. Assembled the device back and when running it, it was making loud knocking noises. Error "coarse plunger travel" was also present. Will disassemble the device and inspect the lead screw as the clutch and worm gear were recently replaced. (TC 11/01/2024 02:14:11 PM, Baghirzade, Nara) Disassembled the device. Removed the lead screw, no issues with the threads on the lead screw. Properly lubricated the lead screw using red grease. Checked the clutch threads; no wearing found. Threaded the lead screw in to the clutch multiple times without removing it and applying grease. Assembled the device back. Conducted a quick functional and found no noises and smooth travelling of the plunger. Completed calibration for: -syringe size and position -force sensor -plunger position -syringe size sensor As per the manual, a full maintenance testing must be completed when motor is replaced. See results in procedure.
| 210518 | RDS | CMPLT | CMPLT | 2024-10-31 14:28 | 2024-10-31 14:32 | Weekly Rounds: Antenatal | (TC 10/31/2024 02:31:53 PM, Baghirzade, Nara) Weekly rounds in Antenatal completed. See procedure for results.
| 210519 | RDS | CMPLT | CMPLT | 2024-10-31 14:33 | 2024-10-31 14:33 | Weekly Rounds: L&D | (TC 10/31/2024 02:33:28 PM, Baghirzade, Nara) Weekly rounds in L&D completed. See procedure for results.
1061682 | 210520 | CM | CMPLT | CMPLT | 2024-10-31 15:19 | 2024-10-31 15:19 | matrix pump hose broken. | (TC 10/31/2024 03:19:24 PM, Alvarado, Karla) Replaced matrix pump hose. Relocated the front case.
1071849 | 210521 | CM | CMPLT | CMPLT | 2024-10-31 15:40 | 2024-10-31 15:42 | Not Working | Call re: not working. Went to check. Waited fro access. FOund ipad not charging. Removed ipad and fixed connection. rebooted all lights and reset wifi. Checked functionality. Tested ok.
1059796 | 210522 | CM | CMPLT | CMPLT | 2024-10-31 16:06 | 2024-10-31 16:06 | No image on the reference monitor display. | Fixed.
1060518 | 210523 | CM | CMPLT | CMPLT | 2024-10-31 16:08 | 2024-10-31 16:08 | No power to table, table top floating. | Fixed.
52033 | 210524 | PM | ACTIV | READY | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
52236 | 210525 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile
64067 | 210526 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM ANNUAL: Power Systems, Uninterruptible
64867 | 210527 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
66508 | 210528 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66509 | 210529 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-26 12:40 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 11/26/2024 12:39:54 PM, Baghirzade, Nara) PM completed. Recorder cleaned and calibrated. Touch calibrated. Functional testing passed; NIBP, SpO2, printer, and 4 transducer ports are functional.
66533 | 210530 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal
66677 | 210531 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-27 10:11 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 11/27/2024 10:10:50 AM, Rhuland, Courtney) Battery replaced Functional check - OK
66678 | 210532 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-29 13:02 | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible | (TC 11/29/2024 01:02:23 PM, Rhuland, Courtney) Replaced battery Vacuumed Added new sticker
66685 | 210533 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
67222 | 210534 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
67225 | 210535 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
67226 | 210536 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
67392 | 210537 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
67411 | 210538 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM
67413 | 210539 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-26 12:52 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 11/26/2024 12:51:55 PM, Baghirzade, Nara) Replaced demand valve and collection ring. PM completed.
67415 | 210540 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM
67416 | 210541 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM
67418 | 210542 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM
67422 | 210543 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-26 12:50 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 11/26/2024 12:49:54 PM, Baghirzade, Nara) PM completed.
67428 | 210544 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-27 07:59 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 11/27/2024 07:58:59 AM, Baghirzade, Nara) Replaced damaged demand valve. PM completed.
67878 | 210545 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67882 | 210546 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
67905 | 210547 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Insufflators
67955 | 210548 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-06 09:18 | PM-MAJOR - Annual PM For: Incubators, Infant, Mobile | (TC 11/05/2024 04:17:11 PM, Rhuland, Courtney) Humidifier was old and cracked, replaced with new one When turned on, humidifier was not registering Reopened bottom and reseated humidifier connections Humidifier registering Replaced: Humidifier Humidity sensor 9V rechargeable battery Calibrated: Scale RH (TC 11/06/2024 09:18:17 AM, Rhuland, Courtney) Functional check - OK Set 35C, 75% humidity Read 35C, 76.4% humidity Electrical safety - OK Ground resistance - 151 mOhms Leakage current - 86.6 uA
68398 | 210549 | PM | ACTIV | READY | 2024-11-01 03:21 |  | PM Major - Delivery Module SEVO
68639 | 210550 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-21 11:00 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/21/2024 11:00:28 AM, Cortez, John) Updated software PM completed successfully, results logged in ASM
68757 | 210551 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69185 | 210552 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
69195 | 210553 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE
67363 | 210554 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
67365 | 210555 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
67366 | 210556 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: WARMERS, BLOOD
32520 | 210557 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM - Semi-Annual
32520 | 210558 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | HARP - Semi Annual
32540 | 210559 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM - Monthly
32560 | 210560 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-24 11:42 | PM - Monthly | PM completed Mechanical PM done TOS phantom a(-996.7,9.2) b(333.4,14.1) c(133.5,13.2) d(95.5,12.6) e(-104.6,12.2) f(0.9,13.5) -passed all phantom tests
32580 | 210561 | PM | CMPLT | CMPLT | 2024-11-01 03:21 | 2024-11-24 11:50 | PM - Monthly | PM completed Mechanical PM done TOS phantom a(-996.3,8.6) b(333.6,13.9) c(133.8,13.6) d(95.8,13.2) e(-104.3,12.2) f(0.8,13.3) -passed all phantom tests
39505 | 210562 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM - Annual
39506 | 210563 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM - Annual
395000 | 210564 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM - Annual
502452 | 210565 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | MSH OPG SPP-3520(Rm E2) VOLUSON PM - Annual |  | MSH OPG SPP-3520(Rm E2) VOLUSON PM - Annual
502453 | 210566 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | MSH OPG SPP-3525(Rm D1) VOLUSON PM - Annual |  | MSH OPG SPP-3525(Rm D1) VOLUSON PM - Annual
1046062 | 210567 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Annual PM
1046442 | 210568 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Annual PM
1046452 | 210569 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Annual PM
1046634 | 210570 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Annual PM
1046635 | 210571 | PM | ACTIV | READY | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1046655 | 210572 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046656 | 210573 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046657 | 210574 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046658 | 210575 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046659 | 210576 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Bladder Scanner
1046781 | 210577 | PM | ACTIV | READY | 2024-11-01 03:21 |  | PM - Semi-Annual
1046781 | 210578 | PM | ACTIV | READY | 2024-11-01 03:21 |  | HARP - Semi Annual
1047359 | 210579 | PM | ACTIV | PV | 2024-11-01 03:21 |  | PM-MAJOR- Annual PM for: Testers
1047360 | 210580 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR- Annual PM for: low flow module
1047361 | 210581 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR- Annual PM for: High Flow module
1049204 | 210582 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1049215 | 210583 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1049217 | 210584 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1049218 | 210585 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1049221 | 210586 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1049222 | 210587 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1049608 | 210588 | PM | PENDG | REQST | 2024-11-01 03:21 |  | MSH OPG SPP-HALL-D-RM6 VOLUSON PM - Annual |  | MSH OPG SPP-HALL-D-RM6 VOLUSON PM - Annual
1049784 | 210589 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitor, Physiologic, Central Station
1049786 | 210590 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station, Main
1049901 | 210591 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | HARP - Semi Annual
1049901 | 210592 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM
1049902 | 210593 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | HARP - Semi Annual
1049902 | 210594 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM
1049903 | 210595 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | HARP - Semi Annual
1049903 | 210596 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM
1049904 | 210597 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | HARP - Semi Annual
1049904 | 210598 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM
1049905 | 210599 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | HARP - Semi Annual
1049905 | 210600 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM
1049940 | 210601 | PM | ACTIV | READY | 2024-11-01 03:21 |  | PM - Annual
1049941 | 210602 | PM | ACTIV | READY | 2024-11-01 03:21 |  | PM - Annual
1050056 | 210603 | PM | ACTIV | READY | 2024-11-01 03:21 |  | PM,
1050904 | 210604 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | MSH OPG SPP-HALL-D-4.5 VOLUSON PM - Annual |  | MSH OPG SPP-HALL-D-4.5 VOLUSON PM - Annual
1051830 | 210605 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PLANNED MAINTENANCE-VACUUM PC
1051831 | 210606 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1051840 | 210607 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR - PM For Central Monitor PC
1053320 | 210608 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Insufflators
1053330 | 210609 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR-Annual PM For: Ventilators, Intensive Care, Neonatal/Pediatric
1053876 | 210610 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR - ANNUAL PM for: ESU
1054227 | 210611 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | MSH OPG SPP SPP-HALL-D-3 VOLUSON PM - Annual |  | MSH OPG SPP SPP-HALL-D-3 VOLUSON PM - Annual
1054591 | 210612 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Annual Philips Monitoring, Central Station
1054599 | 210613 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Annual Philips Monitoring, Central Station
1054615 | 210614 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Power Systems, Uninterruptible
1054616 | 210615 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Power Systems, Uninterruptible
1055707 | 210616 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Power Systems, Uninterruptible
1056190 | 210617 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1057511 | 210618 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057699 | 210619 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057700 | 210620 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057710 | 210621 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057711 | 210622 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057713 | 210623 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057714 | 210624 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057715 | 210625 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057716 | 210626 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057717 | 210627 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057718 | 210628 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057719 | 210629 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057720 | 210630 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057721 | 210631 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057722 | 210632 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057723 | 210633 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057724 | 210634 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057726 | 210635 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057727 | 210636 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057730 | 210637 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057732 | 210638 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057733 | 210639 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057734 | 210640 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057735 | 210641 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057736 | 210642 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057737 | 210643 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057738 | 210644 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057739 | 210645 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057740 | 210646 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1057758 | 210647 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1057760 | 210648 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1057761 | 210649 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1057764 | 210650 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1057842 | 210651 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Electrosurgical Units
1057858 | 210652 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
1057861 | 210653 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Power Systems, Uninterruptible
1058197 | 210654 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Infusion Pumps, Patient-Controlled Analgesic
1059330 | 210655 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1059440 | 210656 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059441 | 210657 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Physiologic Monitor Modules, Exhaled Carbon Dioxide
1059530 | 210658 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059531 | 210659 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059532 | 210660 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059533 | 210661 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059534 | 210662 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1059549 | 210663 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1059550 | 210664 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1059551 | 210665 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1059552 | 210666 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1059555 | 210667 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1059941 | 210668 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Warming Units, Patient, Radiant, Infant, Mobile
1060155 | 210669 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060156 | 210670 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060157 | 210671 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060195 | 210672 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060198 | 210673 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060336 | 210674 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060337 | 210675 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060338 | 210676 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060901 | 210677 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1060902 | 210678 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1061254 | 210679 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1063331 | 210680 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | MSH OPG SPP (Research) VOLUSON PM - Annual |  | MSH OPG SPP (Research) VOLUSON PM - Annual
1063358 | 210681 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1063808 | 210682 | PM | PENDG | REQST | 2024-11-01 03:21 |  | MSH 15 SPP PROCEDURE RM VOLUSON PM - Annual |  | MSH 15 SPP PROCEDURE RM VOLUSON PM - Annual
1063934 | 210683 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM - Annual
1063935 | 210684 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM - Annual
1064496 | 210685 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1064497 | 210686 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1065405 | 210687 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR-Annual PM For: Rehab Chair
1066939 | 210688 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Central Monitor
1066943 | 210689 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM ANNUAL: Power Systems, Uninterruptible
1067968 | 210690 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR-Monitors, Physiologic, Central Station
1069363 | 210691 | PM | PENDG | REQST | 2024-11-01 03:21 |  | MSH OPG SPP VOLUSON PM - Annual |  | MSH OPG SPP VOLUSON PM - Annual
1069387 | 210692 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Units
1069388 | 210693 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Units
1069389 | 210694 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Unit
1069390 | 210695 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Units
1069391 | 210696 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Units
1069392 | 210697 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Units
1069393 | 210698 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Units
1069394 | 210699 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM-MAJOR - Annual PM For: Hemodialysis Units
1069562 | 210700 | PM | PENDG | REQST | 2024-11-01 03:21 |  | PM - Annual
1069563 | 210701 | PM | PENDG | REQST | 2024-11-01 03:21 |  | PM - Annual
1069564 | 210702 | PM | PENDG | REQST | 2024-11-01 03:21 |  | PM - Annual
1069565 | 210703 | PM | PENDG | REQST | 2024-11-01 03:21 |  | PM - Annual
1069566 | 210704 | PM | PENDG | REQST | 2024-11-01 03:21 |  | PM - Annual
1069567 | 210705 | PM | ACTIV | ISSUE | 2024-11-01 03:21 |  | PM MAJOR - ANNUAL PM For: Electrosurgical Unit
1070268 | 210706 | PM | ACTIV | READY | 2024-11-01 03:21 |  | Battery Cycling
1066843 | 210707 | CM | CMPLT | CMPLT | 2024-11-01 08:52 | 2024-10-31 08:53 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
1066934 | 210708 | CM | CMPLT | CMPLT | 2024-11-01 08:54 | 2024-10-31 08:54 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
68021 | 210709 | CM | CMPLT | CMPLT | 2024-11-01 08:55 | 2024-10-31 08:57 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
1066935 | 210710 | CM | CMPLT | CMPLT | 2024-11-01 09:13 | 2024-10-31 09:13 | Quarterly Inspection | Coordinated, scheduled and completed quarterly inspection on unit. No signs of moisture or liquid residue was found. Checked functionality. Tested ok.
32520 | 210711 | CM | CMPLT | CMPLT | 2024-11-01 10:00 | 2024-11-01 10:01 | MSH DR2 DETECTOR BETA NOT LEVE CHEST OR TABLE | (TC 11/01/2024 09:59:59 AM, Whittington, Scott) MSH DR2 DETECTOR BETA NOT LEVE CHEST OR TABLE REZERO DETECTOR BETA | MSH DR2 DETECTOR BETA NOT LEVE CHEST OR TABLE
1053348 | 210712 | CM | CMPLT | NFF | 2024-11-01 10:41 | 2024-11-01 10:42 | CAD/GAID not working | (TC 11/01/2024 10:41:42 AM, Fan, Tak) User reported not working on workstation 2. Confirmed other exams had CAD/GAID. Was specific exam. Resent, no change. Worklist indicates it does not have CAD. Could be spcific to the view. Only 2 exams from this morning didn't have CAD and were the same image type LSCC. Rad actually wanted images from priors but they are no longer processable. No problem found overall.
68545 | 210713 | CM | CMPLT | CMPLT | 2024-11-01 10:43 | 2024-11-12 14:39 | Self test failed - Gas Sensors not available | (TC 11/01/2024 10:42:45 AM, Scero, Tomlyn) Pulled device into OR 1 to troubleshoot. Confirmed Self test failed the O2 and Auto. Giving a gas sensor measurement failure alarm. downloaded the logs for review. Log shows multiple errors related to PGA ILCA2 and IR-Sensorhead (TC 11/11/2024 02:48:40 PM, Scero, Tomlyn) Logs were sent to Draeger and confirmed the issue was a defective ILCA2 sensor. Part was ordered and received from Draeger. ILCA2 replaced and Selftest passed. Confirmed Gas sensor measurements passed with the function test of PGA as per test instructions. (TC 11/12/2024 02:39:18 PM, Scero, Tomlyn) Gas sensor tests complete. PGA test per IPM-L passed. Returned AGM to 15th floor.
62982 | 210714 | CM | CMPLT | CMPLT | 2024-11-01 10:56 | 2024-11-01 10:56 | Blue light is not working | (TC 11/01/2024 10:56:02 AM, Alvarado, Karla) 0.0 hours Intensity light: - High: 31.6 uW/cm2/nm - Low: 13.8 uW/cm2/nm Target light is functional All tests passed. No Fault Found (NFF).
1069388 | 210715 | CM | CMPLT | CMPLT | 2024-11-01 11:01 | 2024-11-01 11:02 | Air detected passed air chamber | (TC 11/01/2024 11:01:38 AM, Canjimavadivel, Kamallesh ) Downloaded logs. Turned on device and ran patient simulation. Ran simulation without issues. All self-tests passed. Returned device back to unit.
1066843 | 210716 | CM | CMPLT | CMPLT | 2024-11-01 11:13 | 2024-11-01 11:16 | Error E315 - Scope not compatible | Call re: Not working. Went to check. Waited for access. Found error E315. checked manual. Found scope not compatible or dirty contacts. Cleaned scope contacts and tried again. Same issue. Requested a different scope. Issue persisted. Took the lightsource out and image processor. Removed lightsource cover. Vacummed everything and removed dust bunnies. Sprayed electrical contact cleaner and cleaned lightsource contact points. Reassembled and checked fucntionality. Issues resolved. Tested ok. Returned to service. Also Found Error E302 internal buffer full. Contacted tech supporet and was advised to remove saved images. Consulted staff and erased old exam images one by one as per tech support advise. Issue resolved. Tested ok.
1065011 | 210717 | CM | CMPLT | CMPLT | 2024-11-01 11:55 | 2024-11-01 11:55 | Functional Testing | (TC 11/01/2024 11:55:01 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065020 | 210718 | CM | CMPLT | CMPLT | 2024-11-01 12:00 | 2024-11-01 12:00 | Functional Testing | (TC 11/01/2024 11:59:44 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065009 | 210719 | CM | CMPLT | CMPLT | 2024-11-01 12:01 | 2024-11-01 12:01 | Functional Testing | (TC 11/01/2024 12:00:40 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065016 | 210720 | CM | CMPLT | CMPLT | 2024-11-01 12:01 | 2024-11-01 12:02 | Functional Testing | (TC 11/01/2024 12:01:39 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065021 | 210721 | CM | CMPLT | CMPLT | 2024-11-01 12:03 | 2024-11-01 12:03 | Functional Testing | (TC 11/01/2024 12:02:50 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065017 | 210722 | CM | CMPLT | CMPLT | 2024-11-01 12:04 | 2024-11-01 12:04 | Functional Testing | (TC 11/01/2024 12:03:39 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065013 | 210723 | CM | CMPLT | CMPLT | 2024-11-01 12:04 | 2024-11-01 12:04 | Functional Testing | (TC 11/01/2024 12:04:26 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065006 | 210724 | CM | ACTIV | ISSUE | 2024-11-01 12:05 |  | Functional Testing | (TC 11/01/2024 12:05:31 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1057536 | 210725 | CM | CMPLT | CMPLT | 2024-11-01 12:42 | 2024-11-01 12:44 | ECG port blocked | (TC 11/01/2024 12:43:43 PM, Scero, Tomlyn) Called to the CCU because they couldn't attached the ECG truck cable to the X3. The trunk cable actually broke off in the port and had to be removed with a screw driver. ECG simulation was complete and passed.
1072086 | 210726 | ACC | CMPLT | CMPLT | 2024-11-01 13:00 | 2024-11-01 15:19 | Incoming Inspection | (TC 11/01/2024 03:19:10 PM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072203 | 210727 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 08:10 | Incoming Inspection | (TC 11/04/2024 08:10:14 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072204 | 210728 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 10:47 | Incoming Inspection | (TC 11/04/2024 10:46:50 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072205 | 210729 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 10:02 | Incoming Inspection | (TC 11/04/2024 10:02:18 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072206 | 210730 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 08:48 | Incoming Inspection | (TC 11/05/2024 08:48:23 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072207 | 210731 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 08:48 | Incoming Inspection | (TC 11/04/2024 08:47:33 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072208 | 210732 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 08:01 | Incoming Inspection | (TC 11/04/2024 08:00:57 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072209 | 210733 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 08:53 | Incoming Inspection | (TC 11/05/2024 08:53:19 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072210 | 210734 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 09:55 | Incoming Inspection | (TC 11/04/2024 09:54:47 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072211 | 210735 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 08:59 | Incoming Inspection | (TC 11/05/2024 08:59:10 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072212 | 210736 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-01 14:33 | Incoming Inspection | (TC 11/01/2024 02:33:11 PM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072213 | 210737 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 08:37 | Incoming Inspection | (TC 11/04/2024 08:37:17 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072214 | 210738 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-01 13:21 | Incoming Inspection | (TC 11/01/2024 01:21:12 PM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072215 | 210739 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 08:36 | Incoming Inspection | (TC 11/04/2024 08:36:19 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072216 | 210740 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 08:18 | Incoming Inspection | (TC 11/04/2024 08:18:17 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072217 | 210741 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 09:01 | Incoming Inspection | (TC 11/04/2024 09:00:57 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072219 | 210742 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 10:55 | Incoming Inspection | (TC 11/04/2024 10:54:36 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072220 | 210743 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 09:37 | Incoming Inspection | (TC 11/04/2024 09:36:58 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072221 | 210744 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 10:30 | Incoming Inspection | (TC 11/04/2024 10:30:26 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072222 | 210745 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 09:47 | Incoming Inspection | (TC 11/04/2024 09:47:17 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072223 | 210746 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 10:24 | Incoming Inspection | (TC 11/04/2024 10:24:03 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072224 | 210747 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 07:54 | Incoming Inspection | (TC 11/04/2024 07:54:01 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072225 | 210748 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 10:39 | Incoming Inspection | (TC 11/04/2024 10:38:32 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072226 | 210749 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-04 09:02 | Incoming Inspection | (TC 11/04/2024 09:02:04 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072227 | 210750 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-01 14:35 | Incoming Inspection | (TC 11/01/2024 02:34:30 PM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072228 | 210751 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 08:37 | Incoming Inspection | (TC 11/05/2024 08:36:47 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072229 | 210752 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 10:53 | Incoming Inspection | (TC 11/05/2024 10:53:28 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072230 | 210753 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:27 | Incoming Inspection | (TC 11/05/2024 09:26:54 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072231 | 210754 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:11 | Incoming Inspection | (TC 11/05/2024 09:11:04 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072232 | 210755 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 10:05 | Incoming Inspection | (TC 11/05/2024 10:04:39 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072233 | 210756 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:59 | Incoming Inspection | (TC 11/05/2024 09:58:54 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072234 | 210757 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-06 08:03 | Incoming Inspection | (TC 11/06/2024 08:02:38 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072235 | 210758 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:32 | Incoming Inspection | (TC 11/05/2024 09:31:40 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072236 | 210759 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 10:41 | Incoming Inspection | (TC 11/05/2024 10:41:14 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072237 | 210760 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 10:46 | Incoming Inspection | (TC 11/05/2024 10:46:20 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072238 | 210761 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-06 07:58 | Incoming Inspection | (TC 11/06/2024 07:57:52 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072239 | 210762 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:54 | Incoming Inspection | (TC 11/05/2024 09:54:02 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072240 | 210763 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 08:43 | Incoming Inspection | (TC 11/05/2024 08:42:44 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072241 | 210764 | ACC | ACTIV | ISSUE | 2024-11-01 13:01 |  | Incoming Inspection | (TC 11/06/2024 08:04:33 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072242 | 210765 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-06 07:56 | Incoming Inspection | (TC 11/06/2024 07:56:12 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072243 | 210766 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:22 | Incoming Inspection | (TC 11/05/2024 09:21:53 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072244 | 210767 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 10:29 | Incoming Inspection | (TC 11/05/2024 10:29:25 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072245 | 210768 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:17 | Incoming Inspection | (TC 11/05/2024 09:17:02 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072246 | 210769 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 10:10 | Incoming Inspection | (TC 11/05/2024 10:10:09 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072247 | 210770 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-06 08:01 | Incoming Inspection | (TC 11/06/2024 08:00:30 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072248 | 210771 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-06 08:07 | Incoming Inspection | (TC 11/06/2024 08:06:38 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072249 | 210772 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 11:07 | Incoming Inspection | (TC 11/05/2024 11:06:34 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072250 | 210773 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 09:48 | Incoming Inspection | (TC 11/05/2024 09:48:19 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
1072251 | 210774 | ACC | CMPLT | CMPLT | 2024-11-01 13:01 | 2024-11-05 10:35 | Incoming Inspection | (TC 11/05/2024 10:35:10 AM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
68371 | 210775 | CM | CMPLT | CMPLT | 2024-11-01 14:34 | 2024-11-04 14:32 | Constantly delivering 5L/min of O2 while in Auto | (TC 11/04/2024 02:31:41 PM, Scero, Tomlyn) Downloaded the logs for the device for review. Found during set up of Auto flow and minimun gas flow was set to 5L/min. I tested what this looked like with Karla and yes, while in automode when a minimun gas flow is selected it will deliver that amount of O2 no matter what. Emailed Peter to let him know that the device is funtional.
1063746 | 210776 | PM | CMPLT | CMPLT | 2024-11-01 14:36 | 2024-11-01 14:37 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/01/2024 02:36:31 PM, Rhuland, Courtney) Drained, disinfected, rinsed and refilled with sterile water
1072202 | 210777 | ACC | CMPLT | CMPLT | 2024-11-01 15:09 | 2024-11-01 15:09 | Incomming inspection | (TC 11/01/2024 03:08:39 PM, Alvarado, Karla) Unboxed. Installed battery. Placed asset label. Completed functional test. See procedure for details. Incoming inspection completed.
| 210778 | CM | CMPLT | CMPLT | 2024-11-01 15:41 | 2024-11-01 15:42 | Suction Tree Broken | CAll re: not working. Went to check. Inspected unit. Founf all tubings and canister had liquid residue and was contaminated. Advised staff to clean the suction and remove the tubing. Also advise to ask Reeba's staff to order the new tubing and reach out once it is delivered.
1054693 | 210779 | CM | CMPLT | NFF | 2024-11-04 09:23 | 2024-11-04 09:23 | System error 33 at 10/16 9:40. | (TC 11/04/2024 09:22:58 AM, Alvarado, Karla) Pump was running for 2 different periods of time (8 hours per day) on one week and 2 different periods of time the following week. No Fault Found (NFF)
1060002 | 210780 | CM | CMPLT | NFF | 2024-11-04 09:25 | 2024-11-04 09:25 | Not reporting | (TC 11/04/2024 09:25:15 AM, Canjimavadivel, Kamallesh ) Turned on and waited for wifi to connect. Was able to time sync. Tested with simulator, all waveforms reading. No issue found.
1056473 | 210781 | CM | CMPLT | CMPLT | 2024-11-04 11:19 | 2024-11-04 11:19 | battery not holding charge | Root has been exchanged by Masimo. received new unit and installed configs. tested. functional.
| 210782 | CON | ACTIV | ISSUE | 2024-11-04 12:21 |  | MSH BI SHIPPING LABELS FOR H1 RETURN PARTS; CHECK SECONDARY TUBE LABELS ON H1 & H3 | (TC 11/04/2024 12:21:37 PM, Whittington, Scott) MSH BI SHIPPING LABELS FOR H1 RETURN PARTS; CHECK SECONDARY TUBE LABELS ON H1 & H3 (TC 11/05/2024 11:14:29 AM, Dolinar, John) MSH BI SHIPPING LABELS FOR H1 RETURN PARTS; CHECK SECONDARY TUBE LABELS ON H1 & H3 | MSH BI SHIPPING LABELS FOR H1 RETURN PARTS; CHECK SECONDARY TUBE LABELS ON H1 & H3
32510 | 210783 | CM | ACTIV | ISSUE | 2024-11-04 13:04 |  | MSH DR1 LATHAN REPORTS CHEST IMAGES DARK |  | MSH DR1 LATHAN REPORTS CHEST IMAGES DARK
32510 | 210784 | CM | ACTIV | ISSUE | 2024-11-04 13:06 |  | MSH DR1 EMRG STOP WARNIG, BUT BUTTONS CLEAR, WIRING? |  | MSH DR1 EMRG STOP WARNIG, BUT BUTTONS CLEAR, WIRING?
1064422 | 210785 | CM | CMPLT | NFF | 2024-11-04 14:03 | 2024-11-04 14:03 | Not detecting signals | (TC 11/04/2024 02:03:14 PM, Canjimavadivel, Kamallesh ) Tested with simulator, all waveforms displaying. Not able to replciate reported issue.
1064356 | 210786 | CM | CMPLT | NFF | 2024-11-04 14:11 | 2024-11-04 14:14 | Not charging | (TC 11/04/2024 02:14:05 PM, Canjimavadivel, Kamallesh ) Plugged into outlet and device charging. Charging rate within expected range. Device left to charge overnight and will be sent back to unit
| 210787 | PA | ACTIV | ISSUE | 2024-11-04 14:25 |  | Altrix parts for stock | (TC 11/04/2024 02:31:59 PM, Canjimavadivel, Kamallesh ) Placed order for parts
1054769 | 210788 | CM | CMPLT | CMPLT | 2024-11-04 15:02 | 2024-11-04 15:02 | Blood pressure not inflating | (TC 11/04/2024 03:02:05 PM, Rhuland, Courtney) Tested and at first, cuffs were inflating After more testing, when BP initiated the reading showed -10/-9 and would not inflate Replaced NIBP module with spare Reassembled device and tested multiple readings with all cuffs and issue was resolved Functional check - OK
1061218 | 210789 | CM | CMPLT | CMPLT | 2024-11-04 15:25 | 2024-11-04 16:43 | Video display Issue | Call re: not able to display microscope on the monitor. Initial call was taken by ER. They called again after ER came back. Spoke with ER and he informed me that he replaced DVI Cable. Went to check. Waited for access. Found DVI connector was broken. Replaced DVI connector and cable. Rebooted VMC and scale OR twice. Rebooted micro scope. Issue resolved. Tested ok.
1048472 | 210790 | CM | CMPLT | CMPLT | 2024-11-04 15:32 | 2024-11-04 11:32 | Detector status always calibrating | (TC 11/04/2024 11:32:00 AM, Fan, Tak) Confirmed issue. Calibrated small detector, passed. Issue resolved. Other detector had bent pins, replaced.
| 210791 | CM | CMPLT | CMPLT | 2024-11-05 07:13 | 2024-11-04 07:14 | Biomed huddle | (TC 11/04/2024 07:13:35 AM, Baghirzade, Nara) Biomed huddle
67359 | 210792 | CM | CMPLT | CMPLT | 2024-11-05 07:16 | 2024-11-04 07:16 | Functional check | (TC 11/04/2024 07:15:46 AM, Baghirzade, Nara) Retrieved this unit that is waiting for PM parts. Kim requested that the OR5 warmer gets removed and swapped with a working unit. Ran a functional check for more than an hour - no issues or errors present. Temperature checker measured 41.5 and warmer temperature was 41. Will swap once OR5 is available.
| 210793 | RDS | CMPLT | CMPLT | 2024-11-05 07:33 | 2024-11-04 07:33 | Weekly Rounds: L&D - L&D OR | (TC 11/04/2024 07:33:23 AM, Baghirzade, Nara) Weekly rounds in L&D and L&D OR completed. See procedure for results.
| 210794 | RDS | CMPLT | CMPLT | 2024-11-05 07:35 | 2024-11-05 07:35 | Weekly Rounds: Antenatal | (TC 11/04/2024 07:35:29 AM, Baghirzade, Nara) Weekly rounds in Antenatal completed. See procedure for results.
66429 | 210795 | CM | CMPLT | CMPLT | 2024-11-05 10:08 | 2024-11-05 15:52 | CCU Telemetry Issue | Call re: No data error. Went to check. Spoke with the nurse and was informed that central station indicated no data error when X3 was assigned. Checked unit. Could not replicate the issue. Assisted staff with reviewinf data. Tested ok.
| 210796 | RDS | CMPLT | CMPLT | 2024-11-05 10:14 | 2024-11-05 10:20 | Weekly check for ED
| 210797 | RDS | CMPLT | CMPLT | 2024-11-05 10:14 | 2024-11-07 14:10 | Weekly check for ICU
| 210798 | PA | CMPLT | CMPLT | 2024-11-05 10:29 | 2024-11-21 13:14 | Batteries for Hamilton vents | (TC 11/05/2024 10:33:48 AM, Canjimavadivel, Kamallesh ) Placed order for batteries (TC 11/21/2024 01:13:51 PM, Canjimavadivel, Kamallesh ) Batteries received
1069833 | 210799 | CM | ACTIV | ISSUE | 2024-11-05 10:40 |  | MSH 5 ULTIMAX SLP TIMS, STORED IN MSH ICU; NETWORK TEST OK, DICOM JDMIMWL NOT PRESENT | (TC 11/05/2024 10:40:26 AM, Whittington, Scott) MSH 5 ULTIMAX SLP TIMS, STORED IN MSH ICU; NETWORK TEST OK, DICOM JDMIMWL NOT PRESENT | MSH 5 ULTIMAX SLP TIMS, STORED IN MSH ICU; NETWORK TEST OK, DICOM JDMIMWL NOT PRESENT
32510 | 210800 | CM | CMPLT | CMPLT | 2024-11-05 11:17 | 2024-11-05 11:18 | msh dr1 tube/detector mvmnt | (TC 11/05/2024 11:17:22 AM, Whittington, Scott) msh dr1 tube/detector mvmnt DISABLE TABLE CONTROLLER TO ELIMINATE FALSE EMRG STOP COMMAND | msh dr1 tube/detector mvmnt
68686 | 210801 | CM | CMPLT | CMPLT | 2024-11-05 11:40 | 2024-11-05 11:42 | Functional Testing | (TC 11/05/2024 11:41:30 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests. Where the GND Leakage = 62.2 uA and the GND Resistance = 163 mOhms All test passed.
1065078 | 210802 | CM | CMPLT | CMPLT | 2024-11-05 12:33 | 2024-11-05 12:34 | Functional Test | (TC 11/05/2024 12:33:33 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests All test passed.
68772 | 210803 | CM | CMPLT | CMPLT | 2024-11-05 12:50 | 2024-11-05 12:50 | Functional Test | (TC 11/05/2024 12:49:59 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests. Where the GND Leakage = 72.1 uA and the GND Resistance = 114 mOhms All test passed.
1072284 | 210804 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:16 | Incoming inspection | (TC 11/05/2024 02:15:51 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072285 | 210805 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:17 | Incoming inspection | (TC 11/05/2024 02:16:46 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072286 | 210806 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:17 | Incoming inspection | (TC 11/05/2024 02:17:15 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072287 | 210807 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:18 | Incoming inspection | (TC 11/05/2024 02:17:50 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072288 | 210808 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:18 | Incoming inspection | (TC 11/05/2024 02:18:21 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072289 | 210809 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:19 | Incoming inspection | (TC 11/05/2024 02:18:53 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072290 | 210810 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:19 | Incoming inspection | (TC 11/05/2024 02:19:25 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072291 | 210811 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:20 | Incoming inspection | (TC 11/05/2024 02:19:52 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072292 | 210812 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:20 | Incoming inspection | (TC 11/05/2024 02:20:17 PM, Rhuland, Courtney) Unboxed and asset tagged Functional check - OK Returned to NICU
1072293 | 210813 | ACC | CMPLT | CMPLT | 2024-11-05 12:57 | 2024-11-05 14:21 | Incoming inspection | (TC 11/05/2024 02:20:41 PM, Rhuland, Courtney) Unboxed and asset tagged Ambient temperature too low error present temperature probe found to be faulty Contacted Hillrom to receive replacement probe Took probe from stock - will replace when new probe arrives Functional check - OK Returned to NICU
69115 | 210814 | CM | CMPLT | CMPLT | 2024-11-05 13:49 | 2024-11-05 13:49 | Functional Test | (TC 11/05/2024 01:49:28 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests Download error log: did not find errors. Performed rate accuracy twice. All test passed.
501071 | 210815 | CM | CMPLT | CMPLT | 2024-11-05 13:57 | 2024-11-05 14:05 | MSH CEOU CEOU-3237(Rm12) PROBE SELECTION ERROR | (TC 11/05/2024 01:58:13 PM, Whittington, Scott) MSH CEOU CEOU-3237(Rm12) PROBE SELECTION ERROR APPEARS TO HAVE CLEARED ISSUE (TC 11/05/2024 03:17:28 PM, Dolinar, John) MSH CEOU CEOU-3237(Rm12) PROBE SELECTION ERROR | MSH CEOU CEOU-3237(Rm12) PROBE SELECTION ERROR
| 210816 | PA | CMPLT | CMPLT | 2024-11-05 14:09 | 2024-11-05 14:13 | SureTemp temp probes with wells | (TC 11/05/2024 02:12:54 PM, Canjimavadivel, Kamallesh ) Placed order for temp probe with wells
65969 | 210817 | CM | CMPLT | CMPLT | 2024-11-05 14:30 | 2024-11-05 14:30 | Broken front case | (TC 11/05/2024 02:29:49 PM, Baghirzade, Nara) Replaced front case. Check-in tests passed and completed. IUIs and device cleaned. Electrical Safety - Ground resistance: 117 mohms Leakage current: 61.3 uA
1054981 | 210818 | CM | CMPLT | CMPLT | 2024-11-05 14:31 | 2024-11-05 14:31 | Incoming, config, and functional check | (TC 11/05/2024 02:31:20 PM, Baghirzade, Nara) Root replacement was received from Masimo; acceptance testing completed. Config was loaded with MICT. Verified that low HR alarm changed after configuration upload. Pushed network configuration to Root via MasimoSketch. Loaded the MSH VSM config to Root and Radical. Updated the radical config with MSH config. Updated the SN on TMS. Printed asset tag and device number. Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
66523 | 210819 | PM | CMPLT | CMPLT | 2024-11-05 14:33 | 2024-11-05 14:33 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 11/05/2024 02:32:54 PM, Baghirzade, Nara) Swapped FM with a spare. PM completed. Cleaned recorder and calibrated. Functional testing passed; NIBP, SpO2, printer, and 4 transducer ports are functional.
| 210820 | CM | CMPLT | CMPLT | 2024-11-05 14:34 | 2024-11-05 14:34 | Meeting with Carol NICU EEG Baby AI study | (TC 11/05/2024 02:34:03 PM, Baghirzade, Nara) Met with Carol. Replaced the EEG cable. Discussed AI studies and Carol voiced her complaints about the omnibed and the difficulty of working with one.
1069391 | 210821 | CM | CMPLT | CMPLT | 2024-11-05 14:41 | 2024-11-05 14:41 | "1)Blood flow error 2)Air leak detection error" | (TC 11/05/2024 02:41:08 PM, Cortez, John) Replacement bag scale not passing scale SST. Calibrated. Ran patient simulation test. Passed with no issues. Unable to replicate described issue.
1069389 | 210822 | CM | CMPLT | CMPLT | 2024-11-05 14:47 | 2024-11-04 14:47 | "problem with effluent scale" | (TC 11/04/2024 02:46:43 PM, Cortez, John) Checked scale calibration. Replacement scale was off (50-100g). Calibrated. Scale SST and patient simulation passed successfully
1054537 | 210823 | CM | CMPLT | CMPLT | 2024-11-05 14:51 | 2024-11-04 14:51 | "water leaked - test tubed" | (TC 11/04/2024 02:51:03 PM, Cortez, John) NFF, ran for 45 minutes. Perhaps they tried to run it without test tubing connected
69042 | 210824 | CM | CMPLT | CMPLT | 2024-11-05 14:55 | 2024-11-06 11:33 | Functional Test | (TC 11/05/2024 02:55:27 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log and Download Event Download error log: pressure sensor broken. Therefore, both pressure sensor were replaced. Performed pressure calibration and rate calibration (twice). (TC 11/06/2024 11:32:59 AM, Alvarado, Karla) Performed Check-in. All test passed.
1063000 | 210825 | CM | CMPLT | CMPLT | 2024-11-05 14:57 | 2024-11-04 14:57 | "malfunction" | (TC 11/04/2024 02:56:47 PM, Cortez, John) Leaked immediately from port 3. O-ring ripped in test tube. O-ring replaced. Ran Altrix for 45 minutes with no issue
1069393 | 210826 | CM | ACTIV | ISSUE | 2024-11-05 15:03 |  | "replacement bag not mixed" | (TC 11/01/2024 03:02:57 PM, Cortez, John) Informed of issue a couple of hours after. Therapy ran without issue since. Service manual suggests to run scale SST. Will try to run SST after machine is free
| 210827 | PA | CMPLT | CMPLT | 2024-11-06 10:02 | 2024-11-11 07:31 | Ordering OR table remote clips | (TC 11/06/2024 10:06:42 AM, Baghirzade, Nara) SC 1003345798 (TC 11/11/2024 07:31:26 AM, Baghirzade, Nara) PO 7202483828
1053391 | 210828 | CM | CMPLT | CMPLT | 2024-11-06 13:21 | 2024-11-06 13:21 | Before and after patient use | (TC 11/06/2024 01:20:56 PM, Baghirzade, Nara) Physical inspection good. Replaced filter. Device operates on battery; battery charge ok. AC indicator light is on when device is connected to AC power. Pressure checks: Set 50 mmHg - Measured 50 mmHg Set 125 mmHg - Measured 125 mmHg Set 200 mmHg - Measured 200 mmHg Set 125 mmHg - Measured 125 mmHg Seal check ok; when opening the patient port, the seal check alarm activates. Adjusted the time on the device due to the time change. Device is functional. Returned to 11S.
1054344 | 210829 | CM | CMPLT | CMPLT | 2024-11-06 13:31 | 2024-11-05 13:31 | "broken" | (TC 11/05/2024 01:30:52 PM, Cortez, John) NFF, Tested ok with calibration key, batteries ok
66068 | 210830 | PS | CMPLT | CMPLT | 2024-11-06 13:36 | 2024-11-06 13:36 | Software upgrade | (TC 11/06/2024 01:36:21 PM, Baghirzade, Nara) Software upgrade completed. Transferred network and dataset files.
1054697 | 210831 | CM | CMPLT | CMPLT | 2024-11-06 13:41 | 2024-11-06 13:41 | Not turning on | (TC 11/06/2024 01:40:48 PM, Baghirzade, Nara) Battery was depleted; recharged battery for a few hours and then ran a feed at varying rates on battery. No issues or errors during feed; battery is functional. No fault found.
1058136 | 210832 | CM | CMPLT | CMPLT | 2024-11-06 13:41 | 2024-11-06 13:42 | Doppler cable coming out | (TC 11/06/2024 01:42:27 PM, Reid, Emmerson ) Replaced the cable, working now.
1057008 | 210833 | CM | CMPLT | CMPLT | 2024-11-06 13:53 | 2024-11-26 10:10 | Created by Failed PM WO # 209375 : PM Major Annual: Monitors, Physiologic, Patient Transport | (TC 11/06/2024 02:36:41 PM, Canjimavadivel, Kamallesh ) Issue is related to MSL board from previous experience. Tried to reseat all connections with MSL board, did not resolve. Placed order for MSL board. (TC 11/26/2024 10:09:33 AM, Canjimavadivel, Kamallesh ) Replaced MSL board. Replaced battery. PM tests pass | (TC 11/06/2024 01:46:19 PM, Canjimavadivel, Kamallesh ) Msmt ext malfunction error when turned on
1064422 | 210834 | CM | CMPLT | CMPLT | 2024-11-07 09:27 | 2024-11-07 09:27 | Not picking up rhythm when leads applied to patient. | (TC 11/07/2024 09:27:18 AM, Alvarado, Karla) Performed functional test with simulator, all waveforms displaying. Not able to replicate reported issue.
| 210835 | CM | CMPLT | CMPLT | 2024-11-07 09:35 | 2024-11-15 12:51 | batteries for incubators
1061449 | 210836 | PM | CMPLT | CMPLT | 2024-11-07 10:45 | 2024-11-07 10:45 | PM MAJOR Annual PM for: Procuity Bed | (TC 11/07/2024 10:44:51 AM, Ferreira, Qwynn) Completed by Stryker
1061999 | 210837 | PM | CMPLT | CMPLT | 2024-11-07 10:47 | 2024-11-07 10:47 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 10:46:51 AM, Ferreira, Qwynn) Completed by Stryker
1062501 | 210838 | PM | CMPLT | CMPLT | 2024-11-07 10:52 | 2024-11-07 10:52 | PM-MAJOR Annual PM for: Procuity Bed | (TC 11/07/2024 10:52:24 AM, Ferreira, Qwynn) Completed by Stryker
1062025 | 210839 | PM | CMPLT | CMPLT | 2024-11-07 10:55 | 2024-11-07 10:55 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 10:55:29 AM, Ferreira, Qwynn) Completed by Stryker
1061957 | 210840 | PM | CMPLT | CMPLT | 2024-11-07 10:56 | 2024-11-07 10:56 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 10:56:29 AM, Ferreira, Qwynn) Completed by Stryker
1061968 | 210841 | PM | CMPLT | CMPLT | 2024-11-07 11:09 | 2024-11-07 11:10 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 11:09:38 AM, Ferreira, Qwynn) Completed by Stryker
1061420 | 210842 | PM | CMPLT | CMPLT | 2024-11-07 11:15 | 2024-11-07 11:15 | PM-MAJOR-Annual PM for: Procuity Bed | (TC 11/07/2024 11:15:03 AM, Ferreira, Qwynn) Completed by Stryker
1061961 | 210843 | PM | CMPLT | CMPLT | 2024-11-07 11:16 | 2024-11-07 11:16 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 11:16:25 AM, Ferreira, Qwynn) Completed by Stryker
1047851 | 210844 | CM | CMPLT | CMPLT | 2024-11-07 11:17 | 2024-11-05 12:27 | System error | (TC 11/05/2024 12:27:20 PM, Rhuland, Courtney) System error A2D BGD test failure Tried replacing board but second board failed Placed original board back and force sensor error appeared Replaced force sensor Replaced plunger cable All errors cleared Full calibration and functional check completed All tests passed
1061428 | 210845 | PM | CMPLT | CMPLT | 2024-11-07 11:29 | 2024-11-07 11:29 | PM-MAJOR-Annual PM for: Procuity Bed | (TC 11/07/2024 11:29:07 AM, Ferreira, Qwynn) Completed by Stryker
1048528 | 210846 | CM | CMPLT | NFF | 2024-11-07 12:22 | 2024-11-07 12:22 | Channel error. | (TC 11/07/2024 12:21:44 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, and Check- in tests. Download error log: did not find errors. All test passed. No fault found.
1061992 | 210847 | PM | CMPLT | CMPLT | 2024-11-07 13:32 | 2024-11-07 13:32 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 01:31:47 PM, Ferreira, Qwynn) Completed by Stryker
1061443 | 210848 | PM | CMPLT | CMPLT | 2024-11-07 13:41 | 2024-11-07 13:41 | PM-MAJOR-Annual PM for: Procuity Bed | (TC 11/07/2024 01:40:45 PM, Ferreira, Qwynn) Completed by Stryker
1062019 | 210849 | PM | CMPLT | CMPLT | 2024-11-07 13:42 | 2024-11-07 13:43 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 01:42:52 PM, Ferreira, Qwynn) Completed by Stryker
1059778 | 210850 | FC | CMPLT | CMPLT | 2024-11-07 13:44 | 2024-11-07 13:44 | QC checklist | (TC 11/07/2024 01:43:49 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
69652 | 210851 | CM | CMPLT | CMPLT | 2024-11-07 13:45 | 2024-11-07 13:45 | Recurring 211.2000 error | (TC 11/07/2024 01:45:10 PM, Baghirzade, Nara) Recurring 211.2000 error pertaining to the logic board; as per manual, the only remediation is to replace the logic board. Replaced logic board. Flashed pump and upgraded software. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
1071913 | 210852 | CM | ACTIV | ISSUE | 2024-11-07 13:48 |  | No signal with needle EMG | (TC 11/07/2024 01:48:19 PM, Baghirzade, Nara) No signal/feedback when needle is inserted in patient during EMG. Dr. Lazarou tried multiple times and after many attempts, then it works. Swapped EMG base unit with an old spare; ran HW programmer and completed successfully. Ran diagnostics and ensured that amplifier is also functional. Will coordinate a remote session with Natus as this has been recurring with the new base units.
1071912 | 210853 | CM | ACTIV | ISSUE | 2024-11-07 13:51 |  | No signal/feedback with needle EMG | (TC 11/07/2024 01:51:02 PM, Baghirzade, Nara) No signal/feedback when needle is inserted in patient during EMG. Dr. Lazarou tried multiple times and after many attempts, then it works. Swapped EMG base unit with an old spare; ran HW programmer and completed successfully. Ran diagnostics and ensured that amplifier is also functional. Will coordinate a remote session with Natus as this has been recurring with the new base units.
69631 | 210854 | PS | CMPLT | CMPLT | 2024-11-07 14:00 | 2024-11-07 14:00 | Software Recall | (TC 11/07/2024 01:59:47 PM, Baghirzade, Nara) Software upgrade completed.
1062026 | 210855 | PM | CMPLT | CMPLT | 2024-11-07 14:06 | 2024-11-07 14:06 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 02:06:19 PM, Ferreira, Qwynn) Completed by Stryker
1062415 | 210856 | PM | CMPLT | CMPLT | 2024-11-07 14:09 | 2024-11-07 14:09 | PM-MAJOR-Annual PM for: Procuity Bed | (TC 11/07/2024 02:08:55 PM, Ferreira, Qwynn) Completed by Stryker
1061899 | 210857 | PM | CMPLT | CMPLT | 2024-11-07 14:10 | 2024-11-07 14:10 | PM-MAJOR_Annual PM For: Med Surg Bed | (TC 11/07/2024 02:10:13 PM, Ferreira, Qwynn) Completed by Stryker
| 210858 | CM | CMPLT | CMPLT | 2024-11-07 14:12 | 2024-11-07 14:12 | Power bar on crash cart broken | (TC 11/07/2024 02:12:15 PM, Canjimavadivel, Kamallesh ) Informed CNS that prong still stuck in outlet and to contact building services. Replaced broken power bar.
| 210859 | FC | CMPLT | CMPLT | 2024-11-07 14:21 | 2024-11-07 09:41 | Functional checks on watchbp devices | (TC 11/07/2024 09:41:17 AM, Scero, Tomlyn) Functional checks complete on all 4 watchbp devices in the heart function clinic. All available BP cuffs tests with our BP simulator. No issues or leaks found. No error codes present.
| 210860 | PA | ACTIV | ISSUE | 2024-11-07 14:22 |  | Power bars for stock | (TC 11/07/2024 02:32:37 PM, Canjimavadivel, Kamallesh ) Placed order for power bars
1064401 | 210861 | CM | ACTIV | ISSUE | 2024-11-07 14:32 |  | Bipolar Cable Order for Endo | Nov 7 - Request from Anita to order bipolar cable for Erbe. Went to check and confirmed the cable that they wanted to order. Contacted vendor and requested part# and pricing. Nov 14 - Followe dup. Rep could not find info. Rep visited SH so met with him and took him to endo to show the cable needed. He took images and informed that he will send info soon. Nov 26 - Followed up. Info received. Communicated to Anita. Was advised to schedule a demo beofre we make the purchase. Contacted rep and requestred a demo for bipolar cables.
| 210862 | CM | ACTIV | ISSUE | 2024-11-07 15:17 |  | MSH CR ROOM CASSETTE CLEAR OUT |  | MSH CR ROOM CASSETTE CLEAR OUT
1054538 | 210863 | CM | CMPLT | CMPLT | 2024-11-07 15:58 | 2024-11-07 17:19 | "screen flickers" | (TC 11/07/2024 05:18:47 PM, Cortez, John) Did notice the screen flicker on power up, but couldn't replicate issue after that Checked ribbon cable from display. No sign of flickering when moved. Keyboard cover ripped, replaced Other functionality ok
68966 | 210864 | CM | CMPLT | CMPLT | 2024-11-08 07:53 | 2024-11-07 13:00 | it is not turning on. | (TC 11/07/2024 01:00:00 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests. Where the GND Leakage = 69.1 uA and the GND Resistance = 106 mOhms The PCU left full charge. All test passed.
68940 | 210865 | CM | CMPLT | CMPLT | 2024-11-08 07:54 | 2024-11-07 13:30 | it is not turning on. | (TC 11/07/2024 01:30:00 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests. Where the GND Leakage =66.1 uA and the GND Resistance = 131 mOhms The PCU left full charge. All test passed.
1065698 | 210866 | CM | CMPLT | CMPLT | 2024-11-08 08:40 | 2024-11-08 08:41 | Power cord was not able to plug in. | (TC 11/08/2024 08:41:01 AM, Alvarado, Karla) Prongs were fixed. Performed tests in ASM: Download Error Log, Download Event, and Check Tests. . Where the GND Leakage = 58.7 uA and the GND Resistance =128 mOhms All test passed.
68998 | 210867 | CM | CMPLT | CMPLT | 2024-11-08 08:46 | 2024-11-08 08:47 | Power cord was not able to plug in | (TC 11/08/2024 08:46:37 AM, Alvarado, Karla) Prongs were fixed Performed tests in ASM: Download Error Log, Download Event, and Check Tests. . Where the GND Leakage = 67.7 uA and the GND Resistance = 208 mOhms All test passed.
69598 | 210868 | CM | CMPLT | CMPLT | 2024-11-08 09:04 | 2024-11-08 09:05 | Power cord was not able to plug in (PCU) | (TC 11/08/2024 09:04:48 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests All test passed.
1055721 | 210869 | CM | CMPLT | CMPLT | 2024-11-08 09:17 | 2024-11-12 09:52 | Cracked MSL connector | (TC 11/08/2024 09:23:28 AM, Canjimavadivel, Kamallesh ) Found crack in msl connector. Placed order for repair part. (TC 11/12/2024 09:51:55 AM, Canjimavadivel, Kamallesh ) Replaced MSL connector, all functions pass
69103 | 210870 | CM | CMPLT | CMPLT | 2024-11-08 09:30 | 2024-11-08 09:31 | Power cord was not able to plug in (PCU) | (TC 11/08/2024 09:30:31 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests All test passed.
69226 | 210871 | CM | CMPLT | CMPLT | 2024-11-08 09:39 | 2024-11-08 09:39 | Power cord was not able to plug in (attached to PCU). | (TC 11/08/2024 09:39:18 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check- in tests All test passed.
1054214 | 210872 | CM | CMPLT | CMPLT | 2024-11-08 09:40 | 2024-11-07 09:40 | "new batteries, not powering on" | (TC 11/07/2024 09:40:03 AM, Cortez, John) Probe well seems to be worn out. Prevents unit from powering on intermittently. Replaced
1071784 | 210873 | CM | CMPLT | CMPLT | 2024-11-08 10:18 | 2024-11-08 10:18 | E3 error | (TC 11/08/2024 10:18:10 AM, Rhuland, Courtney) E3 error reported Tested device and no issues were found Cuff was not being put on the arm tight enough Discussed with reporter and asked her to keep an eye on it for the following few days
32510 | 210874 | CM | ACTIV | ISSUE | 2024-11-08 11:29 |  | MSH DR1 table microcontroller | (TC 11/08/2024 03:34:18 PM, Whittington, Scott) msh dr1 table microcontroller (TC 11/08/2024 03:34:31 PM, Heron, Chris) msh dr1 table microcontroller (TC 11/08/2024 04:13:40 PM, Whittington, Scott) msh dr1 table microcontroller | msh dr1 table microcontroller
66070 | 210875 | CM | CMPLT | CMPLT | 2024-11-08 12:50 | 2024-11-08 12:50 | Returned from SickKids | (TC 11/08/2024 12:50:29 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check Tests. . Where the GND Leakage = 87.7 uA and the GND Resistance = 153 mOhms All test passed.
66002 | 210876 | CM | CMPLT | CMPLT | 2024-11-08 12:51 | 2024-11-08 12:51 | Returned from SickKids | (TC 11/08/2024 12:51:17 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check Tests. . Where the GND Leakage = 87.3 uA and the GND Resistance = 149 mOhms All test passed.
66050 | 210877 | CM | CMPLT | CMPLT | 2024-11-08 12:52 | 2024-11-08 12:52 | Returned from SickKids | (TC 11/08/2024 12:52:12 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check Tests. . Where the GND Leakage = 88.4 uA and the GND Resistance = 159 mOhms All test passed.
1055108 | 210878 | CM | CMPLT | CMPLT | 2024-11-08 12:53 | 2024-11-08 12:53 | Returned from SickKids | (TC 11/08/2024 12:53:01 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log, Download Event, and Check Tests. . Where the GND Leakage = 59.0 uA and the GND Resistance = 153 mOhms All test passed.
1070754 | 210879 | CM | CMPLT | CMPLT | 2024-11-08 12:54 | 2024-11-08 12:54 | Returned from SickKids | (TC 11/08/2024 12:54:06 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
1070753 | 210880 | CM | CMPLT | CMPLT | 2024-11-08 12:55 | 2024-11-08 12:56 | Returned from SickKids | (TC 11/08/2024 12:55:33 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
65931 | 210881 | CM | CMPLT | CMPLT | 2024-11-08 12:57 | 2024-11-08 12:57 | Returned from SickKids | (TC 11/08/2024 12:56:41 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
1050617 | 210882 | CM | CMPLT | CMPLT | 2024-11-08 12:57 | 2024-11-08 12:57 | Returned from SickKids | (TC 11/08/2024 12:57:25 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
65961 | 210883 | CM | CMPLT | CMPLT | 2024-11-08 12:58 | 2024-11-08 12:58 | Returned from SickKids | (TC 11/08/2024 12:58:20 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
65971 | 210884 | CM | CMPLT | CMPLT | 2024-11-08 12:59 | 2024-11-08 12:59 | Returned from SickKids | (TC 11/08/2024 12:59:07 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
65882 | 210885 | CM | CMPLT | CMPLT | 2024-11-08 13:00 | 2024-11-08 13:00 | Returned from SickKids | (TC 11/08/2024 01:00:22 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
65966 | 210886 | CM | CMPLT | CMPLT | 2024-11-08 13:01 | 2024-11-08 13:01 | Returned from SickKids | (TC 11/08/2024 01:01:09 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
65938 | 210887 | CM | CMPLT | CMPLT | 2024-11-08 13:02 | 2024-11-08 13:02 | Returned from SickKids | (TC 11/08/2024 01:02:03 PM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
69404 | 210888 | CM | CMPLT | CMPLT | 2024-11-08 14:10 | 2024-11-08 14:10 | Pump was displaying “channel error”. | (TC 11/08/2024 02:10:18 PM, Alvarado, Karla) Performed tests in ASM: Download Error Log and Download Event. Download error log: upstream pressure sensor failure. Therefore, the upstream pressure sensor was replaced. While the pressure calibration was being performed the downstream pressure sensor failed. Hence, the downstream pressure sensor was replaced too. Performed pressure calibration, rate calibration and check-in tests. All tests passed.
1064418 | 210889 | CM | ACTIV | ISSUE | 2024-11-08 14:10 |  | MRI Incubator Long Hood Assembly + Shipping
1065707 | 210890 | CM | CMPLT | CMPLT | 2024-11-08 14:42 | 2024-11-08 14:44 | Battery does not charge. | (TC 11/07/2024 09:30:00 AM, Alvarado, Karla) The PCU was running in battery to reply the error. Indeed, the battery was unable to keep the charge. Therefore, the battery was replaced. The battery was left charging; then, the PCU was running in normal mode to check that the battery was keeping the charge and working properly. (TC 11/08/2024 02:43:59 PM, Alvarado, Karla) Performed Check Test. Where the GND Leakage = 68.2 uA and the GND Resistance = 170 mOhms All test passed.
| 210891 | MEP | CMPLT | CMPLT | 2024-11-08 15:26 | 2024-11-08 15:26 | Demo setup
1065699 | 210892 | CM | CMPLT | CMPLT | 2024-11-11 09:09 | 2024-11-11 09:09 | battery is not charging. | (TC 11/11/2024 09:09:08 AM, Alvarado, Karla) Performed battery condition test twice. The device kept charging. Performed Check Test. Where the GND Leakage = 63.9 uA and the GND Resistance = 123 mOhms All test passed.
69337 | 210893 | CM | CMPLT | CMPLT | 2024-11-11 09:10 | 2024-11-11 09:10 | Functional Check | (TC 11/11/2024 09:10:12 AM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
69455 | 210894 | CM | CMPLT | CMPLT | 2024-11-11 09:11 | 2024-11-11 09:11 | Functional Check | (TC 11/11/2024 09:10:57 AM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
69386 | 210895 | CM | CMPLT | CMPLT | 2024-11-11 09:12 | 2024-11-11 09:12 | Functional Check | (TC 11/11/2024 09:11:55 AM, Alvarado, Karla) Performed Check-in test in ASM. All test passed.
1055703 | 210896 | CM | CMPLT | CMPLT | 2024-11-11 10:14 | 2024-11-11 10:15 | Not Working | (TC 11/11/2024 10:14:36 AM, Reid, Emmerson ) Printer was not working. Not printing and yellow light was flashing. I needed to reboot the central twice before it reconnected and was able to print again. Working now.
1056001 | 210897 | CM | CMPLT | CMPLT | 2024-11-11 10:45 | 2024-11-11 10:45 | Battery replacement. | (TC 11/11/2024 10:45:22 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055940 | 210898 | CM | CMPLT | CMPLT | 2024-11-11 10:46 | 2024-11-11 10:46 | Temperature is not reading | (TC 11/11/2024 10:46:26 AM, Alvarado, Karla) Replaced temperature probe. Performed tests: Temperature NIBP= set: 120/80, measured: 116/84. SpO2 = set: 95% and 80 b/m, measured: 95% and 80 b/m. All test passed.
1056457 | 210899 | CM | CMPLT | CMPLT | 2024-11-11 10:48 | 2024-11-11 10:48 | Temperature is not reading | (TC 11/11/2024 10:48:22 AM, Alvarado, Karla) Replaced temperature probe. Performed tests: Temperature NIBP = set: 120/80, measured: 117/83. SpO2 = set: 95% and 80 b/m, measured: 95% and 80 b/m. All test passed.
| 210900 | CM | CMPLT | CMPLT | 2024-11-11 11:25 | 2024-11-11 11:25 | New webview install L&D | (TC 11/11/2024 11:24:51 AM, Baghirzade, Nara) Met with Naz and Ellie from RPO to discuss the webview install in L&D.
1072314 | 210901 | ACC | CMPLT | CMPLT | 2024-11-11 11:37 | 2024-11-11 11:45 | Incoming inspection | (TC 11/11/2024 11:45:27 AM, Scero, Tomlyn) Incoming inspection complete. Safety inspection complete and passed. GWR .071 Ohms LC 91.4 mA SF LC 157.5 mA
32510 | 210902 | CM | ACTIV | ISSUE | 2024-11-11 13:16 |  | MSH DR1 DETECTOR HAS QUARTERED EFFECT | (TC 11/11/2024 01:23:07 PM, Whittington, Scott) MSH DR1 DETECTOR HAS QUARTERED EFFECT (TC 11/11/2024 01:23:07 PM, Whittington, Scott) MSH DR1 DETECTOR HAS QUARTERED EFFECT | MSH DR1 DETECTOR HAS QUARTERED EFFECT
65896 | 210903 | CM | CMPLT | CMPLT | 2024-11-11 13:29 | 2024-11-11 13:30 | Returned from SickKids | (TC 11/11/2024 01:29:36 PM, Alvarado, Karla) Performed Check- in test. The pressure Disc Accuracy test was failed. Therefore, the Syringe was calibrated. However, the syringe kept failing. Replaced pressure sensor. Replaced latch. Performed Calibration and Check- in test. All test passed.
1065696 | 210904 | CM | CMPLT | CMPLT | 2024-11-11 14:11 | 2024-11-11 14:12 | Plug removed | (TC 11/11/2024 02:11:37 PM, Alvarado, Karla) Plugged the power cord. Performed Check Test. Where the GND Leakage = 63.8 uA and the GND Resistance = 90 mOhms All test passed.
63220 | 210905 | CM | CMPLT | CMPLT | 2024-11-11 14:20 | 2024-11-11 14:20 | Battery malfunction | (TC 11/11/2024 02:20:03 PM, Baghirzade, Nara) Defective battery. Replaced battery. Ran functional checks (varying rates and volumes); no issues or errors found.
502492 | 210906 | CM | CMPLT | CMPLT | 2024-11-11 14:25 | 2024-11-11 14:25 | Power supply/internal battery problem | Checked connections of power supply. Restarted the system and injector. Injector is now working normally.
1048500 | 210907 | CM | ACTIV | ISSUE | 2024-11-11 14:58 |  | MSH 5 US EPIQ System will not boot up properly. | (TC 11/15/2024 04:05:44 PM, Fan, Tak) Swapped monitor. Problem persists. Checked cabling, reseated. Tested video on external monitor, seems to work. Something about the power not getting to the monitor I guess. Jillian to order more parts, maybe power cable. AIO board had green LEDs so probably providing correct voltage? | System will not boot up properly. POST beeps and main monitor will not power up. Called Philips. Ref# 0124666493
1055691 | 210908 | CM | CMPLT | CMPLT | 2024-11-12 07:47 | 2024-11-11 14:00 | Functional Check | (TC 11/11/2024 02:00:00 PM, Alvarado, Karla) Performed functional and alarm tests. Set temperature: Fan, measured 23.8 C. 32C, measured 34.4 C. 38C, measured 40.2 C. 43C, measured 44.9 C. All test passed.
| 210909 | PA | ACTIV | ISSUE | 2024-11-12 08:29 |  | Batteries for stock | (TC 11/12/2024 08:40:27 AM, Canjimavadivel, Kamallesh ) Placed order for batteries
69021 | 210910 | CM | CMPLT | CMPLT | 2024-11-12 08:32 | 2024-11-12 08:32 | Power cord is damaged | (TC 11/12/2024 08:31:49 AM, Alvarado, Karla) Replaced the power cord. Replaced latch Performed Check Test. Where the GND Leakage = 79.6 uA and the GND Resistance = 86 mOhms All test passed.
1056507 | 210911 | CM | CMPLT | CMPLT | 2024-11-12 08:37 | 2024-11-12 08:37 | Battery replacement. | (TC 11/12/2024 08:37:21 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1069391 | 210912 | FC | CMPLT | NFF | 2024-11-12 10:06 | 2024-11-08 10:08 | functional check requested, "air in line detected" | (TC 11/08/2024 10:05:42 AM, Cortez, John) Amanda (ICU) mentioned that nurses came across air detected message, but didn't notice air in line. Functional check was requested. Check passed successfully
69104 | 210913 | CM | CMPLT | CMPLT | 2024-11-12 10:08 | 2024-11-12 10:08 | Not turning on. | (TC 11/12/2024 10:08:07 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log and Download Event. Download error log: upstream pressure sensor failure several times. Therefore, the upstream pressure sensor was replaced. While the pressure calibration was being performed the downstream pressure sensor failed. Hence, the downstream pressure sensor was replaced. Performed Pressure calibration, Rate Calibration, and Check-in test. All test passed.
1059828 | 210914 | PM | CMPLT | CMPLT | 2024-11-12 10:53 | 2024-11-12 10:54 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/12/2024 10:53:36 AM, Alvarado, Karla) Drained, disinfected, rinsed and refilled with sterile water.
1067824 | 210915 | CM | CMPLT | CMPLT | 2024-11-12 11:17 | 2024-11-12 11:54 | Exp flow sensor failure | (TC 11/12/2024 11:54:05 AM, Scero, Tomlyn) Breathing system was wet after use and failing the self test. They said they tried to clean out the water but it was still failing. When I came to test, the expiratory flow sensor was broken. I replaced it with a flow sensor I had from a spare breathing system and re tested the self test. It passed with no issues.
1049901 | 210916 | CM | CMPLT | CMPLT | 2024-11-12 11:41 | 2024-11-12 11:41 | MSH R61 Revo - Not going up | (TC 11/12/2024 11:41:26 AM, Chau, Ken) Replaced switch on handle bar as broken
1063746 | 210917 | PM | CMPLT | CMPLT | 2024-11-12 12:38 | 2024-11-12 12:38 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/12/2024 12:37:47 PM, Alvarado, Karla) Drained, disinfected, rinsed and refilled with sterile water.
| 210918 | RDS | CMPLT | CMPLT | 2024-11-12 13:52 | 2024-11-12 13:54 | Weekly check for ED
| 210919 | RDS | CMPLT | CMPLT | 2024-11-12 13:53 | 2024-11-20 08:46 | Weekly check for ICU
1049904 | 210920 | CM | CMPLT | CMPLT | 2024-11-12 15:42 | 2024-11-12 15:42 | Complaint of low collimator light output. |  | Checked luminance output. Output was 3500 lux which was higer than R65 and R61 which have an output of 1900 lux.
| 210921 | CM | CMPLT | CMPLT | 2024-11-12 17:00 | 2024-11-12 17:00 | MSH Site coverage. | (TC 11/12/2024 05:02:45 PM, Dolinar, John) MSH Site coverage. | MSH Site coverage.
| 210922 | PA | ACTIV | ISSUE | 2024-11-13 09:59 |  | Ordering PCU logic board for stock | (TC 11/13/2024 09:59:48 AM, Baghirzade, Nara) SC 1003350415
1065040 | 210923 | CM | CMPLT | CMPLT | 2024-11-13 10:14 | 2024-11-13 10:14 | Unable to connect to network | (TC 11/13/2024 10:14:07 AM, Rhuland, Courtney) Error 600.6875 present apon startup Unable to connect to network Tried to push network configuration and it errored out Tried to push network config and would not take Flashed PCU and took much longer than expected Tried comm board from spare PCU and error cleared Pushed network config and tested connectivity All tests passed Functional check - OK
65305 | 210924 | CM | CMPLT | CMPLT | 2024-11-13 11:04 | 2024-11-13 11:07 | Functional check | (TC 11/13/2024 11:06:38 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
| 210925 | CM | CMPLT | CMPLT | 2024-11-13 11:07 | 2024-11-13 11:07 | Pharmacy scale weight verification | (TC 11/13/2024 11:07:24 AM, Rhuland, Courtney) AND EJ-303 scale - SN: 6A5302394 Verification testing as requested by Pharmacy *TEST WEIGHTS USED ARE NOT CERTIFIED* Troemner 311-10S test weights - SN: 86809 Test weight - Actual weight - % difference 10mg 10.000mg 0% 50mg 50.000mg 0% 100mg 100.000mg 0% 500mg 500.000mg 0% 1g 1.001g 0.1% 2g 2.001g 0.05% 5g 5.009g 0.045% 10g 9.999g 0.01% 20g 20.007g 0.035% 50g 50.039g 0.078% 100g 100.030g 0.03%
53009 | 210926 | CM | CMPLT | CMPLT | 2024-11-13 11:08 | 2024-11-13 11:09 | Functional check | (TC 11/13/2024 11:08:33 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065013 | 210927 | CM | CMPLT | CMPLT | 2024-11-13 11:09 | 2024-11-13 11:10 | Functional check | (TC 11/13/2024 11:09:35 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065006 | 210928 | CM | CMPLT | CMPLT | 2024-11-13 11:11 | 2024-11-13 11:11 | Functional Check | (TC 11/13/2024 11:10:40 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065004 | 210929 | CM | CMPLT | CMPLT | 2024-11-13 11:12 | 2024-11-13 11:12 | Functional check | (TC 11/13/2024 11:11:46 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065020 | 210930 | CM | CMPLT | CMPLT | 2024-11-13 11:13 | 2024-11-12 13:12 | Functional check | (TC 11/12/2024 01:12:00 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065011 | 210931 | CM | CMPLT | CMPLT | 2024-11-13 11:14 | 2024-11-12 13:20 | Functional check | (TC 11/12/2024 01:20:00 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065021 | 210932 | CM | ACTIV | ISSUE | 2024-11-13 11:16 |  | Functional check | (TC 11/13/2024 11:16:14 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065016 | 210933 | CM | CMPLT | CMPLT | 2024-11-13 11:17 | 2024-11-13 11:17 | Functional check | (TC 11/13/2024 11:17:04 AM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065005 | 210934 | CM | CMPLT | CMPLT | 2024-11-13 11:19 | 2024-11-12 13:19 | Functional check | (TC 11/12/2024 01:19:00 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
1065017 | 210935 | CM | CMPLT | CMPLT | 2024-11-13 11:20 | 2024-11-12 13:20 | Functional check | (TC 11/12/2024 01:20:00 PM, Alvarado, Karla) No physical damage. Test 1: >= 100 mmHg Test 2: >= 300 mmHg Test 3: >= 500 mmHg Test 4: >= 450 mmHg All test passed.
66044 | 210936 | PS | CMPLT | CMPLT | 2024-11-13 13:56 | 2024-11-13 13:56 | Software recall | (TC 11/13/2024 01:56:24 PM, Baghirzade, Nara) Software upgrade completed.
67229 | 210937 | CM | CMPLT | CMPLT | 2024-11-13 13:57 | 2024-11-21 12:07 | System failure 09 | (TC 11/14/2024 01:03:55 PM, Baghirzade, Nara) Disassembled the top enclosure and rear panel to gain access to the power PCA. Checked both power boards on top of the enclosure for the heater and light as well as put in a new power board to make sure that there is no issue with the power boards; both power boards are functional. Replaced the power PCA; still getting the same system failure error after booting up the warmer. Replaced the ribbon cable between the power PCA and the control board; still getting the same system failure error after booting up the warmer. Measured the resistance on the cable that connects directly from the power PCA to the heating element and found that there is a short in the cable. Went in service mode, got the error codes, and Olcay emailed getinge about this issue. Placed warmer on 18th fl. biomed room. (TC 11/19/2024 07:50:29 AM, Baghirzade, Nara) Swapped the power PCA board as the original one was functional. Received new heater engine. Disassembled the heater, removed all power supply boards, ground connections, and heater itself. Placed new heater and completed alignment of the heater. Placed metal enclosure back on the heater. (TC 11/20/2024 02:21:35 PM, Baghirzade, Nara) Completely assembled the heater, placed all the boards back and assembled the cover back on. Turned on the warmer to 100% and the error did not return; heater engine is functional.. (TC 11/21/2024 12:07:10 PM, Baghirzade, Nara) Functional tests passed. All switches and controllers are functional. Warmer is functional. SpO2 is functional. Resuscitation system is functional.
| 210938 | PA | CMPLT | CMPLT | 2024-11-13 13:58 | 2024-11-13 14:00 | Training AED battery | (TC 11/13/2024 01:59:34 PM, Canjimavadivel, Kamallesh ) Placed order for training AED battery
501648 | 210939 | CM | CMPLT | CMPLT | 2024-11-13 14:30 | 2024-11-13 14:30 | MRT rebooted and hung up. | (TC 11/13/2024 02:29:45 PM, Green, Paul) Travelled to site and inspected. Powered off the SRU unit and rebooted the CRU. Powered up ok. Cusotmer was able to load the syringes and use the injector.
65427 | 210940 | CM | CMPLT | CMPLT | 2024-11-13 14:38 | 2024-11-13 14:39 | Blood pressure cuff inflates but not reading. | (TC 11/13/2024 02:38:39 PM, Alvarado, Karla) Replaced blood pressure cuff. Replaced fan. Performed tests: Temperature NIBP SpO2 = set: 95%, measured: 95% All test passed.
1046480 | 210941 | CM | CMPLT | CMPLT | 2024-11-13 15:17 | 2024-11-13 15:18 | Synch with Conmed cautery | Call re: not working. Went to check. Waited for access. WAs informed unit needs to be synched with the cautery. Sync the unit. Checked functionality. Tested ok.
1059618 | 210942 | CM | CMPLT | CMPLT | 2024-11-13 15:19 | 2024-11-13 15:20 | Repair | Call re: not working. Went to check. Waited for access. Was infomed the microscope does not connect to Olympus monitors. FOund HDMI-DVI cable broken. Repalced cable. Checked functionality. Tested ok.
1061213 | 210943 | MEP | ACTIV | ISSUE | 2024-11-13 15:22 |  | OR 12 Readiness | Nov 13 - Reuqest from OR Clinical staff to prepare OR for surgeries. Was assigned the list of equipment needed. Went to OR and assessed locations and connections. Will be installling required equipment in near future. Nov 14 - Moved and installed reuqested equipment. ESU, Smoke Evac, Cautery. Connected. Checked functionality. Tested ok. Checked Olympus VMC. Found wall monitor not working. Will investigate further. Nov 15 - Rectified connection issue for wall monitor. Chekced fucntionality. Tested ok. Nov 20 - Found one of the monitor handle missing. Reached out to Olympus. Was advised to contact baxter. Reached out to baxter and sent them pictures of the part needed. Requested purchase info. Nov 26 - Followed up. Info received. Created SC. Requested approval. Received. Placed order via MMC.
32510 | 210944 | CM | ACTIV | ISSUE | 2024-11-13 16:49 |  | MSH DR1 OGP MOVEMENTS TUBE STAND RELAY1.1 RE10 | (TC 11/13/2024 04:49:49 PM, Whittington, Scott) MSH DR1 OGP MOVEMENTS TUBE STAND RELAY1.1 RE10 | MSH DR1 OGP MOVEMENTS TUBE STAND RELAY1.1 RE10
1048778 | 210945 | CM | ACTIV | ISSUE | 2024-11-14 08:13 |  | Error message present when turn unit on | (TC 11/14/2024 01:33:49 PM, Scero, Tomlyn) Back and forth emals with tech support and sales staff from Cardialine. they sent me a service manual and guided me in the right direction for troubleshooting. I fixed the problem by replacing the RJ-45 cable from the console to the MCU. Unfortunatly we only had unsheilded patch cables and the unit requires a sheilded cable. I ordered a cable from DigiKey and will replace it on the unit when it comes in.
67880 | 210946 | CM | CMPLT | CMPLT | 2024-11-14 08:38 | 2024-11-14 08:44 | Broken power supply | (TC 11/14/2024 08:43:31 AM, Alvarado, Karla) Replaced batteries and power supply. Performance check completed. Where the accuracy test the intended delivery volume was set to 20 ml, while the actual delivery volume was 20.435 ml. The inaccuracy volume was 0.435 ml; therefore, the accuracy error percentage obtained was 2%. All tests passed
1055772 | 210947 | CM | CMPLT | CMPLT | 2024-11-14 09:07 | 2024-11-14 09:07 | Configuration update | (TC 11/14/2024 09:07:30 AM, Rhuland, Courtney) Transferred all logs to server Pushed updated configuration
66476 | 210948 | CM | CMPLT | CMPLT | 2024-11-14 09:42 | 2024-11-14 09:48 | Keyboard cover replacement | (TC 11/14/2024 09:47:46 AM, Rhuland, Courtney) Brought cover to OPG and replaced worn keyboard cover
| 210949 | CM | CMPLT | CMPLT | 2024-11-14 09:53 | 2024-11-14 09:59 | 2 limb leads LL/LA
1055579 | 210950 | CM | CMPLT | CMPLT | 2024-11-14 10:57 | 2024-11-14 10:57 | Temperature probe not reading. | (TC 11/14/2024 10:57:23 AM, Alvarado, Karla) Replaced probe well. Performed tests: Temperature NIBP=set: 120/80, measured: 115/82. SpO2 =set: 95% and 40 b/m, measured: 95% and 40 b/m. All test passed.
1066462 | 210951 | FC | CMPLT | CMPLT | 2024-11-14 11:20 | 2024-11-14 13:48 | QC checklist | (TC 11/14/2024 01:47:49 PM, Baghirzade, Nara) QC checcklist completed. Returned to 14 South.
66434 | 210952 | CM | CMPLT | CMPLT | 2024-11-14 13:25 | 2024-11-14 13:26 | No data tele | (TC 11/14/2024 01:26:30 PM, Scero, Tomlyn) Used a simulator with the telepak to confirm that it wasn't connecting to the central station. In the Piicix configuration I noticed that the tele equipment label had been removed from the bed label. I reassigned and locked it. Now it's working correctly.
1066462 | 210953 | PS | CMPLT | CMPLT | 2024-11-14 13:46 | 2024-11-14 13:46 | Occlusion sensor recall# 2024110149 | (TC 11/14/2024 01:46:16 PM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall.
69385 | 210954 | CM | CMPLT | CMPLT | 2024-11-14 14:13 | 2024-11-14 14:13 | Channel error | (TC 11/14/2024 02:13:19 PM, Alvarado, Karla) The pump was plugged but was not turning on and the screen displayed error code 210.5020. According to the manual, the logic board should be replaced. First, the pump was flashed but it did not work and the logic board was replaced. The pump was plugged one more time and the screen displayed error code 200.5040. According to the manual, the pump should be reflashed. After flash the pump, the software required to performance PM. PM completed. Door cover and rubber ring were replaced. All test passed.
1047843 | 210955 | CM | CMPLT | CMPLT | 2024-11-14 14:20 | 2024-11-14 14:21 | System failure | (TC 11/14/2024 02:20:37 PM, Rhuland, Courtney) Primary audible alarm post alarm present Could not replicate alarm but was reported before Replaced primary speaker During testing, check clutch error present Will replace clutch in future WO
67229 | 210956 | PA | CMPLT | CMPLT | 2024-11-15 07:18 | 2024-11-18 13:02 | Ordering heater engine for repair | (TC 11/15/2024 07:22:15 AM, Baghirzade, Nara) SC 1003351934 (TC 11/18/2024 01:01:42 PM, Baghirzade, Nara) PO 7202488157
1069789 | 210957 | CM | CMPLT | CMPLT | 2024-11-15 07:53 | 2024-11-22 07:42 | Sending transducer for repair | (TC 11/15/2024 10:28:12 AM, Baghirzade, Nara) SC 1003351961 Shipped by KJV. KJV # 464925 (TC 11/22/2024 07:42:07 AM, Baghirzade, Nara) Received transducer from repair; functional test passed. Dropped off at L&D. (TC 11/22/2024 07:47:53 AM, Baghirzade, Nara) PO# 7500214888 | loud noise during fetal monitoring
1054161 | 210958 | CM | CMPLT | CMPLT | 2024-11-15 07:55 | 2024-11-22 07:49 | Sending transducer for repair | (TC 11/15/2024 10:26:39 AM, Baghirzade, Nara) SC 1003351961 Shipped by KJV. KJV # 464925 (TC 11/22/2024 07:49:24 AM, Baghirzade, Nara) Vendor deemed too much damage to the unit and broken beyond repair. There are no charges applicable and unit will be retired and disposed of. | No sound
1069791 | 210959 | CM | CMPLT | CMPLT | 2024-11-15 07:56 | 2024-11-22 07:49 | Sending transducer for repair | (TC 11/15/2024 10:24:42 AM, Baghirzade, Nara) SC 1003351961 (TC 11/15/2024 10:24:42 AM, Baghirzade, Nara) SC 1003351961 Shipped by KJV. KJV # 464925 (TC 11/22/2024 07:48:27 AM, Baghirzade, Nara) Received transducer from repair; functional test passed. Dropped off at L&D. (TC 11/22/2024 07:48:47 AM, Baghirzade, Nara) PO# 7500214888 | Static noise
| 210960 | OC | CMPLT | CMPLT | 2024-11-15 09:10 | 2024-11-21 08:43 | On-call week Nov 15-21, 2024 | (TC 11/15/2024 09:10:29 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 11/18/2024 08:50:31 AM, Canjimavadivel, Kamallesh ) on-call dailly report (TC 11/19/2024 08:33:22 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 11/20/2024 08:43:12 AM, Canjimavadivel, Kamallesh ) Daily check report (TC 11/21/2024 08:43:12 AM, Canjimavadivel, Kamallesh ) Daily check report
| 210961 | CM | CMPLT | CMPLT | 2024-11-15 09:32 | 2024-11-20 14:51 | Uline Order | Nov 15 - Spoke with staff and determined stationary/parts requirement. Created SC. Sent for approval. Received. Placed order via MMC. Nov 20 - Received. Sorted and organized everything.
1067436 | 210962 | PS | CMPLT | CMPLT | 2024-11-15 10:31 | 2024-11-15 10:31 | Occlusion sensor recall# 2024110149 | (TC 11/15/2024 10:31:18 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall.
1066115 | 210963 | PS | CMPLT | CMPLT | 2024-11-15 10:32 | 2024-11-15 10:33 | Occlusion sensor recall# 2024110149 | (TC 11/15/2024 10:32:34 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall.
1048546 | 210964 | CM | CMPLT | CMPLT | 2024-11-15 11:20 | 2024-11-20 15:21 | Lamp Replacement | Nov 15 - Unit due for lamp replacement. Contacted vendor for part# and pricing. Received. Created shopping cart. Requested approval. Received. Placed order via MMC. Nov 20 - Lamp bulb received. Cleaned unit. Replaced lamp bulb. Checked fuctionality. Tested ok.
1060030 | 210965 | PA | CMPLT | CMPLT | 2024-11-15 12:52 | 2024-11-21 07:10 | No-charge PO for O2 cell | (TC 11/15/2024 01:09:17 PM, Baghirzade, Nara) SC 1003352523 (TC 11/21/2024 07:09:54 AM, Baghirzade, Nara) PO 7202488590
| 210966 | CM | CMPLT | CMPLT | 2024-11-15 13:10 | 2024-11-15 13:10 | Full team meeting | (TC 11/15/2024 01:10:13 PM, Baghirzade, Nara) Full team meeting
1062779 | 210967 | CM | CMPLT | CMPLT | 2024-11-15 13:11 | 2024-11-15 13:16 | Defective O2 cell | (TC 11/15/2024 01:11:12 PM, Baghirzade, Nara) Gaby reported that during calibration, the O2 cell failed. Assessed the calibration graph and saw that the O2 cell did not respond at all (checked that O2 gas was open). Replaced O2 cell with an old cell from 2023. John and I calibrated the gain and offset. Ran the calibration a couple of times and was passing. Sent results to Gaby and she was ok with it. Reached out to Novus regarding the defective cell as it hasn't been a year since it's been installed; will receive a new one at no charge.
| 210968 | RDS | CMPLT | CMPLT | 2024-11-15 13:17 | 2024-11-15 13:17 | Weekly Rounds: 11 North | (TC 11/15/2024 01:17:17 PM, Baghirzade, Nara) Weekly rounds in 11 North completed. See procedure for results.
| 210969 | RDS | CMPLT | CMPLT | 2024-11-15 13:43 | 2024-11-15 13:44 | Weekly Rounds: 11 South | (TC 11/15/2024 01:43:36 PM, Baghirzade, Nara) Weekly rounds in 11 South completed. See procedure for results.
| 210970 | RDS | CMPLT | CMPLT | 2024-11-15 13:56 | 2024-11-15 13:57 | Weekly Rounds: Antenatal | (TC 11/15/2024 01:56:52 PM, Baghirzade, Nara) Weekly rounds in Antenatal completed. See procedure for results.
| 210971 | RDS | CMPLT | CMPLT | 2024-11-15 13:58 | 2024-11-15 13:58 | Weekly Rounds: L&D and L&D OR | (TC 11/15/2024 01:58:16 PM, Baghirzade, Nara) Weekly rounds in L&D and L&D OR completed. See procedure for results.
1046203 | 210972 | CM | CMPLT | CMPLT | 2024-11-15 16:00 | 2024-11-15 16:01 | Test server verification | (TC 11/21/2024 09:20:00 AM, Fan, Tak) SR failed. Eventually came through. User may have turn off system prematurely. Systems mixed up in the morning but sorted out.
1067975 | 210973 | CM | CMPLT | CMPLT | 2024-11-15 16:04 | 2024-11-15 16:04 | Job stuck | (TC 11/15/2024 04:04:12 PM, Fan, Tak) Spooler issue. Cleared, sent OK.
1054991 | 210974 | CM | CMPLT | NFF | 2024-11-18 08:26 | 2024-11-18 08:26 | Oximeter is flashing and does not take reading. | (TC 11/18/2024 08:25:58 AM, Alvarado, Karla) Performed tests: Temperature NIBP = set: 120/80, measured: 114/83. SpO2 = set: 95% and 40 b/m, measured: 95% and 40 b/m. All test passed. NFF
1051357 | 210975 | CM | CMPLT | NFF | 2024-11-18 08:27 | 2024-11-18 08:28 | Light is broken. | (TC 11/18/2024 08:27:31 AM, Alvarado, Karla) 8609.2 hours Intensity light: - High: 30.9 uW/cm2/nm - Low: 13.7 uW/cm2/nm Target light is functional GND Resistance: 120 mohms Leakage current: 50.7 uA All tests passed
| 210976 | RDS | CMPLT | CMPLT | 2024-11-18 08:35 | 2024-11-25 08:47 | Weekly checks in CCU | (TC 11/18/2024 08:46:30 AM, Scero, Tomlyn) Weekly checks complete. (TC 11/25/2024 08:46:51 AM, Scero, Tomlyn) Weekly checks complete.
| 210977 | PA | CMPLT | CMPLT | 2024-11-18 08:45 | 2024-11-20 08:42 | Batteries for servo-u loaners | (TC 11/18/2024 08:49:39 AM, Canjimavadivel, Kamallesh ) Requested no charge PO for batteries (TC 11/20/2024 08:42:22 AM, Canjimavadivel, Kamallesh ) Installed batteries on loaners
62982 | 210978 | CM | CMPLT | CMPLT | 2024-11-18 08:48 | 2024-11-18 08:48 | Light is orange | (TC 11/18/2024 08:47:50 AM, Alvarado, Karla) Light was disassembled to relocate internal connection. 0.0 hours Intensity light: - High: 33.8 uW/cm2/nm - Low: 14.7 uW/cm2/nm Target light is functional GND Resistance: 170 mohms Leakage current: 21.7 uA All tests passed
1069391 | 210979 | CM | CMPLT | CMPLT | 2024-11-18 09:22 | 2024-11-20 14:08 | Extreme positive return, unable to return blood | (TC 11/18/2024 02:19:18 PM, Canjimavadivel, Kamallesh ) Downloaded logs. Reached out to department for additional info. Turned on and set up patient simulation. Simulation ran without issues. Was able to create scenario to replicate reported issue, was able to resolve by following suggestions on screen. (TC 11/20/2024 02:07:36 PM, Cortez, John) Ran pressure sensor SST successfully, pressure readings are accurate Ran patient simulation again amd ran for 30 mins, no issues found
62632 | 210980 | PA | CMPLT | CMPLT | 2024-11-18 12:34 | 2024-11-25 14:30 | Ordering lithium ion batteries for PM | (TC 11/18/2024 12:37:38 PM, Baghirzade, Nara) SC 1003353467 (TC 11/25/2024 02:29:32 PM, Baghirzade, Nara) PO 7202489139
1054228 | 210981 | CM | CMPLT | CMPLT | 2024-11-18 12:39 | 2024-11-18 12:40 | Reported not working | (TC 11/18/2024 12:39:30 PM, Baghirzade, Nara) Unit did not have batteries; placed batteries. Thermometer would default to pediatric and in F; changed the settings to match the other thermometers in antenatal. Functional checks passed.
1055071 | 210982 | CM | CMPLT | CMPLT | 2024-11-18 12:41 | 2024-11-18 12:41 | Reported not working | (TC 11/18/2024 12:40:51 PM, Baghirzade, Nara) Thermometer had security feature activated; deactivated security feature. Functional tests passed.
69651 | 210983 | CM | CMPLT | CMPLT | 2024-11-18 12:42 | 2024-11-18 12:42 | Channel error | (TC 11/18/2024 12:42:05 PM, Baghirzade, Nara) Downloaded error log; did not note any errors. Check-in tests passed and completed. IUI connectors and device cleaned.
69659 | 210984 | CM | CMPLT | CMPLT | 2024-11-18 12:43 | 2024-11-18 12:43 | Channel error | (TC 11/18/2024 12:42:47 PM, Baghirzade, Nara) Downloaded error logs; did not note any errors. During check-in, the rate accuracy was reading high. Rate was calibrated. Pressure was calibrated Check-in tests completed. IUI connectors and device cleaned.
| 210985 | CM | CMPLT | CMPLT | 2024-11-18 12:44 | 2024-11-18 12:44 | 1-1 meeting with Adeel | (TC 11/18/2024 12:44:09 PM, Baghirzade, Nara) 1-1 meeting with Adeel
1057280 | 210986 | CM | ACTIV | ISSUE | 2024-11-18 12:46 |  | Sending for repair | (TC 11/18/2024 12:48:59 PM, Baghirzade, Nara) Module error 51; replaced well and temp probe, error still persists. Completed power cycle after replacing well and probe however it did not resolve the issue. Placed service request with Masimo. (TC 11/20/2024 01:41:45 PM, Baghirzade, Nara) Sent with KJV KJV# 465923 (TC 11/20/2024 01:57:48 PM, Baghirzade, Nara) SC 1003355714 | Module error 51
66972 | 210987 | CM | CMPLT | CMPLT | 2024-11-18 14:33 | 2024-11-18 14:34 | Battery replacement. | (TC 11/18/2024 02:33:40 PM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1057619 | 210988 | CM | CMPLT | CMPLT | 2024-11-18 14:39 | 2024-11-18 14:39 | Relocate unit | (TC 11/18/2024 02:38:54 PM, Alvarado, Karla) Moved unit from MCR room to 16 North (surgery). Updated asset tag. Performed tests: - Temperature - NIBP - SpO2 = set: 95% and 40 b/m, measured: 95% and 40 b/m. All test passed.
1055930 | 210989 | CM | CMPLT | CMPLT | 2024-11-18 14:40 | 2024-11-18 14:41 | Relocate unit | (TC 11/18/2024 02:40:48 PM, Alvarado, Karla) Moved unit from MCR room to 16 North (surgery). Updated asset tag. Performed tests: Temperature NIBP SpO2 = set: 95% and 40 b/m, measured: 95% and 40 b/m. All test passed.
1055906 | 210990 | CM | CMPLT | CMPLT | 2024-11-18 14:43 | 2024-11-18 14:43 | Relocate unit | (TC 11/18/2024 02:43:21 PM, Alvarado, Karla) Moved unit from MCR room to 16 North (surgery). Updated asset tag. Performed tests: Temperature NIBP = set: 120/80, measured: 115/81. SpO2 = set: 95% and 40 b/m, measured: 95% and 40 b/m. All test passed.
| 210991 | RDS | CMPLT | CMPLT | 2024-11-18 14:46 | 2024-11-19 09:57 | Weekly check for ED
| 210992 | CM | CMPLT | CMPLT | 2024-11-19 07:46 | 2024-11-18 14:45 | Biomed huddle | (TC 11/18/2024 02:45:00 PM, Baghirzade, Nara) Biomed huddle
1055899 | 210993 | CM | CMPLT | CMPLT | 2024-11-19 08:03 | 2024-11-18 15:02 | Blood pressure is not working. | (TC 11/18/2024 03:02:00 PM, Alvarado, Karla) Replaced finger clip sensor. Replaced power cord. Performed tests: Temperature NIBP = set: 120/80, measured: 117/84. SpO2 = set: 95% and 40 b/m, measured: 95% and 40 b/m. Electrical safety test: Where the GND Leakage = 92.6 uA and the GND Resistance = 95 mOhms All test passed.
1055933 | 210994 | CM | CMPLT | CMPLT | 2024-11-19 08:13 | 2024-11-19 08:14 | Sensor is not working | (TC 11/19/2024 08:13:32 AM, Alvarado, Karla) Replaced patient cable (SpO2) Performed tests: Temperature NIBP = set: 120/80, measured: 115/84. SpO2 = set: 95% and 40 b/m, measured: 95% and 40 b/m. All test passed.
1061329 | 210995 | CM | ACTIV | ISSUE | 2024-11-19 08:26 |  | Remote screen issues | (TC 11/19/2024 08:26:17 AM, Reid, Emmerson ) Went to OR11 and saw the remote screen was blank but lit up. I was able to unplug and plug back in the remote and this fixed the screen. I then opened up the remote and found red fluid had permiated inside the remote, this is likely causing the issues. I attempted to clean up as much as I could and check the ribbon cable connectors. I put everything back and will see if this issue persists.
1061316 | 210996 | CM | CMPLT | CMPLT | 2024-11-19 10:49 | 2024-11-19 10:49 | Viewpoint test server testing | (TC 11/19/2024 10:49:03 AM, Fan, Tak) WL OK.
1046203 | 210997 | CM | CMPLT | CMPLT | 2024-11-19 10:50 | 2024-11-19 10:50 | Random no archive study prompted by system | (TC 11/19/2024 10:50:03 AM, Fan, Tak) Someone did a short scan without demographics. Worked with Dianne to identify the study. She will reach out to the Rad for archiving.
1067338 | 210998 | CM | CMPLT | CMPLT | 2024-11-19 10:54 | 2024-11-19 10:54 | O2 sensor failure | Collected error logs. reached vendor for technical support. found water in expiratory cassette. put a new cassette. started testing. passed pre-use check 3 times. run the unit one week. no error. returned back to service. updated RT.
1055985 | 210999 | CM | CMPLT | CMPLT | 2024-11-19 15:08 | 2024-11-19 15:08 | Battery replacement. | (TC 11/19/2024 03:08:08 PM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1056435 | 211000 | CM | CMPLT | CMPLT | 2024-11-19 15:11 | 2024-11-19 15:11 | Functional check. | (TC 11/19/2024 03:11:16 PM, Alvarado, Karla) Replaced probe well. Replaced power cord. Update date and time. Provided large BP cuff. Performed tests: Temperature, NIBP, and SpO2 All test passed.
1055925 | 211001 | CM | CMPLT | CMPLT | 2024-11-19 15:13 | 2024-11-19 15:14 | Functional check. | (TC 11/19/2024 03:13:44 PM, Alvarado, Karla) Replaced finger clip (SpO2). Update date and time. Provided large BP cuff. Performed tests: Temperature, NIBP, and SpO2 All test passed.
1055926 | 211002 | CM | CMPLT | CMPLT | 2024-11-19 15:16 | 2024-11-19 15:16 | Functional check. | (TC 11/19/2024 03:16:14 PM, Alvarado, Karla) Replaced probe well. Update date and time. Provided large BP cuff. Performed tests: Temperature, NIBP, and SpO2 All test passed.
1056443 | 211003 | CM | CMPLT | CMPLT | 2024-11-19 15:18 | 2024-11-19 15:18 | Functional check. | (TC 11/19/2024 03:17:55 PM, Alvarado, Karla) Replaced power cord. Update date and time. Performed tests: Temperature, NIBP, and SpO2 All test passed.
1056447 | 211004 | CM | CMPLT | CMPLT | 2024-11-19 15:20 | 2024-11-19 15:20 | Functional check. | (TC 11/19/2024 03:19:47 PM, Alvarado, Karla) Replaced temperature probe. Update date and time. Performed tests: Temperature, NIBP, and SpO2 All test passed.
1048435 | 211005 | CM | ACTIV | ISSUE | 2024-11-19 15:25 |  | MSH 17 NICU ZONARE Battery replacement | (TC 11/19/2024 03:25:23 PM, Fan, Tak) Verified system. Checked battery. Requested quote. | MSH 17 NICU ZONARE Battery replacement
32520 | 211006 | CM | ACTIV | ISSUE | 2024-11-19 16:48 |  | MSH DR2 DETECTOR CLAMSHELL REPAIR |  | MSH DR2 DETECTOR CLAMSHELL REPAIR
500668 | 211007 | CM | CMPLT | CMPLT | 2024-11-19 16:55 | 2024-11-19 16:56 | MSH CEOU RM 6 CRT FAILED; SYSTEM DISMANTLED & HDD REMOVED SCRAPPED & RETIRED | (TC 11/19/2024 04:56:01 PM, Whittington, Scott) MSH CEOU RM 6 CRT FAILED; SYSTEM DISMANTLED & HDD REMOVED SCRAPPED & RETIRED | MSH CEOU RM 6 CRT FAILED; SYSTEM DISMANTLED & HDD REMOVED CRAPPED & RETIRED
1067437 | 211008 | PS | CMPLT | CMPLT | 2024-11-20 07:14 | 2024-11-19 07:14 | Occlusion sensor recall# 2024110149 | (TC 11/19/2024 07:14:05 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall.
1067437 | 211009 | FC | CMPLT | CMPLT | 2024-11-20 07:15 | 2024-11-19 07:15 | QC checklist | (TC 11/19/2024 07:14:51 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
1066460 | 211010 | PS | CMPLT | CMPLT | 2024-11-20 07:16 | 2024-11-19 07:16 | Occlusion sensor recall# 2024110149 | (TC 11/19/2024 07:16:30 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall.
1066460 | 211011 | FC | CMPLT | CMPLT | 2024-11-20 07:17 | 2024-11-19 07:44 | QC checklist | (TC 11/19/2024 07:43:49 AM, Baghirzade, Nara) QC checklist completed. Returned to 14 South.
1070842 | 211012 | CM | CMPLT | CMPLT | 2024-11-20 07:46 | 2024-11-19 07:47 | Reported that temperature does not display | (TC 11/19/2024 07:46:33 AM, Baghirzade, Nara) Ran the omnibed in pre-heat mode; functional. Set the heater to 100% and temperature was being displayed. Both baby temp probe ports are functional. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
1057616 | 211013 | FC | CMPLT | CMPLT | 2024-11-20 07:49 | 2024-11-19 07:49 | Functional check | (TC 11/19/2024 07:48:36 AM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional. Provided power cord as it was damaged/defective.
| 211014 | CM | CMPLT | CMPLT | 2024-11-20 07:53 | 2024-11-21 13:59 | Coordination and organization of the medfusion recalls | (TC 11/19/2024 07:52:34 AM, Baghirzade, Nara) Coordinating the medfusion recall project on behalf of Courtney; brought down medfusions (2 at a time) from storage unit until all were remediated. Connected with Tabish regarding cleaning the pumps that are in the patient rooms for nov. 20th and 21st. Connected with Janet regarding the drug library changes that are to be made. Contacted Patty from medfusion to ask if the on-site tech/FSR can complete these changes. Will wait for Janet to approve the library before uploading it to other pumps.
1056525 | 211015 | CM | CMPLT | CMPLT | 2024-11-20 09:18 | 2024-11-20 09:18 | Battery replacement. | (TC 11/20/2024 09:18:22 AM, Alvarado, Karla) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1055945 | 211016 | CM | CMPLT | CMPLT | 2024-11-20 09:19 | 2024-11-20 09:19 | Blood pressure cuff is leaking | (TC 11/20/2024 09:19:11 AM, Alvarado, Karla) BP cuff was not plug. Performed tests: Temperature, NIBP, and SpO2 All test passed.
1056449 | 211017 | CM | ACTIV | ISSUE | 2024-11-20 09:21 |  | sent out for repair
1061203 | 211018 | CM | CMPLT | CMPLT | 2024-11-20 09:52 | 2024-11-20 09:52 | the ipad was not working | (TC 11/14/2024 09:51:46 AM, Huang, Jessica) received the call from OR 2 the ipad was not working properly reset it, working fine
1061209 | 211019 | CM | CMPLT | CMPLT | 2024-11-20 09:56 | 2024-11-20 09:56 | medical device port on the boom was not activated | (TC 11/14/2024 09:55:49 AM, Huang, Jessica) received the call in OR 8 and spoke to Dr. Vescan the medical port on the foot boom was not activated contacted the help desk to address the issue
1069387 | 211020 | CM | CMPLT | NFF | 2024-11-20 14:14 | 2024-11-19 14:14 | "positive return pressure, unable to return blood" | (TC 11/19/2024 02:14:10 PM, Cortez, John) Ran patient simulation for 30+ mins, no issues found
1055028 | 211021 | CM | CMPLT | CMPLT | 2024-11-20 14:19 | 2024-11-20 14:19 | Battery replacement | (TC 11/20/2024 02:18:59 PM, Baghirzade, Nara) Replaced battery. Performed battery upgrade successfully. Placed the original SD card back in the Radical 7.
1054983 | 211022 | FC | CMPLT | CMPLT | 2024-11-20 14:21 | 2024-11-20 14:21 | Functional check | (TC 11/20/2024 02:21:02 PM, Baghirzade, Nara) Functional testing passed; NIBP, temperature, and SpO2 modules are all functional.
| 211023 | CM | ACTIV | ISSUE | 2024-11-20 14:52 |  | Uline Order - Biomed Cabinets | Nov 19 - Plan to replace biomed cabinets. Measured space. Discussed with SC And AA. Researched storage cabiets available. Sent info to SC and AA. Nov 20 - Discussed with SC and was recommended some changes. Created SC. Sent for approval. Placed order via MMC Nov 26 - call from receiving about the shipment delivery. Was asked to vist the loading dock. Went to check. Found shipment too big to fit in the elevator. Consulted SC for delivery. Will go with him and check tomorrow. Nov 27 - Went with SC to check out shipment. Opened shipment and brought it to biomed on a cart. Assembled one cabinet. Disposed old cabinet and placed new one in its place. Nov 28 - Sorted and organized all testing equipment and accessories in the new cabinet.
1055750 | 211024 | CM | ACTIV | ISSUE | 2024-11-20 14:59 |  | it does not connect to brains | (TC 11/20/2024 02:58:49 PM, Alvarado, Karla) Performed Check- in test in ASM. All test passed. Not Fault Found
1071849 | 211025 | CM | CMPLT | CMPLT | 2024-11-20 15:10 | 2024-11-20 15:11 | Not working | Call re: not working. Went to check. Found lights not synched and wifi off on the ipad. Fixed it. Checked functionality. Tested ok.
1061213 | 211026 | CM | CMPLT | CMPLT | 2024-11-20 15:19 | 2024-11-20 15:19 | Monitor Flickering | Call re: monitors flickering while activating cautery. Went to check. Waited for access. FOund power cord connection at the back of the smoke evac loose. Fixed it. Issue resolved. Tested ok.
1061210 | 211027 | CM | ACTIV | ISSUE | 2024-11-21 09:38 |  | Monitors not holding position | (TC 11/21/2024 09:38:44 AM, Reid, Emmerson ) Monitors were not holding position. I was able to fix one of them but the other two where not accessable. The monitors were also flickering blue, this was due to tension on the sdi cable at the back of the camera processor. I releived the strain and the flickering went away.
| 211028 | PA | CMPLT | CMPLT | 2024-11-21 09:54 | 2024-11-21 10:07 | Hamilton vent PM Kits | (TC 11/21/2024 10:07:11 AM, Canjimavadivel, Kamallesh ) Placed order for PM kits
66504 | 211029 | PM | CMPLT | CMPLT | 2024-11-21 12:08 | 2024-11-21 12:09 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 11/21/2024 12:08:32 PM, Baghirzade, Nara) PM completed. Recorder cleaned and calibrated. Touch calibrated. Functional testing passed; NIBP, SpO2, printer, and 4 transducer ports are functional.
69434 | 211030 | CM | CMPLT | CMPLT | 2024-11-21 15:24 | 2024-11-20 10:39 | LVP display + front case need replacement | (TC 11/20/2024 10:39:22 AM, Cortez, John) Replaced display and front case
1055750 | 211031 | CM | CMPLT | NFF | 2024-11-22 07:20 | 2024-11-21 07:20 | it does not connect to brains. | (TC 11/21/2024 07:19:57 AM, Alvarado, Karla) Performed tests in ASM: Download Error Log, and Check- in tests. Download error log: did not find errors. All test passed. Not Fault Found
1062144 | 211032 | PM | CMPLT | CMPLT | 2024-11-22 08:47 | 2024-11-22 08:47 | Batteries Replacement | (TC 11/22/2024 08:47:17 AM, Baghirzade, Nara) Completed by Stryker.
1062271 | 211033 | PM | CMPLT | CMPLT | 2024-11-22 08:54 | 2024-11-22 08:55 | Batteries Replacement Due | (TC 11/22/2024 08:54:32 AM, Baghirzade, Nara) Completed by Stryker.
1062239 | 211034 | PM | CMPLT | CMPLT | 2024-11-22 08:56 | 2024-11-22 08:56 | Batteries Replacement Due | (TC 11/22/2024 08:56:14 AM, Baghirzade, Nara) Completed by Stryker.
1062248 | 211035 | PM | CMPLT | CMPLT | 2024-11-22 09:07 | 2024-11-22 09:07 | Batteries Replacement Due | (TC 11/22/2024 09:07:19 AM, Baghirzade, Nara) Completed by Stryker.
1062258 | 211036 | PM | CMPLT | CMPLT | 2024-11-22 09:09 | 2024-11-22 09:09 | Batteries Replacement Due | (TC 11/22/2024 09:09:18 AM, Baghirzade, Nara) Completed by Stryker.
1062269 | 211037 | PM | CMPLT | CMPLT | 2024-11-22 09:12 | 2024-11-22 09:12 | Batteries Replacement Due | (TC 11/22/2024 09:12:08 AM, Baghirzade, Nara) Completed by Stryker.
1062390 | 211038 | PM | CMPLT | CMPLT | 2024-11-22 09:19 | 2024-11-22 09:20 | Batteries Replacement Due | (TC 11/22/2024 09:19:32 AM, Baghirzade, Nara) Completed by Stryker.
1062389 | 211039 | PM | CMPLT | CMPLT | 2024-11-22 09:26 | 2024-11-22 09:26 | Batteries Replacement Due | (TC 11/22/2024 09:26:11 AM, Baghirzade, Nara) Completed by Stryker.
1062274 | 211040 | PM | CMPLT | CMPLT | 2024-11-22 09:28 | 2024-11-22 09:28 | Batteries Replacement Due | (TC 11/22/2024 09:28:04 AM, Baghirzade, Nara) Completed by Stryker.
1051379 | 211041 | CM | CMPLT | CMPLT | 2024-11-22 10:13 | 2024-11-22 10:14 | Battery replacement | (TC 11/22/2024 10:13:34 AM, Baghirzade, Nara) Replaced batteries. Set 5kg for accuracy testing, measured 5kg. Scale is functional.
67943 | 211042 | CM | CMPLT | CMPLT | 2024-11-22 10:17 | 2024-11-22 10:17 | Broken door hinge | (TC 11/22/2024 10:17:29 AM, Baghirzade, Nara) Incubator came down without the reported broken door/door hinge; replaced east door with a spare from biomed. Functional tests passed and completed. Scale is functional. Tilt is functional. Set temp to 35C and humidity to 75%; measured 35C across all temperature points and 75%RH Incubator is fully functional.
1065780 | 211043 | PM | CMPLT | CMPLT | 2024-11-22 10:54 | 2024-11-21 10:55 | PM-MAJOR - Annual PM For: INFUSION PUMPS, GENERAL PURPOSE | (TC 11/21/2024 10:54:37 AM, Cortez, John) Updated software PM completed successfully, results logged in ASM
| 211044 | RDS | CMPLT | CMPLT | 2024-11-22 11:34 | 2024-11-22 11:35 | Weekly Rounds: L&D and L&D OR | (TC 11/22/2024 11:34:41 AM, Baghirzade, Nara) Weekly rounds in L&D and L&D OR completed. See procedure for results.
| 211045 | RDS | CMPLT | CMPLT | 2024-11-22 11:37 | 2024-11-22 11:37 | Weekly Rounds: Antenatal | (TC 11/22/2024 11:36:57 AM, Baghirzade, Nara) Weekly rounds in Antenatal completed. See procedure for results.
| 211046 | RDS | CMPLT | CMPLT | 2024-11-22 11:39 | 2024-11-22 11:39 | Weekly Rounds: 11 North | (TC 11/22/2024 11:38:54 AM, Baghirzade, Nara) Weekly rounds in 11 North complete. See procedure for results.
| 211047 | RDS | CMPLT | CMPLT | 2024-11-22 11:41 | 2024-11-22 11:41 | Weekly Rounds: 11 South | (TC 11/22/2024 11:40:40 AM, Baghirzade, Nara) Weekly rounds in 11 South complete. See procedure for results.
1072372 | 211048 | CM | CMPLT | CMPLT | 2024-11-22 16:23 | 2024-11-22 16:24 | Repurposed for 16N surge | (TC 11/22/2024 04:23:36 PM, Cortez, John) Monitor previously used for testing, TNC1TEST Loaded default config, added spare BP hose, cuffs, charged spo2 sensor+extension cable
67955 | 211049 | CM | CMPLT | CMPLT | 2024-11-25 07:37 | 2024-11-25 07:38 | Broken latch | (TC 11/25/2024 07:37:35 AM, Rhuland, Courtney) Replaced broken latch Functional check - OK
| 211050 | RDS | CMPLT | CMPLT | 2024-11-25 08:34 | 2024-11-25 09:42 | Weekly rounds for ED
| 211051 | RDS | CMPLT | CMPLT | 2024-11-25 08:35 | 2024-11-25 14:39 | Weekly rounds for ICU
1049048 | 211052 | CM | ACTIV | ISSUE | 2024-11-25 09:06 |  | warranty exchange
1054521 | 211053 | CM | CMPLT | CMPLT | 2024-11-25 09:11 | 2024-11-25 09:11 | LA lead broken | (TC 11/25/2024 09:11:04 AM, Canjimavadivel, Kamallesh ) Replaced LA lead, tested okay with simulator. Charging batteries before sending back to unit.
| 211054 | CM | ACTIV | ISSUE | 2024-11-25 09:16 |  | ordering a couple of suction regulators
1050607 | 211055 | CM | CMPLT | CMPLT | 2024-11-25 09:39 | 2024-11-25 09:39 | the scale shuts off while in use. | (TC 11/25/2024 09:39:10 AM, Alvarado, Karla) Replaced batteries. Tested 25lb, 35lb, and 50 lb. All tests passed.
1050087 | 211056 | CM | CMPLT | CMPLT | 2024-11-25 09:41 | 2024-11-25 09:41 | check battery | (TC 11/25/2024 09:41:19 AM, Alvarado, Karla) Replaced batteries. Tested 5kg and 10 kg. All tests passed.
1048040 | 211057 | PS | CMPLT | CMPLT | 2024-11-25 09:42 | 2024-11-25 09:43 | Occlusion sensor recall# 2024110149 | (TC 11/25/2024 09:42:56 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall. Checked by Jessica.
1048041 | 211058 | PS | CMPLT | CMPLT | 2024-11-25 09:42 | 2024-11-25 09:44 | Occlusion sensor recall# 2024110149 | (TC 11/25/2024 09:43:45 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall. Checked by Jessica.
1048042 | 211059 | PS | CMPLT | CMPLT | 2024-11-25 09:42 | 2024-11-25 09:45 | Occlusion sensor recall# 2024110149 | (TC 11/25/2024 09:44:36 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall. Checked by Jessica.
| 211060 | CM | ACTIV | ISSUE | 2024-11-25 09:52 |  | SCD devices for L&D | (TC 11/25/2024 09:52:10 AM, Baghirzade, Nara) Picking up all SCD devices for L&D from 6th fl. storage. (TC 11/25/2024 11:25:55 AM, Baghirzade, Nara) Retrieved old SCD devices to swap with the new ones. Dropped off new SCD devices with Kim. Will retire old devices and store them on the 6th fl. storage area. Followed up with Becky.
1060974 | 211061 | CM | ACTIV | ISSUE | 2024-11-25 10:20 |  | Screen not working | (TC 11/25/2024 10:29:56 AM, Reid, Emmerson ) Saw when the team was using this device the right-most screen was not lit up. I switched the device with another one from the hallway. Testing it now the screen is turning on. I will leave it on for a bit to see if anything changes.
| 211062 | CM | CMPLT | CMPLT | 2024-11-25 11:27 | 2024-11-25 11:27 | PFT upgrade overview meeting | (TC 11/25/2024 11:27:28 AM, Baghirzade, Nara) PFT upgrade overview meeting
1070738 | 211063 | CM | CMPLT | CMPLT | 2024-11-25 12:29 | 2024-11-29 12:04 | Not turning on | (TC 11/25/2024 12:29:42 PM, Rhuland, Courtney) Device would not boot up passed the Giraffe screen Opened back of device and manually put omnibed into service mode Device turned on in service mode and error log showed system failure 70 - unexpected software error Put device back into normal mode and device powered on normally Calibrated RH Missing SE flange corner - waiting on parts (TC 11/29/2024 12:03:53 PM, Rhuland, Courtney) Replaced missing part
1065521 | 211064 | CM | CMPLT | CMPLT | 2024-11-25 14:07 | 2024-11-25 14:07 | Removing from service | (TC 11/25/2024 02:07:07 PM, Baghirzade, Nara) Retrieved device from L&D. Will retire and store on 6th fl. Replaced with new SCD device. Followed up with Becky.
1065528 | 211065 | CM | CMPLT | CMPLT | 2024-11-25 14:09 | 2024-11-25 14:09 | Removing from service | (TC 11/25/2024 02:08:37 PM, Baghirzade, Nara) Retrieved device from L&D. Will retire and store on 6th fl. Replaced with new SCD device. Followed up with Becky.
1065525 | 211066 | CM | CMPLT | CMPLT | 2024-11-25 14:10 | 2024-11-25 14:10 | Removing from service | (TC 11/25/2024 02:09:37 PM, Baghirzade, Nara) Retrieved device from L&D. Will retire and store on 6th fl. Replaced with new SCD device. Followed up with Becky.
| 211067 | CM | CMPLT | CMPLT | 2024-11-25 14:14 | 2024-11-25 14:14 | Biomed PM huddle | (TC 11/25/2024 02:13:37 PM, Baghirzade, Nara) Biomed PM huddle
1072376 | 211068 | ACC | ACTIV | ISSUE | 2024-11-25 14:53 |  | Acceptance | (TC 11/25/2024 02:53:34 PM, Fan, Tak) MIF and NAC req.
1046480 | 211069 | CM | ACTIV | ISSUE | 2024-11-25 15:03 |  | Sent for PM under consignment | Unit was stored in biomed as a spare for a while and need t get a PM done before we put it back in service. Contacted rep. Was advised issue no charge PO and send for PM to conmed. Createed SC. Requested approval. Received.
1060334 | 211070 | CM | ACTIV | ISSUE | 2024-11-25 15:12 |  | Sent for PM under consignment | Nov 25 - Unit was stored in biomed as a spare for a while and need t get a PM done before we put it back in service. Contacted rep. Was advised issue no charge PO and send for PM to conmed. Createed SC. Requested approval. Received. Nov 28 - Advised MMC to issue the PO. Coordinate and shipped unit with the sales rep
1066482 | 211071 | CM | CMPLT | CMPLT | 2024-11-25 23:35 | 2024-11-25 17:35 | MSH 16 CATH LAB SITE RITE 8 POWER CORD REQ CORD CAP REPLACED | (TC 11/25/2024 05:35:00 PM, Whittington, Scott) MSH 16 CATH LAB SITE RITE 8 POWER CORD REQ CORD CAP REPLACED | MSH 16 CATH LAB SITE RITE 8 POWER CORD REQ CORD CAP REPLACED
| 211072 | CM | CMPLT | CMPLT | 2024-11-26 07:42 | 2024-11-29 12:04 | Omnibed parts
66681 | 211073 | CM | CMPLT | CMPLT | 2024-11-26 08:40 | 2024-11-26 11:14 | Doors will not latch | (TC 11/26/2024 11:13:33 AM, Rhuland, Courtney) Incubator sent down with two of the same doors Replaced wrong door with spare Other side door was missing spring on latch Replaced North latch During functional check, device failed RH calibration multiple times RH sensor found to be corroded Replaced sensor and calibrated RH Functional check - OK
1061210 | 211074 | CM | CMPLT | CMPLT | 2024-11-26 10:20 | 2024-11-26 17:15 | Monitor Tilting Issue | Call re: Monitors are tilting back when adjusted to an angle. Went tot check. Waited for access. Found two monitors with same issue. Removed side arm cover and tightned screws. Checked functionality. Tested ok.
68537 | 211075 | CM | ACTIV | ISSUE | 2024-11-26 10:51 |  | Pulled from case, not going into Auto mode
1067437 | 211076 | FC | CMPLT | CMPLT | 2024-11-26 12:36 | 2024-11-26 12:37 | QC checklist | (TC 11/26/2024 12:36:37 PM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
66517 | 211077 | PM | CMPLT | CMPLT | 2024-11-26 12:42 | 2024-11-26 12:42 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 11/26/2024 12:41:49 PM, Baghirzade, Nara) Touch was completely out of calibration, unable to enter service mode. Connected FM to Philips Support Tool and calibrated touch from there. Touch is fully functional. AC LED was not lit up when connected to AC; switch board found defective. Replaced switch board. PM completed. Recorder cleaned and calibrated. Functional testing passed; NIBP, SpO2, printer, and 4 transducer ports are functional.
67423 | 211078 | PM | CMPLT | CMPLT | 2024-11-26 12:51 | 2024-11-26 12:51 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 11/26/2024 12:51:23 PM, Baghirzade, Nara) PM completed.
67412 | 211079 | PM | CMPLT | CMPLT | 2024-11-26 13:04 | 2024-11-26 13:04 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 11/26/2024 01:04:26 PM, Baghirzade, Nara) Replaced demand valve and collection ring. PM completed.
| 211080 | CM | CMPLT | CMPLT | 2024-11-26 14:00 | 2024-11-26 14:00 | Scale check | (TC 11/26/2024 01:59:44 PM, Rhuland, Courtney) 2 portable infant scales used by midwives for home visits were reported to be off Tested both, one found to be off by 20g Brought 2 5kg weights over and calibrated off scale Functional check - OK for both
32520 | 211081 | CM | ACTIV | ISSUE | 2024-11-26 15:42 |  | MSH DR2 OGP MOVEMENT ISSUES | (TC 11/26/2024 03:42:14 PM, Whittington, Scott) MSH DR2 OGP MOVEMENT ISSUES (TC 11/26/2024 08:33:00 PM, Whittington, Scott) MSH DR2 OGP MOVEMENT ISSUES | MSH DR2 OGP MOVEMENT ISSUES
1058166 | 211082 | CM | CMPLT | CMPLT | 2024-11-26 16:08 | 2024-11-26 16:09 | Battery replacement | (TC 11/26/2024 04:09:12 PM, Canjimavadivel, Kamallesh ) Replaced 2x batteries
1058528 | 211083 | CM | CMPLT | CMPLT | 2024-11-26 16:11 | 2024-11-27 14:30 | Battery replacement | (TC 11/27/2024 02:30:18 PM, Canjimavadivel, Kamallesh ) 2x batteries replaced
| 211084 | CM | ACTIV | ISSUE | 2024-11-26 16:33 |  | Formlabs Order
| 211085 | CM | CMPLT | CMPLT | 2024-11-26 17:20 | 2024-11-26 17:01 | Exam Table Issue in IVF Clinic Room 1 | Call re; exam table not working and was asked to visit IVF urgently. Visited IVF. Checked unit. Found table automatically going down without pressing the foot paddle. Found foot paddle malfunction. Found foot switch corroded and top and bottom plate screw not able to hold them together. Found down switch spring detached. Rectified it. The footswitch will need to be replaced eventually. Temporarily duct taped top and bottom plates together to make it functional. Checked functionality. Tested ok.
1072380 | 211086 | CM | CMPLT | CMPLT | 2024-11-27 08:49 | 2024-11-27 08:50 | Loose lighthead | (TC 11/27/2024 08:49:33 AM, Rhuland, Courtney) Tightened screw under cap Light is functional and stays in place
1072381 | 211087 | CM | CMPLT | CMPLT | 2024-11-27 08:49 | 2024-11-27 08:50 | Loose lighthead | (TC 11/27/2024 08:50:16 AM, Rhuland, Courtney) Tightened screw under cap Light is functional and stays in place
1072382 | 211088 | CM | CMPLT | CMPLT | 2024-11-27 08:49 | 2024-11-27 08:51 | Loose lighthead | (TC 11/27/2024 08:50:37 AM, Rhuland, Courtney) Tightened screw under cap Light is functional and stays in place
1072384 | 211089 | CM | CMPLT | CMPLT | 2024-11-27 08:49 | 2024-11-27 08:51 | Loose lighthead | (TC 11/27/2024 08:50:55 AM, Rhuland, Courtney) Tightened screw under cap Light is functional and stays in place
1062661 | 211090 | CM | ACTIV | ISSUE | 2024-11-27 09:23 |  | shipping the device to ConMed for PM
1053349 | 211091 | CM | ACTIV | ISSUE | 2024-11-27 11:16 |  | MSH H2 - Hologic Selenia - Artifact In Image | (TC 11/27/2024 11:16:31 AM, Chau, Ken) Image artifact present even after gain calibration Removed covers and cleaned system for any potential dust -the artifact still present but smaller size and barely noticable, however since its still present Hologic to remote in to do pixel manipulation to see if it removes artifact. If artifact still present the tube requires replacement as a burnt spot is present
502088 | 211092 | CM | CMPLT | CMPLT | 2024-11-27 12:05 | 2024-11-27 12:05 | Persistant error | (TC 11/27/2024 12:04:45 PM, Green, Paul) Persistent error. Uncommanded motion on SLD1. Will monitor this. Was able to clear this.
1059828 | 211093 | PM | CMPLT | CMPLT | 2024-11-27 12:45 | 2024-11-27 12:45 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/27/2024 12:45:17 PM, Alvarado, Karla) Drained, disinfected, rinsed and refilled with sterile water.
1063746 | 211094 | PM | CMPLT | CMPLT | 2024-11-27 12:46 | 2024-11-27 12:46 | Cleaning-HYPO/HYPERTHERMIA UNITS | (TC 11/27/2024 12:46:17 PM, Alvarado, Karla) Drained, disinfected, rinsed and refilled with sterile water.
1058622 | 211095 | ACC | CMPLT | CMPLT | 2024-11-27 13:50 | 2024-11-27 13:51 | Incoming inspection | (TC 11/27/2024 01:50:35 PM, Rhuland, Courtney) Device was faulty - under warranty Received new device Functional check - OK New SN: 1890756 Updated warranty dates
1065350 | 211096 | ACC | CMPLT | CMPLT | 2024-11-27 13:58 | 2024-11-27 13:58 | Incoming inspection | (TC 11/27/2024 01:57:44 PM, Rhuland, Courtney) Device was faulty - under warranty Received new device Functional check - OK New SN: 1890757 Updated warranty dates
500198 | 211097 | CM | CMPLT | CMPLT | 2024-11-27 14:24 | 2024-11-27 14:24 | No network | (TC 11/27/2024 02:23:48 PM, Green, Paul) Swapped NIC cards between systems. Network now ok.
| 211098 | CM | CMPLT | CMPLT | 2024-11-27 14:30 | 2024-11-27 14:30 | ConMed Training
1064366 | 211099 | CM | ACTIV | ISSUE | 2024-11-28 06:54 |  | MSH OPG SPP GE VOLUSON TV PROBE MEMBRANE COMPROMISED | (TC 11/27/2024 06:55:00 PM, Whittington, Scott) MSH OPG SPP GE VOLUSON TV PROBE MEMBRANE COMPROMISED (TC 11/29/2024 02:25:26 PM, Dolinar, John) MSH OPG SPP GE VOLUSON TV PROBE MEMBRANE COMPROMISED | MSH OPG SPP GE VOLUSON TV PROBE MEMBRANE COMPROMISED Ref# 1-5532-3554-4295
66525 | 211100 | PM | CMPLT | CMPLT | 2024-11-28 07:52 | 2024-11-27 07:52 | PM-MAJOR - Annual PM For: Monitors, Bedside, Fetal | (TC 11/27/2024 07:52:21 AM, Baghirzade, Nara) PM completed. Recorder cleaned and calibrated. Touch calibrated. Functional testing passed; NIBP, SpO2, printer, and 4 transducer ports are functional.
67417 | 211101 | PM | CMPLT | CMPLT | 2024-11-28 07:54 | 2024-11-27 07:54 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 11/27/2024 07:53:56 AM, Baghirzade, Nara) PM completed.
67410 | 211102 | PM | CMPLT | CMPLT | 2024-11-28 07:56 | 2024-11-28 07:58 | PM-MAJOR - Annual PM For: NITRONOX INHALATION ANALGESIA SYSTEM | (TC 11/27/2024 07:57:47 AM, Baghirzade, Nara) Replaced diaphragm and scavenging tubing. PM completed.
| 211103 | TEDU | CMPLT | CMPLT | 2024-11-28 08:48 | 2024-11-27 08:48 | Dialysis troubleshooting overview | (TC 11/27/2024 08:48:29 AM, Baghirzade, Nara) Dialysis troubleshooting overview
67311 | 211104 | CM | ACTIV | ISSUE | 2024-11-28 08:52 |  | Not warming bags, stays cold | (TC 11/27/2024 08:51:51 AM, Baghirzade, Nara) Warmer stays cold, at 22C. Power cycled the warmer and came back after 30 minutes; warmer still at 22C. Removed from OR3, notified Gen and Kim. Will order bag sensor. (TC 11/28/2024 07:22:36 AM, Baghirzade, Nara) When disassembling the device, the connector that connects the main board and control/display board broke off. Samir soldered it back on. Replaced the bag sensor. Tested device and it is warming the bag; fully functional. Will assemble and mount it back on the wall in OR3.
67965 | 211105 | CM | CMPLT | CMPLT | 2024-11-28 10:15 | 2024-11-28 10:16 | Broken doors | (TC 11/28/2024 10:15:40 AM, Rhuland, Courtney) Doors came down on the wrong side Resulted in 2 broken south latches Replaced broken latches and 2 side door hinges Reached out to SA manager & NICU coordinator to reiterate that doors cannot be interchanged Functional check - OK
67311 | 211106 | PA | ACTIV | ISSUE | 2024-11-28 10:17 |  | Ordering bag sensors | (TC 11/28/2024 12:12:33 PM, Baghirzade, Nara) SC 1003361442
1059517 | 211107 | CM | CMPLT | CMPLT | 2024-11-28 11:10 | 2024-11-28 11:10 | Power cord fell into device | (TC 11/28/2024 11:09:53 AM, Reid, Emmerson ) Was able to pull the cord out of the device, from the bottom. Working now.
1061204 | 211108 | CM | CMPLT | CMPLT | 2024-11-28 11:12 | 2024-11-28 11:12 | Monitor #2 was loose | (TC 11/28/2024 11:11:47 AM, Reid, Emmerson ) tightened the monitor. Also helped with the nurse PC in OR 5
| 211109 | CM | ACTIV | ISSUE | 2024-11-28 11:33 |  | OR HUG-U-VAC Repair
| 211110 | PA | ACTIV | ISSUE | 2024-11-28 12:09 |  | Ordering rollstands for dopplers | (TC 11/28/2024 12:11:26 PM, Baghirzade, Nara) SC 1003361605 | Kim Ngo requested 3 rollstands for 3 dopplers
| 211111 | CM | CMPLT | CMPLT | 2024-11-28 12:14 | 2024-11-28 12:14 | Capital meeting | (TC 11/28/2024 12:14:15 PM, Baghirzade, Nara) Capital meeting
502492 | 211112 | PM | ACTIV | READY | 2024-11-28 14:18 |  | PM-Major, annual PM
1072394 | 211113 | ACC | CMPLT | CMPLT | 2024-11-28 14:28 | 2024-11-28 14:49 | Incoming inspection | (TC 11/28/2024 02:49:01 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1072395 | 211114 | CM | CMPLT | CMPLT | 2024-11-28 14:31 | 2024-11-28 14:49 | Incoming inspection | (TC 11/28/2024 02:48:43 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1072396 | 211115 | ACC | CMPLT | CMPLT | 2024-11-28 14:34 | 2024-11-28 14:48 | Incoming inspection | (TC 11/28/2024 02:48:16 PM, Canjimavadivel, Kamallesh ) Added to database and asset tagged
1061415 | 211116 | CM | CMPLT | CMPLT | 2024-11-28 15:01 | 2024-11-28 15:02 | Monitor tilting issue | Call re: monitor tilting issue. Went to check. Waited for access. found monitor 1 with tilting issue. Fixed it. Tested ok.
1061209 | 211117 | CM | CMPLT | CMPLT | 2024-11-28 15:03 | 2024-11-28 15:04 | Monitors Flickering | Call re: monitors flickering. Went to check. Waited for access. Rectified power cord connections on FT10 and smoke evac. Checked functionality. Tested ok.
1059536 | 211118 | CM | ACTIV | ISSUE | 2024-11-28 17:17 |  | MSH OB ANAES 15th Floor - Murray Side C probe has dropped crystals. |  | MSH OB ANAES 15th Floor - Murray Side C probe has dropped crystals. Called problem in to SonoSite. System SN# Q5GOCW Probe SN# 04XNQ7 Call Steve Arai Ext 6546 to update.
1061333 | 211119 | PS | CMPLT | CMPLT | 2024-11-29 07:15 | 2024-11-28 07:15 | Occlusion sensor recall# 2024110149 | (TC 11/28/2024 07:14:59 AM, Baghirzade, Nara) Occlusion sensor passed. Unit not affected by recall.
1061333 | 211120 | FC | CMPLT | CMPLT | 2024-11-29 07:16 | 2024-11-28 07:16 | QC checklist | (TC 11/28/2024 07:16:07 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 North.
1059778 | 211121 | PS | ACTIV | ISSUE | 2024-11-29 07:17 |  | Occlusion sensor recall# 2024110149 | (TC 11/28/2024 07:17:07 AM, Baghirzade, Nara) Unit is affected by recall - when testing the occlusion sensor, the plunger separates from the collar. Contacted Solventum to report this and they have mentioned that they currently do not have parts to service the VAC and that the VAC is fully operational even with the defective sensor; once Solventum has parts available, VAC will be shipped for repair. Contacted Sarabeth to let her know and she had mentioned that they very rarely use the bottle side of the VAC, she was ok with putting the VAC back in use. Will keep this WO open once the VAC is remediated.
1059778 | 211122 | FC | CMPLT | CMPLT | 2024-11-29 07:22 | 2024-11-28 07:22 | QC checklist | (TC 11/28/2024 07:21:47 AM, Baghirzade, Nara) QC checklist completed. Returned to 11 South.
67422 | 211123 | PA | ACTIV | ISSUE | 2024-11-29 07:28 |  | Ordering Demand Valve Exhalation Valves | (TC 11/29/2024 07:42:16 AM, Baghirzade, Nara) SC 1003362083
68628 | 211124 | CM | CMPLT | CMPLT | 2024-11-29 07:46 | 2024-11-28 07:46 | Not working. | (TC 11/28/2024 07:45:41 AM, Alvarado, Karla) Replaced power cord. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage = 79.1uA and the GND Resistance = 51 mOhms All test passed.
66014 | 211125 | CM | CMPLT | CMPLT | 2024-11-29 07:48 | 2024-11-28 07:48 | it is not charging | (TC 11/28/2024 07:47:42 AM, Alvarado, Karla) Replaced power cord. Performed tests in ASM: Download Error Log, Download Event Log and Check- in tests. Where the GND Leakage = 84.6uA and the GND Resistance =54 mOhms All test passed.
1064355 | 211126 | CM | CMPLT | CMPLT | 2024-11-29 08:23 | 2024-11-29 09:29 | V5 lead not working | (TC 11/29/2024 09:28:56 AM, Canjimavadivel, Kamallesh ) Replaced V4-V6 chest lead set. Replaced LA, RA, RL leads. Replaced keyboard cover. Tested okay with simulator
1067975 | 211127 | CM | ACTIV | ISSUE | 2024-11-29 08:26 |  | MSH ER ROOM G7 DATABASE NOT PROVIDING SUFFICIENT SPACE, DIFFICULTY REMAKING DB |  | MSH ER ROOM G7 DATABASE NOT PROVIDING SUFFICIENT SPACE, DIFFICULTY REMAKING DB
68375 | 211128 | CM | CMPLT | CMPLT | 2024-11-29 08:53 | 2024-11-29 14:29 | Intermittently losing CO2 trace on | (TC 11/29/2024 02:29:16 PM, Scero, Tomlyn) Replaced the Intellibridge EC5 on the AGM. AAs were complaining that intermittently the CO2 trace would disappear until they "wiggled" the wire at the EC5. I replaced it with a new one to see if it solves the problem.
1064851 | 211129 | CM | CMPLT | CMPLT | 2024-11-29 09:25 | 2024-11-29 14:26 | Broken APL valve | (TC 11/29/2024 02:25:43 PM, Scero, Tomlyn) APL valve was broken by the Dr. Replaced and gave the breathing system back to the AAs
1054523 | 211130 | CM | CMPLT | CMPLT | 2024-11-29 09:55 | 2024-11-29 15:14 | V4 not working
1072401 | 211131 | ACC | CMPLT | CMPLT | 2024-11-29 10:28 | 2024-11-29 10:37 | Incoming inspection | (TC 11/29/2024 10:36:37 AM, Rhuland, Courtney) Unboxed and tagged treatment cabinet Glass atomizer arrived broken - reached out to company for replacement Moved to ENT and set up Functional check - OK
1072402 | 211132 | ACC | CMPLT | CMPLT | 2024-11-29 10:28 | 2024-11-29 10:29 | Incoming inspection | (TC 11/29/2024 10:28:49 AM, Rhuland, Courtney) Unboxed and tagged treatment cabinet Moved to ENT and set up Functional check - OK
502088 | 211133 | RQ | ACTIV | ISSUE | 2024-11-29 10:34 |  | persistent error |  | occurred during collimator exchange
69318 | 211134 | CM | CMPLT | CMPLT | 2024-11-29 11:36 | 2024-11-29 11:36 | Not working. | (TC 11/29/2024 11:35:44 AM, Alvarado, Karla) The pump came to biomed department with a PCU that was not charging. During the functional test, the rate accuracy failed. Therefore, the rate calibration was performed twice and the pump still failed. The bezel was replaced. Performed pressure calibration, rate calibration, and check-in test. All test passed
1069389 | 211135 | CM | CMPLT | CMPLT | 2024-11-29 13:11 | 2024-11-29 13:11 | Deaeration chamber not filling up during priming | (TC 11/29/2024 01:10:58 PM, Canjimavadivel, Kamallesh ) Set up and ran patient simulation. Deaeration chamber filled up as expected during first prime cycle. Ran continuously for 1 hour with no alarms triggered. | Attempted to prime 4 times and did not fill up
66436 | 211136 | CM | CMPLT | CMPLT | 2024-11-29 14:06 | 2024-11-29 14:25 | not all leads working | (TC 11/29/2024 02:24:36 PM, Scero, Tomlyn) Confirmed on simulator RL and V not working. Replaced with new leads, tested on simulator and returned the working unit to CCU.
1072405 | 211137 | PA | CMPLT | CMPLT | 2024-11-29 15:05 | 2024-11-29 15:05 | Incoming Inspection | (TC 11/29/2024 03:05:12 PM, Reid, Emmerson ) No electrical Safety. Tag created. Incoming Complete. Copied from asset this one is replacing (1065273)
1056435 | 211138 | CM | CMPLT | NFF | 2024-11-29 15:27 | 2024-11-29 15:27 | Screen not lighting up even when plugged in. | (TC 11/29/2024 03:27:29 PM, Alvarado, Karla) Performed tests: Temperature NIBP= set: 120/80, measured: 116/83 SpO2 = set: 95% and 40 b/m, measured: 96% and 40 b/m. All test passed. No Fault Found (NFF)
1053347 | 211139 | CM | ACTIV | ISSUE | 2024-11-29 16:11 |  | 8X10 comperssion paddle not being recognized. |  | 8X10 comperssion paddle not being recognized. Error GCB(14:89) System reboot cleared the problem. Ref# 01656035